## Data Flags Legend
High internal standard recovery, corresponding analyte in sample. Reviewed and not disqualifying | IS1 | Internal standard recovery was above the method acceptance limits of 150%. Target analytes were detected in the sample. Impact on results is not expected based on analyst review. See comments. | Analyst and validator concurrence required to release data. If this is not met, reanalysis and/or repreparation is required.
High internal standard recovery, no analyte in sample | IS2 | Internal standard recovery was above the method acceptance limits. No target analytes were detected in the sample. | Analyst and validator concurrence required to release data. If this is not met, reanalysis and/or repreparation is required.
Low internal standard recovery, corresponding analyte in sample. Reviewed and not disqualifying | IS3 | Internal standard recovery was below the method acceptance limit of 50%. Target analytes were detected in the sample. Impact on results is not expected based on analyst review. See comments. | Analyst and validator concurrence required to release data. If this is not met, reanalysis and/or repreparation is required.
Low internal standard recovery, no analyte in sample. | IS4 | Internal standard recovery was below the method acceptance limits. No target analytes were detected in the sample. | Analyst and validator concurrence required to release data. If this is not met, reanalysis and/or repreparation is required.
Internal standards outside limit, disqualifying. | IS5 | Internal standard recovery was outside of method limits of 50-150%, and upon review is expected to impact data quality. Data not useable. [Margie would rather we said things like 'data biased low', but my impression is you prefer a go/no-go for data, right?] | Sample re-analysis required; If failure persists, re-prepare sample and re-analyze.
Laboratory duplicate precision failure | D | RPD/RSD for duplicate analysis exceeded method acceptance. | Sample re-analysis required; If failure persists, further troubleshooting is required to improve precision and all samples analyzed since previous passing duplicate must be re-analyzed.
Laboratory Reagent Blank (LRB) failure | B | Result is >50% of MRL. | Flag LRB failure and sample results within 10-fold of the LRB concentration as biased high.
Continuing Calibration Check (CCC) failure | V1 | Continuing calibration check (CCC) recovery was outside the method acceptance limits (70-130%). | Samples analyzed since previous passing CCC must be re-analyzed. If failure persists, re-calibration is required. Upon recalibration, all other QC must be re-analyzed as well.
Quality Control Sample (QCS) failure | V2 | Quality control sample (QCS) recovery was outside of the method acceptance limits (70-130%). | Re-analysis of QCS required; if failure persists, re-preparation of QCS is required and re-calibration may be necessary. Upon recalibration, all other QC must be re-analyzed as well. No sample results may be reported if QCS failure occurs in batch.
Inter-Batch Consistency check (IBC) failure | V3 | The inter-batch consistency check (IBC) sample was outside of acceptance criteria (80-120%). This suggests a systematic bias for the analytical batch relative to prior sample batches. | Compare peak areas of calibration standards to those of previous batches to determine if source of failure is the calibration curve or the IBC itself. For IBC failure, re-preparation and/or re-analysis of IBC is required. For calibration curve failure, re-preparation and/or re-analysis of calibration curve is required. Upon recalibration, all other QC must be re-analyzed as well. No sample results may be reported if IBC failure occurs in batch.
High Laboratory Fortified Blank (LFB) recovery | L1 | LFB results >150% for LFB-Low or >130% for LFB-High. Affected sample results biased high. | Re-analysis of LFB required; If failure persists, re-prepare LFB and re-analyze. Re-preparation of spiking solution (PDS/SDS) may also be necessary. Affected samples, i.e. all samples analyzed since previous passing LFB, must be flagged.
Low Laboratory Fortified Blank (LFB) recovery | L2 | LFB results <50% for LFB-Low or <70% for LFB-High. Affected sample results biased low. | Re-analysis of LFB required; If failure persists, re-prepare LFB and re-analyze. Re-preparation of spiking solution (PDS/SDS) may also be necessary. Affected samples, i.e. all samples analyzed since previous passing LFB, must be flagged.
Laboratory Fortified Sample Matrix (LFSM) failure | M1 | LFSM recovery outside | Absent LFB issues:
|  | 70-130% for LFSM-High or | Re-analysis of LFSM is required; Passing LFB confirms proper preparation and analysis so no further troubleshooting is necessary. If failure persists, samples analyzed since previous passing LFSM of same type (-Low/-High) must be flagged.
50-150% for LFSM-Low. Sample results may be biased proportionally due to matrix interferences.
Coordinating LFSM/LFB failures:
|  | Note: Simultaneous failure of LFB and LFSM indicates systematic bias in the preparation or analysis processes. | Re-preparation and re-analysis of LFB/LFSM required. Re-preparation of spiking solution (PDS/SDS) may also be necessary. If failure persists, samples analyzed since previous passing set of LFSM and LFB must be flagged.
LFSM interference | M2 | Spike recovery unable to be calculated due to other factors, e.g. spiking concentration too low. | Re-preparation and re-analysis of LFSM required. If failure persists, samples analyzed since previous passing LFSM of same type (-Low/-High) must be flagged. Analyst must describe in comments.
Upper calibration limit exceeded | UL | Sample results are above the upper calibration limit. | Sample must be re-analyzed at a higher dilution factor to obtain results within calibration range. Adjacent samples must be inspected for signs of carryover (see flag “C”).
Lower calibration limit exceeded for diluted sample | LL | Sample dilution factor is greater than “2” but results are <MRL. Sample was therefore over-diluted. | Sample must be re-analyzed at a lower dilution factor, undiluted if necessary, to obtain results within calibration range.
Carryover/Contamination observed | C | Carryover and/or contamination observed in sample. Sample results are suspected to be biased high. | Sample re-analysis is required.

## Template
Batch Name: | [5 empty] | SA: | Area: | SAH: | AreaH: | Last Acquisition Date: |  |  | Sample Count: |  | Analytes | PFBA | L-PFBS | HFPO-DA | PFHxA | L-PFHxS | PFOA | PFNA | L-PFOS | PFDA | HCal Std |  | Sheet1_1 | [15 empty] | Verifier Count | Validator Count | MPFBA | M3PFBS | MPHFPO-DA | M5PFHxA | M3PFHxS | M8PFOA | M9PFNA | M8PFOS | M6PFDA |  | Type | [7 empty] | D (LD) | K (LRB) | M (Cal) | CCC/QCS | S (LFSM) | Z (LFB) | IBC | <MRL | <MRL OVRD
2022-07-15-AB2 | [5 empty] | Sheet2 | Sheet4 | 8 | 8 | 2022-07-26 04:21 |  |  | 342 |  | MRL> | 200 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 0 |  | Sheet3 | [15 empty] | 0 | 0 | 63255 | 1143888 | 365038 | 3150523 | 2633040 | 15825420 | 20097918 | 3718322 | 31298744 |  | U |  |  |  |  | Quality Control Maximum Recovery %: |  |  | 30% | 50% | 30% | 30% | 30% | 30% | 20% | 50% | ON
[6 empty] | Measured Concentration, ng/L | [9 empty] | Concentration, ng/L, Corrected for Sample Dilution | [10 empty] | Sheet4 |  |  | Flags | [8 empty] | Re-Analysis Flag |  |  | Verifier/Validator Initials |  |  | Area (Internal Standard) | [17 empty] | Measured Concentration, ng/L | [9 empty] | Quality Control Checks | [18 empty] | These fields are editable with password: 'fixit'
# | Time Stamp: | Sample ID: | # | Type | QC? | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | Dilution Factor: | C-PFBA | C-PFBS | C-HFPO-DA | C-PFHxA | C-PFHxS | C-PFOA | C-PFNA | C-PFOS | C-PFDA | True Value | True Value 2 | True Value 3 | Sample Type | Flags (342) | F-PFBA | F-PFBS | F-HFPO-DA | F-PFHxA | F-PFHxS | F-PFOA | F-PFNA | F-PFOS | F-PFDA |  | List for drop down flags | Re-Analysis Flags | Sample Comments | Verifier Comments | Validator Comments | MPFBA | MPFBS | MPHFPO-DA | MPFHxA | MPFHxS | MPFOA | MPFNA | MPFOS | MPFDA | IS Comments | M&Q | K&Z | D | S | # | QC Samples | Type | Dilution Factor: | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | True Value | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | PFBA | PFBS | GenX | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | Pairing
1 | 2022-07-16 18:08 | AB-220715-01 Cal Std 1 1 | 0 | M | QC | 0 | 1.235 | 0 | 0 | 1.65 | 1.244 | 1.19 | 1.566 | 1.209 | 1 | 0 | 1.235 | 0 | 0 | 1.65 | 1.244 | 1.19 | 1.566 | 1.209 | Std 1 1 | Std 1 1 | NA | Calib Std | No TrueV! | [10 empty] | IS1 | R1 |  |  |  | 61078 | 1500040 | 480815 | 4123113 | 3726379 | 22704030 | 27045316 | 4705476 | 40182026 |  | 1 |  |  |  | 1 | AB-220715-01 Cal Std 1 1 | M | 1 | 0 | 1.235 | 0 | 0 | 1.65 | 1.244 | 1.19 | 1.566 | 1.209 | Std 1 1
2 | 2022-07-16 18:44 | AB-220715-01 Cal Std 2 1 | 1 | M | QC | 0 | 1.809 | 2.756 | 2.915 | 2.259 | 2.18 | 2.12 | 2.219 | 2.178 | 1 | 0 | 1.809 | 2.756 | 2.915 | 2.259 | 2.18 | 2.12 | 2.219 | 2.178 | Std 2 1 | Std 2 1 | NA | Calib Std | No TrueV! | [10 empty] | IS2 | R2 |  |  |  | 35290 | 1148968 | 340592 | 3078995 | 2883380 | 17285957 | 22129693 | 4017638 | 33735557 |  | 2 |  |  |  | 2 | AB-220715-01 Cal Std 2 1 | M | 1 | 0 | 1.809 | 2.756 | 2.915 | 2.259 | 2.18 | 2.12 | 2.219 | 2.178 | Std 2 1
3 | 2022-07-16 19:20 | AB-220715-01 Cal Std 3 1 | 2 | M | QC | 0 | 3.71 | 4.106 | 4.587 | 3.685 | 3.99 | 3.898 | 3.587 | 4.036 | 1 | 0 | 3.71 | 4.106 | 4.587 | 3.685 | 3.99 | 3.898 | 3.587 | 4.036 | Std 3 1 | Std 3 1 | NA | Calib Std | No TrueV! | [10 empty] | IS3 | R3 |  |  |  | 61302 | 1320162 | 406880 | 3674530 | 3410846 | 20100191 | 25080673 | 4518567 | 39237316 |  | 3 |  |  |  | 3 | AB-220715-01 Cal Std 3 1 | M | 1 | 0 | 3.71 | 4.106 | 4.587 | 3.685 | 3.99 | 3.898 | 3.587 | 4.036 | Std 3 1
4 | 2022-07-16 19:56 | AB-220715-01 Cal Std 4 1 | 3 | M | QC | 0 | 8.981 | 7.465 | 9.617 | 7.953 | 9.57 | 9.647 | 8.162 | 9.592 | 1 | 0 | 8.981 | 7.465 | 9.617 | 7.953 | 9.57 | 9.647 | 8.162 | 9.592 | Std 4 1 | Std 4 1 | NA | Calib Std | No TrueV! | [10 empty] | IS4 |  |  |  |  | 73743 | 1275747 | 433529 | 3599898 | 3262781 | 19694899 | 25586736 | 4610237 | 40566614 |  | 4 |  |  |  | 4 | AB-220715-01 Cal Std 4 1 | M | 1 | 0 | 8.981 | 7.465 | 9.617 | 7.953 | 9.57 | 9.647 | 8.162 | 9.592 | Std 4 1
5 | 2022-07-16 20:32 | AB-220715-01 Cal Std 5 1 | 4 | M | QC | 0 | 15.471 | 20.512 | 17.849 | 14.817 | 18.982 | 19.29 | 15.556 | 19.349 | 1 | 0 | 15.471 | 20.512 | 17.849 | 14.817 | 18.982 | 19.29 | 15.556 | 19.349 | Std 5 1 | Std 5 1 | NA | Calib Std | No TrueV! | [10 empty] | IS5 |  |  |  |  | 61896 | 1277711 | 394758 | 3438958 | 3040930 | 18617801 | 23616240 | 4289270 | 36873395 |  | 5 |  |  |  | 5 | AB-220715-01 Cal Std 5 1 | M | 1 | 0 | 15.471 | 20.512 | 17.849 | 14.817 | 18.982 | 19.29 | 15.556 | 19.349 | Std 5 1
6 | 2022-07-16 21:08 | AB-220715-01 Cal Std 6 1 | 5 | M | QC | 0 | 32.123 | 39.383 | 35.309 | 29.069 | 38.825 | 38.945 | 31.304 | 38.803 | 1 | 0 | 32.123 | 39.383 | 35.309 | 29.069 | 38.825 | 38.945 | 31.304 | 38.803 | Std 6 1 | Std 6 1 | NA | Calib Std | No TrueV! | [10 empty] | D |  |  |  |  | 63255 | 1138115 | 364862 | 3150523 | 2895638 | 17085621 | 22429940 | 4084789 | 35342251 |  | 6 |  |  |  | 6 | AB-220715-01 Cal Std 6 1 | M | 1 | 0 | 32.123 | 39.383 | 35.309 | 29.069 | 38.825 | 38.945 | 31.304 | 38.803 | Std 6 1
7 | 2022-07-16 21:44 | AB-220715-01 Cal Std 7 1 | 6 | M | QC | 0 | 78.091 | 88.145 | 90.964 | 72.445 | 96.896 | 96.207 | 76.777 | 96.261 | 1 | 0 | 78.091 | 88.145 | 90.964 | 72.445 | 96.896 | 96.207 | 76.777 | 96.261 | Std 7 1 | Std 7 1 | NA | Calib Std | No TrueV! | [10 empty] | B |  |  |  |  | 59957 | 1063643 | 334185 | 2802723 | 2332530 | 13707584 | 19642767 | 3718322 | 31298744 |  | 7 |  |  |  | 7 | AB-220715-01 Cal Std 7 1 | M | 1 | 0 | 78.091 | 88.145 | 90.964 | 72.445 | 96.896 | 96.207 | 76.777 | 96.261 | Std 7 1
8 | 2022-07-16 22:20 | AB-220715-01 Cal Std 8 1 | 7 | M | QC | 192.335 | 165.757 | 187.477 | 189.927 | 158.86 | 195 | 197.798 | 164.702 | 195.086 | 1 | 192.335 | 165.757 | 187.477 | 189.927 | 158.86 | 195 | 197.798 | 164.702 | 195.086 | Std 8 1 | Std 8 1 | NA | Calib Std | No TrueV! | [10 empty] | V1 |  |  |  |  | 67846 | 1192670 | 413789 | 3471177 | 2633040 | 15825420 | 20097918 | 3565013 | 29182620 |  | 8 |  |  |  | 8 | AB-220715-01 Cal Std 8 1 | M | 1 | 192.335 | 165.757 | 187.477 | 189.927 | 158.86 | 195 | 197.798 | 164.702 | 195.086 | Std 8 1
9 | 2022-07-16 22:56 | AB-220715-01 Cal Std 9 1 | 8 | M | QC | 410.126 | 345.579 | 413.164 | 381.334 | 319.1 | 385.596 | 403.989 | 341.327 | 385.574 | 1 | 410.126 | 345.579 | 413.164 | 381.334 | 319.1 | 385.596 | 403.989 | 341.327 | 385.574 | Std 9 1 | Std 9 1 | NA | Calib Std | No TrueV! | [10 empty] | V2 |  |  |  |  | 67795 | 1143888 | 396794 | 3374919 | 2346729 | 13421321 | 17013245 | 3123532 | 24297545 |  | 9 |  |  |  | 9 | AB-220715-01 Cal Std 9 1 | M | 1 | 410.126 | 345.579 | 413.164 | 381.334 | 319.1 | 385.596 | 403.989 | 341.327 | 385.574 | Std 9 1
10 | 2022-07-16 23:32 | AB-220715-01 Cal Std 10 1 | 9 | M | QC | 930.971 | 830.668 | 983.792 | 867.467 | 810.873 | 924.669 | 919.292 | 841.7 | 924.724 | 1 | 930.971 | 830.668 | 983.792 | 867.467 | 810.873 | 924.669 | 919.292 | 841.7 | 924.724 | Std 10 1 | Std 10 1 | NA | Calib Std | No TrueV! | [10 empty] | V3 |  |  |  |  | 72321 | 1004844 | 365038 | 2835104 | 1657654 | 9637873 | 12925865 | 2277056 | 17020941 |  | 10 |  |  |  | 10 | AB-220715-01 Cal Std 10 1 | M | 1 | 930.971 | 830.668 | 983.792 | 867.467 | 810.873 | 924.669 | 919.292 | 841.7 | 924.724 | Std 10 1
11 | 2022-07-17 00:08 | AB-220715-01 Cal Std 11 1 | 10 | M | QC | 2075.212 | 1747.461 | 1959.323 | 1844.46 | 1770.213 | 1893.058 | 1936.193 | 1829.534 | 1916.836 | 1 | 2075.212 | 1747.461 | 1959.323 | 1844.46 | 1770.213 | 1893.058 | 1936.193 | 1829.534 | 1916.836 | Std 11 1 | Std 11 1 | NA | Calib Std | No TrueV! |  |  |  |  | IS3 | IS3 | IS3 | IS3 | IS3 |  | L1 |  |  |  |  | 65082 | 763538 | 318901 | 2281360 | 1116701 | 6720492 | 8862935 | 1590397 | 11373297 |  | 11 |  |  |  | 11 | AB-220715-01 Cal Std 11 1 | M | 1 | 2075.212 | 1747.461 | 1959.323 | 1844.46 | 1770.213 | 1893.058 | 1936.193 | 1829.534 | 1916.836 | Std 11 1
12 | 2022-07-17 00:44 | AB-220715-01 Cal Std 12 1 | 11 | M | QC | 4301.969 | 3719.585 | 4112.801 | 4155.832 | 4004.568 | 3945.156 | 4063.315 | 3973.032 | 3952.891 | 1 | 4301.969 | 3719.585 | 4112.801 | 4155.832 | 4004.568 | 3945.156 | 4063.315 | 3973.032 | 3952.891 | Std 12 1 | Std 12 1 | NA | Calib Std | No TrueV! |  |  |  | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  | L2 |  |  |  |  | 69192 | 580441 | 237441 | 1574116 | 694484 | 4250243 | 5789422 | 1079609 | 7246388 |  | 12 |  |  |  | 12 | AB-220715-01 Cal Std 12 1 | M | 1 | 4301.969 | 3719.585 | 4112.801 | 4155.832 | 4004.568 | 3945.156 | 4063.315 | 3973.032 | 3952.891 | Std 12 1
13 | 2022-07-17 01:20 | AB-220715-01 Cal Std 13 1 | 12 | M | QC | 9689.387 | 8816.874 | 9957.078 | 10175.739 | 9042.279 | 10261.834 | 10085.115 | 9202.601 | 10230.461 | 1 | 9689.387 | 8816.874 | 9957.078 | 10175.739 | 9042.279 | 10261.834 | 10085.115 | 9202.601 | 10230.461 | Std 13 1 | Std 13 1 | NA | Calib Std | No TrueV! |  | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  | M1 |  |  |  |  | 54188 | 356739 | 145507 | 906084 | 387630 | 2058190 | 3012684 | 637062 | 3649764 |  | 13 |  |  |  | 13 | AB-220715-01 Cal Std 13 1 | M | 1 | 9689.387 | 8816.874 | 9957.078 | 10175.739 | 9042.279 | 10261.834 | 10085.115 | 9202.601 | 10230.461 | Std 13 1
14 | 2022-07-17 01:56 | AB-220715-01 LFB-H 1 | 13 | Z | QC | 184.824 | 167.941 | 199.413 | 198.905 | 154.586 | 195.518 | 197.021 | 168.032 | 196.5 | 1 | <MRL | 167.941 | 199.413 | 198.905 | 154.586 | 195.518 | 197.021 | 168.032 | 196.5 | H 1 | H 1 | NA | Blank Spike | No TrueV! | [10 empty] | M2 |  |  |  |  | 72806 | 1258466 | 433139 | 3578022 | 2609141 | 15840003 | 20165202 | 3363767 | 28986281 |  |  | 14 |  |  | 17 | AB-220715-01 CCC-H 1 | Q | 1 | 402.94 | 352.033 | 402.071 | 385.132 | 325.628 | 391.959 | 401.704 | 344.494 | 398.521 | 1 | 40294% | 35203% | 40207% | 38513% | 32563% | 39196% | 40170% | 34449% | 39852% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
15 | 2022-07-17 02:32 | AB-220715-01 LRB 1_INJ001 | 14 | K | QC | 0 | 0 | 0 | 1.749 | 0 | 1.349 | 0.917 | 1.258 | 0.984 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [10 empty] | UL |  |  |  |  | 56360 | 1346465 | 442707 | 3827595 | 3154476 | 19792477 | 24084153 | 4180363 | 37971765 |  |  | 15 |  |  | 18 | AB-220715-01 IBC 1 | Q | 1 | 191.996 | 173.785 | 197.324 | 192.575 | 163.267 | 196.228 | 202.398 | 171.927 | 195.534 | 1 | 19200% | 17378% | 19732% | 19258% | 16327% | 19623% | 20240% | 17193% | 19553% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
16 | 2022-07-17 03:08 | AB-220715-01 LRB 1 | 15 | K | QC | 0 | 0 | 0 | 3.07 | 0 | 1.317 | 0.967 | 1.494 | 1.116 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | [9 empty] | LL |  |  |  |  | 30308 | 889053 | 286866 | 2587113 | 2553652 | 15886293 | 19918783 | 3533089 | 31479184 |  |  | 16 |  |  | 19 | AB-220715-01 QCS 1 | Q | 1 | 180.3 | 176.792 | 153.347 | 187.09 | 161.888 | 193.982 | 198.336 | 172.792 | 191.245 | 1 | 18030% | 17679% | 15335% | 18709% | 16189% | 19398% | 19834% | 17279% | 19124% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
17 | 2022-07-17 03:44 | AB-220715-01 CCC-H 1 | 16 | Q | QC | 402.94 | 352.033 | 402.071 | 385.132 | 325.628 | 391.959 | 401.704 | 344.494 | 398.521 | 1 | 402.94 | 352.033 | 402.071 | 385.132 | 325.628 | 391.959 | 401.704 | 344.494 | 398.521 | 1 | 01 | NA | CCC | UL! | [10 empty] | C |  |  |  |  | 62643 | 1075641 | 403694 | 3215096 | 2144318 | 12833415 | 16388396 | 2836570 | 23483308 |  | 17 |  |  |  | 46 | AB-220715-01 CCC-H 2 | Q | 1 | 364.718 | 347.787 | 399.928 | 380.048 | 320.885 | 378.815 | 387.617 | 336.485 | 386.616 | 1 | 36472% | 34779% | 39993% | 38005% | 32088% | 37882% | 38762% | 33648% | 38662% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
18 | 2022-07-17 04:20 | AB-220715-01 IBC 1 | 17 | Q | QC | 191.996 | 173.785 | 197.324 | 192.575 | 163.267 | 196.228 | 202.398 | 171.927 | 195.534 | 1 | <MRL | 173.785 | 197.324 | 192.575 | 163.267 | 196.228 | 202.398 | 171.927 | 195.534 | 1 | NA | 1 | IBC | UL! | [15 empty] | 68645 | 1220172 | 434362 | 3566735 | 2572751 | 15612554 | 20318022 | 3393225 | 31908879 |  | 18 |  |  |  | 74 | AB-220715-01 CCC-H 3 | Q | 1 | 361.667 | 340.157 | 397.811 | 382.755 | 317.243 | 391.053 | 391.511 | 332.226 | 389.63 | 1 | 36167% | 34016% | 39781% | 38276% | 31724% | 39105% | 39151% | 33223% | 38963% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
19 | 2022-07-17 04:56 | AB-220715-01 QCS 1 | 18 | Q | QC | 180.3 | 176.792 | 153.347 | 187.09 | 161.888 | 193.982 | 198.336 | 172.792 | 191.245 | 1 | <MRL | 176.792 | 153.347 | 187.09 | 161.888 | 193.982 | 198.336 | 172.792 | 191.245 | 1 | NA | 1 | QCS | UL! | [15 empty] | 66107 | 1209903 | 424097 | 3427631 | 2669141 | 15921519 | 20459537 | 3375356 | 29652036 |  | 19 |  |  |  | 102 | AB-220715-01 CCC-H 4 | Q | 1 | 336.451 | 332.142 | 388.596 | 377.704 | 321.727 | 387.419 | 386.028 | 331.024 | 382.399 | 1 | 33645% | 33214% | 38860% | 37770% | 32173% | 38742% | 38603% | 33102% | 38240% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
20 | 2022-07-17 05:32 | WO-22270-01-A-1 | 19 | U |  | 0 | 0 | 0 | 0 | 0 | 2.455 | 1.284 | 1.107 | 4.515 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8248 | 192834 | 67698 | 647220 | 908010 | 5930160 | 10683728 | 2506050 | 21056090 | [5 empty] | 130 | AB-220715-01 CCC-H 5 | Q | 1 | 356.628 | 338.354 | 389.461 | 384.448 | 315.397 | 376.568 | 395.37 | 332.343 | 386.242 | 1 | 35663% | 33835% | 38946% | 38445% | 31540% | 37657% | 39537% | 33234% | 38624% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
21 | 2022-07-17 06:08 | WO-22270-02-A-1 | 20 | U |  | 0 | 0 | 0 | 0 | 0 | 2.629 | 0.801 | 1.166 | 4.114 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 9251 | 236621 | 83138 | 786145 | 1029128 | 6874699 | 12208610 | 2904969 | 23902714 | [5 empty] | 158 | AB-220715-01 CCC-H 6 | Q | 1 | 377.628 | 350.691 | 375.438 | 368.549 | 302.006 | 372.428 | 377.013 | 313.329 | 378.911 | 1 | 37763% | 35069% | 37544% | 36855% | 30201% | 37243% | 37701% | 31333% | 37891% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
22 | 2022-07-17 06:44 | WO-22270-03-A-1 | 21 | U |  | 0 | 0 | 0 | 0 | 0 | 1.389 | 0.734 | 0.982 | 3.879 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 10607 | 238397 | 81746 | 797719 | 1021694 | 6865251 | 12613090 | 2919310 | 24275846 | [5 empty] | 186 | AB-220715-01 CCC-H 7 | Q | 1 | 369.407 | 335.708 | 385.941 | 375.749 | 303.514 | 387.221 | 384.409 | 317.825 | 372.392 | 1 | 36941% | 33571% | 38594% | 37575% | 30351% | 38722% | 38441% | 31782% | 37239% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
23 | 2022-07-17 07:20 | WO-22270-04-A-1 | 22 | U |  | 0 | 0 | 0 | 0 | 0 | 1.569 | 0.854 | 1.082 | 3.959 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8158 | 213958 | 78456 | 700096 | 971558 | 6380733 | 11465262 | 2739134 | 22475000 | [5 empty] | 214 | AB-220715-01 CCC-H 8 | Q | 1 | 380.715 | 336.433 | 385.036 | 378.446 | 316.872 | 374.878 | 379.099 | 325.053 | 375.361 | 1 | 38072% | 33643% | 38504% | 37845% | 31687% | 37488% | 37910% | 32505% | 37536% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
24 | 2022-07-17 07:56 | WO-22270-05-A-1 | 23 | U |  | 0 | 0 | 0 | 0 | 0 | 3.116 | 0.677 | 1.161 | 4.702 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 9053 | 256964 | 89911 | 832899 | 1055212 | 6786453 | 12838486 | 2931235 | 24352740 | [5 empty] | 242 | AB-220715-01 CCC-H 9 | Q | 1 | 361.131 | 341.763 | 369.546 | 391.441 | 304.674 | 386.79 | 391.626 | 311.608 | 363.614 | 1 | 36113% | 34176% | 36955% | 39144% | 30467% | 38679% | 39163% | 31161% | 36361% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
25 | 2022-07-17 08:32 | WO-22270-06-A-1 | 24 | U |  | 0 | 0 | 3.727 | 0 | 0 | 3.167 | 1.073 | 1.226 | 3.878 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 10949 | 217242 | 73672 | 709091 | 979994 | 6504078 | 12100289 | 2718954 | 23800787 | [5 empty] | 272 | AB-220715-01 CCC-H 10 | Q | 1 | 368.792 | 339.369 | 386.392 | 388.865 | 312.752 | 398.474 | 383.256 | 345.56 | 383.207 | 1 | 36879% | 33937% | 38639% | 38886% | 31275% | 39847% | 38326% | 34556% | 38321% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
26 | 2022-07-17 09:08 | WO-22270-07-A-1 | 25 | U |  | 0 | 0 | 0 | 0 | 0 | 2.314 | 0.509 | 0.834 | 3.862 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8642 | 226617 | 85596 | 774955 | 997478 | 6580585 | 12143878 | 2727936 | 23786530 | [5 empty] | 301 | AB-220715-01 CCC-H 11 | Q | 1 | 343.684 | 351.242 | 422.132 | 372.456 | 319.013 | 389.571 | 397.692 | 330.388 | 407.738 | 1 | 34368% | 35124% | 42213% | 37246% | 31901% | 38957% | 39769% | 33039% | 40774% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
27 | 2022-07-17 09:44 | WO-22270-08-A-1 | 26 | U |  | 0 | 0 | 0 | 0 | 0 | 2.862 | 0.848 | 0.938 | 4.404 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 6088 | 174942 | 64185 | 574949 | 860282 | 5709385 | 10567001 | 2416625 | 20956659 | [5 empty] | 331 | AB-220715-01 CCC-H 12 | Q | 1 | 337.382 | 330.585 | 418.533 | 384.909 | 310.509 | 386.724 | 390.787 | 319.683 | 381.938 | 1 | 33738% | 33058% | 41853% | 38491% | 31051% | 38672% | 39079% | 31968% | 38194% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
28 | 2022-07-17 10:20 | WO-22270-09-A-1 | 27 | U |  | 0 | 0 | 0 | 0 | 0 | 1.798 | 1.009 | 2.294 | 5.182 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 5.182 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8956 | 252013 | 91451 | 841657 | 1014121 | 6892188 | 12991287 | 2912696 | 25413825 | [5 empty] | 341 | AB-220715-01 CCC-H 13 | Q | 1 | 332.365 | 341.506 | 376.546 | 374.334 | 303.614 | 388.839 | 386.732 | 323.323 | 387.108 | 1 | 33236% | 34151% | 37655% | 37433% | 30361% | 38884% | 38673% | 32332% | 38711% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
29 | 2022-07-17 10:56 | WO-22270-10-A-1 | 28 | U |  | 0 | 0 | 0 | 0 | 0 | 2.843 | 0.804 | 0.933 | 6.309 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 6.309 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7559 | 241047 | 84932 | 763305 | 1015479 | 6737953 | 12581059 | 2862004 | 24903677 | [5 empty] | 342 | AB-220715-01 QCS 2 | Q | 1 | 161.011 | 171.1 | 199.036 | 193.156 | 158.561 | 196.615 | 209.797 | 162.451 | 191.063 | 2 | 8051% | 8555% | 9952% | 9658% | 7928% | 9831% | 10490% | 8123% | 9553% | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL | FAIL
30 | 2022-07-17 11:32 | AB-220715-01 LRB 2_INJ001 | 29 | K | QC | 0 | 0 | 0 | 0 | 1.014 | 1.595 | 0.708 | 0.918 | 0.443 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 55088 | 1198731 | 399197 | 3489215 | 2987314 | 18792228 | 23446218 | 4077662 | 36852981 |  |  | 30
31 | 2022-07-17 12:08 | AB-220715-01 LRB 2 | 30 | K | QC | 0 | 0 | 0 | 0 | 0 | 1.568 | 0.802 | 1.055 | 0.367 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 63071 | 1203801 | 408133 | 3479659 | 2966182 | 18880621 | 23536400 | 3931466 | 36367310 |  |  | 31
32 | 2022-07-17 12:44 | WO-22270-11-A-1 | 31 | U |  | 0 | 0 | 0 | 0 | 0 | 1.094 | 0.482 | 2.047 | 4.526 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7880 | 243227 | 86169 | 752337 | 973170 | 6499004 | 12415592 | 2756593 | 24750178
33 | 2022-07-17 13:20 | WO-22270-12-A-1 | 32 | U |  | 0 | 0 | 0 | 0 | 0 | 2.504 | 0.455 | 0.921 | 3.046 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 14360 | 224022 | 78773 | 751211 | 943192 | 6244197 | 11980648 | 2676764 | 23411615
34 | 2022-07-17 13:56 | WO-22270-13-A-1 | 33 | U |  | 0 | 0 | 1.644 | 0 | 0.942 | 0 | 0.516 | 0.867 | 4.196 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 11567 | 163457 | 60443 | 552437 | 806418 | 5559627 | 10373051 | 2342145 | 20831902
35 | 2022-07-17 14:32 | WO-22270-14-A-1 | 34 | U |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.372 | 0.847 | 4.113 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 9159 | 224289 | 86858 | 756615 | 957616 | 6661242 | 12409492 | 2682247 | 24688316
36 | 2022-07-17 15:08 | WO-22270-15-A-1 | 35 | U |  | 0 | 0 | 0 | 0 | 0 | 2.884 | 0.761 | 1.973 | 2.456 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7200 | 215373 | 77998 | 722193 | 945971 | 6359552 | 12095536 | 2761557 | 23345937
37 | 2022-07-17 15:44 | WO-22270-16-A-1 | 36 | U |  | 0 | 0 | 0 | 0 | 0 | 2.347 | 0.681 | 0.965 | 4.631 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 9043 | 253091 | 93335 | 866499 | 1007387 | 6896612 | 12993054 | 2856034 | 25899008
38 | 2022-07-17 16:20 | WO-22270-17-A-1 | 37 | U |  | 0 | 0 | 0 | 0 | 0 | 3.796 | 0.916 | 0.832 | 4.525 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8750 | 212397 | 70790 | 702222 | 874969 | 6171068 | 11191368 | 2464575 | 22950166
39 | 2022-07-17 16:56 | WO-22270-18-A-1 | 38 | U |  | 0 | 0 | 0 | 0 | 1.431 | 1.384 | 0.394 | 0.83 | 1.851 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7583 | 209350 | 71362 | 684538 | 862988 | 5952927 | 11429261 | 2473151 | 23325324
40 | 2022-07-17 17:32 | WO-22270-19-A-1 | 39 | U |  | 0 | 0 | 1.644 | 0 | 0 | 1.812 | 0.482 | 0.866 | 3.432 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7571 | 215765 | 72915 | 726015 | 877577 | 6127998 | 11774606 | 2551859 | 24026812
41 | 2022-07-17 18:08 | WO-22270-20-A-1 | 40 | U |  | 0 | 0 | 2.717 | 0 | 0 | 2.423 | 0.447 | 0.99 | 5.105 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 5.105 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 9344 | 187792 | 75300 | 709476 | 807537 | 6092727 | 11580470 | 2279685 | 22800078
42 | 2022-07-17 18:44 | AB-220715-01 LFB-H 2 | 41 | Z | QC | 159.874 | 161.04 | 194.921 | 187.047 | 152.172 | 190.583 | 192.039 | 160.442 | 193.606 | 1 | <MRL | 161.04 | 194.921 | 187.047 | 152.172 | 190.583 | 192.039 | 160.442 | 193.606 | H 2 | H 2 | NA | Blank Spike | No TrueV! | [15 empty] | 64930 | 1098509 | 394898 | 3271639 | 2424591 | 15597174 | 20010453 | 3134823 | 28517778 |  |  | 42
43 | 2022-07-17 19:20 | AB-220715-01 LRB 3_INJ001 | 42 | K | QC | 0 | 0 | 0 | 1.04 | 1.095 | 1.137 | 0.487 | 1.016 | 0.521 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 73454 | 1180945 | 413998 | 3688857 | 2944018 | 20033002 | 24842549 | 4024004 | 38568092 |  |  | 43 |  |  | LRB & LFB Samples |  |  |  | Measured Concentration, ng/L | [9 empty] | Quality Control Checks
44 | 2022-07-17 19:56 | AB-220715-01 LRB 3 | 43 | K | QC | 0 | 0 | 0 | 0.977 | 0.969 | 1.107 | 0.453 | 0.842 | 0.4 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 36802 | 1175836 | 416397 | 3571226 | 2801679 | 18187751 | 22675639 | 3696160 | 34668358 |  |  | 44 |  |  | # | QC Samples | Type | Dilution Factor: | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | True Value | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA
45 | 2022-07-17 20:32 | AB-220715-01 LD 1 | 44 | D | QC | 0 | 0 | 2.029 | 0 | 1.804 | 2.867 | 0.621 | 0.834 | 3.428 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Duplicate | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 9334 | 229663 | 89095 | 798127 | 936028 | 6343479 | 12213570 | 2637392 | 24291523 |  |  |  | 45
46 | 2022-07-17 21:08 | AB-220715-01 CCC-H 2 | 45 | Q | QC | 364.718 | 347.787 | 399.928 | 380.048 | 320.885 | 378.815 | 387.617 | 336.485 | 386.616 | 1 | 364.718 | 347.787 | 399.928 | 380.048 | 320.885 | 378.815 | 387.617 | 336.485 | 386.616 | 1 | 01 | NA | CCC | UL! | [15 empty] | 50312 | 936774 | 368167 | 3073192 | 1927568 | 12833118 | 16922867 | 2652217 | 22522730 |  | 46 |  |  |  | 14 | AB-220715-01 LFB-H 1 | Z | 1 | 184.824 | 167.941 | 199.413 | 198.905 | 154.586 | 195.518 | 197.021 | 168.032 | 196.5 | H 1
47 | 2022-07-17 21:44 | AB-220715-01 LFSM-H 1 | 46 | S | QC | 0 | 164.433 | 188.888 | 194.271 | 151.916 | 194.121 | 194.736 | 149.476 | 196.881 | 1 | <MRL | 164.433 | 188.888 | 194.271 | 151.916 | 194.121 | 194.736 | 149.476 | 196.881 | H 1 | NA | H 1 | Sample Spike | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | [9 empty] | 11665 | 216376 | 82309 | 709783 | 821320 | 5718369 | 10492766 | 2247881 | 20039375 |  |  |  |  | 47 | 15 | AB-220715-01 LRB 1_INJ001 | K | 1 | 0 | 0 | 0 | 1.749 | 0 | 1.349 | 0.917 | 1.258 | 0.984 | NA
48 | 2022-07-17 22:20 | WO-22270-21-A-1 | 47 | U |  | 0 | 0 | 2.659 | 0 | 0 | 4.374 | 0.654 | 1.043 | 3.786 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 6573 | 195198 | 74382 | 654006 | 830131 | 6014678 | 10956068 | 2358870 | 21910933 | [5 empty] | 16 | AB-220715-01 LRB 1 | K | 1 | 0 | 0 | 0 | 3.07 | 0 | 1.317 | 0.967 | 1.494 | 1.116 | NA | [12 empty] | FAIL
49 | 2022-07-17 22:56 | WO-22270-22-A-1 | 48 | U |  | 0 | 0 | 0 | 0 | 0 | 1.838 | 0.608 | 0.808 | 3.208 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 10036 | 232969 | 85811 | 782289 | 951987 | 6548169 | 12533477 | 2605651 | 24552358 | [5 empty] | 30 | AB-220715-01 LRB 2_INJ001 | K | 1 | 0 | 0 | 0 | 0 | 1.014 | 1.595 | 0.708 | 0.918 | 0.443 | NA
50 | 2022-07-17 23:32 | WO-22270-23-A-1 | 49 | U |  | 0 | 0 | 0 | 2.005 | 0 | 3.327 | 0.641 | 0.896 | 4.061 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 10071 | 228669 | 83242 | 774613 | 951131 | 6553624 | 12070045 | 2536246 | 23873966 | [5 empty] | 31 | AB-220715-01 LRB 2 | K | 1 | 0 | 0 | 0 | 0 | 0 | 1.568 | 0.802 | 1.055 | 0.367 | NA
51 | 2022-07-18 00:08 | WO-22270-24-A-1 | 50 | U |  | 0 | 0 | 0 | 0 | 0 | 2.263 | 0.582 | 1.016 | 3.347 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8155 | 205118 | 75539 | 697914 | 864524 | 6063512 | 11585880 | 2380645 | 22565786 | [5 empty] | 42 | AB-220715-01 LFB-H 2 | Z | 1 | 159.874 | 161.04 | 194.921 | 187.047 | 152.172 | 190.583 | 192.039 | 160.442 | 193.606 | H 2
52 | 2022-07-18 00:44 | WO-22270-25-A-1 | 51 | U |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.552 | 1.775 | 4.386 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8796 | 198886 | 80841 | 762129 | 926178 | 6436247 | 12191895 | 2578343 | 24376114 | [5 empty] | 43 | AB-220715-01 LRB 3_INJ001 | K | 1 | 0 | 0 | 0 | 1.04 | 1.095 | 1.137 | 0.487 | 1.016 | 0.521 | NA
53 | 2022-07-18 01:20 | WO-22270-26-A-1 | 52 | U |  | 0 | 0 | 0 | 0 | 0 | 2.693 | 0.619 | 2.491 | 3.774 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7011 | 162614 | 63660 | 581597 | 803516 | 5598039 | 10809945 | 2253364 | 21456251 | [5 empty] | 44 | AB-220715-01 LRB 3 | K | 1 | 0 | 0 | 0 | 0.977 | 0.969 | 1.107 | 0.453 | 0.842 | 0.4 | NA
54 | 2022-07-18 01:56 | WO-22270-27-A-1 | 53 | U |  | 0 | 0 | 0 | 0 | 0 | 1.492 | 0.47 | 0.946 | 3.312 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 10699 | 212558 | 88315 | 780466 | 871868 | 6437946 | 12334450 | 2467084 | 24393292 | [5 empty] | 58 | AB-220715-01 LRB 4_INJ001 | K | 1 | 0 | 0 | 0 | 1.03 | 0.968 | 1.15 | 0.597 | 1.507 | 0.328 | NA
55 | 2022-07-18 02:32 | WO-22270-28-A-1 | 54 | U |  | 0 | 0 | 0 | 0 | 0 | 0 | 0.471 | 0.892 | 4.549 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7360 | 178048 | 77701 | 711838 | 868396 | 6198911 | 11956537 | 2520075 | 23251013 | [5 empty] | 59 | AB-220715-01 LRB 4 | K | 1 | 0 | 0 | 0 | 1.05 | 0.951 | 1.046 | 0.398 | 0.839 | 0.33 | NA
56 | 2022-07-18 03:08 | WO-22270-29-A-1 | 55 | U |  | 0 | 0 | 0 | 0 | 0 | 0.799 | 0.539 | 0.826 | 4.642 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 9079 | 220808 | 89823 | 844415 | 958138 | 6971406 | 13663824 | 2761177 | 25150085 | [5 empty] | 70 | AB-220715-01 LFB-H 3 | Z | 1 | 157.99 | 168.028 | 196.235 | 190.785 | 153.757 | 199.713 | 199.696 | 164.912 | 198.894 | H 3
57 | 2022-07-18 03:44 | WO-22270-30-A-1 | 56 | U |  | 0 | 0 | 2.817 | 1.854 | 0 | 2.664 | 0.836 | 1.847 | 4.086 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8079 | 204675 | 83667 | 768432 | 918954 | 6664518 | 12720413 | 2658465 | 23382076 | [5 empty] | 71 | AB-220715-01 LRB 5_INJ001 | K | 1 | 0 | 0 | -0.174 | 0.976 | 1.63 | 0.766 | 0.45 | 0.988 | 0.496 | NA
58 | 2022-07-18 04:20 | AB-220715-01 LRB 4_INJ001 | 57 | K | QC | 0 | 0 | 0 | 1.03 | 0.968 | 1.15 | 0.597 | 1.507 | 0.328 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 69085 | 1024715 | 389709 | 3438556 | 2806334 | 19118635 | 24146421 | 3894796 | 37059342 |  |  | 58 |  |  | 72 | AB-220715-01 LRB 5 | K | 1 | 0 | 0.683 | -0.157 | 1.012 | 1.69 | 0.861 | 0.41 | 1.616 | 0.431 | NA
59 | 2022-07-18 04:56 | AB-220715-01 LRB 4 | 58 | K | QC | 0 | 0 | 0 | 1.05 | 0.951 | 1.046 | 0.398 | 0.839 | 0.33 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 40255 | 1052684 | 414308 | 3423926 | 2720683 | 18050520 | 22767222 | 3495357 | 35065625 |  |  | 59 |  |  | 92 | AB-220715-01 LRB 6_INJ001 | K | 1 | 0 | 0 | 0.438 | 0.966 | 1.676 | 0.485 | 0.346 | 1.236 | 0.36 | NA
60 | 2022-07-18 05:32 | WO-22270-31-A-1 | 59 | U |  | 0 | 0 | 0.624 | 0 | 0 | 1.614 | 1.049 | 0.864 | 2.993 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8390 | 214161 | 76706 | 809761 | 928348 | 6415707 | 12305829 | 2588435 | 24969566 | [5 empty] | 93 | AB-220715-01 LRB 6 | K | 1 | 0 | 0.557 | 1.195 | 1.007 | 1.456 | 0.535 | 0.342 | 0.934 | 0.365 | NA
61 | 2022-07-18 06:08 | WO-22270-32-A-1 | 60 | U |  | 0 | 0 | 0 | 0 | 0 | 1.812 | 0.671 | 2.278 | 3.971 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 4741 | 84477 | 36053 | 328226 | 572347 | 3991210 | 7756948 | 1568949 | 15464622 | [5 empty] | 98 | AB-220715-01 LFB-H 4 | Z | 1 | 162.744 | 175.872 | 197.68 | 199.009 | 157.175 | 199.897 | 204.168 | 165.163 | 198.772 | H 4
62 | 2022-07-18 06:44 | WO-22270-33-A-1 | 61 | U |  | 0 | 0 | 0.752 | 0 | 0 | 2.319 | 0.42 | 0.899 | 3.255 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7519 | 140821 | 52877 | 565262 | 760076 | 5467538 | 10249490 | 2116313 | 20700734 | [5 empty] | 99 | AB-220715-01 LRB 7_INJ001 | K | 1 | 0 | 0.669 | -0.005 | 1.05 | 1.998 | 1.489 | 0.567 | 1.365 | 0.506 | NA
63 | 2022-07-18 07:20 | WO-22270-34-A-1 | 62 | U |  | 0 | 0 | 0 | 0 | 0 | 3.094 | 0.998 | 1.188 | 4.461 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [8 empty] | 4927 | 126717 | 54088 | 507631 | 692041 | 5012065 | 9767518 | 1973366 | 19171427 | [5 empty] | 100 | AB-220715-01 LRB 7 | K | 1 | 0 | 0 | 0.5 | 0.992 | 2.109 | 1.449 | 0.776 | 0.94 | 0.397 | NA
64 | 2022-07-18 07:56 | WO-22270-35-A-1 | 63 | U |  | 0 | 0 | 0 | 0 | 0 | 0.962 | 1.591 | 1.696 | 2.965 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7517 | 176728 | 79629 | 725329 | 846287 | 6213151 | 11749688 | 2480709 | 23300048 | [5 empty] | 114 | AB-220715-01 LRB 8_INJ001 | K | 1 | 0 | 0.733 | -0.188 | 0.985 | 2.078 | 1.451 | 0.692 | 0.885 | 0.33 | NA
65 | 2022-07-18 08:32 | WO-22270-36-A-1 | 64 | U |  | 0 | 0 | 0.078 | 0 | 0 | 1.501 | 0.517 | 1.719 | 3.607 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 9213 | 268439 | 105468 | 1022167 | 1045386 | 7413062 | 14258175 | 2946160 | 27583948 | [5 empty] | 115 | AB-220715-01 LRB 8 | K | 1 | 0 | 0.678 | 0.346 | 0.981 | 2.103 | 1.359 | 0.846 | 1.442 | 0.371 | NA
66 | 2022-07-18 09:08 | WO-22270-37-A-1 | 65 | U |  | 0 | 0.915 | 0 | 0 | 0 | 1.387 | 0.447 | 1.024 | 3.229 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7128 | 161519 | 73754 | 651708 | 787237 | 5639306 | 11143067 | 2273046 | 21995949 | [5 empty] | 126 | AB-220715-01 LFB-H 5 | Z | 1 | 190.537 | 174.405 | 211.05 | 205.523 | 155.005 | 201.716 | 206.05 | 164.39 | 199.853 | H 5
67 | 2022-07-18 09:44 | WO-22270-38-A-1 | 66 | U |  | 0 | 1.299 | 0 | 0 | 0 | 1.752 | 0.48 | 0.958 | 2.328 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7073 | 167345 | 79210 | 670769 | 789416 | 5871523 | 11099051 | 2368565 | 23508771 | [5 empty] | 127 | AB-220715-01 LRB 9_INJ001 | K | 1 | 0 | 0.709 | 0.111 | 1.066 | 2.124 | 1.724 | 0.986 | 1.658 | 0.497 | NA
68 | 2022-07-18 10:20 | WO-22270-39-A-1 | 67 | U |  | 0 | 0 | 0 | 0 | 0 | 1.962 | 0.429 | 2.016 | 2.578 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7799 | 162958 | 74250 | 658063 | 774833 | 5599197 | 10961914 | 2214400 | 21946613 | [5 empty] | 128 | AB-220715-01 LRB 9 | K | 1 | 0 | 0.742 | 0.494 | 1.009 | 2.23 | 1.482 | 1.115 | 0.933 | 0.364 | NA
69 | 2022-07-18 10:56 | WO-22270-40-A-1 | 68 | U |  | 0 | 0 | 1.034 | 0 | 0 | 0.761 | 1.428 | 0.952 | 3.462 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7929 | 158309 | 71697 | 641009 | 777921 | 5748344 | 10864387 | 2242114 | 22015251 | [5 empty] | 142 | AB-220715-01 LRB 10_INJ001 | K | 1 | 0 | 0.855 | 0.364 | 0 | 0 | 1.382 | 1.144 | 0.955 | 0.39 | NA
70 | 2022-07-18 11:32 | AB-220715-01 LFB-H 3 | 69 | Z | QC | 157.99 | 168.028 | 196.235 | 190.785 | 153.757 | 199.713 | 199.696 | 164.912 | 198.894 | 1 | <MRL | 168.028 | 196.235 | 190.785 | 153.757 | 199.713 | 199.696 | 164.912 | 198.894 | H 3 | H 3 | NA | Blank Spike | No TrueV! | [15 empty] | 58803 | 875068 | 354933 | 3161735 | 2211219 | 14776874 | 19755156 | 2985029 | 28117518 |  |  | 70 |  |  | 143 | AB-220715-01 LRB 10 | K | 1 | 0 | 0.656 | 0.581 | 1.06 | 0 | 1.482 | 1.02 | 1.082 | 0.442 | NA
71 | 2022-07-18 12:08 | AB-220715-01 LRB 5_INJ001 | 70 | K | QC | 0 | 0 | -0.174 | 0.976 | 1.63 | 0.766 | 0.45 | 0.988 | 0.496 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 69278 | 929593 | 409478 | 3521842 | 2745820 | 19081863 | 24113664 | 3774492 | 37502477 |  |  | 71 |  |  | 154 | AB-220715-01 LFB-H 6 | Z | 1 | 156.313 | 166.991 | 182.753 | 191.319 | 148.72 | 191.836 | 188.34 | 161.946 | 188.731 | H 6
72 | 2022-07-18 12:44 | AB-220715-01 LRB 5 | 71 | K | QC | 0 | 0.683 | -0.157 | 1.012 | 1.69 | 0.861 | 0.41 | 1.616 | 0.431 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 41800 | 931713 | 370035 | 3314814 | 2582623 | 17477041 | 21661304 | 3390735 | 33609010 |  |  | 72 |  |  | 155 | AB-220715-01 LRB 11_INJ001 | K | 1 | 0 | 0.684 | 0.603 | 0 | 0 | 1.569 | 0.798 | 2.04 | 0.503 | NA
73 | 2022-07-18 13:20 | AB-220715-01 LD 2 | 72 | D | QC | 0 | 1.048 | 0 | 0 | 2.332 | 4.203 | 1.059 | 1.012 | 2.589 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Duplicate | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7489 | 164230 | 67770 | 637486 | 782533 | 5642949 | 10589898 | 2166130 | 22153123 |  |  |  | 73 |  | 156 | AB-220715-01 LRB 11 | K | 1 | 0 | 0.596 | 0.293 | 0.998 | 0 | 1.468 | 0.811 | 1.677 | 0.383 | NA
74 | 2022-07-18 13:56 | AB-220715-01 CCC-H 3 | 73 | Q | QC | 361.667 | 340.157 | 397.811 | 382.755 | 317.243 | 391.053 | 391.511 | 332.226 | 389.63 | 1 | 361.667 | 340.157 | 397.811 | 382.755 | 317.243 | 391.053 | 391.511 | 332.226 | 389.63 | 1 | 01 | NA | CCC | UL! | [15 empty] | 55846 | 789700 | 349251 | 2920353 | 1820963 | 12414542 | 16675537 | 2451267 | 22653173 |  | 74 |  |  |  | 170 | AB-220715-01 LRB 12_INJ001 | K | 1 | 0 | 0.705 | 0.578 | 0 | 0 | 0.51 | 0.335 | 1.732 | 0.332 | NA
75 | 2022-07-18 14:32 | AB-220715-01 LFSM-H 2 | 74 | S | QC | 0 | 168.055 | 184.578 | 183.486 | 150.282 | 198.046 | 196.189 | 152.35 | 196.473 | 1 | <MRL | 168.055 | 184.578 | 183.486 | 150.282 | 198.046 | 196.189 | 152.35 | 196.473 | H 2 | NA | H 2 | Sample Spike | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | [7 empty] | 8264 | 168355 | 81085 | 704358 | 752646 | 5276322 | 9470779 | 1847622 | 18025841 |  |  |  |  | 75 | 171 | AB-220715-01 LRB 12 | K | 1 | 0 | 0.946 | 0.319 | 0 | 0 | 0.517 | 0.321 | 1.105 | 0.346 | NA
76 | 2022-07-18 15:08 | WO-22270-41-A-1 | 75 | U |  | 0 | 0 | 0 | 0 | 0 | 0.959 | 0.66 | 0.97 | 2.578 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8915 | 150874 | 66208 | 601965 | 743759 | 5644774 | 11178511 | 2172604 | 21698381 | [5 empty] | 182 | AB-220715-01 LFB-H 7 | Z | 1 | 172.757 | 171.667 | 189.674 | 192.398 | 156.023 | 192.64 | 195.667 | 158.717 | 190.922 | H 7
77 | 2022-07-18 15:44 | WO-22270-42-A-1 | 76 | U |  | 0 | 1.077 | 2.162 | 0 | 0 | 1.255 | 0.428 | 0.841 | 1.633 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7287 | 183430 | 72490 | 748662 | 829097 | 6032911 | 11485926 | 2347910 | 23151283 | [5 empty] | 183 | AB-220715-01 LRB 13_INJ001 | K | 1 | 0 | 0.708 | 0.574 | 1.13 | 0 | 0.579 | 0.447 | 1.31 | 0.522 | NA
78 | 2022-07-18 16:20 | WO-22270-43-A-1 | 77 | U |  | 0 | 0 | 0 | 0 | 0 | 4.642 | 0.663 | 0.917 | 2.633 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7526 | 169870 | 73276 | 684181 | 810068 | 5623487 | 10691478 | 2210459 | 21964913 | [5 empty] | 184 | AB-220715-01 LRB 13 | K | 1 | 0 | 0.858 | 0.337 | 0.956 | 0 | 0.599 | 0.354 | 1.057 | 0.374 | NA
79 | 2022-07-18 16:56 | WO-22270-44-A-1 | 78 | U |  | 0 | 0 | 1.368 | 0 | 0 | 1.306 | 0.442 | 0.928 | 1.905 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 10678 | 187714 | 80655 | 757318 | 867535 | 6108574 | 12037762 | 2357439 | 24105423 | [5 empty] | 198 | AB-220715-01 LRB 14_INJ001 | K | 1 | 0 | 0.835 | 0.484 | 1.109 | 0 | 0.587 | 0.321 | 1.409 | 0.356 | NA
80 | 2022-07-18 17:32 | WO-22270-45-A-1 | 79 | U |  | 0 | 1.289 | 3.345 | 0 | 0 | 0 | 1.441 | 0.858 | 1.785 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8379 | 206195 | 86272 | 801893 | 926799 | 6624417 | 12634452 | 2543666 | 24106814 | [5 empty] | 199 | AB-220715-01 LRB 14 | K | 1 | 0 | 0.797 | 0.65 | 0 | 0 | 0.56 | 0.339 | 1.439 | 0.373 | NA
81 | 2022-07-18 18:08 | WO-22270-52-A-1 | 80 | U |  | 0 | 1.25 | 3.526 | 0 | 0 | 1.301 | 0.646 | 2.063 | 2.789 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 4926 | 86810 | 40552 | 351720 | 593628 | 4543967 | 8766585 | 1681002 | 17035994 | [5 empty] | 210 | AB-220715-01 LFB-H 8 | Z | 1 | 184.848 | 169.13 | 182.476 | 190.233 | 153.04 | 200.802 | 207.329 | 153.851 | 183.032 | H 8
82 | 2022-07-18 18:44 | WO-22270-46-A-1 | 81 | U |  | 0 | 1.039 | 2.103 | 0 | 0 | 1.241 | 0.561 | 1.511 | 3.148 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8136 | 189053 | 86130 | 752980 | 865596 | 6296925 | 12022917 | 2375957 | 24586598 | [5 empty] | 211 | AB-220715-01 LRB 15_INJ001 | K | 1 | 0 | 1.064 | 0.909 | 1.636 | 1.4 | 1.407 | 1.026 | 1.582 | 1.466 | NA
83 | 2022-07-18 19:20 | WO-22270-47-A-1 | 82 | U |  | 0 | 0 | 0.861 | 0 | 0 | 1.16 | 0.441 | 0.922 | 1.894 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 12054 | 156945 | 71626 | 630850 | 810284 | 5846099 | 11703113 | 2286942 | 22479422 | [5 empty] | 212 | AB-220715-01 LRB 15 | K | 1 | 0 | 0.813 | 0.951 | 1.008 | 0 | 0.881 | 0.374 | 1.271 | 0.494 | NA
84 | 2022-07-18 19:56 | WO-22270-53-A-1 | 83 | U |  | 0 | 1.507 | 1.472 | 0 | 0 | 1.634 | 0.685 | 0.826 | 3.553 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 9140 | 200515 | 90937 | 829180 | 915032 | 7017773 | 14229850 | 2757852 | 27991159 | [5 empty] | 226 | AB-220715-01 LRB 16_INJ001 | K | 1 | 0 | 0.682 | 2.056 | 0 | 0 | 0.858 | 0.56 | 1.009 | 0.43 | NA
85 | 2022-07-18 20:32 | WO-22270-48-A-1 | 84 | U |  | 0 | 0 | 1.281 | 0 | 0 | 0.488 | 0.378 | 1.93 | 3.977 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 6969 | 182187 | 81609 | 767247 | 873671 | 6441919 | 12535265 | 2495139 | 24901790 | [5 empty] | 227 | AB-220715-01 LRB 16 | K | 1 | 0 | 0.881 | 1.823 | 0 | 0 | 0.865 | 0.378 | 0.899 | 0.429 | NA
86 | 2022-07-18 21:08 | WO-22270-54-A-1 | 85 | U |  | 0 | 0 | 3.071 | 1.304 | 0 | 0.828 | 0.375 | 1.31 | 3.216 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | [11 empty] | 11737 | 298970 | 129241 | 1206695 | 1426202 | 10975723 | 21571853 | 3872401 | 36975477 | [5 empty] | 238 | AB-220715-01 LFB-H 9 | Z | 1 | 225.666 | 160.929 | 180.037 | 195.058 | 144.71 | 199.266 | 209.15 | 147.606 | 201.143 | H 9
87 | 2022-07-18 21:44 | WO-22270-49-A-1 | 86 | U |  | 0 | 0 | 0 | 0 | 0 | 0.776 | 0.497 | 0.988 | 4.442 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7700 | 156686 | 65892 | 598947 | 805874 | 5872741 | 11118381 | 2184481 | 22711009
88 | 2022-07-18 22:20 | WO-22270-55-A-1 | 87 | U |  | 0 | 1.289 | 2.916 | 0 | 0 | 0 | 3.128 | 3.016 | 2.961 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 7978 | 115484 | 53133 | 423319 | 548155 | 4001568 | 6579510 | 1054358 | 11295160 | [5 empty] | LD Samples |  |  |  | Measured Concentration, ng/L | [9 empty] | Quality Control Checks
89 | 2022-07-18 22:56 | WO-22270-50-A-1 | 88 | U |  | 0 | 0.742 | 0 | 0 | 0 | 0.833 | 1.212 | 0.949 | 4.275 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 9505 | 194929 | 83917 | 796369 | 837207 | 6125561 | 11990486 | 2303412 | 23587410 | [5 empty] | # | QC Samples | Type | Dilution Factor: | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | True Value | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA
90 | 2022-07-18 23:32 | WO-22270-56-A-1 | 89 | U |  | 0 | 1.308 | 0 | 0 | 0 | 1.763 | 0.793 | 0.891 | 4.414 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8377 | 191057 | 79209 | 739990 | 818452 | 6054762 | 11905052 | 2354598 | 24097262
91 | 2022-07-19 00:08 | WO-22270-57-A-1 | 90 | U |  | 0 | 1.272 | 2.437 | 0 | 0 | 0.912 | 0.685 | 1.085 | 4.114 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | [10 empty] | 12298 | 250975 | 102772 | 986250 | 1125337 | 8768564 | 18264750 | 3228336 | 33342342
92 | 2022-07-19 00:44 | AB-220715-01 LRB 6_INJ001 | 91 | K | QC | 0 | 0 | 0.438 | 0.966 | 1.676 | 0.485 | 0.346 | 1.236 | 0.36 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 68749 | 926924 | 380763 | 3234487 | 2633926 | 18386672 | 23132124 | 3520878 | 35663352 |  |  | 92 |  |  | 45 | AB-220715-01 LD 1 | D | 1 | 0 | 0 | 2.029 | 0 | 1.804 | 2.867 | 0.621 | 0.834 | 3.428 | NA | [18 empty] | FAIL
93 | 2022-07-19 01:20 | AB-220715-01 LRB 6 | 92 | K | QC | 0 | 0.557 | 1.195 | 1.007 | 1.456 | 0.535 | 0.342 | 0.934 | 0.365 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 37349 | 938199 | 384775 | 3297290 | 2505670 | 17770400 | 22861691 | 3461344 | 33771748 |  |  | 93
94 | 2022-07-19 01:56 | WO-22270-58-A-1 | 93 | U |  | 0 | 2.421 | 0 | 0 | 0 | 1.42 | 0.383 | 1.217 | 3.988 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 3990 | 81253 | 38250 | 366671 | 530893 | 4361498 | 9103504 | 1713680 | 17839994 | [5 empty] | 73 | AB-220715-01 LD 2 | D | 1 | 0 | 1.048 | 0 | 0 | 2.332 | 4.203 | 1.059 | 1.012 | 2.589 | NA | [18 empty] | FAIL
95 | 2022-07-19 02:32 | WO-22270-51-A-1 | 94 | U |  | 0 | 0 | 0 | 0 | 0 | 4.097 | 0.944 | 1.713 | 2.866 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7845 | 224738 | 91409 | 795112 | 922328 | 6789733 | 13508564 | 2656259 | 26825806
96 | 2022-07-19 03:08 | WO-22270-59-A-1 | 95 | U |  | 0 | 1.069 | 0 | 0 | 0 | 0.765 | 0.473 | 1.563 | 3.171 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8397 | 180885 | 82960 | 738467 | 772260 | 6053205 | 12371284 | 2337281 | 23459572 | [5 empty] | 101 | AB-220715-01 LD 3 | D | 1 | 0 | 1.254 | 0 | 0 | 0 | 1.999 | 0.487 | 0.955 | 3.264 | NA | [18 empty] | FAIL
97 | 2022-07-19 03:44 | WO-22270-60-A-1 | 96 | U |  | 0 | 2.623 | 0 | 0 | 0 | 3.544 | 1.577 | 0.999 | 3.518 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 7086 | 145756 | 62466 | 626941 | 774035 | 5882129 | 11885493 | 2250888 | 23786204
98 | 2022-07-19 04:20 | AB-220715-01 LFB-H 4 | 97 | Z | QC | 162.744 | 175.872 | 197.68 | 199.009 | 157.175 | 199.897 | 204.168 | 165.163 | 198.772 | 1 | <MRL | 175.872 | 197.68 | 199.009 | 157.175 | 199.897 | 204.168 | 165.163 | 198.772 | H 4 | H 4 | NA | Blank Spike | No TrueV! | [15 empty] | 60308 | 823884 | 364744 | 3179654 | 2084483 | 14682995 | 19333922 | 2784852 | 27461333 |  |  | 98 |  |  | 129 | AB-220715-01 LD 4 | D | 1 | 0 | 2.589 | 3.027 | 2.379 | 0 | 3.311 | 1.152 | 2.064 | 3.69 | NA | [18 empty] | FAIL
99 | 2022-07-19 04:56 | AB-220715-01 LRB 7_INJ001 | 98 | K | QC | 0 | 0.669 | -0.005 | 1.05 | 1.998 | 1.489 | 0.567 | 1.365 | 0.506 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 67495 | 922069 | 400382 | 3371818 | 2444509 | 17384755 | 21251175 | 3432322 | 33799083 |  |  | 99
100 | 2022-07-19 05:32 | AB-220715-01 LRB 7 | 99 | K | QC | 0 | 0 | 0.5 | 0.992 | 2.109 | 1.449 | 0.776 | 0.94 | 0.397 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | [14 empty] | 30987 | 836401 | 347450 | 3069010 | 2057854 | 14208774 | 17045869 | 2763779 | 26605293 |  |  | 100 |  |  | 157 | AB-220715-01 LD 5 | D | 1 | 0 | 2.34 | 1.745 | 2.824 | 0 | 3.298 | 1.161 | 3.174 | 4.628 | NA | [18 empty] | FAIL
101 | 2022-07-19 06:08 | AB-220715-01 LD 3 | 100 | D | QC | 0 | 1.254 | 0 | 0 | 0 | 1.999 | 0.487 | 0.955 | 3.264 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Duplicate | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 9064 | 178949 | 76012 | 695541 | 729556 | 5371566 | 10928132 | 2242228 | 22038283 |  |  |  | 101
102 | 2022-07-19 06:44 | AB-220715-01 CCC-H 4 | 101 | Q | QC | 336.451 | 332.142 | 388.596 | 377.704 | 321.727 | 387.419 | 386.028 | 331.024 | 382.399 | 1 | 336.451 | 332.142 | 388.596 | 377.704 | 321.727 | 387.419 | 386.028 | 331.024 | 382.399 | 1 | 01 | NA | CCC | UL! | [15 empty] | 42768 | 688206 | 309516 | 2491474 | 1527722 | 9983058 | 12885821 | 2197810 | 18661633 |  | 102 |  |  |  | 185 | AB-220715-01 LD 6 | D | 1 | 0 | 1.889 | 0 | 1.984 | 0 | 1.785 | 0.814 | 0.93 | 3.868 | NA | [18 empty] | FAIL
103 | 2022-07-19 07:20 | AB-220715-01 LFSM-H 3 | 102 | S | QC | 0 | 176.858 | 191.02 | 182.049 | 151.448 | 197.926 | 198.045 | 154.815 | 200.285 | 1 | <MRL | 176.858 | 191.02 | 182.049 | 151.448 | 197.926 | 198.045 | 154.815 | 200.285 | H 3 | NA | H 3 | Sample Spike | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | [9 empty] | 8005 | 171966 | 80393 | 760210 | 840934 | 5587522 | 10882993 | 2356133 | 20032416 |  |  |  |  | 103
104 | 2022-07-19 07:56 | WO-22270-61-A-1 | 103 | U |  | 0 | 1.818 | 0 | 0 | 0 | 0.83 | 0.482 | 2.014 | 3.348 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8100 | 159381 | 76889 | 629715 | 784412 | 5632908 | 11059184 | 2417223 | 23853721 | [5 empty] | 213 | AB-220715-01 LD 7 | D | 1 | 0 | 1.639 | 7.487 | 0 | 0 | 1.55 | 0.653 | 1.203 | 4.949 | NA | [18 empty] | FAIL |  | True duplicate failure. Flagging
105 | 2022-07-19 08:32 | WO-22270-62-A-1 | 104 | U |  | 0 | 1.636 | 1.249 | 0 | 0 | 1.122 | 0.586 | 2.345 | 3.215 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [8 empty] | 7449 | 142858 | 70819 | 558706 | 704598 | 5277855 | 8851652 | 1890861 | 21352912
106 | 2022-07-19 09:08 | WO-22270-63-A-1 | 105 | U |  | 0 | 2.378 | 0 | 0 | 0 | 1.792 | 0.534 | 1.061 | 3.361 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 5932 | 97773 | 45185 | 424919 | 562587 | 4194261 | 7870391 | 1618097 | 16824197 | [5 empty] | 241 | AB-220715-01 LD 8 | D | 1 | 0 | 1.145 | 8.326 | 2.782 | 0 | 3.612 | 0.66 | 1.327 | 3.377 | NA | [18 empty] | FAIL
107 | 2022-07-19 09:44 | WO-22270-64-A-1 | 106 | U |  | 0 | 2.725 | 2.491 | 0 | 0 | 0 | 3.413 | 0.884 | 3.534 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 5328 | 85349 | 39411 | 355301 | 525246 | 3735298 | 7084669 | 1461250 | 15389981
108 | 2022-07-19 10:20 | WO-22270-65-A-1 | 107 | U |  | 0 | 0.788 | 0 | 1.145 | 0 | 0.741 | 0.555 | 1.601 | 3.516 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 10197 | 210409 | 94969 | 889358 | 928365 | 6960461 | 13445178 | 2805077 | 27040132 | [5 empty] | 271 | AB-220715-01 LD 9 | D | 1 | 0 | 1.746 | 7.444 | 0 | 0 | 0.625 | 0.708 | 2.715 | 2.987 | NA | [18 empty] | FAIL
109 | 2022-07-19 10:56 | WO-22270-66-A-1 | 108 | U |  | 0 | 3.244 | 3.031 | 0 | 0 | 1.099 | 0.722 | 2.321 | 3.2 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [8 empty] | 9090 | 156805 | 74964 | 693512 | 677177 | 5210331 | 9956516 | 2073449 | 21093445
110 | 2022-07-19 11:32 | WO-22270-67-A-1 | 109 | U |  | 0 | 2.024 | 2.759 | 0 | 0 | 2.636 | 20.342 | 2.084 | 4.157 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 20.342 | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | [7 empty] | 6707 | 132533 | 59699 | 540893 | 642445 | 4549314 | 8858423 | 1789267 | 19195880 | [5 empty] | 300 | AB-220715-01 LD 10 | D | 1 | 0 | 1.452 | 5472.988 | 2201.609 | 0 | 1.37 | 0.723 | 2.393 | 2.957 | NA | [18 empty] | FAIL |  | LD isn't calculating right because the LD was made at 4-fold dilution, not 2-fold.
111 | 2022-07-19 12:08 | WO-22270-68-A-1 | 110 | U |  | 0 | 3.843 | 2.365 | 0 | 0 | 1.27 | 0.457 | 2.14 | 3.686 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 6750 | 110015 | 52466 | 493220 | 600447 | 4280878 | 8346155 | 1755705 | 18173454 | [39 empty] | 0.11823 | 0.22245 | 0.12377 | Decimal LD errors
112 | 2022-07-19 12:44 | WO-22270-69-A-1 | 111 | U |  | 0 | 3.852 | 0 | 0 | 0 | 0.793 | 0.518 | 2.992 | 4.846 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 4206 | 56380 | 28093 | 270748 | 412816 | 3107400 | 5936933 | 1202969 | 12658951 | [5 empty] | 330 | AB-220715-01 LD 11 | D | 1 | 0 | 1.08 | 2421.487 | 1091.661 | 4.165 | 1.364 | 0.555 | 0.907 | 1.575 | NA | [18 empty] | FAIL |  | 0.13523
113 | 2022-07-19 13:20 | WO-22270-70-A-1 | 112 | U |  | 0 | 1.922 | 1.632 | 0 | 0 | 3.853 | 1.354 | 1.318 | 4.055 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 5216 | 101322 | 47787 | 420469 | 563025 | 4183401 | 7799028 | 1639413 | 17148751
114 | 2022-07-19 13:56 | AB-220715-01 LRB 8_INJ001 | 113 | K | QC | 0 | 0.733 | -0.188 | 0.985 | 2.078 | 1.451 | 0.692 | 0.885 | 0.33 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 79005 | 902082 | 434366 | 3527356 | 2411403 | 17191201 | 20522310 | 3354216 | 31665544 |  |  | 114
115 | 2022-07-19 14:32 | AB-220715-01 LRB 8 | 114 | K | QC | 0 | 0.678 | 0.346 | 0.981 | 2.103 | 1.359 | 0.846 | 1.442 | 0.371 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 35119 | 816716 | 347082 | 3015102 | 2081068 | 14127326 | 17494277 | 2866857 | 25841839 |  |  | 115
116 | 2022-07-19 15:08 | WO-22270-71-A-1 | 115 | U |  | 0 | 1.464 | 2.742 | 0 | 0 | 2.626 | 5.465 | 0.896 | 8.605 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 5.465 | <MRL | 8.605 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [8 empty] | 8916 | 140433 | 70935 | 606222 | 688630 | 5108915 | 9865279 | 2019373 | 20858411
117 | 2022-07-19 15:44 | WO-22270-72-A-1 | 116 | U |  | 0 | 1.316 | 1.019 | 0 | 0 | 1.363 | 0.393 | 0.981 | 3.996 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [8 empty] | 7587 | 154375 | 70231 | 618722 | 676054 | 5013576 | 9800270 | 1987691 | 20577747
118 | 2022-07-19 16:20 | WO-22270-73-A-1 | 117 | U |  | 0 | 1.954 | 3.521 | 0 | 0 | 0.576 | 0.46 | 0.921 | 3.362 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 6110 | 124508 | 56268 | 507276 | 615152 | 4622658 | 8907123 | 1778582 | 18861487
119 | 2022-07-19 16:56 | WO-22270-74-A-1 | 118 | U |  | 0 | 1.892 | 2.192 | 0 | 0 | 2.129 | 2.505 | 2.049 | 6.787 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 6.787 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 7115 | 122767 | 56226 | 511395 | 598123 | 4366331 | 8820378 | 1752307 | 18839817
120 | 2022-07-19 17:32 | WO-22270-75-A-1 | 119 | U |  | 0 | 1.166 | 2.943 | 0 | 0 | 4.871 | 1.793 | 1.729 | 3.346 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8457 | 189479 | 80966 | 754436 | 743176 | 5477571 | 10411323 | 2099428 | 21679778
121 | 2022-07-19 18:08 | WO-22270-76-A-1 | 120 | U |  | 0 | 2.297 | 3.26 | 0 | 0 | 1.654 | 0.522 | 3.292 | 4.155 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 4541 | 61407 | 28247 | 264233 | 414464 | 2991368 | 5806209 | 1075264 | 11850992
122 | 2022-07-19 18:44 | WO-22270-77-A-1 | 121 | U |  | 0 | 1.209 | 0.62 | 1.755 | 0 | 0.829 | 0.445 | 2.161 | 19.796 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 19.796 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 9382 | 193017 | 89528 | 763299 | 724842 | 5720353 | 10925884 | 2155191 | 22956505
123 | 2022-07-19 19:20 | WO-22270-78-A-1 | 122 | U |  | 0 | 0.789 | 2.357 | 1.926 | 0 | 0.94 | 0.575 | 1.883 | 3.982 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 11120 | 213636 | 97909 | 831775 | 775811 | 5776537 | 11367131 | 2165997 | 23688616
124 | 2022-07-19 19:56 | WO-22270-79-A-1 | 123 | U |  | 0 | 1.09 | 1.421 | 1.998 | 0 | 1.095 | 0.403 | 1.036 | 3.306 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 11439 | 204186 | 92966 | 816834 | 756442 | 5728536 | 11034958 | 2258374 | 23328471
125 | 2022-07-19 20:32 | WO-22270-80-A-1 | 124 | U |  | 0 | 1.521 | 2.245 | 2.097 | 0 | 1.857 | 0.983 | 0.867 | 3.667 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 11307 | 235939 | 103221 | 876904 | 784201 | 5798240 | 11573539 | 2245614 | 24003414
126 | 2022-07-19 21:08 | AB-220715-01 LFB-H 5 | 125 | Z | QC | 190.537 | 174.405 | 211.05 | 205.523 | 155.005 | 201.716 | 206.05 | 164.39 | 199.853 | 1 | <MRL | 174.405 | 211.05 | 205.523 | 155.005 | 201.716 | 206.05 | 164.39 | 199.853 | H 5 | H 5 | NA | Blank Spike | No TrueV! | [15 empty] | 50507 | 707518 | 307686 | 2509330 | 1746852 | 11001705 | 13828390 | 2265980 | 20292643 |  |  | 126
127 | 2022-07-19 21:44 | AB-220715-01 LRB 9_INJ001 | 126 | K | QC | 0 | 0.709 | 0.111 | 1.066 | 2.124 | 1.724 | 0.986 | 1.658 | 0.497 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 69981 | 838718 | 351091 | 2983127 | 2252011 | 15068280 | 18439011 | 2921618 | 29943930 |  |  | 127
128 | 2022-07-19 22:20 | AB-220715-01 LRB 9 | 127 | K | QC | 0 | 0.742 | 0.494 | 1.009 | 2.23 | 1.482 | 1.115 | 0.933 | 0.364 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 48600 | 809751 | 345555 | 2903771 | 1976682 | 13921341 | 17044426 | 2677379 | 26409809 |  |  | 128
129 | 2022-07-19 22:56 | AB-220715-01 LD 4 | 128 | D | QC | 0 | 2.589 | 3.027 | 2.379 | 0 | 3.311 | 1.152 | 2.064 | 3.69 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Duplicate | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 8799 | 178794 | 82518 | 716808 | 672550 | 5339548 | 10202374 | 2022776 | 21008614 |  |  |  | 129
130 | 2022-07-19 23:32 | AB-220715-01 CCC-H 5 | 129 | Q | QC | 356.628 | 338.354 | 389.461 | 384.448 | 315.397 | 376.568 | 395.37 | 332.343 | 386.242 | 1 | 356.628 | 338.354 | 389.461 | 384.448 | 315.397 | 376.568 | 395.37 | 332.343 | 386.242 | 1 | 01 | NA | CCC | UL! | [15 empty] | 62264 | 727228 | 335321 | 2693442 | 1486649 | 10341770 | 12768315 | 2060162 | 18565659 |  | 130
131 | 2022-07-20 00:08 | AB-220715-01 LFSM-H 4 | 130 | S | QC | 0 | 170.383 | 190.142 | 193.605 | 149.526 | 196.17 | 204.594 | 151.141 | 201.625 | 1 | <MRL | 170.383 | 190.142 | 193.605 | 149.526 | 196.17 | 204.594 | 151.141 | 201.625 | H 4 | NA | H 4 | Sample Spike | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | [8 empty] | 12178 | 197978 | 92993 | 796480 | 709155 | 5148850 | 9525681 | 1943838 | 18544192 |  |  |  |  | 131
132 | 2022-07-20 00:44 | WO-22270-81-A-1 | 131 | U |  | 0 | 1.934 | 3.201 | 3.35 | 0 | 1.584 | 0.552 | 3.031 | 3.721 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 6348 | 123221 | 56564 | 501260 | 585225 | 4423039 | 8639728 | 1690896 | 18293922
133 | 2022-07-20 01:20 | WO-22270-82-A-1 | 132 | U |  | 0 | 1.181 | 2.934 | 3.299 | 0 | 2.325 | 0.474 | 0.86 | 3.553 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 10423 | 197317 | 89715 | 800165 | 816846 | 6277768 | 12553357 | 2372708 | 24731305 | [5 empty] | LFSM Samples |  |  |  | Measured Concentration, ng/L | [9 empty] | Quality Control Checks
134 | 2022-07-20 01:56 | WO-22270-83-A-1 | 133 | U |  | 0 | 1.222 | 2.189 | 3.92 | 0 | 0.903 | 0.568 | 1.866 | 3.516 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 10377 | 205795 | 94527 | 844460 | 789113 | 6020994 | 11784212 | 2323868 | 24460084 | [5 empty] | # | QC Samples | Type | Dilution Factor: | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | True Value | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA
135 | 2022-07-20 02:32 | WO-22270-84-A-1 | 134 | U |  | 0 | 2.623 | 0 | 10.43 | 0 | 1.798 | 0.444 | 0.926 | 2.515 | 1 | <MRL | <MRL | <MRL | 10.43 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 10508 | 181820 | 91522 | 772173 | 750904 | 5720144 | 11441580 | 2220971 | 23008855
136 | 2022-07-20 03:08 | WO-22270-85-A-1 | 135 | U |  | 0 | 1.233 | 1.331 | 8.256 | 0 | 1.847 | 0.342 | 2.173 | 3.296 | 1 | <MRL | <MRL | <MRL | 8.256 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 10741 | 203653 | 91493 | 780033 | 758643 | 5866912 | 11662114 | 2293834 | 24050284
137 | 2022-07-20 03:44 | WO-22270-86-A-1 | 136 | U |  | 0 | 1.149 | 1.607 | 8.317 | 0 | 3.486 | 0.451 | 0.937 | 3.299 | 1 | <MRL | <MRL | <MRL | 8.317 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 10626 | 189036 | 86556 | 757330 | 766801 | 6104626 | 11998524 | 2335104 | 23675098 | [5 empty] | 47 | AB-220715-01 LFSM-H 1 | S | 1 | 0 | 164.433 | 188.888 | 194.271 | 151.916 | 194.121 | 194.736 | 149.476 | 196.881 | H 1 | [18 empty] | FAIL
138 | 2022-07-20 04:20 | WO-22270-87-A-1 | 137 | U |  | 0 | 0 | 0 | 29.138 | 0 | 0.651 | 0.522 | 1.704 | 3.777 | 1 | <MRL | <MRL | <MRL | 29.138 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 10305 | 202143 | 86474 | 795368 | 753415 | 6274669 | 12290455 | 2416998 | 22623919
139 | 2022-07-20 04:56 | WO-22270-88-A-1 | 138 | U |  | 0 | 1.205 | 2.168 | 29.571 | 0 | 0.822 | 0.458 | 2.057 | 4.311 | 1 | <MRL | <MRL | <MRL | 29.571 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 9324 | 181172 | 83236 | 709023 | 741640 | 5546647 | 11225127 | 2202924 | 23474682 | [5 empty] | 75 | AB-220715-01 LFSM-H 2 | S | 1 | 0 | 168.055 | 184.578 | 183.486 | 150.282 | 198.046 | 196.189 | 152.35 | 196.473 | H 2 | [18 empty] | FAIL
140 | 2022-07-20 05:32 | WO-22270-89-A-1 | 139 | U |  | 0 | 1.45 | 1.877 | 34.603 | 0 | 1.466 | 0.764 | 2.282 | 3.765 | 1 | <MRL | <MRL | <MRL | 34.603 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 7368 | 113973 | 47149 | 420888 | 574467 | 4370898 | 8539778 | 1660729 | 18307018
141 | 2022-07-20 06:08 | WO-22270-90-A-1 | 140 | U |  | 0 | 1.248 | 1.717 | 34.536 | 0 | 0 | 0 | 0.996 | 3.436 | 1 | <MRL | <MRL | <MRL | 34.536 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 7240 | 130434 | 59977 | 524788 | 587176 | 4568482 | 7750832 | 1514821 | 18276354 | [5 empty] | 103 | AB-220715-01 LFSM-H 3 | S | 1 | 0 | 176.858 | 191.02 | 182.049 | 151.448 | 197.926 | 198.045 | 154.815 | 200.285 | H 3 | [18 empty] | FAIL
142 | 2022-07-20 06:44 | AB-220715-01 LRB 10_INJ001 | 141 | K | QC | 0 | 0.855 | 0.364 | 0 | 0 | 1.382 | 1.144 | 0.955 | 0.39 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 51251 | 746001 | 294329 | 2761186 | 1996278 | 14133799 | 16210796 | 2386365 | 26542230 |  |  | 142
143 | 2022-07-20 07:20 | AB-220715-01 LRB 10 | 142 | K | QC | 0 | 0.656 | 0.581 | 1.06 | 0 | 1.482 | 1.02 | 1.082 | 0.442 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | IS4 | [13 empty] | 18497 | 504579 | 210863 | 1855802 | 1417805 | 10464064 | 13478997 | 2159872 | 20692741 |  |  | 143 |  |  | 131 | AB-220715-01 LFSM-H 4 | S | 1 | 0 | 170.383 | 190.142 | 193.605 | 149.526 | 196.17 | 204.594 | 151.141 | 201.625 | H 4 | [18 empty] | FAIL
144 | 2022-07-20 07:56 | WO-22270-91-A-1 | 143 | U |  | 0 | 0.964 | 0 | 89.122 | 0 | 2.673 | 0.341 | 1.001 | 2.934 | 1 | <MRL | <MRL | <MRL | 89.122 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 11147 | 199719 | 84669 | 748555 | 746753 | 5809917 | 11308729 | 2198104 | 23113498
145 | 2022-07-20 08:32 | WO-22270-92-A-1 | 144 | U |  | 0 | 1.274 | 1.92 | 118.451 | 0 | 1.678 | 0.488 | 1.027 | 3.82 | 1 | <MRL | <MRL | <MRL | 118.451 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 11620 | 211782 | 89135 | 832541 | 746938 | 5952482 | 12500813 | 2382566 | 24576698 | [5 empty] | 159 | AB-220715-01 LFSM-H 5 | S | 1 | 0 | 171.217 | 184.913 | 184.057 | 144.456 | 196.1 | 197.638 | 149.42 | 198.493 | H 5 | [18 empty] | FAIL
146 | 2022-07-20 09:08 | WO-22270-93-A-1 | 145 | U |  | 0 | 2.193 | 1.009 | 95.675 | 0 | 2.634 | 0.505 | 1.807 | 4.301 | 1 | <MRL | <MRL | <MRL | 95.675 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 12938 | 224370 | 99668 | 863242 | 805398 | 6257337 | 12289407 | 2352554 | 24694734
147 | 2022-07-20 09:44 | WO-22270-94-A-1 | 146 | U |  | 0 | 1.385 | 1.973 | 513.964 | 0 | 4.213 | 0.553 | 0.892 | 3.73 | 1 | <MRL | <MRL | <MRL | 513.964 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 10572 | 169459 | 80065 | 666687 | 715334 | 5846242 | 12149788 | 2195247 | 22499349 | [5 empty] | 187 | AB-220715-01 LFSM-H 6 | S | 1 | 0 | 178.901 | 202.898 | 189.642 | 145.053 | 200.076 | 195.386 | 146.784 | 201.023 | H 6 | [18 empty] | FAIL
148 | 2022-07-20 10:20 | WO-22270-95-A-1 | 147 | U |  | 0 | 1.772 | 1.934 | 296.758 | 0 | 1.156 | 0.443 | 0.933 | 3.59 | 1 | <MRL | <MRL | <MRL | 296.758 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | [8 empty] | 7724 | 141093 | 64144 | 565736 | 616178 | 4906830 | 9458542 | 1862901 | 20794691
149 | 2022-07-20 10:56 | WO-22270-96-A-1 | 148 | U |  | 0 | 0.84 | 0 | 429.411 | 0 | 2.135 | 0.395 | 1.081 | 3.898 | 1 | <MRL | <MRL | <MRL | 429.411 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 12916 | 211412 | 97388 | 813660 | 748224 | 5844232 | 11541020 | 2230558 | 23826961 | [5 empty] | 215 | AB-220715-01 LFSM-H 7 | S | 1 | 0 | 176.03 | 220.478 | 188.25 | 153.039 | 196.063 | 196.726 | 150.528 | 201.943 | H 7 | [18 empty] | FAIL
150 | 2022-07-20 11:32 | WO-22270-97-A-1 | 149 | U |  | 0 | 0.982 | 1.411 | 2.168 | 0 | 0.986 | 0.561 | 0.924 | 3.657 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 12260 | 210196 | 92828 | 843725 | 743371 | 5748496 | 11492106 | 2214819 | 23791887
151 | 2022-07-20 12:08 | WO-22270-98-A-1 | 150 | U |  | 0 | 0.913 | 3.012 | 1.918 | 0 | 0.611 | 0.427 | 0.86 | 4.501 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 10973 | 215239 | 90782 | 805018 | 750296 | 5652528 | 11383458 | 2104134 | 23127370 | [5 empty] | 243 | AB-220715-01 LFSM-H 8 | S | 1 | 0 | 178.667 | 214.18 | 184.277 | 144.849 | 202.262 | 199.472 | 151.908 | 187.394 | H 8 | [18 empty] | FAIL
152 | 2022-07-20 12:44 | WO-22270-99-A-1 | 151 | U |  | 0 | 1.609 | 2.588 | 0 | 0 | 0 | 1.034 | 2.825 | 4.353 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 6270 | 66716 | 30747 | 287145 | 412706 | 3190425 | 5970633 | 984422 | 12534084
153 | 2022-07-20 13:20 | WO-22270-100-A-1 | 152 | U |  | 0 | 2.049 | 0 | 1.921 | 0 | 1.687 | 0.447 | 0.868 | 4.313 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 6952 | 122929 | 54327 | 505037 | 565199 | 4410993 | 8549173 | 1587015 | 17880141 | [5 empty] | 273 | AB-220715-01 LFSM-H 9 | S | 1 | 0 | 169.264 | 197.643 | 200.41 | 146.015 | 195.88 | 197.609 | 153.973 | 191.971 | H 9 | [18 empty] | FAIL
154 | 2022-07-20 13:56 | AB-220715-01 LFB-H 6 | 153 | Z | QC | 156.313 | 166.991 | 182.753 | 191.319 | 148.72 | 191.836 | 188.34 | 161.946 | 188.731 | 1 | <MRL | 166.991 | 182.753 | 191.319 | 148.72 | 191.836 | 188.34 | 161.946 | 188.731 | H 6 | H 6 | NA | Blank Spike | No TrueV! | [15 empty] | 39912 | 590622 | 265343 | 2239385 | 1526333 | 10675315 | 14485110 | 2205900 | 21637634 |  |  | 154
155 | 2022-07-20 14:32 | AB-220715-01 LRB 11_INJ001 | 154 | K | QC | 0 | 0.684 | 0.603 | 0 | 0 | 1.569 | 0.798 | 2.04 | 0.503 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 76916 | 880847 | 377121 | 3189079 | 2134600 | 15551575 | 19371096 | 2808866 | 29425554 |  |  | 155 |  |  | 303 | AB-220715-01 LFSM-H 10 | S | 1 | 0 | 167.184 | 5413.034 | 2486.335 | 149.14 | 188.316 | 197.927 | 147.507 | 187.19 | H 10 | [18 empty] | FAIL |  | Same issue with 2-fold vs 4-fold dilution as in the LD.
156 | 2022-07-20 15:08 | AB-220715-01 LRB 11 | 155 | K | QC | 0 | 0.596 | 0.293 | 0.998 | 0 | 1.468 | 0.811 | 1.677 | 0.383 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 53079 | 853202 | 347007 | 3153962 | 2004282 | 15029356 | 18225009 | 2767757 | 27909548 |  |  | 156
157 | 2022-07-20 15:44 | AB-220715-01 LD 5 | 156 | D | QC | 0 | 2.34 | 1.745 | 2.824 | 0 | 3.298 | 1.161 | 3.174 | 4.628 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Duplicate | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 4897 | 66288 | 30142 | 270600 | 416502 | 3149240 | 6088699 | 1090966 | 12616466 |  |  |  | 157 |  | 332 | AB-220715-01 LFSM-H 11 | S | 1 | 0 | 165.797 | 2977.339 | 1328.5 | 144.818 | 194.494 | 194.899 | 151.821 | 194.326 | H 11 | [18 empty] | FAIL |  | This is just a low spike amount vs the ambient FPAS.
158 | 2022-07-20 16:20 | AB-220715-01 CCC-H 6 | 157 | Q | QC | 377.628 | 350.691 | 375.438 | 368.549 | 302.006 | 372.428 | 377.013 | 313.329 | 378.911 | 1 | 377.628 | 350.691 | 375.438 | 368.549 | 302.006 | 372.428 | 377.013 | 313.329 | 378.911 | 1 | 01 | NA | CCC | UL! | [15 empty] | 52567 | 618538 | 286156 | 2363626 | 1392965 | 10085369 | 13396215 | 2118416 | 18987306 |  | 158
159 | 2022-07-20 16:56 | AB-220715-01 LFSM-H 5 | 158 | S | QC | 0 | 171.217 | 184.913 | 184.057 | 144.456 | 196.1 | 197.638 | 149.42 | 198.493 | 1 | <MRL | 171.217 | 184.913 | 184.057 | 144.456 | 196.1 | 197.638 | 149.42 | 198.493 | H 5 | NA | H 5 | Sample Spike | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | [6 empty] | 6486 | 107413 | 49652 | 440231 | 510075 | 3657926 | 7062576 | 1348555 | 14363915 |  |  |  |  | 159
160 | 2022-07-20 17:32 | WO-22270-101-A-1 | 159 | U |  | 0 | 0.982 | 3.681 | 2.944 | 0 | 0.731 | 0.434 | 1.145 | 4.04 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 12883 | 214921 | 94777 | 875364 | 839518 | 6443524 | 13267462 | 2482525 | 26517096
161 | 2022-07-20 18:08 | WO-22270-102-A-1 | 160 | U |  | 0 | 1.622 | 1.641 | 2.814 | 0 | 1.723 | 0.503 | 2.04 | 4.044 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 11977 | 206879 | 88647 | 821723 | 723586 | 5790764 | 11273981 | 2132395 | 23878042
162 | 2022-07-20 18:44 | WO-22270-103-A-1 | 161 | U |  | 0 | 1.09 | 1.6 | 3.613 | 0 | 1.333 | 0.843 | 1.683 | 4.974 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | [10 empty] | 11421 | 258852 | 109029 | 1025468 | 1140323 | 9321004 | 18248804 | 3157958 | 32941178
163 | 2022-07-20 19:20 | WO-22270-104-A-1 | 162 | U |  | 0 | 1.437 | 1.636 | 8.259 | 0 | 0.625 | 0.365 | 1.001 | 3.97 | 1 | <MRL | <MRL | <MRL | 8.259 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | [8 empty] | 7143 | 146485 | 68480 | 594156 | 627835 | 4923165 | 9773223 | 1894019 | 20829860
164 | 2022-07-20 19:56 | WO-22270-105-A-1 | 163 | U |  | 0 | 1.197 | 1.889 | 7.025 | 0 | 0.608 | 0.782 | 1.069 | 4.763 | 1 | <MRL | <MRL | <MRL | 7.025 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 7358 | 149845 | 65824 | 594750 | 667558 | 5264972 | 10364697 | 1880352 | 20703722
165 | 2022-07-20 20:32 | WO-22270-106-A-1 | 164 | U |  | 0 | 0.963 | 0 | 7.749 | 0 | 0.523 | 0.619 | 2.235 | 4.019 | 1 | <MRL | <MRL | <MRL | 7.749 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 9382 | 169669 | 83047 | 757298 | 693100 | 5855602 | 11728390 | 2110216 | 23012653
166 | 2022-07-20 21:08 | WO-22270-107-A-1 | 165 | U |  | 0 | 1.928 | 2.217 | 26.028 | 0 | 0.535 | 0.499 | 1.906 | 3.585 | 1 | <MRL | <MRL | <MRL | 26.028 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 12559 | 221668 | 96248 | 868187 | 776790 | 6231741 | 12348570 | 2249895 | 24780276
167 | 2022-07-20 21:44 | WO-22270-108-A-1 | 166 | U |  | 0 | 1.173 | 3.237 | 24.938 | 0 | 0.557 | 0.489 | 1.045 | 3.661 | 1 | <MRL | <MRL | <MRL | 24.938 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 10918 | 184261 | 86995 | 789601 | 738575 | 6080203 | 12446477 | 2179844 | 22216407
168 | 2022-07-20 22:20 | WO-22270-109-A-1 | 167 | U |  | 0 | 1.642 | 1.92 | 36.818 | 0 | 0.617 | 1.347 | 0.961 | 4.195 | 1 | <MRL | <MRL | <MRL | 36.818 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 9062 | 176680 | 83904 | 714474 | 691625 | 5583680 | 11210476 | 2009241 | 23125269
169 | 2022-07-20 22:56 | WO-22270-110-A-1 | 168 | U |  | 0 | 0.734 | 2.75 | 35.931 | 0 | 1.446 | 1.286 | 0.98 | 3.898 | 1 | <MRL | <MRL | <MRL | 35.931 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 9122 | 167364 | 76826 | 717074 | 639255 | 5485504 | 10807877 | 2044642 | 22851904
170 | 2022-07-20 23:32 | AB-220715-01 LRB 12_INJ001 | 169 | K | QC | 0 | 0.705 | 0.578 | 0 | 0 | 0.51 | 0.335 | 1.732 | 0.332 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 55209 | 648203 | 289548 | 2598668 | 1718292 | 13146927 | 17583313 | 2596459 | 27455193 |  |  | 170
171 | 2022-07-21 00:08 | AB-220715-01 LRB 12 | 170 | K | QC | 0 | 0.946 | 0.319 | 0 | 0 | 0.517 | 0.321 | 1.105 | 0.346 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | [14 empty] | 24568 | 653605 | 272721 | 2474200 | 1613514 | 12028132 | 15855749 | 2430888 | 24969885 |  |  | 171
172 | 2022-07-21 00:44 | WO-22270-111-A-1 | 171 | U |  | 0 | 1.267 | 1.185 | 72.024 | 0 | 0.756 | 0.505 | 0.886 | 3.531 | 1 | <MRL | <MRL | <MRL | 72.024 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 8953 | 159741 | 72226 | 629552 | 607635 | 4971085 | 9759566 | 1801407 | 20700911
173 | 2022-07-21 01:20 | WO-22270-112-A-1 | 172 | U |  | 0 | 1.418 | 2.563 | 72.186 | 0 | 0.638 | 0.855 | 1.181 | 3.652 | 1 | <MRL | <MRL | <MRL | 72.186 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 8481 | 168350 | 82708 | 724965 | 659295 | 5706817 | 11568769 | 2029749 | 22186783
174 | 2022-07-21 01:56 | WO-22270-113-A-1 | 173 | U |  | 0 | 1.485 | 1.551 | 94.178 | 0 | 0.999 | 0.372 | 2.63 | 3.505 | 1 | <MRL | <MRL | <MRL | 94.178 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 10972 | 187370 | 88396 | 828045 | 671160 | 5569755 | 10515390 | 1882586 | 21945106
175 | 2022-07-21 02:32 | WO-22270-114-A-1 | 174 | U |  | 0 | 1.335 | 0 | 326.568 | 0 | 1.482 | 0.57 | 2.402 | 4.205 | 1 | <MRL | <MRL | <MRL | 326.568 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 7648 | 133810 | 58578 | 509911 | 588942 | 4792766 | 9550675 | 1760651 | 20575282
176 | 2022-07-21 03:08 | WO-22270-115-A-1 | 175 | U |  | 0 | 1.218 | 3.524 | 222.047 | 0 | 2.163 | 2.753 | 1.171 | 4.431 | 1 | <MRL | <MRL | <MRL | 222.047 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 10171 | 186150 | 86584 | 760625 | 684544 | 6071450 | 11883155 | 2120464 | 21189412
177 | 2022-07-21 03:44 | WO-22270-116-A-1 | 176 | U |  | 0 | 2.36 | 0 | 239.793 | 0 | 1.595 | 0.482 | 0.987 | 3.651 | 1 | <MRL | <MRL | <MRL | 239.793 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 10969 | 203753 | 83753 | 764118 | 705787 | 5794834 | 12331419 | 2244457 | 23720312
178 | 2022-07-21 04:20 | WO-22270-117-A-1 | 177 | U |  | 0 | 1.803 | 2.873 | 2.136 | 0 | 1.126 | 0.507 | 0.964 | 27.663 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 27.663 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 9589 | 179288 | 77324 | 724122 | 633138 | 5166087 | 9898041 | 1775753 | 18305780
179 | 2022-07-21 04:56 | WO-22270-118-A-1 | 178 | U |  | 0 | 1.197 | 0 | 2.136 | 0 | 0.666 | 0.423 | 0.913 | 2.864 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 9560 | 172905 | 75347 | 695421 | 660948 | 5271343 | 10991257 | 1922212 | 22323639
180 | 2022-07-21 05:32 | WO-22270-119-A-1 | 179 | U |  | 0 | 1.762 | 0 | 1.872 | 0 | 1.779 | 0.726 | 2.894 | 4.091 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 6734 | 130696 | 57455 | 526616 | 579807 | 4337529 | 8736161 | 1501090 | 18653659
181 | 2022-07-21 06:08 | WO-22270-120-A-1 | 180 | U |  | 0 | 1.069 | 0 | 1.811 | 0 | 1.868 | 0.958 | 1.033 | 4.189 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  | IS4 | [7 empty] | 9573 | 164834 | 71437 | 700236 | 625207 | 5097827 | 10252495 | 1830436 | 19981202
182 | 2022-07-21 06:44 | AB-220715-01 LFB-H 7 | 181 | Z | QC | 172.757 | 171.667 | 189.674 | 192.398 | 156.023 | 192.64 | 195.667 | 158.717 | 190.922 | 1 | <MRL | 171.667 | 189.674 | 192.398 | 156.023 | 192.64 | 195.667 | 158.717 | 190.922 | H 7 | H 7 | NA | Blank Spike | No TrueV! | [15 empty] | 52308 | 588835 | 271714 | 2320535 | 1348651 | 10225391 | 14186538 | 2121381 | 19325253 |  |  | 182
183 | 2022-07-21 07:20 | AB-220715-01 LRB 13_INJ001 | 182 | K | QC | 0 | 0.708 | 0.574 | 1.13 | 0 | 0.579 | 0.447 | 1.31 | 0.522 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 51800 | 677789 | 298463 | 2577659 | 1653649 | 12666893 | 16585527 | 2487994 | 25459165 |  |  | 183
184 | 2022-07-21 07:56 | AB-220715-01 LRB 13 | 183 | K | QC | 0 | 0.858 | 0.337 | 0.956 | 0 | 0.599 | 0.354 | 1.057 | 0.374 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | [14 empty] | 18202 | 602438 | 246703 | 2317303 | 1417471 | 10266887 | 13284930 | 1896730 | 19935037 |  |  | 184
185 | 2022-07-21 08:32 | AB-220715-01 LD 6 | 184 | D | QC | 0 | 1.889 | 0 | 1.984 | 0 | 1.785 | 0.814 | 0.93 | 3.868 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Duplicate | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 6918 | 121477 | 56706 | 527612 | 520854 | 4320638 | 8631675 | 1486179 | 18592197 |  |  |  | 185
186 | 2022-07-21 09:08 | AB-220715-01 CCC-H 7 | 185 | Q | QC | 369.407 | 335.708 | 385.941 | 375.749 | 303.514 | 387.221 | 384.409 | 317.825 | 372.392 | 1 | 369.407 | 335.708 | 385.941 | 375.749 | 303.514 | 387.221 | 384.409 | 317.825 | 372.392 | 1 | 01 | NA | CCC | UL! |  |  |  |  | IS3 | [10 empty] | 57256 | 599962 | 285666 | 2451488 | 1272681 | 9574526 | 13051808 | 1962335 | 16990268 |  | 186
187 | 2022-07-21 09:44 | AB-220715-01 LFSM-H 6 | 186 | S | QC | 0 | 178.901 | 202.898 | 189.642 | 145.053 | 200.076 | 195.386 | 146.784 | 201.023 | 1 | <MRL | 178.901 | 202.898 | 189.642 | 145.053 | 200.076 | 195.386 | 146.784 | 201.023 | H 6 | NA | H 6 | Sample Spike | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | [7 empty] | 7883 | 143272 | 67886 | 623203 | 547779 | 4488239 | 9111106 | 1676334 | 17823352 |  |  |  |  | 187
188 | 2022-07-21 10:20 | WO-22270-121-A-1 | 187 | U |  | 0 | 1.007 | 1.054 | 2.834 | 0 | 0 | 0.406 | 1.38 | 3.691 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 12286 | 224847 | 94550 | 872626 | 663452 | 5454144 | 10044035 | 1834652 | 18940059
189 | 2022-07-21 10:56 | WO-22270-122-A-1 | 188 | U |  | 0 | 2.476 | 0 | 4.153 | 0 | 1.193 | 0.714 | 0.971 | 3.841 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 5276 | 76465 | 37659 | 341178 | 432530 | 3497743 | 6324605 | 1091594 | 13693187
190 | 2022-07-21 11:32 | WO-22270-123-A-1 | 189 | U |  | 0 | 0.914 | 0 | 3.574 | 0 | 0.544 | 0.454 | 2.967 | 4.068 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 9628 | 187488 | 73169 | 716815 | 585518 | 5333152 | 11064215 | 1912900 | 21031285
191 | 2022-07-21 12:08 | WO-22270-124-A-1 | 190 | U |  | 0 | 1.17 | 0 | 8.903 | 0 | 3.316 | 0.583 | 1.115 | 4.517 | 1 | <MRL | <MRL | <MRL | 8.903 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 11548 | 194237 | 89594 | 796284 | 698798 | 5699289 | 11270126 | 1990011 | 20479034
192 | 2022-07-21 12:44 | WO-22270-125-A-1 | 191 | U |  | 0 | 2.165 | 0 | 8.16 | 0 | 1.426 | 0.925 | 0.877 | 4.939 | 1 | <MRL | <MRL | <MRL | 8.16 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 12097 | 188027 | 76656 | 780730 | 675536 | 6037692 | 10993672 | 1941943 | 19296311
193 | 2022-07-21 13:20 | WO-22270-126-A-1 | 192 | U |  | 0 | 2.014 | 0 | 8.852 | 0 | 2.024 | 0.878 | 0.952 | 3.937 | 1 | <MRL | <MRL | <MRL | 8.852 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 7859 | 117206 | 54211 | 528099 | 531254 | 4521694 | 8799705 | 1525797 | 18499001
194 | 2022-07-21 13:56 | WO-22270-127-A-1 | 193 | U |  | 0 | 1.381 | 0 | 28.051 | 0 | 2.717 | 0.484 | 0.903 | 3.877 | 1 | <MRL | <MRL | <MRL | 28.051 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 10110 | 145717 | 70416 | 625802 | 589442 | 4825909 | 8553778 | 1437646 | 17297313
195 | 2022-07-21 14:32 | WO-22270-128-A-1 | 194 | U |  | 0 | 2.627 | 0 | 33.692 | 0 | 1.561 | 0.52 | 1.558 | 3.491 | 1 | <MRL | <MRL | <MRL | 33.692 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 8183 | 145476 | 63632 | 626511 | 602592 | 4834988 | 9499258 | 1634387 | 19982121
196 | 2022-07-21 15:08 | WO-22270-129-A-1 | 195 | U |  | 0 | 0.913 | 1.124 | 38.125 | 0 | 2.706 | 0.837 | 0.979 | 4.154 | 1 | <MRL | <MRL | <MRL | 38.125 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 9689 | 187501 | 77786 | 745632 | 625168 | 5481662 | 11240230 | 1991819 | 23228761
197 | 2022-07-21 15:44 | WO-22270-130-A-1 | 196 | U |  | 0 | 1.542 | 0 | 39.499 | 0 | 21.233 | 1.533 | 0.829 | 5.842 | 1 | <MRL | <MRL | <MRL | 39.499 | <MRL | 21.233 | <MRL | <MRL | 5.842 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS3 | IS4 | [7 empty] | 7022 | 91354 | 50714 | 416689 | 470343 | 4352324 | 8791448 | 1412516 | 17402469
198 | 2022-07-21 16:20 | AB-220715-01 LRB 14_INJ001 | 197 | K | QC | 0 | 0.835 | 0.484 | 1.109 | 0 | 0.587 | 0.321 | 1.409 | 0.356 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | [15 empty] | 49023 | 605013 | 265948 | 2390547 | 1479328 | 11851395 | 15909323 | 2218538 | 24648982 |  |  | 198
199 | 2022-07-21 16:56 | AB-220715-01 LRB 14 | 198 | K | QC | 0 | 0.797 | 0.65 | 0 | 0 | 0.56 | 0.339 | 1.439 | 0.373 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | IS4 |  |  | IS4 |  |  | IS4 | [7 empty] | 17287 | 439049 | 184050 | 1665274 | 1164324 | 9277931 | 12816098 | 1765251 | 19358553 |  |  | 199
200 | 2022-07-21 17:32 | WO-22270-131-A-1 | 199 | U |  | 0 | 1.28 | 0 | 142.268 | 0 | 1.455 | 0.796 | 0.743 | 3.517 | 1 | <MRL | <MRL | <MRL | 142.268 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 7442 | 123781 | 56755 | 517798 | 550218 | 4487682 | 9084723 | 1494266 | 17528152
201 | 2022-07-21 18:08 | WO-22270-132-A-1 | 200 | U |  | 0 | 1.683 | 0 | 112.938 | 0 | 3.691 | 0.725 | 0 | 3.637 | 1 | <MRL | <MRL | <MRL | 112.938 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 10844 | 180112 | 82526 | 766261 | 653695 | 5903689 | 11965642 | 1980270 | 21597146
202 | 2022-07-21 18:44 | WO-22270-133-A-1 | 201 | U |  | 0 | 1.435 | 0 | 83.818 | 0 | 2.816 | 2.182 | 3.056 | 4.59 | 1 | <MRL | <MRL | <MRL | 83.818 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 7134 | 120387 | 53059 | 537745 | 538697 | 4592238 | 9484061 | 1634970 | 18901668
203 | 2022-07-21 19:20 | WO-22270-134-A-1 | 202 | U |  | 0 | 2.327 | 0 | 2231.973 | 0 | 2.168 | 3.39 | 1.377 | 4.942 | 1 | <MRL | <MRL | <MRL | 2231.973 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 8036 | 126845 | 62531 | 503867 | 566135 | 5122490 | 9798107 | 1689109 | 19266677
204 | 2022-07-21 19:56 | WO-22270-135-A-1 | 203 | U |  | 0 | 1.205 | 0 | 222.327 | 0 | 0.598 | 0.704 | 2.956 | 3.738 | 1 | <MRL | <MRL | <MRL | 222.327 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 6475 | 99331 | 49385 | 420742 | 494603 | 4251316 | 8965474 | 1445158 | 18105377
205 | 2022-07-21 20:32 | WO-22270-136-A-1 | 204 | U |  | 0 | 1.742 | 3.41 | 51.868 | 0 | 1.873 | 0.774 | 0.913 | 4.151 | 1 | <MRL | <MRL | <MRL | 51.868 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [9 empty] | 3981 | 121399 | 55484 | 591592 | 516680 | 5855521 | 13194586 | 2165432 | 22119472
206 | 2022-07-21 21:08 | WO-22270-137-A-1 | 205 | U |  | 0 | 1.387 | 11.283 | 1.438 | 0 | 1.092 | 0.447 | 0.983 | 2.975 | 1 | <MRL | <MRL | 11.283 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 6665 | 136130 | 66307 | 659323 | 477072 | 4941408 | 9263209 | 1297714 | 14943000
207 | 2022-07-21 21:44 | WO-22270-138-A-1 | 206 | U |  | 0 | 1.31 | 16.461 | 0 | 0 | 0.526 | 0.365 | 0.896 | 2.294 | 1 | <MRL | <MRL | 16.461 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 3857 | 101071 | 53890 | 520661 | 416637 | 4391642 | 9369027 | 1403595 | 13390151
208 | 2022-07-21 22:20 | WO-22270-139-A-1 | 207 | U |  | 0 | 1.429 | 15.189 | 0 | 0 | 0.633 | 0.538 | 0.876 | 4.91 | 1 | <MRL | <MRL | 15.189 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 3358 | 53525 | 30142 | 323209 | 350065 | 3271597 | 6537586 | 1022183 | 14070998
209 | 2022-07-21 22:56 | WO-22270-140-A-1 | 208 | U |  | 0 | 1.545 | 10.723 | 0 | 0 | 0.612 | 0.56 | 0.971 | 3.954 | 1 | <MRL | <MRL | 10.723 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 6284 | 104719 | 53287 | 495760 | 463804 | 4299628 | 8340094 | 1339615 | 17201517
210 | 2022-07-21 23:32 | AB-220715-01 LFB-H 8 | 209 | Z | QC | 184.848 | 169.13 | 182.476 | 190.233 | 153.04 | 200.802 | 207.329 | 153.851 | 183.032 | 1 | <MRL | 169.13 | 182.476 | 190.233 | 153.04 | 200.802 | 207.329 | 153.851 | 183.032 | H 8 | H 8 | NA | Blank Spike | No TrueV! |  | IS3 |  |  | IS3 | [10 empty] | 52700 | 523805 | 240908 | 2042640 | 1162800 | 8931373 | 12567876 | 2019840 | 20685732 |  |  | 210
211 | 2022-07-22 00:08 | AB-220715-01 LRB 15_INJ001 | 210 | K | QC | 0 | 1.064 | 0.909 | 1.636 | 1.4 | 1.407 | 1.026 | 1.582 | 1.466 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! |  | IS4 | [13 empty] | 33463 | 524588 | 231321 | 2322692 | 1528681 | 12555336 | 15948317 | 2366627 | 25839013 |  |  | 211
212 | 2022-07-22 00:44 | AB-220715-01 LRB 15 | 211 | K | QC | 0 | 0.813 | 0.951 | 1.008 | 0 | 0.881 | 0.374 | 1.271 | 0.494 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | IS4 | [13 empty] | 23381 | 525317 | 251665 | 2179518 | 1521114 | 11929136 | 15116895 | 2262813 | 26677815 |  |  | 212
213 | 2022-07-22 01:20 | AB-220715-01 LD 7 | 212 | D | QC | 0 | 1.639 | 7.487 | 0 | 0 | 1.55 | 0.653 | 1.203 | 4.949 | 1 | <MRL | <MRL | 7.487 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Duplicate | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 6424 | 77348 | 45449 | 370515 | 442798 | 3675123 | 7328787 | 1199582 | 14886295 |  |  |  | 213
214 | 2022-07-22 01:56 | AB-220715-01 CCC-H 8 | 213 | Q | QC | 380.715 | 336.433 | 385.036 | 378.446 | 316.872 | 374.878 | 379.099 | 325.053 | 375.361 | 1 | 380.715 | 336.433 | 385.036 | 378.446 | 316.872 | 374.878 | 379.099 | 325.053 | 375.361 | 1 | 01 | NA | CCC | UL! |  | IS3 |  |  | IS3 |  |  | IS3 | [7 empty] | 41276 | 489858 | 237285 | 2077923 | 1056800 | 7967798 | 11021968 | 1667999 | 15691752 |  | 214
215 | 2022-07-22 02:32 | AB-220715-01 LFSM-H 7 | 214 | S | QC | 0 | 176.03 | 220.478 | 188.25 | 153.039 | 196.063 | 196.726 | 150.528 | 201.943 | 1 | <MRL | 176.03 | 220.478 | 188.25 | 153.039 | 196.063 | 196.726 | 150.528 | 201.943 | H 7 | NA | H 7 | Sample Spike | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | [7 empty] | 8678 | 123238 | 60200 | 621772 | 494490 | 4360705 | 8084466 | 1414281 | 15804340 |  |  |  |  | 215
216 | 2022-07-22 03:08 | WO-22270-141-A-1 | 215 | U |  | 0 | 1.527 | 34.124 | 1.641 | 0 | 1.229 | 0.808 | 1.175 | 3.221 | 1 | <MRL | <MRL | 34.124 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 8651 | 145861 | 71295 | 683065 | 573003 | 4969822 | 9793481 | 1717122 | 19782291
217 | 2022-07-22 03:44 | WO-22270-142-A-1 | 216 | U |  | 0 | 1.463 | 33.507 | 0 | 0 | 1.132 | 0.486 | 1.051 | 2.606 | 1 | <MRL | <MRL | 33.507 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 6102 | 125794 | 63843 | 546595 | 502274 | 4796678 | 9330821 | 1508916 | 19080449
218 | 2022-07-22 04:20 | WO-22270-143-A-1 | 217 | U |  | 0 | 2.015 | 26.498 | 0 | 0 | 0.601 | 0.608 | 0.984 | 3.584 | 1 | <MRL | <MRL | 26.498 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 5425 | 61997 | 31973 | 297470 | 362414 | 3388263 | 6503170 | 1086537 | 14059759
219 | 2022-07-22 04:56 | WO-22270-144-A-1 | 218 | U |  | 0 | 1.164 | 102.136 | 0 | 0 | 1.771 | 0.83 | 1.097 | 3.23 | 1 | <MRL | <MRL | 102.136 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 |  | IS4 | [7 empty] | 10473 | 145314 | 70659 | 686914 | 573430 | 5437062 | 10268993 | 1655493 | 20363449
220 | 2022-07-22 05:32 | WO-22270-145-A-1 | 219 | U |  | 0 | 0.689 | 88.027 | 0 | 0 | 1.335 | 0.525 | 0.844 | 2.649 | 1 | <MRL | <MRL | 88.027 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 10870 | 143294 | 77164 | 643791 | 547070 | 4886714 | 9773329 | 1600079 | 21179565
221 | 2022-07-22 06:08 | WO-22270-146-A-1 | 220 | U |  | 0 | 0.882 | 92.579 | 0 | 0 | 1.95 | 0.591 | 0.965 | 2.223 | 1 | <MRL | <MRL | 92.579 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 8906 | 90916 | 48029 | 460155 | 433524 | 4034522 | 7903729 | 1225828 | 16305276
222 | 2022-07-22 06:44 | WO-22270-147-A-1 | 221 | U |  | 0 | 1.916 | 324.339 | 0 | 0 | 2.685 | 0.489 | 1.445 | 3.744 | 1 | <MRL | <MRL | 324.339 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 5007 | 47744 | 27133 | 246120 | 339378 | 3225824 | 6775621 | 1042155 | 11583296
223 | 2022-07-22 07:20 | WO-22270-148-A-1 | 222 | U |  | 0 | 1.139 | 304.539 | 2.76 | 0 | 4.589 | 0.533 | 1.57 | 2.35 | 1 | <MRL | <MRL | 304.539 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 |  | IS4 | [7 empty] | 9707 | 165268 | 78227 | 774954 | 569266 | 5704854 | 10865694 | 1659164 | 21295071
224 | 2022-07-22 07:56 | WO-22270-149-A-1 | 223 | U |  | 0 | 0.908 | 312.784 | 2.979 | 0 | 1.944 | 0.671 | 0.928 | 4.147 | 1 | <MRL | <MRL | 312.784 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 |  | IS4 | [7 empty] | 9394 | 169894 | 72788 | 737088 | 594572 | 5450588 | 10149306 | 1681935 | 20994098
225 | 2022-07-22 08:32 | WO-22270-150-A-1 | 224 | U |  | 0 | 1.28 | 267.728 | 0 | 0 | 0.552 | 0.73 | 0.854 | 3.079 | 1 | <MRL | <MRL | 267.728 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 8914 | 123721 | 64735 | 588086 | 499943 | 4809651 | 8941046 | 1482777 | 18687443
226 | 2022-07-22 09:08 | AB-220715-01 LRB 16_INJ001 | 225 | K | QC | 0 | 0.682 | 2.056 | 0 | 0 | 0.858 | 0.56 | 1.009 | 0.43 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | IS4 | d |  | IS4 | [10 empty] | 31362 | 412519 | 192514 | 1839962 | 1165125 | 10435240 | 14346854 | 2040059 | 24724590 |  |  | 226
227 | 2022-07-22 09:44 | AB-220715-01 LRB 16 | 226 | K | QC | 0 | 0.881 | 1.823 | 0 | 0 | 0.865 | 0.378 | 0.899 | 0.429 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | IS4 | d |  | IS4 | [10 empty] | 28700 | 398216 | 183473 | 1841664 | 1207377 | 10238552 | 13706642 | 2094631 | 23758425 |  |  | 227
228 | 2022-07-22 10:20 | WO-22270-151-A-1 | 227 | U |  | 0 | 0.918 | 993.029 | 1.571 | 0 | 0.623 | 1.754 | 0.829 | 2.406 | 1 | <MRL | <MRL | 993.029 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 6017 | 115917 | 57239 | 565315 | 501542 | 4398704 | 8451782 | 1396366 | 17809073
229 | 2022-07-22 10:56 | WO-22270-152-A-1 | 228 | U |  | 0 | 1.351 | 1050.632 | 1.672 | 0 | 0.75 | 0.517 | 2.257 | 2.648 | 1 | <MRL | <MRL | 1050.632 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 9329 | 147061 | 63729 | 658354 | 535569 | 4858332 | 9492805 | 1549125 | 19950926
230 | 2022-07-22 11:32 | WO-22270-153-A-1 | 229 | U |  | 0 | 1.915 | 958.71 | 2.461 | 0 | 1.145 | 0.645 | 1.34 | 3.295 | 1 | <MRL | <MRL | 958.71 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 4068 | 43772 | 22075 | 226704 | 309984 | 2952713 | 6427292 | 1051780 | 12714269
231 | 2022-07-22 12:08 | WO-22270-154-A-1 | 230 | U |  | 0 | 0.893 | 2069.312 | 2.409 | 0 | 0.773 | 5.502 | 1.161 | 4.213 | 1 | <MRL | <MRL | 2069.312 | <MRL | <MRL | <MRL | 5.502 | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS3 | [7 empty] | 7703 | 130265 | 65835 | 633498 | 527595 | 4786316 | 9308754 | 1516944 | 18892054
232 | 2022-07-22 12:44 | WO-22270-155-A-1 | 231 | U |  | 0 | 0.924 | 3164.557 | 2.974 | 0 | 0.664 | 0.599 | 0.936 | 3.675 | 1 | <MRL | <MRL | 3164.557 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | [10 empty] | 9634 | 212998 | 92967 | 1072043 | 1124693 | 10742136 | 17632076 | 2483686 | 29629816
233 | 2022-07-22 13:20 | WO-22270-156-A-1 | 232 | U |  | 0 | 1.656 | 2732.931 | 2.804 | 0 | 1.519 | 0.575 | 1.285 | 3.488 | 1 | <MRL | <MRL | 2732.931 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 6618 | 113207 | 48029 | 545367 | 477127 | 4447850 | 9283976 | 1419354 | 19334524
234 | 2022-07-22 13:56 | WO-22270-157-A-1 | 233 | U |  | 0 | 1.469 | 17.397 | 3.054 | 0 | 1.923 | 4.66 | 1.133 | 3.856 | 1 | <MRL | <MRL | 17.397 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 7115 | 127295 | 65854 | 593474 | 544228 | 5196545 | 8946909 | 1504434 | 19537435
235 | 2022-07-22 14:32 | WO-22270-158-A-1 | 234 | U |  | 0 | 1.107 | 7.677 | 1.143 | 0 | 1.129 | 1.068 | 2.294 | 2.164 | 1 | <MRL | <MRL | 7.677 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 7034 | 124087 | 59343 | 611183 | 496865 | 4676702 | 9566553 | 1445449 | 18063766
236 | 2022-07-22 15:08 | WO-22270-159-A-1 | 235 | U |  | 0 | 1.589 | 8.152 | 0 | 0 | 1.534 | 0.726 | 1.054 | 3.551 | 1 | <MRL | <MRL | 8.152 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | [9 empty] | 5910 | 115004 | 63918 | 591800 | 701554 | 7154057 | 13331964 | 2006936 | 25609527
237 | 2022-07-22 15:44 | WO-22270-160-A-1 | 236 | U |  | 0 | 1.25 | 3.988 | 0 | 0 | 0.554 | 0.636 | 2.993 | 3.234 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | d | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 5944 | 72027 | 37585 | 378715 | 366427 | 3403835 | 6645223 | 1030482 | 12729966
238 | 2022-07-22 16:20 | AB-220715-01 LFB-H 9 | 237 | Z | QC | 225.666 | 160.929 | 180.037 | 195.058 | 144.71 | 199.266 | 209.15 | 147.606 | 201.143 | 1 | 225.666 | 160.929 | 180.037 | 195.058 | 144.71 | 199.266 | 209.15 | 147.606 | 201.143 | H 9 | H 9 | NA | Blank Spike | No TrueV! | IS3 | IS3 |  |  | IS3 |  |  | IS3 | [7 empty] | 24022 | 437559 | 211075 | 1841312 | 1000400 | 8235689 | 11264146 | 1663172 | 17287027 |  |  | 238
239 | 2022-07-22 16:56 | AB-220715-01 LRB 17_INJ001 | 238 | K | QC | 0 | 1.121 | 0.812 | 1.817 | 1.497 | 1.693 | 1.235 | 1.555 | 1.538 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | IS4 |  |  | IS4 | [10 empty] | 30471 | 440566 | 206198 | 1982393 | 1283922 | 10618169 | 14436278 | 2018735 | 24594162 |  |  | 239
240 | 2022-07-22 17:32 | AB-220715-01 LRB 17 | 239 | K | QC | 0 | 0.641 | 3.442 | 1.134 | 0 | 0.893 | 0.388 | 1.002 | 0.466 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | IS4 | [13 empty] | 26613 | 482924 | 195002 | 1894534 | 1362703 | 10012992 | 14326661 | 2164904 | 23348606 |  |  | 240
241 | 2022-07-22 18:08 | AB-220715-01 LD 8 | 240 | D | QC | 0 | 1.145 | 8.326 | 2.782 | 0 | 3.612 | 0.66 | 1.327 | 3.377 | 1 | <MRL | <MRL | 8.326 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Duplicate | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 6367 | 78067 | 38257 | 349396 | 392269 | 3429914 | 6126156 | 1034197 | 13275971 |  |  |  | 241
242 | 2022-07-22 18:44 | AB-220715-01 CCC-H 9 | 241 | Q | QC | 361.131 | 341.763 | 369.546 | 391.441 | 304.674 | 386.79 | 391.626 | 311.608 | 363.614 | 1 | 361.131 | 341.763 | 369.546 | 391.441 | 304.674 | 386.79 | 391.626 | 311.608 | 363.614 | 1 | 01 | NA | CCC | UL! |  | IS3 |  |  | IS3 |  |  | IS3 | [7 empty] | 51690 | 510230 | 245512 | 1971418 | 1072569 | 8055316 | 10457366 | 1695888 | 16032009 |  | 242
243 | 2022-07-22 19:20 | AB-220715-01 LFSM-H 8 | 242 | S | QC | 0 | 178.667 | 214.18 | 184.277 | 144.849 | 202.262 | 199.472 | 151.908 | 187.394 | 1 | <MRL | 178.667 | 214.18 | 184.277 | 144.849 | 202.262 | 199.472 | 151.908 | 187.394 | H 8 | NA | H 8 | Sample Spike | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | [6 empty] | 7833 | 97378 | 48644 | 479856 | 462711 | 3756214 | 7389918 | 1294954 | 14810303 |  |  |  |  | 243
244 | 2022-07-22 19:56 | WO-22270-161-A-1 | 243 | U |  | 0 | 1.741 | 12.151 | 1.612 | 0 | 0.627 | 1.19 | 1.224 | 2.606 | 1 | <MRL | <MRL | 12.151 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 4753 | 96858 | 49045 | 412731 | 421825 | 3674705 | 7362581 | 1294928 | 16116943
245 | 2022-07-22 20:32 | WO-22270-162-A-1 | 244 | U |  | 0 | 1.197 | 15.364 | 0 | 0 | 0.958 | 0.54 | 1.038 | 3.833 | 1 | <MRL | <MRL | 15.364 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 5357 | 114259 | 54481 | 487440 | 500477 | 4109872 | 8525979 | 1468628 | 16928868
246 | 2022-07-22 21:08 | WO-22270-163-A-1 | 245 | U |  | 0 | 1.477 | 14.745 | 0 | 0 | 1.607 | 0.599 | 1.276 | 2.784 | 1 | <MRL | <MRL | 14.745 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 3741 | 54525 | 30667 | 273405 | 345906 | 2898261 | 5966876 | 1032507 | 12857812
247 | 2022-07-22 21:44 | WO-22270-164-A-1 | 246 | U |  | 0 | 1.355 | 41.544 | 1.282 | 0 | 1.813 | 1.221 | 2.595 | 3.054 | 1 | <MRL | <MRL | 41.544 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 8580 | 145041 | 75186 | 666467 | 564007 | 5072580 | 9497668 | 1585522 | 19781444
248 | 2022-07-22 22:20 | WO-22270-165-A-1 | 247 | U |  | 0 | 1.596 | 41.894 | 0 | 0 | 0.662 | 0.495 | 0.961 | 3.212 | 1 | <MRL | <MRL | 41.894 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 4607 | 131824 | 67796 | 556366 | 505746 | 4585177 | 8954291 | 1487528 | 17130650
249 | 2022-07-22 22:56 | WO-22270-166-A-1 | 248 | U |  | 0 | 1.124 | 43.64 | 0 | 0 | 0.702 | 0.432 | 1.05 | 3.827 | 1 | <MRL | <MRL | 43.64 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 6403 | 132070 | 64993 | 591614 | 571861 | 4721342 | 9307376 | 1578553 | 20176247
250 | 2022-07-22 23:32 | WO-22270-167-A-1 | 249 | U |  | 0 | 1.721 | 148.683 | 0 | 0 | 1.69 | 1.157 | 1.072 | 5.225 | 1 | <MRL | <MRL | 148.683 | <MRL | <MRL | <MRL | <MRL | <MRL | 5.225 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 3531 | 74108 | 38149 | 344472 | 416676 | 3611791 | 8053607 | 1214396 | 16788719
251 | 2022-07-23 00:08 | WO-22270-168-A-1 | 250 | U |  | 0 | 1.091 | 139.298 | 0 | 0 | 1.087 | 0.534 | 2.366 | 2.786 | 1 | <MRL | <MRL | 139.298 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 7327 | 133297 | 62015 | 576787 | 509001 | 4653025 | 9125726 | 1540673 | 18385424
252 | 2022-07-23 00:44 | WO-22270-169-A-1 | 251 | U |  | 0 | 1.6 | 211.612 | 1.952 | 0 | 1.623 | 0.82 | 1.398 | 3.554 | 1 | <MRL | <MRL | 211.612 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 5124 | 72228 | 31504 | 318093 | 387787 | 3405037 | 7248648 | 1161526 | 14113607
253 | 2022-07-23 01:20 | WO-22270-170-A-1 | 252 | U |  | 0 | 1.209 | 223.747 | 0 | 0 | 0.549 | 0.459 | 2.394 | 2.996 | 1 | <MRL | <MRL | 223.747 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 8010 | 107664 | 53943 | 491147 | 501441 | 4397285 | 8815748 | 1480671 | 17762304
254 | 2022-07-23 01:56 | AB-220715-01 LRB 18_INJ001 | 253 | K | QC | 0 | 0.741 | 2.299 | 1.219 | 2.417 | 0.906 | 0.361 | 0.953 | 0.499 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | IS4 | [13 empty] | 29718 | 482731 | 214031 | 1918687 | 1416177 | 11009099 | 15008419 | 2214346 | 24380115 |  |  | 254
255 | 2022-07-23 02:32 | WO-22209-62-A-1B | 254 | U |  | 0 | 1.04 | 0 | 0 | 0 | 3.26 | 0.453 | 0.892 | 114.95 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 114.95 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | Out of holding time, electively reanalyzed |  |  | 5468 | 81465 | 43172 | 407490 | 388031 | 4145766 | 8313996 | 1129293 | 15793284
256 | 2022-07-23 03:08 | AB-220715-01 LRB 18 | 255 | K | QC | 0 | 0.778 | 1.268 | 1.199 | 0 | 0.945 | 0.363 | 0.933 | 0.455 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | IS4 |  |  | IS4 | [10 empty] | 15606 | 438976 | 200118 | 1770614 | 1260583 | 9609305 | 13137166 | 2051760 | 21369570 |  |  | 256
257 | 2022-07-23 03:44 | WO-22270-171-A-1 | 256 | U |  | 0 | 0.883 | 448.566 | 0 | 0 | 1.172 | 1.36 | 0.835 | 2.703 | 1 | <MRL | <MRL | 448.566 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 8750 | 146873 | 67583 | 642829 | 549928 | 4696955 | 9109177 | 1549442 | 18517951
258 | 2022-07-23 04:20 | WO-22270-172-A-1 | 257 | U |  | 0 | 0.83 | 490.756 | 1.676 | 0 | 1.077 | 0.561 | 2.218 | 3.409 | 1 | <MRL | <MRL | 490.756 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 7624 | 140379 | 69748 | 624566 | 534822 | 4712013 | 9120655 | 1540644 | 18818518
259 | 2022-07-23 04:56 | WO-22270-173-A-1 | 258 | U |  | 0 | 1.345 | 652.269 | 1.572 | 0 | 1.491 | 0.645 | 0.901 | 2.969 | 1 | <MRL | <MRL | 652.269 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 7022 | 152186 | 69738 | 663709 | 535398 | 4820123 | 9704164 | 1539161 | 20252525
260 | 2022-07-23 05:32 | WO-22270-174-A-1 | 259 | U |  | 0 | 1.375 | 2324.998 | 3.302 | 0 | 1.675 | 2.072 | 1.104 | 2.746 | 1 | <MRL | <MRL | 2324.998 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 5824 | 124965 | 55322 | 519485 | 521543 | 4975969 | 9471485 | 1533210 | 17439954
261 | 2022-07-23 06:08 | WO-22270-175-A-1 | 260 | U |  | 0 | 0.928 | 1553.208 | 1.554 | 0 | 0.839 | 0.439 | 1.139 | 2.174 | 1 | <MRL | <MRL | 1553.208 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 6686 | 129497 | 64679 | 594416 | 543406 | 4637843 | 9280381 | 1545338 | 18830491
262 | 2022-07-23 06:44 | WO-22270-176-A-1 | 261 | U |  | 0 | 3.32 | 2159.757 | 3.578 | 0 | 0.633 | 0.586 | 1.438 | 2.838 | 1 | <MRL | <MRL | 2159.757 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 4826 | 64699 | 33535 | 308150 | 376408 | 3435999 | 6876162 | 1113208 | 14235878
263 | 2022-07-23 07:20 | WO-22270-177-A-1 | 262 | U |  | 0 | 0.98 | 14.533 | 0 | 0 | 0.592 | 1.46 | 0.977 | 1.97 | 1 | <MRL | <MRL | 14.533 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 |  | IS4 | [7 empty] | 6742 | 142206 | 73075 | 688421 | 645123 | 5828726 | 11134728 | 1831898 | 21958673
264 | 2022-07-23 07:56 | WO-22270-178-A-1 | 263 | U |  | 0 | 1.21 | 11.442 | 0 | 0 | 1.784 | 0.663 | 2.858 | 2.336 | 1 | <MRL | <MRL | 11.442 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 3037 | 62119 | 31875 | 311949 | 328193 | 2867754 | 6148355 | 1060428 | 13680992
265 | 2022-07-23 08:32 | WO-22270-179-A-1 | 264 | U |  | 0 | 1.082 | 9.067 | 0 | 0 | 0.588 | 0.545 | 0.941 | 3.117 | 1 | <MRL | <MRL | 9.067 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 7275 | 91873 | 44654 | 411207 | 449059 | 3815560 | 7580320 | 1270501 | 16153915
266 | 2022-07-23 09:08 | WO-22270-180-A-1 | 265 | U |  | 0 | 1.39 | 7.807 | 1.518 | 0 | 1.33 | 0.622 | 2.126 | 3.526 | 1 | <MRL | <MRL | 7.807 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 6772 | 115193 | 59021 | 543370 | 509927 | 4484080 | 9321170 | 1482883 | 18072983
267 | 2022-07-23 09:44 | WO-22209-64-A-1B | 266 | U |  | 0 | 1.237 | 0 | 0 | 0 | 7.981 | 0.769 | 0.812 | 567.441 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | 7.981 | <MRL | <MRL | 567.441 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 |  |  |  | Out of holding time, electively reanalyzed |  |  | 6654 | 74984 | 41538 | 393360 | 380524 | 4102904 | 8234356 | 1135938 | 13756687
268 | 2022-07-23 10:20 | AB-220715-01 LFB-H 10 | 267 | Z | QC | 148.473 | 165.749 | 181.985 | 193.018 | 154.851 | 200.351 | 196.216 | 164.645 | 192.246 | 1 | <MRL | 165.749 | 181.985 | 193.018 | 154.851 | 200.351 | 196.216 | 164.645 | 192.246 | H 10 | H 10 | NA | Blank Spike | No TrueV! |  | IS3 |  |  | IS3 |  |  | IS3 | [7 empty] | 47073 | 502091 | 231281 | 2024045 | 1142118 | 8528689 | 12295535 | 1788550 | 19219979 |  |  | 268
269 | 2022-07-23 10:56 | AB-220715-01 LRB 19_INJ001 | 268 | K | QC | 0 | 1.233 | 2.874 | 1.801 | 2.675 | 1.349 | 0.997 | 1.411 | 1.205 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! |  | IS4 | [13 empty] | 32898 | 523247 | 228759 | 2083916 | 1433477 | 12104147 | 16362010 | 2324384 | 27205917 |  |  | 269
270 | 2022-07-23 11:32 | AB-220715-01 LRB 19 | 269 | K | QC | 0 | 0.991 | 2.143 | 1.098 | 2.013 | 0.961 | 0.459 | 1.059 | 0.481 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | IS4 |  |  | IS4 | [10 empty] | 18754 | 455894 | 226224 | 1989883 | 1226736 | 10277703 | 13901883 | 1992401 | 23794591 |  |  | 270
271 | 2022-07-23 12:08 | AB-220715-01 LD 9 | 270 | D | QC | 0 | 1.746 | 7.444 | 0 | 0 | 0.625 | 0.708 | 2.715 | 2.987 | 1 | <MRL | <MRL | 7.444 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Duplicate | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 6853 | 87728 | 44848 | 427754 | 437075 | 4044534 | 7983247 | 1344566 | 16173618 |  |  |  | 271
272 | 2022-07-23 12:44 | AB-220715-01 CCC-H 10 | 271 | Q | QC | 368.792 | 339.369 | 386.392 | 388.865 | 312.752 | 398.474 | 383.256 | 345.56 | 383.207 | 1 | 368.792 | 339.369 | 386.392 | 388.865 | 312.752 | 398.474 | 383.256 | 345.56 | 383.207 | 1 | 01 | NA | CCC | UL! |  | IS3 |  |  | IS3 | IS3 |  | IS3 | [7 empty] | 48911 | 496252 | 222916 | 1937687 | 1001249 | 7871819 | 11186151 | 1533437 | 16156602 |  | 272
273 | 2022-07-23 13:20 | AB-220715-01 LFSM-H 9 | 272 | S | QC | 0 | 169.264 | 197.643 | 200.41 | 146.015 | 195.88 | 197.609 | 153.973 | 191.971 | 1 | <MRL | 169.264 | 197.643 | 200.41 | 146.015 | 195.88 | 197.609 | 153.973 | 191.971 | H 9 | NA | H 9 | Sample Spike | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | [6 empty] | 4407 | 87609 | 44911 | 388732 | 399748 | 3360966 | 6758566 | 1117379 | 13002454 |  |  |  |  | 273
274 | 2022-07-23 13:56 | WO-22270-181-A-1 | 273 | U |  | 0 | 1.563 | 15.258 | 1.517 | 0 | 2.224 | 1.717 | 1.208 | 5.86 | 1 | <MRL | <MRL | 15.258 | <MRL | <MRL | <MRL | <MRL | <MRL | 5.86 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 7099 | 128274 | 66730 | 637872 | 530007 | 4816138 | 8673978 | 1431554 | 18956485
275 | 2022-07-23 14:32 | WO-22270-182-A-1 | 274 | U |  | 0 | 1.902 | 13.876 | 0 | 0 | 1.034 | 0.692 | 1.079 | 4.081 | 1 | <MRL | <MRL | 13.876 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 6042 | 77988 | 43052 | 398398 | 404719 | 3981976 | 8167641 | 1188611 | 16459906
276 | 2022-07-23 15:08 | WO-22270-183-A-1 | 275 | U |  | 0 | 0.822 | 16.358 | 0 | 0 | 1.842 | 0.573 | 1.153 | 3.847 | 1 | <MRL | <MRL | 16.358 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 8715 | 138842 | 62339 | 608931 | 528202 | 4724255 | 9552472 | 1473230 | 19253640
277 | 2022-07-23 15:44 | WO-22270-184-A-1 | 276 | U |  | 0 | 1.233 | 35.552 | 0 | 0 | 1.077 | 1.819 | 0.882 | 3.932 | 1 | <MRL | <MRL | 35.552 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 8644 | 134816 | 70329 | 639985 | 517852 | 4647348 | 9434538 | 1462737 | 20139215
278 | 2022-07-23 16:20 | WO-22270-185-A-1 | 277 | U |  | 0 | 1.791 | 45.211 | 4.181 | 0 | 0.856 | 0.582 | 0.861 | 3.544 | 1 | <MRL | <MRL | 45.211 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | [10 empty] | 8066 | 173171 | 90282 | 814638 | 741466 | 8056361 | 14661485 | 2206331 | 23150911
279 | 2022-07-23 16:56 | WO-22209-50-A-1B | 278 | U |  | 0 | 1.689 | 0 | 0 | 0 | 32.495 | 1.284 | 2.502 | 1038.324 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | 32.495 | <MRL | <MRL | 1038.324 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 |  |  |  | Out of holding time, electively reanalyzed |  |  | 6448 | 96543 | 50146 | 514912 | 438961 | 4449372 | 9071610 | 1291088 | 13086565
280 | 2022-07-23 17:32 | WO-22270-186-A-1 | 279 | U |  | 0 | 1.268 | 49.661 | 2.227 | 0 | 1.14 | 0.7 | 3.391 | 3.091 | 1 | <MRL | <MRL | 49.661 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 5478 | 69596 | 34076 | 326720 | 358339 | 3342417 | 6822586 | 1068896 | 13512307
281 | 2022-07-23 18:08 | WO-22270-187-A-1 | 280 | U |  | 0 | 1.112 | 190.082 | 1.457 | 0 | 1.973 | 0.67 | 1.227 | 4.366 | 1 | <MRL | <MRL | 190.082 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | [9 empty] | 20413 | 247399 | 122710 | 1155541 | 678660 | 6336298 | 12985380 | 1965195 | 23880077
282 | 2022-07-23 18:44 | WO-22270-188-A-1 | 281 | U |  | 0 | 1.364 | 166.615 | 1.564 | 0 | 4.744 | 0.578 | 0.929 | 7.36 | 1 | <MRL | <MRL | 166.615 | <MRL | <MRL | <MRL | <MRL | <MRL | 7.36 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | [10 empty] | 20853 | 269362 | 125616 | 1236410 | 761956 | 7920627 | 14474228 | 2152651 | 24832194
283 | 2022-07-23 19:20 | WO-22270-189-A-1 | 282 | U |  | 0 | 1.217 | 163.398 | 1.988 | 0 | 1.948 | 1.001 | 1.014 | 6.961 | 1 | <MRL | <MRL | 163.398 | <MRL | <MRL | <MRL | <MRL | <MRL | 6.961 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | [9 empty] | 16292 | 230129 | 120679 | 1042060 | 719147 | 6857499 | 14163833 | 2026103 | 26099177
284 | 2022-07-23 19:56 | WO-22270-190-A-1 | 283 | U |  | 0 | 1.053 | 157.924 | 1.455 | 0 | 1.469 | 1.21 | 0.88 | 7.471 | 1 | <MRL | <MRL | 157.924 | <MRL | <MRL | <MRL | <MRL | <MRL | 7.471 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 |  | IS4 | [7 empty] | 10518 | 173696 | 90012 | 765946 | 555196 | 5309754 | 10823491 | 1732972 | 22583744
285 | 2022-07-23 20:32 | AB-220715-01 LRB 20_INJ001 | 284 | K | QC | 0 | 1.062 | 2.254 | 0 | 0 | 0.643 | 0.395 | 1.871 | 0.413 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! |  | IS4 | [13 empty] | 46383 | 495819 | 252520 | 2217661 | 1365573 | 11680942 | 15617144 | 2111009 | 25699741 |  |  | 285
286 | 2022-07-23 21:08 | AB-220715-01 LRB 20 | 285 | K | QC | 0 | 0.784 | 1.802 | 0 | 2.185 | 0.678 | 0.65 | 0.947 | 0.406 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | IS4 |  |  | IS4 |  |  | IS4 | [7 empty] | 19932 | 425025 | 204594 | 1842133 | 1047795 | 9592561 | 12643900 | 1689166 | 19889324 |  |  | 286
287 | 2022-07-23 21:44 | WO-22270-196-A-1 | 286 | U |  | 0 | 0.888 | 1495.148 | 2.292 | 0 | 1.261 | 0.658 | 2.261 | 7.764 | 1 | <MRL | <MRL | 1495.148 | <MRL | <MRL | <MRL | <MRL | <MRL | 7.764 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | [9 empty] | 20721 | 240146 | 122240 | 1139282 | 643972 | 6522938 | 13141993 | 1933863 | 25598630
288 | 2022-07-23 22:20 | WO-22270-191-A-1 | 287 | U |  | 0 | 0.892 | 590.481 | 1.491 | 0 | 1.588 | 1.144 | 2.233 | 7.844 | 1 | <MRL | <MRL | 590.481 | <MRL | <MRL | <MRL | <MRL | <MRL | 7.844 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | [9 empty] | 21024 | 266195 | 126505 | 1247643 | 718456 | 7343352 | 14534261 | 2119930 | 25991201
289 | 2022-07-23 22:56 | WO-22270-197-A-1 | 288 | U |  | 0 | 1.04 | 1403.188 | 2.479 | 0 | 1.949 | 1.063 | 1.726 | 3.32 | 1 | <MRL | <MRL | 1403.188 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 |  | IS4 | [7 empty] | 10025 | 140258 | 71453 | 685938 | 532702 | 5169965 | 10400087 | 1546636 | 20959273
290 | 2022-07-23 23:32 | WO-22270-192-A-1 | 289 | U |  | 0 | 0.761 | 714.487 | 1.811 | 0 | 1.947 | 0.578 | 2.008 | 7.309 | 1 | <MRL | <MRL | 714.487 | <MRL | <MRL | <MRL | <MRL | <MRL | 7.309 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | [9 empty] | 24153 | 303874 | 136209 | 1275267 | 712567 | 6666580 | 13334156 | 2093905 | 27126592
291 | 2022-07-24 00:08 | WO-22270-198-A-1 | 290 | U |  | 0 | 3.573 | 11190.616 | 3333.538 | 0 | 0.887 | 0.556 | 2.524 | 4.209 | 1 | <MRL | <MRL | 11190.616 | 3333.538 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | UL | IS3 | IS3 | IS4 |  | IS4 |  | R2 | [5 empty] | 11645 | 169954 | 67161 | 634144 | 588901 | 5584033 | 11065748 | 1639334 | 22458245
292 | 2022-07-24 00:44 | WO-22270-193-A-1 | 291 | U |  | 0 | 0.826 | 629.943 | 1.486 | 0 | 0.916 | 0.646 | 1.053 | 8.219 | 1 | <MRL | <MRL | 629.943 | <MRL | <MRL | <MRL | <MRL | <MRL | 8.219 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 |  | IS4 | [7 empty] | 16422 | 209770 | 106735 | 956387 | 600249 | 5772170 | 11873065 | 1812329 | 22468485
293 | 2022-07-24 01:20 | WO-22270-199-A-1 | 292 | U |  | 0 | 2.963 | 12321.404 | 3521.747 | 0 | 1.283 | 0.678 | 2.154 | 5.232 | 1 | <MRL | <MRL | 12321.404 | 3521.747 | <MRL | <MRL | <MRL | <MRL | 5.232 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | UL | IS3 | IS3 | IS4 |  | IS4 |  | R2 | [5 empty] | 13829 | 162087 | 64619 | 604404 | 567704 | 5411755 | 10627487 | 1606146 | 21670250
294 | 2022-07-24 01:56 | WO-22209-26-A-1B | 293 | U |  | 0 | 1.409 | 0 | 1.719 | 0 | 27.337 | 0.873 | 0.907 | 29.291 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | 27.337 | <MRL | <MRL | 29.291 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | [7 empty] | 8466 | 85097 | 46123 | 424919 | 385010 | 3955236 | 8137461 | 1107166 | 16095196
295 | 2022-07-24 02:32 | WO-22270-200-A-1 | 294 | U |  | 0 | 2.631 | 11615.642 | 3344.616 | 0 | 1.155 | 1.058 | 1.4 | 3.552 | 1 | <MRL | <MRL | 11615.642 | 3344.616 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | UL | IS3 | IS3 | IS4 | IS4 | IS4 |  | R2 | [5 empty] | 9850 | 151874 | 61754 | 599146 | 541995 | 5338460 | 9989664 | 1516460 | 19825367
296 | 2022-07-24 03:08 | WO-22270-194-A-1 | 295 | U |  | 0 | 1.068 | 1279.045 | 1.773 | 0 | 1.515 | 0.623 | 0.834 | 5.876 | 1 | <MRL | <MRL | 1279.045 | <MRL | <MRL | <MRL | <MRL | <MRL | 5.876 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | [10 empty] | 23998 | 299005 | 139349 | 1414189 | 815770 | 8333244 | 15929215 | 2263229 | 29615948
297 | 2022-07-24 03:44 | WO-22270-195-A-1 | 296 | U |  | 0 | 0.896 | 911.43 | 2.016 | 0 | 1.486 | 0.519 | 2.281 | 4.423 | 1 | <MRL | <MRL | 911.43 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 |  | IS4 | [7 empty] | 12327 | 172405 | 89450 | 813259 | 569531 | 5337636 | 10417731 | 1629132 | 21702220
298 | 2022-07-24 04:20 | AB-220715-01 LFB-H 11 | 297 | Z | QC | 187.116 | 167.433 | 231.801 | 198.179 | 151.857 | 190.277 | 199.559 | 160.42 | 192.084 | 1 | <MRL | 167.433 | 231.801 | 198.179 | 151.857 | 190.277 | 199.559 | 160.42 | 192.084 | H 11 | H 11 | NA | Blank Spike | No TrueV! |  | IS3 |  |  | IS3 |  |  | IS3 | IS3 | [6 empty] | 59029 | 551945 | 268344 | 2370944 | 1211979 | 10268475 | 12889642 | 1658269 | 12123754 |  |  | 298
299 | 2022-07-24 04:56 | AB-220715-01 LRB 21 | 298 | K | QC | 0 | 0.881 | 1.799 | 1.323 | 0 | 0.894 | 0.413 | 2.293 | 0.455 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | IS4 |  |  | IS4 | [10 empty] | 22186 | 459063 | 241550 | 2093452 | 1190488 | 10613237 | 14466924 | 1873834 | 22951723 |  |  | 299
300 | 2022-07-24 05:32 | AB-220715-01 LD 10 | 299 | D | QC | 0 | 1.452 | 5472.988 | 2201.609 | 0 | 1.37 | 0.723 | 2.393 | 2.957 | 1 | <MRL | <MRL | 5472.988 | 2201.609 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Duplicate | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 8432 | 140679 | 64412 | 627072 | 538589 | 5347087 | 11200934 | 1698249 | 21373739 |  |  |  | 300
301 | 2022-07-24 06:08 | AB-220715-01 CCC-H 11 | 300 | Q | QC | 343.684 | 351.242 | 422.132 | 372.456 | 319.013 | 389.571 | 397.692 | 330.388 | 407.738 | 1 | 343.684 | 351.242 | 422.132 | 372.456 | 319.013 | 389.571 | 397.692 | 330.388 | 407.738 | 1 | 01 | NA | CCC | UL! |  | IS3 |  |  | IS3 |  |  | IS3 | [7 empty] | 63416 | 533110 | 286852 | 2516157 | 1057160 | 9109447 | 12641679 | 1602482 | 16211907 |  | 301
302 | 2022-07-24 06:44 | WO-22209-16-A-1B | 301 | U |  | 0 | 2.387 | 37.659 | 9.482 | 0 | 2.417 | 1.101 | 1.251 | 38.635 | 1 | <MRL | <MRL | 37.659 | 9.482 | <MRL | <MRL | <MRL | <MRL | 38.635 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 |  |  |  |  | Out of holding time, electively reanalyzed |  |  | 11790 | 108593 | 78051 | 652479 | 467768 | 5062016 | 9339887 | 1271826 | 18631992
303 | 2022-07-24 07:20 | AB-220715-01 LFSM-H 10 | 302 | S | QC | 0 | 167.184 | 5413.034 | 2486.335 | 149.14 | 188.316 | 197.927 | 147.507 | 187.19 | 1 | <MRL | 167.184 | 5413.034 | 2486.335 | 149.14 | 188.316 | 197.927 | 147.507 | 187.19 | H 10 | NA | H 10 | Sample Spike | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | [7 empty] | 6755 | 97878 | 51976 | 423311 | 425695 | 4128931 | 7908652 | 1217569 | 15765296 |  |  |  |  | 303
304 | 2022-07-24 07:56 | WO-22270-201-A-1 | 303 | U |  | 0 | 2.236 | 11785.487 | 3175.149 | 0 | 1.51 | 1.074 | 1.709 | 3.658 | 1 | <MRL | <MRL | 11785.487 | 3175.149 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | UL | IS3 | IS3 | IS4 | IS4 | IS4 |  | R2 | [5 empty] | 7649 | 115076 | 53468 | 503451 | 479193 | 4412096 | 8828642 | 1387673 | 18531576
305 | 2022-07-24 08:32 | WO-22270-202-A-1 | 304 | U |  | 0 | 2.206 | 12663.908 | 3534.884 | 0 | 0.69 | 0.541 | 1.114 | 2.071 | 1 | <MRL | <MRL | 12663.908 | 3534.884 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | UL | IS3 | IS3 | IS4 |  | IS4 |  | R2 | [5 empty] | 10213 | 110319 | 49345 | 453729 | 528127 | 4885475 | 11059800 | 1603564 | 20517245
306 | 2022-07-24 09:08 | WO-22270-203-A-1 | 305 | U |  | 0 | 2.256 | 12265.672 | 3405.809 | 0 | 0.655 | 0.56 | 1.007 | 3.444 | 1 | <MRL | <MRL | 12265.672 | 3405.809 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | UL | IS3 | IS3 | IS4 |  | IS4 |  | R2 | [5 empty] | 11191 | 159799 | 61876 | 637910 | 587666 | 5152350 | 11351161 | 1567207 | 20235970
307 | 2022-07-24 09:44 | AB-220715-01 LRB 21_INJ001 | 306 | K | QC | 0 | 1.046 | 2.938 | 2.423 | 0 | 1.379 | 0.881 | 2.29 | 0.736 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! |  | IS4 | B | [12 empty] | 37697 | 476769 | 247394 | 2165647 | 1347453 | 11944023 | 16532369 | 2116050 | 25827315 |  |  | 307
308 | 2022-07-24 10:20 | WO-22270-204-A-1 | 307 | U |  | 0 | 2.686 | 11125.024 | 3483.05 | 0 | 3.116 | 0.626 | 1.336 | 1.971 | 1 | <MRL | <MRL | 11125.024 | 3483.05 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 8216 | 72212 | 40094 | 310477 | 369779 | 3656935 | 7737605 | 1206656 | 16437401
309 | 2022-07-24 10:56 | WO-22270-205-A-1 | 308 | U |  | 0 | 2.122 | 11343.019 | 3455.628 | 0 | 0.526 | 0.737 | 1.674 | 2.288 | 1 | <MRL | <MRL | 11343.019 | 3455.628 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 9178 | 135280 | 63613 | 553318 | 524284 | 5052642 | 9676819 | 1438636 | 19113363
310 | 2022-07-24 11:32 | WO-22270-206-A-1 | 309 | U |  | 0 | 2.199 | 11381.093 | 3284.437 | 0 | 1.254 | 0.995 | 3.293 | 4.095 | 1 | <MRL | <MRL | 11381.093 | 3284.437 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 6202 | 89706 | 51225 | 430233 | 405642 | 4363613 | 8297687 | 1161523 | 17110554
311 | 2022-07-24 12:08 | WO-22270-207-A-1 | 310 | U |  | 0 | 2.467 | 10832.086 | 3592.031 | 0 | 2.265 | 1.019 | 1.242 | 5.134 | 1 | <MRL | <MRL | 10832.086 | 3592.031 | <MRL | <MRL | <MRL | <MRL | 5.134 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 7888 | 62211 | 39905 | 288327 | 331950 | 3735302 | 7687885 | 981687 | 14515480
312 | 2022-07-24 12:44 | WO-22270-208-A-1 | 311 | U |  | 0 | 2.764 | 7278.211 | 2847.722 | 0 | 1.277 | 0.571 | 1.031 | 2.486 | 1 | <MRL | <MRL | 7278.211 | 2847.722 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 7325 | 121904 | 63483 | 530208 | 466090 | 4888982 | 9756775 | 1331510 | 19465500
313 | 2022-07-24 13:20 | WO-22270-209-A-1 | 312 | U |  | 0 | 2.036 | 7487.394 | 2860.632 | 0 | 2.189 | 0.812 | 6.395 | 6.272 | 1 | <MRL | <MRL | 7487.394 | 2860.632 | <MRL | <MRL | <MRL | 6.395 | 6.272 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 11136 | 172114 | 77582 | 673820 | 537979 | 5839527 | 11268998 | 1542083 | 22370756
314 | 2022-07-24 13:56 | WO-22270-210-A-1 | 313 | U |  | 0 | 2.025 | 7969.456 | 2869.4 | 0 | 1.151 | 0.526 | 1.962 | 6.031 | 1 | <MRL | <MRL | 7969.456 | 2869.4 | <MRL | <MRL | <MRL | <MRL | 6.031 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 13982 | 164860 | 81550 | 719747 | 571580 | 5975251 | 12077200 | 1600983 | 23175947
315 | 2022-07-24 14:32 | AB-220715-01 LRB 22_INJ001 | 314 | K | QC | 0 | 1.076 | 22.399 | 8.449 | 0 | 0.813 | 0.737 | 1.786 | 0.462 | 1 | <MRL | <MRL | 22.399 | 8.449 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! |  | IS4 |  |  | IS3 | [10 empty] | 44556 | 477532 | 280454 | 2460376 | 1304177 | 12532133 | 17969396 | 2087106 | 27996557 |  |  | 315
316 | 2022-07-24 15:08 | AB-220715-01 LRB 22 | 315 | K | QC | 0 | 0.799 | 6.275 | 3.437 | 0 | 1.067 | 0.701 | 2.425 | 0.496 | 1 | <MRL | <MRL | 6.275 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | IS4 | IS4 |  | IS4 |  |  | IS4 | [7 empty] | 20648 | 343182 | 179572 | 1697926 | 922690 | 9138883 | 12911343 | 1463968 | 19412130 |  |  | 316
317 | 2022-07-24 15:44 | WO-22270-211-A-1 | 316 | U |  | 0 | 2.3 | 7781.393 | 2748.222 | 0 | 0.993 | 0.51 | 1.99 | 3.187 | 1 | <MRL | <MRL | 7781.393 | 2748.222 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 9728 | 155620 | 71482 | 688156 | 546469 | 5715933 | 11448530 | 1597403 | 22430735
318 | 2022-07-24 16:20 | WO-22270-212-A-1 | 317 | U |  | 0 | 3.041 | 7105.418 | 2676.981 | 0 | 1.958 | 0.59 | 5.603 | 4.309 | 1 | <MRL | <MRL | 7105.418 | 2676.981 | <MRL | <MRL | <MRL | 5.603 | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 11841 | 124086 | 66935 | 604138 | 494734 | 5130742 | 10811392 | 1497004 | 21237452
319 | 2022-07-24 16:56 | WO-22270-213-A-1 | 318 | U |  | 0 | 2.359 | 7173.014 | 2778.734 | 0 | 1.614 | 0.677 | 2.514 | 4.681 | 1 | <MRL | <MRL | 7173.014 | 2778.734 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 14404 | 156739 | 85676 | 718723 | 566561 | 5825835 | 12507360 | 1722576 | 22977323
320 | 2022-07-24 17:32 | WO-22270-214-A-1 | 319 | U |  | 0 | 2.749 | 12554.794 | 3346.883 | 0 | 1.669 | 1.638 | 1.131 | 14.656 | 1 | <MRL | <MRL | 12554.794 | 3346.883 | <MRL | <MRL | <MRL | <MRL | 14.656 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | [10 empty] | 21435 | 225661 | 89100 | 926666 | 682200 | 7941049 | 15341442 | 2028914 | 26389497
321 | 2022-07-24 18:08 | WO-22270-215-A-1 | 320 | U |  | 0 | 4.025 | 12151.767 | 3235.902 | 0 | 1.903 | 0.805 | 0.993 | 2.926 | 1 | <MRL | <MRL | 12151.767 | 3235.902 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 7520 | 99597 | 56715 | 530274 | 480468 | 5368983 | 10532768 | 1424552 | 19718899
322 | 2022-07-24 18:44 | WO-22270-216-A-1 | 321 | U |  | 0 | 3.068 | 7184.12 | 3348.427 | 0 | 2.498 | 1.048 | 0.951 | 8.092 | 1 | <MRL | <MRL | 7184.12 | 3348.427 | <MRL | <MRL | <MRL | <MRL | 8.092 | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 11818 | 145052 | 85626 | 695961 | 545555 | 5124682 | 11166318 | 1591366 | 21788958
323 | 2022-07-24 19:20 | WO-22270-217-A-1 | 322 | U |  | 0 | 2.209 | 7307.135 | 3450.62 | 0 | 2.127 | 0.919 | 2.409 | 4.849 | 1 | <MRL | <MRL | 7307.135 | 3450.62 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 9822 | 128228 | 77522 | 628020 | 515257 | 5442785 | 10309220 | 1379006 | 20589884
324 | 2022-07-24 19:56 | WO-22270-218-A-1 | 323 | U |  | 0 | 1.585 | 2491.878 | 1141.632 | 0 | 0.719 | 0.649 | 1.886 | 2.803 | 1 | <MRL | <MRL | 2491.878 | 1141.632 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 12976 | 167267 | 103059 | 832272 | 596399 | 6062329 | 10781459 | 1592824 | 22102196
325 | 2022-07-24 20:32 | WO-22270-219-A-1 | 324 | U |  | 0 | 1.36 | 2772.695 | 1133.005 | 0 | 1.084 | 0.554 | 1.051 | 1.586 | 1 | <MRL | <MRL | 2772.695 | 1133.005 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 14625 | 200937 | 109178 | 985580 | 635989 | 6853538 | 12721918 | 1608310 | 25175785
326 | 2022-07-24 21:08 | WO-22270-220-A-1 | 325 | U |  | 0 | 1.044 | 2530.797 | 1166.087 | 0 | 0.644 | 0.508 | 1.058 | 2.469 | 1 | <MRL | <MRL | 2530.797 | 1166.087 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 11184 | 159008 | 90086 | 794512 | 548267 | 5814175 | 11275577 | 1490656 | 21250524
327 | 2022-07-24 21:44 | AB-220715-01 LFB-H 12 | 326 | Z | QC | 144.503 | 165.072 | 207.993 | 182.251 | 145.325 | 186.411 | 192.991 | 152.715 | 184.98 | 1 | <MRL | 165.072 | 207.993 | 182.251 | 145.325 | 186.411 | 192.991 | 152.715 | 184.98 | H 12 | H 12 | NA | Blank Spike | No TrueV! |  | IS3 |  |  | IS3 |  |  | IS3 | [7 empty] | 50643 | 463823 | 262976 | 2498642 | 1177803 | 11098232 | 15111322 | 1741780 | 21976360 |  |  | 327
328 | 2022-07-24 22:20 | AB-220715-01 LRB 23_INJ001 | 327 | K | QC | 0 | 1.162 | 1.639 | 2.016 | 1.448 | 1.26 | 1.078 | 1.597 | 1.255 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! |  | IS4 | b |  | IS4 | [10 empty] | 49716 | 476218 | 267466 | 2491489 | 1315732 | 12901297 | 17624904 | 2037344 | 27014804 |  |  | 328
329 | 2022-07-24 22:56 | AB-220715-01 LRB 23 | 328 | K | QC | 0 | 0.726 | 6.134 | 2.845 | 0 | 0.779 | 0.425 | 2.333 | 0.472 | 1 | <MRL | <MRL | 6.134 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | IS4 | b |  | IS4 |  |  | IS4 | [7 empty] | 16664 | 453195 | 236105 | 2186411 | 1061182 | 10513831 | 14217666 | 1617515 | 20842471 |  |  | 329
330 | 2022-07-24 23:32 | AB-220715-01 LD 11 | 329 | D | QC | 0 | 1.08 | 2421.487 | 1091.661 | 4.165 | 1.364 | 0.555 | 0.907 | 1.575 | 1 | <MRL | <MRL | 2421.487 | 1091.661 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Duplicate | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 13637 | 192560 | 108695 | 968559 | 610756 | 6647475 | 12323242 | 1667184 | 23851746 |  |  |  | 330
331 | 2022-07-25 00:08 | AB-220715-01 CCC-H 12 | 330 | Q | QC | 337.382 | 330.585 | 418.533 | 384.909 | 310.509 | 386.724 | 390.787 | 319.683 | 381.938 | 1 | 337.382 | 330.585 | 418.533 | 384.909 | 310.509 | 386.724 | 390.787 | 319.683 | 381.938 | 1 | 01 | NA | CCC | UL! |  | IS3 |  |  | IS3 |  |  | IS3 | [7 empty] | 52383 | 464159 | 267099 | 2381097 | 1021654 | 9563423 | 13445640 | 1520143 | 18249143 |  | 331
332 | 2022-07-25 00:44 | AB-220715-01 LFSM-H 11 | 331 | S | QC | 0 | 165.797 | 2977.339 | 1328.5 | 144.818 | 194.494 | 194.899 | 151.821 | 194.326 | 1 | <MRL | 165.797 | 2977.339 | 1328.5 | 144.818 | 194.494 | 194.899 | 151.821 | 194.326 | H 11 | NA | H 11 | Sample Spike | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | [7 empty] | 10404 | 155733 | 90253 | 794392 | 533729 | 5431109 | 10005229 | 1303000 | 18420698 |  |  |  |  | 332
333 | 2022-07-25 01:20 | WO-22270-221-A-1 | 332 | U |  | 0 | 1.511 | 2681.881 | 1085.099 | 0 | 0.693 | 0.525 | 1.66 | 1.787 | 1 | <MRL | <MRL | 2681.881 | 1085.099 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 7905 | 197478 | 108003 | 945716 | 643498 | 6540357 | 11936740 | 1709687 | 21968010
334 | 2022-07-25 01:56 | WO-22270-222-A-1 | 333 | U |  | 0 | 1.5 | 2794.788 | 1067.661 | 0 | 2.007 | 0.698 | 1.203 | 1.519 | 1 | <MRL | <MRL | 2794.788 | 1067.661 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 10046 | 153941 | 81361 | 802223 | 547460 | 5659326 | 10703960 | 1394304 | 19031010
335 | 2022-07-25 02:32 | WO-22270-223-A-1 | 334 | U |  | 0 | 0.92 | 2436.532 | 1051.989 | 0 | 1.524 | 0.721 | 2.83 | 1.186 | 1 | <MRL | <MRL | 2436.532 | 1051.989 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 14152 | 195432 | 105189 | 956808 | 601342 | 6492293 | 11896518 | 1625551 | 21877461
336 | 2022-07-25 03:08 | WO-22270-224-A-1 | 335 | U |  | 0 | 1.157 | 2567.789 | 1132.739 | 0 | 1.089 | 0.741 | 0.94 | 2.39 | 1 | <MRL | <MRL | 2567.789 | 1132.739 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 14602 | 197599 | 106397 | 912321 | 617437 | 6240995 | 11117897 | 1546054 | 20876279
337 | 2022-07-25 03:44 | WO-22270-225-A-1 | 336 | U |  | 0 | 0.964 | 2686.931 | 1066.892 | 0 | 1.694 | 0.6 | 0.968 | 2.002 | 1 | <MRL | <MRL | 2686.931 | 1066.892 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 10087 | 178774 | 97477 | 939533 | 581314 | 6278527 | 11814640 | 1586849 | 22968960
338 | 2022-07-25 04:20 | WO-22270-226-A-1 | 337 | U |  | 0 | 1.551 | 2806.729 | 1116.773 | 0 | 1.358 | 0.686 | 2.501 | 1.745 | 1 | <MRL | <MRL | 2806.729 | 1116.773 | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Unknown | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4 | [7 empty] | 15159 | 206238 | 105271 | 982839 | 607152 | 6307692 | 11399572 | 1582656 | 21528776
339 | 2022-07-25 04:56 | AB-220715-01 LRB 24_INJ001 | 338 | K | QC | 0 | 0.766 | 8.854 | 4.418 | 2.309 | 0.64 | 0.348 | 0.875 | 0.375 | 1 | <MRL | <MRL | 8.854 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! |  | IS4 | b | b | [11 empty] | 52145 | 491306 | 262966 | 2594506 | 1362996 | 13271086 | 18250216 | 2089775 | 27937892 |  |  | 339
340 | 2022-07-25 05:32 | AB-220715-01 LRB 24 | 339 | K | QC | 0 | 0.849 | 4.17 | 2.162 | 2.368 | 0.668 | 0.373 | 2.338 | 0.408 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | NA | NA | NA | Blank | UL! | IS3 | IS4 | b | IS4 | IS4 |  |  | IS4 | [7 empty] | 12666 | 281158 | 159622 | 1527220 | 993605 | 9800208 | 13529513 | 1581659 | 20882693 |  |  | 340
341 | 2022-07-25 06:08 | AB-220715-01 CCC-H 13 | 340 | Q | QC | 332.365 | 341.506 | 376.546 | 374.334 | 303.614 | 388.839 | 386.732 | 323.323 | 387.108 | 1 | 332.365 | 341.506 | 376.546 | 374.334 | 303.614 | 388.839 | 386.732 | 323.323 | 387.108 | 1 | 01 | NA | CCC | UL! |  | IS3 |  |  | IS3 |  |  | IS3 | [7 empty] | 49744 | 469969 | 288393 | 2465554 | 1024841 | 10213630 | 13836837 | 1505355 | 18796639 |  | 341
342 | 2022-07-25 06:44 | AB-220715-01 QCS 2 | 341 | Q | QC | 161.011 | 171.1 | 199.036 | 193.156 | 158.561 | 196.615 | 209.797 | 162.451 | 191.063 | 1 | <MRL | 171.1 | 199.036 | 193.156 | 158.561 | 196.615 | 209.797 | 162.451 | 191.063 | 2 | NA | 2 | QCS | UL! | [7 empty] | IS3 | [7 empty] | 62716 | 604229 | 360443 | 3117107 | 1333423 | 12800782 | 16710888 | 1829471 | 25271347 |  | 342
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA

## Processing Narrative
1. Creeping background of PFBA in the mobile phases in this batch resulted in extremely high background signal swamping any unknown reports < 100-200 ng/L. I have raised the MRL as a result to 200 ng/L for PFBA.

## Report Template
Type
U
|  |  |  | Concentration (Raw) | [9 empty] | Concentration | [9 empty] | Flags | [15 empty] | Raw Concentrations, ng/L | [9 empty] | Concentration, ng/L, Corrected for Sample Dilution | [9 empty] | Flags
# | Time Stamp: | Sample ID: | Type | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | Dilution Factor: | C-PFBA | C-PFBS | C-HFPO-DA | C-PFHxA | C-PFHxS | C-PFOA | C-PFNA | C-PFOS | C-PFDA | Flags | F-PFBA | F-PFBS | F-HFPO-DA | F-PFHxA | F-PFHxS | F-PFOA | F-PFNA | F-PFOS | F-PFDA | Re-Analysis Flag | Sample Comments | IS Comments | # | Time Stamp: | Sample ID: | Type | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | Dilution Factor: | C-PFBA | C-PFBS | C-HFPO-DA | C-PFHxA | C-PFHxS | C-PFOA | C-PFNA | C-PFOS | C-PFDA | Flags | F-PFBA | F-PFBS | F-HFPO-DA | F-PFHxA | F-PFHxS | F-PFOA | F-PFNA | F-PFOS | F-PFDA | Re-Analysis Flag | Sample Comments | IS Comments | Final Notes
1 | 2022-07-16 18:08 | AB-220715-01 Cal Std 1 1 | M | 0 | 1.235 | 0 | 0 | 1.65 | 1.244 | 1.19 | 1.566 | 1.209 | 1 | 0 | 1.235 | 0 | 0 | 1.65 | 1.244 | 1.19 | 1.566 | 1.209 | No TrueV! | [12 empty] | 20 | 2022-07-17 05:32 | 211827-d2 | U | 0 | 0 | 0 | 0 | 0 | 2.455 | 1.284 | 1.107 | 4.515 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
2 | 2022-07-16 18:44 | AB-220715-01 Cal Std 2 1 | M | 0 | 1.809 | 2.756 | 2.915 | 2.259 | 2.18 | 2.12 | 2.219 | 2.178 | 1 | 0 | 1.809 | 2.756 | 2.915 | 2.259 | 2.18 | 2.12 | 2.219 | 2.178 | No TrueV! | [12 empty] | 21 | 2022-07-17 06:08 | 211828-d2 | U | 0 | 0 | 0 | 0 | 0 | 2.629 | 0.801 | 1.166 | 4.114 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
3 | 2022-07-16 19:20 | AB-220715-01 Cal Std 3 1 | M | 0 | 3.71 | 4.106 | 4.587 | 3.685 | 3.99 | 3.898 | 3.587 | 4.036 | 1 | 0 | 3.71 | 4.106 | 4.587 | 3.685 | 3.99 | 3.898 | 3.587 | 4.036 | No TrueV! | [12 empty] | 22 | 2022-07-17 06:44 | 211829-d2 | U | 0 | 0 | 0 | 0 | 0 | 1.389 | 0.734 | 0.982 | 3.879 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
4 | 2022-07-16 19:56 | AB-220715-01 Cal Std 4 1 | M | 0 | 8.981 | 7.465 | 9.617 | 7.953 | 9.57 | 9.647 | 8.162 | 9.592 | 1 | 0 | 8.981 | 7.465 | 9.617 | 7.953 | 9.57 | 9.647 | 8.162 | 9.592 | No TrueV! | [12 empty] | 23 | 2022-07-17 07:20 | 211830-d2 | U | 0 | 0 | 0 | 0 | 0 | 1.569 | 0.854 | 1.082 | 3.959 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
5 | 2022-07-16 20:32 | AB-220715-01 Cal Std 5 1 | M | 0 | 15.471 | 20.512 | 17.849 | 14.817 | 18.982 | 19.29 | 15.556 | 19.349 | 1 | 0 | 15.471 | 20.512 | 17.849 | 14.817 | 18.982 | 19.29 | 15.556 | 19.349 | No TrueV! | [12 empty] | 24 | 2022-07-17 07:56 | 211831-d2 | U | 0 | 0 | 0 | 0 | 0 | 3.116 | 0.677 | 1.161 | 4.702 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
6 | 2022-07-16 21:08 | AB-220715-01 Cal Std 6 1 | M | 0 | 32.123 | 39.383 | 35.309 | 29.069 | 38.825 | 38.945 | 31.304 | 38.803 | 1 | 0 | 32.123 | 39.383 | 35.309 | 29.069 | 38.825 | 38.945 | 31.304 | 38.803 | No TrueV! | [12 empty] | 25 | 2022-07-17 08:32 | 211832-d2 | U | 0 | 0 | 3.727 | 0 | 0 | 3.167 | 1.073 | 1.226 | 3.878 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
7 | 2022-07-16 21:44 | AB-220715-01 Cal Std 7 1 | M | 0 | 78.091 | 88.145 | 90.964 | 72.445 | 96.896 | 96.207 | 76.777 | 96.261 | 1 | 0 | 78.091 | 88.145 | 90.964 | 72.445 | 96.896 | 96.207 | 76.777 | 96.261 | No TrueV! | [12 empty] | 26 | 2022-07-17 09:08 | 211833-d2 | U | 0 | 0 | 0 | 0 | 0 | 2.314 | 0.509 | 0.834 | 3.862 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
8 | 2022-07-16 22:20 | AB-220715-01 Cal Std 8 1 | M | 192.335 | 165.757 | 187.477 | 189.927 | 158.86 | 195 | 197.798 | 164.702 | 195.086 | 1 | 192.335 | 165.757 | 187.477 | 189.927 | 158.86 | 195 | 197.798 | 164.702 | 195.086 | No TrueV! | [12 empty] | 27 | 2022-07-17 09:44 | 211834-d2 | U | 0 | 0 | 0 | 0 | 0 | 2.862 | 0.848 | 0.938 | 4.404 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
9 | 2022-07-16 22:56 | AB-220715-01 Cal Std 9 1 | M | 410.126 | 345.579 | 413.164 | 381.334 | 319.1 | 385.596 | 403.989 | 341.327 | 385.574 | 1 | 410.126 | 345.579 | 413.164 | 381.334 | 319.1 | 385.596 | 403.989 | 341.327 | 385.574 | No TrueV! | [12 empty] | 28 | 2022-07-17 10:20 | 211835-d2 | U | 0 | 0 | 0 | 0 | 0 | 1.798 | 1.009 | 2.294 | 5.182 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 10.364
10 | 2022-07-16 23:32 | AB-220715-01 Cal Std 10 1 | M | 930.971 | 830.668 | 983.792 | 867.467 | 810.873 | 924.669 | 919.292 | 841.7 | 924.724 | 1 | 930.971 | 830.668 | 983.792 | 867.467 | 810.873 | 924.669 | 919.292 | 841.7 | 924.724 | No TrueV! | [12 empty] | 29 | 2022-07-17 10:56 | 211836-d2 | U | 0 | 0 | 0 | 0 | 0 | 2.843 | 0.804 | 0.933 | 6.309 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 12.618
11 | 2022-07-17 00:08 | AB-220715-01 Cal Std 11 1 | M | 2075.212 | 1747.461 | 1959.323 | 1844.46 | 1770.213 | 1893.058 | 1936.193 | 1829.534 | 1916.836 | 1 | 2075.212 | 1747.461 | 1959.323 | 1844.46 | 1770.213 | 1893.058 | 1936.193 | 1829.534 | 1916.836 | No TrueV! |  |  |  |  | IS3 | IS3 | IS3 | IS3 | IS3 |  |  |  | 32 | 2022-07-17 12:44 | 211837-d2 | U | 0 | 0 | 0 | 0 | 0 | 1.094 | 0.482 | 2.047 | 4.526 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
12 | 2022-07-17 00:44 | AB-220715-01 Cal Std 12 1 | M | 4301.969 | 3719.585 | 4112.801 | 4155.832 | 4004.568 | 3945.156 | 4063.315 | 3973.032 | 3952.891 | 1 | 4301.969 | 3719.585 | 4112.801 | 4155.832 | 4004.568 | 3945.156 | 4063.315 | 3973.032 | 3952.891 | No TrueV! |  |  |  | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  |  |  | 33 | 2022-07-17 13:20 | 211838-d2 | U | 0 | 0 | 0 | 0 | 0 | 2.504 | 0.455 | 0.921 | 3.046 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
13 | 2022-07-17 01:20 | AB-220715-01 Cal Std 13 1 | M | 9689.387 | 8816.874 | 9957.078 | 10175.739 | 9042.279 | 10261.834 | 10085.115 | 9202.601 | 10230.461 | 1 | 9689.387 | 8816.874 | 9957.078 | 10175.739 | 9042.279 | 10261.834 | 10085.115 | 9202.601 | 10230.461 | No TrueV! |  | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  |  |  | 34 | 2022-07-17 13:56 | 211847-d2 | U | 0 | 0 | 1.644 | 0 | 0.942 | 0 | 0.516 | 0.867 | 4.196 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
14 | 2022-07-17 01:56 | AB-220715-01 LFB-H 1 | Z | 184.824 | 167.941 | 199.413 | 198.905 | 154.586 | 195.518 | 197.021 | 168.032 | 196.5 | 1 | <MRL | 167.941 | 199.413 | 198.905 | 154.586 | 195.518 | 197.021 | 168.032 | 196.5 | No TrueV! | [12 empty] | 35 | 2022-07-17 14:32 | 211848-d2 | U | 0 | 0 | 0 | 0 | 0 | 0 | 0.372 | 0.847 | 4.113 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
15 | 2022-07-17 02:32 | AB-220715-01 LRB 1_INJ001 | K | 0 | 0 | 0 | 1.749 | 0 | 1.349 | 0.917 | 1.258 | 0.984 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 36 | 2022-07-17 15:08 | 211849-d2 | U | 0 | 0 | 0 | 0 | 0 | 2.884 | 0.761 | 1.973 | 2.456 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
16 | 2022-07-17 03:08 | AB-220715-01 LRB 1 | K | 0 | 0 | 0 | 3.07 | 0 | 1.317 | 0.967 | 1.494 | 1.116 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | [11 empty] | 37 | 2022-07-17 15:44 | 211850-d2 | U | 0 | 0 | 0 | 0 | 0 | 2.347 | 0.681 | 0.965 | 4.631 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
17 | 2022-07-17 03:44 | AB-220715-01 CCC-H 1 | Q | 402.94 | 352.033 | 402.071 | 385.132 | 325.628 | 391.959 | 401.704 | 344.494 | 398.521 | 1 | 402.94 | 352.033 | 402.071 | 385.132 | 325.628 | 391.959 | 401.704 | 344.494 | 398.521 | UL! | [12 empty] | 38 | 2022-07-17 16:20 | 211851-d2 | U | 0 | 0 | 0 | 0 | 0 | 3.796 | 0.916 | 0.832 | 4.525 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
18 | 2022-07-17 04:20 | AB-220715-01 IBC 1 | Q | 191.996 | 173.785 | 197.324 | 192.575 | 163.267 | 196.228 | 202.398 | 171.927 | 195.534 | 1 | <MRL | 173.785 | 197.324 | 192.575 | 163.267 | 196.228 | 202.398 | 171.927 | 195.534 | UL! | [12 empty] | 39 | 2022-07-17 16:56 | 211852-d2 | U | 0 | 0 | 0 | 0 | 1.431 | 1.384 | 0.394 | 0.83 | 1.851 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
19 | 2022-07-17 04:56 | AB-220715-01 QCS 1 | Q | 180.3 | 176.792 | 153.347 | 187.09 | 161.888 | 193.982 | 198.336 | 172.792 | 191.245 | 1 | <MRL | 176.792 | 153.347 | 187.09 | 161.888 | 193.982 | 198.336 | 172.792 | 191.245 | UL! | [12 empty] | 40 | 2022-07-17 17:32 | 211853-d2 | U | 0 | 0 | 1.644 | 0 | 0 | 1.812 | 0.482 | 0.866 | 3.432 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
20 | 2022-07-17 05:32 | WO-22270-01-A-1 | U | 0 | 0 | 0 | 0 | 0 | 2.455 | 1.284 | 1.107 | 4.515 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 41 | 2022-07-17 18:08 | 211854-d2 | U | 0 | 0 | 2.717 | 0 | 0 | 2.423 | 0.447 | 0.99 | 5.105 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 10.21
21 | 2022-07-17 06:08 | WO-22270-02-A-1 | U | 0 | 0 | 0 | 0 | 0 | 2.629 | 0.801 | 1.166 | 4.114 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 48 | 2022-07-17 22:20 | 211855-d2 | U | 0 | 0 | 2.659 | 0 | 0 | 4.374 | 0.654 | 1.043 | 3.786 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
22 | 2022-07-17 06:44 | WO-22270-03-A-1 | U | 0 | 0 | 0 | 0 | 0 | 1.389 | 0.734 | 0.982 | 3.879 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 49 | 2022-07-17 22:56 | 211856-d2 | U | 0 | 0 | 0 | 0 | 0 | 1.838 | 0.608 | 0.808 | 3.208 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
23 | 2022-07-17 07:20 | WO-22270-04-A-1 | U | 0 | 0 | 0 | 0 | 0 | 1.569 | 0.854 | 1.082 | 3.959 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 50 | 2022-07-17 23:32 | 211857-d2 | U | 0 | 0 | 0 | 2.005 | 0 | 3.327 | 0.641 | 0.896 | 4.061 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
24 | 2022-07-17 07:56 | WO-22270-05-A-1 | U | 0 | 0 | 0 | 0 | 0 | 3.116 | 0.677 | 1.161 | 4.702 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 51 | 2022-07-18 00:08 | 211858-d2 | U | 0 | 0 | 0 | 0 | 0 | 2.263 | 0.582 | 1.016 | 3.347 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
25 | 2022-07-17 08:32 | WO-22270-06-A-1 | U | 0 | 0 | 3.727 | 0 | 0 | 3.167 | 1.073 | 1.226 | 3.878 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 52 | 2022-07-18 00:44 | 211859-d2 | U | 0 | 0 | 0 | 0 | 0 | 0 | 0.552 | 1.775 | 4.386 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
26 | 2022-07-17 09:08 | WO-22270-07-A-1 | U | 0 | 0 | 0 | 0 | 0 | 2.314 | 0.509 | 0.834 | 3.862 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 53 | 2022-07-18 01:20 | 211860-d2 | U | 0 | 0 | 0 | 0 | 0 | 2.693 | 0.619 | 2.491 | 3.774 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
27 | 2022-07-17 09:44 | WO-22270-08-A-1 | U | 0 | 0 | 0 | 0 | 0 | 2.862 | 0.848 | 0.938 | 4.404 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 54 | 2022-07-18 01:56 | 211861-d2 | U | 0 | 0 | 0 | 0 | 0 | 1.492 | 0.47 | 0.946 | 3.312 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
28 | 2022-07-17 10:20 | WO-22270-09-A-1 | U | 0 | 0 | 0 | 0 | 0 | 1.798 | 1.009 | 2.294 | 5.182 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 5.182 | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 55 | 2022-07-18 02:32 | 211862-d2 | U | 0 | 0 | 0 | 0 | 0 | 0 | 0.471 | 0.892 | 4.549 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
29 | 2022-07-17 10:56 | WO-22270-10-A-1 | U | 0 | 0 | 0 | 0 | 0 | 2.843 | 0.804 | 0.933 | 6.309 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 6.309 | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 56 | 2022-07-18 03:08 | 211863-d2 | U | 0 | 0 | 0 | 0 | 0 | 0.799 | 0.539 | 0.826 | 4.642 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
30 | 2022-07-17 11:32 | AB-220715-01 LRB 2_INJ001 | K | 0 | 0 | 0 | 0 | 1.014 | 1.595 | 0.708 | 0.918 | 0.443 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 57 | 2022-07-18 03:44 | 211864-d2 | U | 0 | 0 | 2.817 | 1.854 | 0 | 2.664 | 0.836 | 1.847 | 4.086 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
31 | 2022-07-17 12:08 | AB-220715-01 LRB 2 | K | 0 | 0 | 0 | 0 | 0 | 1.568 | 0.802 | 1.055 | 0.367 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 60 | 2022-07-18 05:32 | 211865-d2 | U | 0 | 0 | 0.624 | 0 | 0 | 1.614 | 1.049 | 0.864 | 2.993 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
32 | 2022-07-17 12:44 | WO-22270-11-A-1 | U | 0 | 0 | 0 | 0 | 0 | 1.094 | 0.482 | 2.047 | 4.526 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 61 | 2022-07-18 06:08 | 211866-d2 | U | 0 | 0 | 0 | 0 | 0 | 1.812 | 0.671 | 2.278 | 3.971 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
33 | 2022-07-17 13:20 | WO-22270-12-A-1 | U | 0 | 0 | 0 | 0 | 0 | 2.504 | 0.455 | 0.921 | 3.046 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 62 | 2022-07-18 06:44 | 211867-d2 | U | 0 | 0 | 0.752 | 0 | 0 | 2.319 | 0.42 | 0.899 | 3.255 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
34 | 2022-07-17 13:56 | WO-22270-13-A-1 | U | 0 | 0 | 1.644 | 0 | 0.942 | 0 | 0.516 | 0.867 | 4.196 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 63 | 2022-07-18 07:20 | 211868-d2 | U | 0 | 0 | 0 | 0 | 0 | 3.094 | 0.998 | 1.188 | 4.461 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
35 | 2022-07-17 14:32 | WO-22270-14-A-1 | U | 0 | 0 | 0 | 0 | 0 | 0 | 0.372 | 0.847 | 4.113 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 64 | 2022-07-18 07:56 | 211869-d2 | U | 0 | 0 | 0 | 0 | 0 | 0.962 | 1.591 | 1.696 | 2.965 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
36 | 2022-07-17 15:08 | WO-22270-15-A-1 | U | 0 | 0 | 0 | 0 | 0 | 2.884 | 0.761 | 1.973 | 2.456 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 65 | 2022-07-18 08:32 | 211870-d2 | U | 0 | 0 | 0.078 | 0 | 0 | 1.501 | 0.517 | 1.719 | 3.607 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
37 | 2022-07-17 15:44 | WO-22270-16-A-1 | U | 0 | 0 | 0 | 0 | 0 | 2.347 | 0.681 | 0.965 | 4.631 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 66 | 2022-07-18 09:08 | 211871-d2 | U | 0 | 0.915 | 0 | 0 | 0 | 1.387 | 0.447 | 1.024 | 3.229 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
38 | 2022-07-17 16:20 | WO-22270-17-A-1 | U | 0 | 0 | 0 | 0 | 0 | 3.796 | 0.916 | 0.832 | 4.525 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 67 | 2022-07-18 09:44 | 211872-d2 | U | 0 | 1.299 | 0 | 0 | 0 | 1.752 | 0.48 | 0.958 | 2.328 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
39 | 2022-07-17 16:56 | WO-22270-18-A-1 | U | 0 | 0 | 0 | 0 | 1.431 | 1.384 | 0.394 | 0.83 | 1.851 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 68 | 2022-07-18 10:20 | 211873-d2 | U | 0 | 0 | 0 | 0 | 0 | 1.962 | 0.429 | 2.016 | 2.578 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
40 | 2022-07-17 17:32 | WO-22270-19-A-1 | U | 0 | 0 | 1.644 | 0 | 0 | 1.812 | 0.482 | 0.866 | 3.432 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 69 | 2022-07-18 10:56 | 211874-d2 | U | 0 | 0 | 1.034 | 0 | 0 | 0.761 | 1.428 | 0.952 | 3.462 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
41 | 2022-07-17 18:08 | WO-22270-20-A-1 | U | 0 | 0 | 2.717 | 0 | 0 | 2.423 | 0.447 | 0.99 | 5.105 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 5.105 | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 76 | 2022-07-18 15:08 | 211875-d2 | U | 0 | 0 | 0 | 0 | 0 | 0.959 | 0.66 | 0.97 | 2.578 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
42 | 2022-07-17 18:44 | AB-220715-01 LFB-H 2 | Z | 159.874 | 161.04 | 194.921 | 187.047 | 152.172 | 190.583 | 192.039 | 160.442 | 193.606 | 1 | <MRL | 161.04 | 194.921 | 187.047 | 152.172 | 190.583 | 192.039 | 160.442 | 193.606 | No TrueV! | [12 empty] | 77 | 2022-07-18 15:44 | 211876-d2 | U | 0 | 1.077 | 2.162 | 0 | 0 | 1.255 | 0.428 | 0.841 | 1.633 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
43 | 2022-07-17 19:20 | AB-220715-01 LRB 3_INJ001 | K | 0 | 0 | 0 | 1.04 | 1.095 | 1.137 | 0.487 | 1.016 | 0.521 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 78 | 2022-07-18 16:20 | 211877-d2 | U | 0 | 0 | 0 | 0 | 0 | 4.642 | 0.663 | 0.917 | 2.633 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
44 | 2022-07-17 19:56 | AB-220715-01 LRB 3 | K | 0 | 0 | 0 | 0.977 | 0.969 | 1.107 | 0.453 | 0.842 | 0.4 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 79 | 2022-07-18 16:56 | 211878-d2 | U | 0 | 0 | 1.368 | 0 | 0 | 1.306 | 0.442 | 0.928 | 1.905 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
45 | 2022-07-17 20:32 | AB-220715-01 LD 1 | D | 0 | 0 | 2.029 | 0 | 1.804 | 2.867 | 0.621 | 0.834 | 3.428 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 80 | 2022-07-18 17:32 | 211879-d2 | U | 0 | 1.289 | 3.345 | 0 | 0 | 0 | 1.441 | 0.858 | 1.785 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
46 | 2022-07-17 21:08 | AB-220715-01 CCC-H 2 | Q | 364.718 | 347.787 | 399.928 | 380.048 | 320.885 | 378.815 | 387.617 | 336.485 | 386.616 | 1 | 364.718 | 347.787 | 399.928 | 380.048 | 320.885 | 378.815 | 387.617 | 336.485 | 386.616 | UL! | [12 empty] | 81 | 2022-07-18 18:08 | 211886-d2 | U | 0 | 1.25 | 3.526 | 0 | 0 | 1.301 | 0.646 | 2.063 | 2.789 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
47 | 2022-07-17 21:44 | AB-220715-01 LFSM-H 1 | S | 0 | 164.433 | 188.888 | 194.271 | 151.916 | 194.121 | 194.736 | 149.476 | 196.881 | 1 | <MRL | 164.433 | 188.888 | 194.271 | 151.916 | 194.121 | 194.736 | 149.476 | 196.881 | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | [6 empty] | 82 | 2022-07-18 18:44 | 211880-d2 | U | 0 | 1.039 | 2.103 | 0 | 0 | 1.241 | 0.561 | 1.511 | 3.148 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
48 | 2022-07-17 22:20 | WO-22270-21-A-1 | U | 0 | 0 | 2.659 | 0 | 0 | 4.374 | 0.654 | 1.043 | 3.786 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 83 | 2022-07-18 19:20 | 211881-d2 | U | 0 | 0 | 0.861 | 0 | 0 | 1.16 | 0.441 | 0.922 | 1.894 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
49 | 2022-07-17 22:56 | WO-22270-22-A-1 | U | 0 | 0 | 0 | 0 | 0 | 1.838 | 0.608 | 0.808 | 3.208 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 84 | 2022-07-18 19:56 | 211887-d2 | U | 0 | 1.507 | 1.472 | 0 | 0 | 1.634 | 0.685 | 0.826 | 3.553 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
50 | 2022-07-17 23:32 | WO-22270-23-A-1 | U | 0 | 0 | 0 | 2.005 | 0 | 3.327 | 0.641 | 0.896 | 4.061 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 85 | 2022-07-18 20:32 | 211882-d2 | U | 0 | 0 | 1.281 | 0 | 0 | 0.488 | 0.378 | 1.93 | 3.977 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
51 | 2022-07-18 00:08 | WO-22270-24-A-1 | U | 0 | 0 | 0 | 0 | 0 | 2.263 | 0.582 | 1.016 | 3.347 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 86 | 2022-07-18 21:08 | 211888-d2 | U | 0 | 0 | 3.071 | 1.304 | 0 | 0.828 | 0.375 | 1.31 | 3.216 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
52 | 2022-07-18 00:44 | WO-22270-25-A-1 | U | 0 | 0 | 0 | 0 | 0 | 0 | 0.552 | 1.775 | 4.386 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 87 | 2022-07-18 21:44 | 211883-d2 | U | 0 | 0 | 0 | 0 | 0 | 0.776 | 0.497 | 0.988 | 4.442 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
53 | 2022-07-18 01:20 | WO-22270-26-A-1 | U | 0 | 0 | 0 | 0 | 0 | 2.693 | 0.619 | 2.491 | 3.774 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 88 | 2022-07-18 22:20 | 211889-d2 | U | 0 | 1.289 | 2.916 | 0 | 0 | 0 | 3.128 | 3.016 | 2.961 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
54 | 2022-07-18 01:56 | WO-22270-27-A-1 | U | 0 | 0 | 0 | 0 | 0 | 1.492 | 0.47 | 0.946 | 3.312 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 89 | 2022-07-18 22:56 | 211884-d2 | U | 0 | 0.742 | 0 | 0 | 0 | 0.833 | 1.212 | 0.949 | 4.275 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
55 | 2022-07-18 02:32 | WO-22270-28-A-1 | U | 0 | 0 | 0 | 0 | 0 | 0 | 0.471 | 0.892 | 4.549 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 90 | 2022-07-18 23:32 | 211890-d2 | U | 0 | 1.308 | 0 | 0 | 0 | 1.763 | 0.793 | 0.891 | 4.414 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
56 | 2022-07-18 03:08 | WO-22270-29-A-1 | U | 0 | 0 | 0 | 0 | 0 | 0.799 | 0.539 | 0.826 | 4.642 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 91 | 2022-07-19 00:08 | 211891-d2 | U | 0 | 1.272 | 2.437 | 0 | 0 | 0.912 | 0.685 | 1.085 | 4.114 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
57 | 2022-07-18 03:44 | WO-22270-30-A-1 | U | 0 | 0 | 2.817 | 1.854 | 0 | 2.664 | 0.836 | 1.847 | 4.086 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 94 | 2022-07-19 01:56 | 211892-d2 | U | 0 | 2.421 | 0 | 0 | 0 | 1.42 | 0.383 | 1.217 | 3.988 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
58 | 2022-07-18 04:20 | AB-220715-01 LRB 4_INJ001 | K | 0 | 0 | 0 | 1.03 | 0.968 | 1.15 | 0.597 | 1.507 | 0.328 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 95 | 2022-07-19 02:32 | 211885-d2 | U | 0 | 0 | 0 | 0 | 0 | 4.097 | 0.944 | 1.713 | 2.866 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
59 | 2022-07-18 04:56 | AB-220715-01 LRB 4 | K | 0 | 0 | 0 | 1.05 | 0.951 | 1.046 | 0.398 | 0.839 | 0.33 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 96 | 2022-07-19 03:08 | 211893-d2 | U | 0 | 1.069 | 0 | 0 | 0 | 0.765 | 0.473 | 1.563 | 3.171 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
60 | 2022-07-18 05:32 | WO-22270-31-A-1 | U | 0 | 0 | 0.624 | 0 | 0 | 1.614 | 1.049 | 0.864 | 2.993 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 97 | 2022-07-19 03:44 | 211894-d2 | U | 0 | 2.623 | 0 | 0 | 0 | 3.544 | 1.577 | 0.999 | 3.518 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
61 | 2022-07-18 06:08 | WO-22270-32-A-1 | U | 0 | 0 | 0 | 0 | 0 | 1.812 | 0.671 | 2.278 | 3.971 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  | 104 | 2022-07-19 07:56 | 211895-d2 | U | 0 | 1.818 | 0 | 0 | 0 | 0.83 | 0.482 | 2.014 | 3.348 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
62 | 2022-07-18 06:44 | WO-22270-33-A-1 | U | 0 | 0 | 0.752 | 0 | 0 | 2.319 | 0.42 | 0.899 | 3.255 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 105 | 2022-07-19 08:32 | 211896-d2 | U | 0 | 1.636 | 1.249 | 0 | 0 | 1.122 | 0.586 | 2.345 | 3.215 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
63 | 2022-07-18 07:20 | WO-22270-34-A-1 | U | 0 | 0 | 0 | 0 | 0 | 3.094 | 0.998 | 1.188 | 4.461 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [5 empty] | 106 | 2022-07-19 09:08 | 211897-d2 | U | 0 | 2.378 | 0 | 0 | 0 | 1.792 | 0.534 | 1.061 | 3.361 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
64 | 2022-07-18 07:56 | WO-22270-35-A-1 | U | 0 | 0 | 0 | 0 | 0 | 0.962 | 1.591 | 1.696 | 2.965 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 107 | 2022-07-19 09:44 | 211898-d2 | U | 0 | 2.725 | 2.491 | 0 | 0 | 0 | 3.413 | 0.884 | 3.534 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
65 | 2022-07-18 08:32 | WO-22270-36-A-1 | U | 0 | 0 | 0.078 | 0 | 0 | 1.501 | 0.517 | 1.719 | 3.607 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 108 | 2022-07-19 10:20 | 211899-d2 | U | 0 | 0.788 | 0 | 1.145 | 0 | 0.741 | 0.555 | 1.601 | 3.516 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
66 | 2022-07-18 09:08 | WO-22270-37-A-1 | U | 0 | 0.915 | 0 | 0 | 0 | 1.387 | 0.447 | 1.024 | 3.229 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 109 | 2022-07-19 10:56 | 211900-d2 | U | 0 | 3.244 | 3.031 | 0 | 0 | 1.099 | 0.722 | 2.321 | 3.2 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
67 | 2022-07-18 09:44 | WO-22270-38-A-1 | U | 0 | 1.299 | 0 | 0 | 0 | 1.752 | 0.48 | 0.958 | 2.328 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 110 | 2022-07-19 11:32 | 211901-d2 | U | 0 | 2.024 | 2.759 | 0 | 0 | 2.636 | 20.342 | 2.084 | 4.157 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 40.684 | <MRL | <MRL
68 | 2022-07-18 10:20 | WO-22270-39-A-1 | U | 0 | 0 | 0 | 0 | 0 | 1.962 | 0.429 | 2.016 | 2.578 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 111 | 2022-07-19 12:08 | 211902-d2 | U | 0 | 3.843 | 2.365 | 0 | 0 | 1.27 | 0.457 | 2.14 | 3.686 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
69 | 2022-07-18 10:56 | WO-22270-40-A-1 | U | 0 | 0 | 1.034 | 0 | 0 | 0.761 | 1.428 | 0.952 | 3.462 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 112 | 2022-07-19 12:44 | 211903-d2 | U | 0 | 3.852 | 0 | 0 | 0 | 0.793 | 0.518 | 2.992 | 4.846 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
70 | 2022-07-18 11:32 | AB-220715-01 LFB-H 3 | Z | 157.99 | 168.028 | 196.235 | 190.785 | 153.757 | 199.713 | 199.696 | 164.912 | 198.894 | 1 | <MRL | 168.028 | 196.235 | 190.785 | 153.757 | 199.713 | 199.696 | 164.912 | 198.894 | No TrueV! | [12 empty] | 113 | 2022-07-19 13:20 | 211904-d2 | U | 0 | 1.922 | 1.632 | 0 | 0 | 3.853 | 1.354 | 1.318 | 4.055 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
71 | 2022-07-18 12:08 | AB-220715-01 LRB 5_INJ001 | K | 0 | 0 | -0.174 | 0.976 | 1.63 | 0.766 | 0.45 | 0.988 | 0.496 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 116 | 2022-07-19 15:08 | 211905-d2 | U | 0 | 1.464 | 2.742 | 0 | 0 | 2.626 | 5.465 | 0.896 | 8.605 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 10.93 | <MRL | 17.21
72 | 2022-07-18 12:44 | AB-220715-01 LRB 5 | K | 0 | 0.683 | -0.157 | 1.012 | 1.69 | 0.861 | 0.41 | 1.616 | 0.431 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 117 | 2022-07-19 15:44 | 211906-d2 | U | 0 | 1.316 | 1.019 | 0 | 0 | 1.363 | 0.393 | 0.981 | 3.996 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
73 | 2022-07-18 13:20 | AB-220715-01 LD 2 | D | 0 | 1.048 | 0 | 0 | 2.332 | 4.203 | 1.059 | 1.012 | 2.589 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 118 | 2022-07-19 16:20 | 211907-d2 | U | 0 | 1.954 | 3.521 | 0 | 0 | 0.576 | 0.46 | 0.921 | 3.362 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
74 | 2022-07-18 13:56 | AB-220715-01 CCC-H 3 | Q | 361.667 | 340.157 | 397.811 | 382.755 | 317.243 | 391.053 | 391.511 | 332.226 | 389.63 | 1 | 361.667 | 340.157 | 397.811 | 382.755 | 317.243 | 391.053 | 391.511 | 332.226 | 389.63 | UL! | [12 empty] | 119 | 2022-07-19 16:56 | 211908-d2 | U | 0 | 1.892 | 2.192 | 0 | 0 | 2.129 | 2.505 | 2.049 | 6.787 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 13.574
75 | 2022-07-18 14:32 | AB-220715-01 LFSM-H 2 | S | 0 | 168.055 | 184.578 | 183.486 | 150.282 | 198.046 | 196.189 | 152.35 | 196.473 | 1 | <MRL | 168.055 | 184.578 | 183.486 | 150.282 | 198.046 | 196.189 | 152.35 | 196.473 | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  |  |  |  | 120 | 2022-07-19 17:32 | 211909-d2 | U | 0 | 1.166 | 2.943 | 0 | 0 | 4.871 | 1.793 | 1.729 | 3.346 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
76 | 2022-07-18 15:08 | WO-22270-41-A-1 | U | 0 | 0 | 0 | 0 | 0 | 0.959 | 0.66 | 0.97 | 2.578 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 121 | 2022-07-19 18:08 | 211910-d2 | U | 0 | 2.297 | 3.26 | 0 | 0 | 1.654 | 0.522 | 3.292 | 4.155 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
77 | 2022-07-18 15:44 | WO-22270-42-A-1 | U | 0 | 1.077 | 2.162 | 0 | 0 | 1.255 | 0.428 | 0.841 | 1.633 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 122 | 2022-07-19 18:44 | 211911-d2 | U | 0 | 1.209 | 0.62 | 1.755 | 0 | 0.829 | 0.445 | 2.161 | 19.796 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 39.592
78 | 2022-07-18 16:20 | WO-22270-43-A-1 | U | 0 | 0 | 0 | 0 | 0 | 4.642 | 0.663 | 0.917 | 2.633 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 123 | 2022-07-19 19:20 | 211912-d2 | U | 0 | 0.789 | 2.357 | 1.926 | 0 | 0.94 | 0.575 | 1.883 | 3.982 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
79 | 2022-07-18 16:56 | WO-22270-44-A-1 | U | 0 | 0 | 1.368 | 0 | 0 | 1.306 | 0.442 | 0.928 | 1.905 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 124 | 2022-07-19 19:56 | 211913-d2 | U | 0 | 1.09 | 1.421 | 1.998 | 0 | 1.095 | 0.403 | 1.036 | 3.306 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
80 | 2022-07-18 17:32 | WO-22270-45-A-1 | U | 0 | 1.289 | 3.345 | 0 | 0 | 0 | 1.441 | 0.858 | 1.785 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 125 | 2022-07-19 20:32 | 211914-d2 | U | 0 | 1.521 | 2.245 | 2.097 | 0 | 1.857 | 0.983 | 0.867 | 3.667 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
81 | 2022-07-18 18:08 | WO-22270-52-A-1 | U | 0 | 1.25 | 3.526 | 0 | 0 | 1.301 | 0.646 | 2.063 | 2.789 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 132 | 2022-07-20 00:44 | 211915-d2 | U | 0 | 1.934 | 3.201 | 3.35 | 0 | 1.584 | 0.552 | 3.031 | 3.721 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
82 | 2022-07-18 18:44 | WO-22270-46-A-1 | U | 0 | 1.039 | 2.103 | 0 | 0 | 1.241 | 0.561 | 1.511 | 3.148 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 133 | 2022-07-20 01:20 | 211916-d2 | U | 0 | 1.181 | 2.934 | 3.299 | 0 | 2.325 | 0.474 | 0.86 | 3.553 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
83 | 2022-07-18 19:20 | WO-22270-47-A-1 | U | 0 | 0 | 0.861 | 0 | 0 | 1.16 | 0.441 | 0.922 | 1.894 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 134 | 2022-07-20 01:56 | 211917-d2 | U | 0 | 1.222 | 2.189 | 3.92 | 0 | 0.903 | 0.568 | 1.866 | 3.516 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
84 | 2022-07-18 19:56 | WO-22270-53-A-1 | U | 0 | 1.507 | 1.472 | 0 | 0 | 1.634 | 0.685 | 0.826 | 3.553 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 135 | 2022-07-20 02:32 | 211918-d2 | U | 0 | 2.623 | 0 | 10.43 | 0 | 1.798 | 0.444 | 0.926 | 2.515 | 2 | <MRL | <MRL | <MRL | 20.86 | <MRL | <MRL | <MRL | <MRL | <MRL
85 | 2022-07-18 20:32 | WO-22270-48-A-1 | U | 0 | 0 | 1.281 | 0 | 0 | 0.488 | 0.378 | 1.93 | 3.977 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 136 | 2022-07-20 03:08 | 211919-d2 | U | 0 | 1.233 | 1.331 | 8.256 | 0 | 1.847 | 0.342 | 2.173 | 3.296 | 2 | <MRL | <MRL | <MRL | 16.512 | <MRL | <MRL | <MRL | <MRL | <MRL
86 | 2022-07-18 21:08 | WO-22270-54-A-1 | U | 0 | 0 | 3.071 | 1.304 | 0 | 0.828 | 0.375 | 1.31 | 3.216 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | [8 empty] | 137 | 2022-07-20 03:44 | 211920-d2 | U | 0 | 1.149 | 1.607 | 8.317 | 0 | 3.486 | 0.451 | 0.937 | 3.299 | 2 | <MRL | <MRL | <MRL | 16.634 | <MRL | <MRL | <MRL | <MRL | <MRL
87 | 2022-07-18 21:44 | WO-22270-49-A-1 | U | 0 | 0 | 0 | 0 | 0 | 0.776 | 0.497 | 0.988 | 4.442 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 138 | 2022-07-20 04:20 | 211921-d2 | U | 0 | 0 | 0 | 29.138 | 0 | 0.651 | 0.522 | 1.704 | 3.777 | 2 | <MRL | <MRL | <MRL | 58.276 | <MRL | <MRL | <MRL | <MRL | <MRL
88 | 2022-07-18 22:20 | WO-22270-55-A-1 | U | 0 | 1.289 | 2.916 | 0 | 0 | 0 | 3.128 | 3.016 | 2.961 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  | 139 | 2022-07-20 04:56 | 211922-d2 | U | 0 | 1.205 | 2.168 | 29.571 | 0 | 0.822 | 0.458 | 2.057 | 4.311 | 2 | <MRL | <MRL | <MRL | 59.142 | <MRL | <MRL | <MRL | <MRL | <MRL
89 | 2022-07-18 22:56 | WO-22270-50-A-1 | U | 0 | 0.742 | 0 | 0 | 0 | 0.833 | 1.212 | 0.949 | 4.275 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 140 | 2022-07-20 05:32 | 211923-d2 | U | 0 | 1.45 | 1.877 | 34.603 | 0 | 1.466 | 0.764 | 2.282 | 3.765 | 2 | <MRL | <MRL | <MRL | 69.206 | <MRL | <MRL | <MRL | <MRL | <MRL
90 | 2022-07-18 23:32 | WO-22270-56-A-1 | U | 0 | 1.308 | 0 | 0 | 0 | 1.763 | 0.793 | 0.891 | 4.414 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 141 | 2022-07-20 06:08 | 211924-d2 | U | 0 | 1.248 | 1.717 | 34.536 | 0 | 0 | 0 | 0.996 | 3.436 | 2 | <MRL | <MRL | <MRL | 69.072 | <MRL | <MRL | <MRL | <MRL | <MRL
91 | 2022-07-19 00:08 | WO-22270-57-A-1 | U | 0 | 1.272 | 2.437 | 0 | 0 | 0.912 | 0.685 | 1.085 | 4.114 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 144 | 2022-07-20 07:56 | 211925-d2 | U | 0 | 0.964 | 0 | 89.122 | 0 | 2.673 | 0.341 | 1.001 | 2.934 | 2 | <MRL | <MRL | <MRL | 178.244 | <MRL | <MRL | <MRL | <MRL | <MRL
92 | 2022-07-19 00:44 | AB-220715-01 LRB 6_INJ001 | K | 0 | 0 | 0.438 | 0.966 | 1.676 | 0.485 | 0.346 | 1.236 | 0.36 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 145 | 2022-07-20 08:32 | 211926-d2 | U | 0 | 1.274 | 1.92 | 118.451 | 0 | 1.678 | 0.488 | 1.027 | 3.82 | 2 | <MRL | <MRL | <MRL | 236.902 | <MRL | <MRL | <MRL | <MRL | <MRL
93 | 2022-07-19 01:20 | AB-220715-01 LRB 6 | K | 0 | 0.557 | 1.195 | 1.007 | 1.456 | 0.535 | 0.342 | 0.934 | 0.365 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 146 | 2022-07-20 09:08 | 211927-d2 | U | 0 | 2.193 | 1.009 | 95.675 | 0 | 2.634 | 0.505 | 1.807 | 4.301 | 2 | <MRL | <MRL | <MRL | 191.35 | <MRL | <MRL | <MRL | <MRL | <MRL
94 | 2022-07-19 01:56 | WO-22270-58-A-1 | U | 0 | 2.421 | 0 | 0 | 0 | 1.42 | 0.383 | 1.217 | 3.988 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 147 | 2022-07-20 09:44 | 211928-d2 | U | 0 | 1.385 | 1.973 | 513.964 | 0 | 4.213 | 0.553 | 0.892 | 3.73 | 2 | <MRL | <MRL | <MRL | 1027.928 | <MRL | <MRL | <MRL | <MRL | <MRL
95 | 2022-07-19 02:32 | WO-22270-51-A-1 | U | 0 | 0 | 0 | 0 | 0 | 4.097 | 0.944 | 1.713 | 2.866 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 148 | 2022-07-20 10:20 | 211929-d2 | U | 0 | 1.772 | 1.934 | 296.758 | 0 | 1.156 | 0.443 | 0.933 | 3.59 | 2 | <MRL | <MRL | <MRL | 593.516 | <MRL | <MRL | <MRL | <MRL | <MRL
96 | 2022-07-19 03:08 | WO-22270-59-A-1 | U | 0 | 1.069 | 0 | 0 | 0 | 0.765 | 0.473 | 1.563 | 3.171 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 149 | 2022-07-20 10:56 | 211930-d2 | U | 0 | 0.84 | 0 | 429.411 | 0 | 2.135 | 0.395 | 1.081 | 3.898 | 2 | <MRL | <MRL | <MRL | 858.822 | <MRL | <MRL | <MRL | <MRL | <MRL
97 | 2022-07-19 03:44 | WO-22270-60-A-1 | U | 0 | 2.623 | 0 | 0 | 0 | 3.544 | 1.577 | 0.999 | 3.518 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 150 | 2022-07-20 11:32 | 211931-d2 | U | 0 | 0.982 | 1.411 | 2.168 | 0 | 0.986 | 0.561 | 0.924 | 3.657 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
98 | 2022-07-19 04:20 | AB-220715-01 LFB-H 4 | Z | 162.744 | 175.872 | 197.68 | 199.009 | 157.175 | 199.897 | 204.168 | 165.163 | 198.772 | 1 | <MRL | 175.872 | 197.68 | 199.009 | 157.175 | 199.897 | 204.168 | 165.163 | 198.772 | No TrueV! | [12 empty] | 151 | 2022-07-20 12:08 | 211932-d2 | U | 0 | 0.913 | 3.012 | 1.918 | 0 | 0.611 | 0.427 | 0.86 | 4.501 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
99 | 2022-07-19 04:56 | AB-220715-01 LRB 7_INJ001 | K | 0 | 0.669 | -0.005 | 1.05 | 1.998 | 1.489 | 0.567 | 1.365 | 0.506 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 152 | 2022-07-20 12:44 | 211933-d2 | U | 0 | 1.609 | 2.588 | 0 | 0 | 0 | 1.034 | 2.825 | 4.353 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
100 | 2022-07-19 05:32 | AB-220715-01 LRB 7 | K | 0 | 0 | 0.5 | 0.992 | 2.109 | 1.449 | 0.776 | 0.94 | 0.397 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | [11 empty] | 153 | 2022-07-20 13:20 | 211934-d2 | U | 0 | 2.049 | 0 | 1.921 | 0 | 1.687 | 0.447 | 0.868 | 4.313 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
101 | 2022-07-19 06:08 | AB-220715-01 LD 3 | D | 0 | 1.254 | 0 | 0 | 0 | 1.999 | 0.487 | 0.955 | 3.264 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 160 | 2022-07-20 17:32 | 211935-d2 | U | 0 | 0.982 | 3.681 | 2.944 | 0 | 0.731 | 0.434 | 1.145 | 4.04 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
102 | 2022-07-19 06:44 | AB-220715-01 CCC-H 4 | Q | 336.451 | 332.142 | 388.596 | 377.704 | 321.727 | 387.419 | 386.028 | 331.024 | 382.399 | 1 | 336.451 | 332.142 | 388.596 | 377.704 | 321.727 | 387.419 | 386.028 | 331.024 | 382.399 | UL! | [12 empty] | 161 | 2022-07-20 18:08 | 211936-d2 | U | 0 | 1.622 | 1.641 | 2.814 | 0 | 1.723 | 0.503 | 2.04 | 4.044 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
103 | 2022-07-19 07:20 | AB-220715-01 LFSM-H 3 | S | 0 | 176.858 | 191.02 | 182.049 | 151.448 | 197.926 | 198.045 | 154.815 | 200.285 | 1 | <MRL | 176.858 | 191.02 | 182.049 | 151.448 | 197.926 | 198.045 | 154.815 | 200.285 | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | [6 empty] | 162 | 2022-07-20 18:44 | 211937-d2 | U | 0 | 1.09 | 1.6 | 3.613 | 0 | 1.333 | 0.843 | 1.683 | 4.974 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
104 | 2022-07-19 07:56 | WO-22270-61-A-1 | U | 0 | 1.818 | 0 | 0 | 0 | 0.83 | 0.482 | 2.014 | 3.348 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 163 | 2022-07-20 19:20 | 211938-d2 | U | 0 | 1.437 | 1.636 | 8.259 | 0 | 0.625 | 0.365 | 1.001 | 3.97 | 2 | <MRL | <MRL | <MRL | 16.518 | <MRL | <MRL | <MRL | <MRL | <MRL
105 | 2022-07-19 08:32 | WO-22270-62-A-1 | U | 0 | 1.636 | 1.249 | 0 | 0 | 1.122 | 0.586 | 2.345 | 3.215 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [5 empty] | 164 | 2022-07-20 19:56 | 211939-d2 | U | 0 | 1.197 | 1.889 | 7.025 | 0 | 0.608 | 0.782 | 1.069 | 4.763 | 2 | <MRL | <MRL | <MRL | 14.05 | <MRL | <MRL | <MRL | <MRL | <MRL
106 | 2022-07-19 09:08 | WO-22270-63-A-1 | U | 0 | 2.378 | 0 | 0 | 0 | 1.792 | 0.534 | 1.061 | 3.361 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 165 | 2022-07-20 20:32 | 211940-d2 | U | 0 | 0.963 | 0 | 7.749 | 0 | 0.523 | 0.619 | 2.235 | 4.019 | 2 | <MRL | <MRL | <MRL | 15.498 | <MRL | <MRL | <MRL | <MRL | <MRL
107 | 2022-07-19 09:44 | WO-22270-64-A-1 | U | 0 | 2.725 | 2.491 | 0 | 0 | 0 | 3.413 | 0.884 | 3.534 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  | 166 | 2022-07-20 21:08 | 211941-d2 | U | 0 | 1.928 | 2.217 | 26.028 | 0 | 0.535 | 0.499 | 1.906 | 3.585 | 2 | <MRL | <MRL | <MRL | 52.056 | <MRL | <MRL | <MRL | <MRL | <MRL
108 | 2022-07-19 10:20 | WO-22270-65-A-1 | U | 0 | 0.788 | 0 | 1.145 | 0 | 0.741 | 0.555 | 1.601 | 3.516 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 167 | 2022-07-20 21:44 | 211942-d2 | U | 0 | 1.173 | 3.237 | 24.938 | 0 | 0.557 | 0.489 | 1.045 | 3.661 | 2 | <MRL | <MRL | <MRL | 49.876 | <MRL | <MRL | <MRL | <MRL | <MRL
109 | 2022-07-19 10:56 | WO-22270-66-A-1 | U | 0 | 3.244 | 3.031 | 0 | 0 | 1.099 | 0.722 | 2.321 | 3.2 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [5 empty] | 168 | 2022-07-20 22:20 | 211943-d2 | U | 0 | 1.642 | 1.92 | 36.818 | 0 | 0.617 | 1.347 | 0.961 | 4.195 | 2 | <MRL | <MRL | <MRL | 73.636 | <MRL | <MRL | <MRL | <MRL | <MRL
110 | 2022-07-19 11:32 | WO-22270-67-A-1 | U | 0 | 2.024 | 2.759 | 0 | 0 | 2.636 | 20.342 | 2.084 | 4.157 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 20.342 | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 |  |  |  |  | 169 | 2022-07-20 22:56 | 211944-d2 | U | 0 | 0.734 | 2.75 | 35.931 | 0 | 1.446 | 1.286 | 0.98 | 3.898 | 2 | <MRL | <MRL | <MRL | 71.862 | <MRL | <MRL | <MRL | <MRL | <MRL
111 | 2022-07-19 12:08 | WO-22270-68-A-1 | U | 0 | 3.843 | 2.365 | 0 | 0 | 1.27 | 0.457 | 2.14 | 3.686 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 172 | 2022-07-21 00:44 | 211945-d2 | U | 0 | 1.267 | 1.185 | 72.024 | 0 | 0.756 | 0.505 | 0.886 | 3.531 | 2 | <MRL | <MRL | <MRL | 144.048 | <MRL | <MRL | <MRL | <MRL | <MRL
112 | 2022-07-19 12:44 | WO-22270-69-A-1 | U | 0 | 3.852 | 0 | 0 | 0 | 0.793 | 0.518 | 2.992 | 4.846 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  | 173 | 2022-07-21 01:20 | 211946-d2 | U | 0 | 1.418 | 2.563 | 72.186 | 0 | 0.638 | 0.855 | 1.181 | 3.652 | 2 | <MRL | <MRL | <MRL | 144.372 | <MRL | <MRL | <MRL | <MRL | <MRL
113 | 2022-07-19 13:20 | WO-22270-70-A-1 | U | 0 | 1.922 | 1.632 | 0 | 0 | 3.853 | 1.354 | 1.318 | 4.055 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 174 | 2022-07-21 01:56 | 211947-d2 | U | 0 | 1.485 | 1.551 | 94.178 | 0 | 0.999 | 0.372 | 2.63 | 3.505 | 2 | <MRL | <MRL | <MRL | 188.356 | <MRL | <MRL | <MRL | <MRL | <MRL
114 | 2022-07-19 13:56 | AB-220715-01 LRB 8_INJ001 | K | 0 | 0.733 | -0.188 | 0.985 | 2.078 | 1.451 | 0.692 | 0.885 | 0.33 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 175 | 2022-07-21 02:32 | 211948-d2 | U | 0 | 1.335 | 0 | 326.568 | 0 | 1.482 | 0.57 | 2.402 | 4.205 | 2 | <MRL | <MRL | <MRL | 653.136 | <MRL | <MRL | <MRL | <MRL | <MRL
115 | 2022-07-19 14:32 | AB-220715-01 LRB 8 | K | 0 | 0.678 | 0.346 | 0.981 | 2.103 | 1.359 | 0.846 | 1.442 | 0.371 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 176 | 2022-07-21 03:08 | 211949-d2 | U | 0 | 1.218 | 3.524 | 222.047 | 0 | 2.163 | 2.753 | 1.171 | 4.431 | 2 | <MRL | <MRL | <MRL | 444.094 | <MRL | <MRL | <MRL | <MRL | <MRL
116 | 2022-07-19 15:08 | WO-22270-71-A-1 | U | 0 | 1.464 | 2.742 | 0 | 0 | 2.626 | 5.465 | 0.896 | 8.605 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 5.465 | <MRL | 8.605 | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [5 empty] | 177 | 2022-07-21 03:44 | 211950-d2 | U | 0 | 2.36 | 0 | 239.793 | 0 | 1.595 | 0.482 | 0.987 | 3.651 | 2 | <MRL | <MRL | <MRL | 479.586 | <MRL | <MRL | <MRL | <MRL | <MRL
117 | 2022-07-19 15:44 | WO-22270-72-A-1 | U | 0 | 1.316 | 1.019 | 0 | 0 | 1.363 | 0.393 | 0.981 | 3.996 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [5 empty] | 178 | 2022-07-21 04:20 | 211951-d2 | U | 0 | 1.803 | 2.873 | 2.136 | 0 | 1.126 | 0.507 | 0.964 | 27.663 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 55.326
118 | 2022-07-19 16:20 | WO-22270-73-A-1 | U | 0 | 1.954 | 3.521 | 0 | 0 | 0.576 | 0.46 | 0.921 | 3.362 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 179 | 2022-07-21 04:56 | 211952-d2 | U | 0 | 1.197 | 0 | 2.136 | 0 | 0.666 | 0.423 | 0.913 | 2.864 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
119 | 2022-07-19 16:56 | WO-22270-74-A-1 | U | 0 | 1.892 | 2.192 | 0 | 0 | 2.129 | 2.505 | 2.049 | 6.787 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 6.787 | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 180 | 2022-07-21 05:32 | 211953-d2 | U | 0 | 1.762 | 0 | 1.872 | 0 | 1.779 | 0.726 | 2.894 | 4.091 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
120 | 2022-07-19 17:32 | WO-22270-75-A-1 | U | 0 | 1.166 | 2.943 | 0 | 0 | 4.871 | 1.793 | 1.729 | 3.346 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 181 | 2022-07-21 06:08 | 211954-d2 | U | 0 | 1.069 | 0 | 1.811 | 0 | 1.868 | 0.958 | 1.033 | 4.189 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
121 | 2022-07-19 18:08 | WO-22270-76-A-1 | U | 0 | 2.297 | 3.26 | 0 | 0 | 1.654 | 0.522 | 3.292 | 4.155 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  | 188 | 2022-07-21 10:20 | 211955-d2 | U | 0 | 1.007 | 1.054 | 2.834 | 0 | 0 | 0.406 | 1.38 | 3.691 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
122 | 2022-07-19 18:44 | WO-22270-77-A-1 | U | 0 | 1.209 | 0.62 | 1.755 | 0 | 0.829 | 0.445 | 2.161 | 19.796 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 19.796 | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 189 | 2022-07-21 10:56 | 211956-d2 | U | 0 | 2.476 | 0 | 4.153 | 0 | 1.193 | 0.714 | 0.971 | 3.841 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
123 | 2022-07-19 19:20 | WO-22270-78-A-1 | U | 0 | 0.789 | 2.357 | 1.926 | 0 | 0.94 | 0.575 | 1.883 | 3.982 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 190 | 2022-07-21 11:32 | 211957-d2 | U | 0 | 0.914 | 0 | 3.574 | 0 | 0.544 | 0.454 | 2.967 | 4.068 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
124 | 2022-07-19 19:56 | WO-22270-79-A-1 | U | 0 | 1.09 | 1.421 | 1.998 | 0 | 1.095 | 0.403 | 1.036 | 3.306 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 191 | 2022-07-21 12:08 | 211958-d2 | U | 0 | 1.17 | 0 | 8.903 | 0 | 3.316 | 0.583 | 1.115 | 4.517 | 2 | <MRL | <MRL | <MRL | 17.806 | <MRL | <MRL | <MRL | <MRL | <MRL
125 | 2022-07-19 20:32 | WO-22270-80-A-1 | U | 0 | 1.521 | 2.245 | 2.097 | 0 | 1.857 | 0.983 | 0.867 | 3.667 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 192 | 2022-07-21 12:44 | 211959-d2 | U | 0 | 2.165 | 0 | 8.16 | 0 | 1.426 | 0.925 | 0.877 | 4.939 | 2 | <MRL | <MRL | <MRL | 16.32 | <MRL | <MRL | <MRL | <MRL | <MRL
126 | 2022-07-19 21:08 | AB-220715-01 LFB-H 5 | Z | 190.537 | 174.405 | 211.05 | 205.523 | 155.005 | 201.716 | 206.05 | 164.39 | 199.853 | 1 | <MRL | 174.405 | 211.05 | 205.523 | 155.005 | 201.716 | 206.05 | 164.39 | 199.853 | No TrueV! | [12 empty] | 193 | 2022-07-21 13:20 | 211960-d2 | U | 0 | 2.014 | 0 | 8.852 | 0 | 2.024 | 0.878 | 0.952 | 3.937 | 2 | <MRL | <MRL | <MRL | 17.704 | <MRL | <MRL | <MRL | <MRL | <MRL
127 | 2022-07-19 21:44 | AB-220715-01 LRB 9_INJ001 | K | 0 | 0.709 | 0.111 | 1.066 | 2.124 | 1.724 | 0.986 | 1.658 | 0.497 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 194 | 2022-07-21 13:56 | 211961-d2 | U | 0 | 1.381 | 0 | 28.051 | 0 | 2.717 | 0.484 | 0.903 | 3.877 | 2 | <MRL | <MRL | <MRL | 56.102 | <MRL | <MRL | <MRL | <MRL | <MRL
128 | 2022-07-19 22:20 | AB-220715-01 LRB 9 | K | 0 | 0.742 | 0.494 | 1.009 | 2.23 | 1.482 | 1.115 | 0.933 | 0.364 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 195 | 2022-07-21 14:32 | 211962-d2 | U | 0 | 2.627 | 0 | 33.692 | 0 | 1.561 | 0.52 | 1.558 | 3.491 | 2 | <MRL | <MRL | <MRL | 67.384 | <MRL | <MRL | <MRL | <MRL | <MRL
129 | 2022-07-19 22:56 | AB-220715-01 LD 4 | D | 0 | 2.589 | 3.027 | 2.379 | 0 | 3.311 | 1.152 | 2.064 | 3.69 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 196 | 2022-07-21 15:08 | 211963-d2 | U | 0 | 0.913 | 1.124 | 38.125 | 0 | 2.706 | 0.837 | 0.979 | 4.154 | 2 | <MRL | <MRL | <MRL | 76.25 | <MRL | <MRL | <MRL | <MRL | <MRL
130 | 2022-07-19 23:32 | AB-220715-01 CCC-H 5 | Q | 356.628 | 338.354 | 389.461 | 384.448 | 315.397 | 376.568 | 395.37 | 332.343 | 386.242 | 1 | 356.628 | 338.354 | 389.461 | 384.448 | 315.397 | 376.568 | 395.37 | 332.343 | 386.242 | UL! | [12 empty] | 197 | 2022-07-21 15:44 | 211964-d2 | U | 0 | 1.542 | 0 | 39.499 | 0 | 21.233 | 1.533 | 0.829 | 5.842 | 2 | <MRL | <MRL | <MRL | 78.998 | <MRL | 42.466 | <MRL | <MRL | 11.684
131 | 2022-07-20 00:08 | AB-220715-01 LFSM-H 4 | S | 0 | 170.383 | 190.142 | 193.605 | 149.526 | 196.17 | 204.594 | 151.141 | 201.625 | 1 | <MRL | 170.383 | 190.142 | 193.605 | 149.526 | 196.17 | 204.594 | 151.141 | 201.625 | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | [5 empty] | 200 | 2022-07-21 17:32 | 211965-d2 | U | 0 | 1.28 | 0 | 142.268 | 0 | 1.455 | 0.796 | 0.743 | 3.517 | 2 | <MRL | <MRL | <MRL | 284.536 | <MRL | <MRL | <MRL | <MRL | <MRL
132 | 2022-07-20 00:44 | WO-22270-81-A-1 | U | 0 | 1.934 | 3.201 | 3.35 | 0 | 1.584 | 0.552 | 3.031 | 3.721 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 201 | 2022-07-21 18:08 | 211966-d2 | U | 0 | 1.683 | 0 | 112.938 | 0 | 3.691 | 0.725 | 0 | 3.637 | 2 | <MRL | <MRL | <MRL | 225.876 | <MRL | <MRL | <MRL | <MRL | <MRL
133 | 2022-07-20 01:20 | WO-22270-82-A-1 | U | 0 | 1.181 | 2.934 | 3.299 | 0 | 2.325 | 0.474 | 0.86 | 3.553 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 202 | 2022-07-21 18:44 | 211967-d2 | U | 0 | 1.435 | 0 | 83.818 | 0 | 2.816 | 2.182 | 3.056 | 4.59 | 2 | <MRL | <MRL | <MRL | 167.636 | <MRL | <MRL | <MRL | <MRL | <MRL
134 | 2022-07-20 01:56 | WO-22270-83-A-1 | U | 0 | 1.222 | 2.189 | 3.92 | 0 | 0.903 | 0.568 | 1.866 | 3.516 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 203 | 2022-07-21 19:20 | 211968-d2 | U | 0 | 2.327 | 0 | 2231.973 | 0 | 2.168 | 3.39 | 1.377 | 4.942 | 2 | <MRL | <MRL | <MRL | 4463.946 | <MRL | <MRL | <MRL | <MRL | <MRL
135 | 2022-07-20 02:32 | WO-22270-84-A-1 | U | 0 | 2.623 | 0 | 10.43 | 0 | 1.798 | 0.444 | 0.926 | 2.515 | 1 | <MRL | <MRL | <MRL | 10.43 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 204 | 2022-07-21 19:56 | 211969-d2 | U | 0 | 1.205 | 0 | 222.327 | 0 | 0.598 | 0.704 | 2.956 | 3.738 | 2 | <MRL | <MRL | <MRL | 444.654 | <MRL | <MRL | <MRL | <MRL | <MRL
136 | 2022-07-20 03:08 | WO-22270-85-A-1 | U | 0 | 1.233 | 1.331 | 8.256 | 0 | 1.847 | 0.342 | 2.173 | 3.296 | 1 | <MRL | <MRL | <MRL | 8.256 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 205 | 2022-07-21 20:32 | 211970-d2 | U | 0 | 1.742 | 3.41 | 51.868 | 0 | 1.873 | 0.774 | 0.913 | 4.151 | 2 | <MRL | <MRL | <MRL | 103.736 | <MRL | <MRL | <MRL | <MRL | <MRL
137 | 2022-07-20 03:44 | WO-22270-86-A-1 | U | 0 | 1.149 | 1.607 | 8.317 | 0 | 3.486 | 0.451 | 0.937 | 3.299 | 1 | <MRL | <MRL | <MRL | 8.317 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 206 | 2022-07-21 21:08 | 211971-d2 | U | 0 | 1.387 | 11.283 | 1.438 | 0 | 1.092 | 0.447 | 0.983 | 2.975 | 2 | <MRL | <MRL | 22.566 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
138 | 2022-07-20 04:20 | WO-22270-87-A-1 | U | 0 | 0 | 0 | 29.138 | 0 | 0.651 | 0.522 | 1.704 | 3.777 | 1 | <MRL | <MRL | <MRL | 29.138 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 207 | 2022-07-21 21:44 | 211972-d2 | U | 0 | 1.31 | 16.461 | 0 | 0 | 0.526 | 0.365 | 0.896 | 2.294 | 2 | <MRL | <MRL | 32.922 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
139 | 2022-07-20 04:56 | WO-22270-88-A-1 | U | 0 | 1.205 | 2.168 | 29.571 | 0 | 0.822 | 0.458 | 2.057 | 4.311 | 1 | <MRL | <MRL | <MRL | 29.571 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 208 | 2022-07-21 22:20 | 211973-d2 | U | 0 | 1.429 | 15.189 | 0 | 0 | 0.633 | 0.538 | 0.876 | 4.91 | 2 | <MRL | <MRL | 30.378 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
140 | 2022-07-20 05:32 | WO-22270-89-A-1 | U | 0 | 1.45 | 1.877 | 34.603 | 0 | 1.466 | 0.764 | 2.282 | 3.765 | 1 | <MRL | <MRL | <MRL | 34.603 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  |  |  | 209 | 2022-07-21 22:56 | 211974-d2 | U | 0 | 1.545 | 10.723 | 0 | 0 | 0.612 | 0.56 | 0.971 | 3.954 | 2 | <MRL | <MRL | 21.446 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
141 | 2022-07-20 06:08 | WO-22270-90-A-1 | U | 0 | 1.248 | 1.717 | 34.536 | 0 | 0 | 0 | 0.996 | 3.436 | 1 | <MRL | <MRL | <MRL | 34.536 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  |  |  | 216 | 2022-07-22 03:08 | 211975-d2 | U | 0 | 1.527 | 34.124 | 1.641 | 0 | 1.229 | 0.808 | 1.175 | 3.221 | 2 | <MRL | <MRL | 68.248 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
142 | 2022-07-20 06:44 | AB-220715-01 LRB 10_INJ001 | K | 0 | 0.855 | 0.364 | 0 | 0 | 1.382 | 1.144 | 0.955 | 0.39 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 217 | 2022-07-22 03:44 | 211976-d2 | U | 0 | 1.463 | 33.507 | 0 | 0 | 1.132 | 0.486 | 1.051 | 2.606 | 2 | <MRL | <MRL | 67.014 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
143 | 2022-07-20 07:20 | AB-220715-01 LRB 10 | K | 0 | 0.656 | 0.581 | 1.06 | 0 | 1.482 | 1.02 | 1.082 | 0.442 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | [10 empty] | 218 | 2022-07-22 04:20 | 211977-d2 | U | 0 | 2.015 | 26.498 | 0 | 0 | 0.601 | 0.608 | 0.984 | 3.584 | 2 | <MRL | <MRL | 52.996 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
144 | 2022-07-20 07:56 | WO-22270-91-A-1 | U | 0 | 0.964 | 0 | 89.122 | 0 | 2.673 | 0.341 | 1.001 | 2.934 | 1 | <MRL | <MRL | <MRL | 89.122 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 219 | 2022-07-22 04:56 | 211978-d2 | U | 0 | 1.164 | 102.136 | 0 | 0 | 1.771 | 0.83 | 1.097 | 3.23 | 2 | <MRL | <MRL | 204.272 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
145 | 2022-07-20 08:32 | WO-22270-92-A-1 | U | 0 | 1.274 | 1.92 | 118.451 | 0 | 1.678 | 0.488 | 1.027 | 3.82 | 1 | <MRL | <MRL | <MRL | 118.451 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 220 | 2022-07-22 05:32 | 211979-d2 | U | 0 | 0.689 | 88.027 | 0 | 0 | 1.335 | 0.525 | 0.844 | 2.649 | 2 | <MRL | <MRL | 176.054 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
146 | 2022-07-20 09:08 | WO-22270-93-A-1 | U | 0 | 2.193 | 1.009 | 95.675 | 0 | 2.634 | 0.505 | 1.807 | 4.301 | 1 | <MRL | <MRL | <MRL | 95.675 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 221 | 2022-07-22 06:08 | 211980-d2 | U | 0 | 0.882 | 92.579 | 0 | 0 | 1.95 | 0.591 | 0.965 | 2.223 | 2 | <MRL | <MRL | 185.158 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
147 | 2022-07-20 09:44 | WO-22270-94-A-1 | U | 0 | 1.385 | 1.973 | 513.964 | 0 | 4.213 | 0.553 | 0.892 | 3.73 | 1 | <MRL | <MRL | <MRL | 513.964 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 222 | 2022-07-22 06:44 | 211981-d2 | U | 0 | 1.916 | 324.339 | 0 | 0 | 2.685 | 0.489 | 1.445 | 3.744 | 2 | <MRL | <MRL | 648.678 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
148 | 2022-07-20 10:20 | WO-22270-95-A-1 | U | 0 | 1.772 | 1.934 | 296.758 | 0 | 1.156 | 0.443 | 0.933 | 3.59 | 1 | <MRL | <MRL | <MRL | 296.758 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | [5 empty] | 223 | 2022-07-22 07:20 | 211982-d2 | U | 0 | 1.139 | 304.539 | 2.76 | 0 | 4.589 | 0.533 | 1.57 | 2.35 | 2 | <MRL | <MRL | 609.078 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
149 | 2022-07-20 10:56 | WO-22270-96-A-1 | U | 0 | 0.84 | 0 | 429.411 | 0 | 2.135 | 0.395 | 1.081 | 3.898 | 1 | <MRL | <MRL | <MRL | 429.411 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 224 | 2022-07-22 07:56 | 211983-d2 | U | 0 | 0.908 | 312.784 | 2.979 | 0 | 1.944 | 0.671 | 0.928 | 4.147 | 2 | <MRL | <MRL | 625.568 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
150 | 2022-07-20 11:32 | WO-22270-97-A-1 | U | 0 | 0.982 | 1.411 | 2.168 | 0 | 0.986 | 0.561 | 0.924 | 3.657 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 225 | 2022-07-22 08:32 | 211984-d2 | U | 0 | 1.28 | 267.728 | 0 | 0 | 0.552 | 0.73 | 0.854 | 3.079 | 2 | <MRL | <MRL | 535.456 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
151 | 2022-07-20 12:08 | WO-22270-98-A-1 | U | 0 | 0.913 | 3.012 | 1.918 | 0 | 0.611 | 0.427 | 0.86 | 4.501 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 228 | 2022-07-22 10:20 | 211985-d2 | U | 0 | 0.918 | 993.029 | 1.571 | 0 | 0.623 | 1.754 | 0.829 | 2.406 | 2 | <MRL | <MRL | 1986.058 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
152 | 2022-07-20 12:44 | WO-22270-99-A-1 | U | 0 | 1.609 | 2.588 | 0 | 0 | 0 | 1.034 | 2.825 | 4.353 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  | 229 | 2022-07-22 10:56 | 211986-d2 | U | 0 | 1.351 | 1050.632 | 1.672 | 0 | 0.75 | 0.517 | 2.257 | 2.648 | 2 | <MRL | <MRL | 2101.264 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
153 | 2022-07-20 13:20 | WO-22270-100-A-1 | U | 0 | 2.049 | 0 | 1.921 | 0 | 1.687 | 0.447 | 0.868 | 4.313 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 230 | 2022-07-22 11:32 | 211987-d2 | U | 0 | 1.915 | 958.71 | 2.461 | 0 | 1.145 | 0.645 | 1.34 | 3.295 | 2 | <MRL | <MRL | 1917.42 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
154 | 2022-07-20 13:56 | AB-220715-01 LFB-H 6 | Z | 156.313 | 166.991 | 182.753 | 191.319 | 148.72 | 191.836 | 188.34 | 161.946 | 188.731 | 1 | <MRL | 166.991 | 182.753 | 191.319 | 148.72 | 191.836 | 188.34 | 161.946 | 188.731 | No TrueV! | [12 empty] | 231 | 2022-07-22 12:08 | 211988-d2 | U | 0 | 0.893 | 2069.312 | 2.409 | 0 | 0.773 | 5.502 | 1.161 | 4.213 | 2 | <MRL | <MRL | 4138.624 | <MRL | <MRL | <MRL | 11.004 | <MRL | <MRL |  |  |  | d
155 | 2022-07-20 14:32 | AB-220715-01 LRB 11_INJ001 | K | 0 | 0.684 | 0.603 | 0 | 0 | 1.569 | 0.798 | 2.04 | 0.503 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 232 | 2022-07-22 12:44 | 211989-d2 | U | 0 | 0.924 | 3164.557 | 2.974 | 0 | 0.664 | 0.599 | 0.936 | 3.675 | 2 | <MRL | <MRL | 6329.114 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
156 | 2022-07-20 15:08 | AB-220715-01 LRB 11 | K | 0 | 0.596 | 0.293 | 0.998 | 0 | 1.468 | 0.811 | 1.677 | 0.383 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 233 | 2022-07-22 13:20 | 211990-d2 | U | 0 | 1.656 | 2732.931 | 2.804 | 0 | 1.519 | 0.575 | 1.285 | 3.488 | 2 | <MRL | <MRL | 5465.862 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
157 | 2022-07-20 15:44 | AB-220715-01 LD 5 | D | 0 | 2.34 | 1.745 | 2.824 | 0 | 3.298 | 1.161 | 3.174 | 4.628 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  | 234 | 2022-07-22 13:56 | 211991-d2 | U | 0 | 1.469 | 17.397 | 3.054 | 0 | 1.923 | 4.66 | 1.133 | 3.856 | 2 | <MRL | <MRL | 34.794 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
158 | 2022-07-20 16:20 | AB-220715-01 CCC-H 6 | Q | 377.628 | 350.691 | 375.438 | 368.549 | 302.006 | 372.428 | 377.013 | 313.329 | 378.911 | 1 | 377.628 | 350.691 | 375.438 | 368.549 | 302.006 | 372.428 | 377.013 | 313.329 | 378.911 | UL! | [12 empty] | 235 | 2022-07-22 14:32 | 211992-d2 | U | 0 | 1.107 | 7.677 | 1.143 | 0 | 1.129 | 1.068 | 2.294 | 2.164 | 2 | <MRL | <MRL | 15.354 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
159 | 2022-07-20 16:56 | AB-220715-01 LFSM-H 5 | S | 0 | 171.217 | 184.913 | 184.057 | 144.456 | 196.1 | 197.638 | 149.42 | 198.493 | 1 | <MRL | 171.217 | 184.913 | 184.057 | 144.456 | 196.1 | 197.638 | 149.42 | 198.493 | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  |  |  | 236 | 2022-07-22 15:08 | 211993-d2 | U | 0 | 1.589 | 8.152 | 0 | 0 | 1.534 | 0.726 | 1.054 | 3.551 | 2 | <MRL | <MRL | 16.304 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
160 | 2022-07-20 17:32 | WO-22270-101-A-1 | U | 0 | 0.982 | 3.681 | 2.944 | 0 | 0.731 | 0.434 | 1.145 | 4.04 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 237 | 2022-07-22 15:44 | 211994-d2 | U | 0 | 1.25 | 3.988 | 0 | 0 | 0.554 | 0.636 | 2.993 | 3.234 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  |  |  | d
161 | 2022-07-20 18:08 | WO-22270-102-A-1 | U | 0 | 1.622 | 1.641 | 2.814 | 0 | 1.723 | 0.503 | 2.04 | 4.044 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 244 | 2022-07-22 19:56 | 211995-d2 | U | 0 | 1.741 | 12.151 | 1.612 | 0 | 0.627 | 1.19 | 1.224 | 2.606 | 2 | <MRL | <MRL | 24.302 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
162 | 2022-07-20 18:44 | WO-22270-103-A-1 | U | 0 | 1.09 | 1.6 | 3.613 | 0 | 1.333 | 0.843 | 1.683 | 4.974 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 245 | 2022-07-22 20:32 | 211996-d2 | U | 0 | 1.197 | 15.364 | 0 | 0 | 0.958 | 0.54 | 1.038 | 3.833 | 2 | <MRL | <MRL | 30.728 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
163 | 2022-07-20 19:20 | WO-22270-104-A-1 | U | 0 | 1.437 | 1.636 | 8.259 | 0 | 0.625 | 0.365 | 1.001 | 3.97 | 1 | <MRL | <MRL | <MRL | 8.259 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | [5 empty] | 246 | 2022-07-22 21:08 | 211997-d2 | U | 0 | 1.477 | 14.745 | 0 | 0 | 1.607 | 0.599 | 1.276 | 2.784 | 2 | <MRL | <MRL | 29.49 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
164 | 2022-07-20 19:56 | WO-22270-105-A-1 | U | 0 | 1.197 | 1.889 | 7.025 | 0 | 0.608 | 0.782 | 1.069 | 4.763 | 1 | <MRL | <MRL | <MRL | 7.025 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 247 | 2022-07-22 21:44 | 211998-d2 | U | 0 | 1.355 | 41.544 | 1.282 | 0 | 1.813 | 1.221 | 2.595 | 3.054 | 2 | <MRL | <MRL | 83.088 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
165 | 2022-07-20 20:32 | WO-22270-106-A-1 | U | 0 | 0.963 | 0 | 7.749 | 0 | 0.523 | 0.619 | 2.235 | 4.019 | 1 | <MRL | <MRL | <MRL | 7.749 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 248 | 2022-07-22 22:20 | 211999-d2 | U | 0 | 1.596 | 41.894 | 0 | 0 | 0.662 | 0.495 | 0.961 | 3.212 | 2 | <MRL | <MRL | 83.788 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
166 | 2022-07-20 21:08 | WO-22270-107-A-1 | U | 0 | 1.928 | 2.217 | 26.028 | 0 | 0.535 | 0.499 | 1.906 | 3.585 | 1 | <MRL | <MRL | <MRL | 26.028 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 249 | 2022-07-22 22:56 | 212000-d2 | U | 0 | 1.124 | 43.64 | 0 | 0 | 0.702 | 0.432 | 1.05 | 3.827 | 2 | <MRL | <MRL | 87.28 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
167 | 2022-07-20 21:44 | WO-22270-108-A-1 | U | 0 | 1.173 | 3.237 | 24.938 | 0 | 0.557 | 0.489 | 1.045 | 3.661 | 1 | <MRL | <MRL | <MRL | 24.938 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 250 | 2022-07-22 23:32 | 212001-d2 | U | 0 | 1.721 | 148.683 | 0 | 0 | 1.69 | 1.157 | 1.072 | 5.225 | 2 | <MRL | <MRL | 297.366 | <MRL | <MRL | <MRL | <MRL | <MRL | 10.45
168 | 2022-07-20 22:20 | WO-22270-109-A-1 | U | 0 | 1.642 | 1.92 | 36.818 | 0 | 0.617 | 1.347 | 0.961 | 4.195 | 1 | <MRL | <MRL | <MRL | 36.818 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 251 | 2022-07-23 00:08 | 212002-d2 | U | 0 | 1.091 | 139.298 | 0 | 0 | 1.087 | 0.534 | 2.366 | 2.786 | 2 | <MRL | <MRL | 278.596 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
169 | 2022-07-20 22:56 | WO-22270-110-A-1 | U | 0 | 0.734 | 2.75 | 35.931 | 0 | 1.446 | 1.286 | 0.98 | 3.898 | 1 | <MRL | <MRL | <MRL | 35.931 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 252 | 2022-07-23 00:44 | 212003-d2 | U | 0 | 1.6 | 211.612 | 1.952 | 0 | 1.623 | 0.82 | 1.398 | 3.554 | 2 | <MRL | <MRL | 423.224 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
170 | 2022-07-20 23:32 | AB-220715-01 LRB 12_INJ001 | K | 0 | 0.705 | 0.578 | 0 | 0 | 0.51 | 0.335 | 1.732 | 0.332 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 253 | 2022-07-23 01:20 | 212004-d2 | U | 0 | 1.209 | 223.747 | 0 | 0 | 0.549 | 0.459 | 2.394 | 2.996 | 2 | <MRL | <MRL | 447.494 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
171 | 2022-07-21 00:08 | AB-220715-01 LRB 12 | K | 0 | 0.946 | 0.319 | 0 | 0 | 0.517 | 0.321 | 1.105 | 0.346 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | [11 empty] | 255 | 2022-07-23 02:32 | 210451-d2-reprep-shaken | U | 0 | 1.04 | 0 | 0 | 0 | 3.26 | 0.453 | 0.892 | 114.95 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 229.9 | [11 empty] | Out of holding time, electively reanalyzed
172 | 2022-07-21 00:44 | WO-22270-111-A-1 | U | 0 | 1.267 | 1.185 | 72.024 | 0 | 0.756 | 0.505 | 0.886 | 3.531 | 1 | <MRL | <MRL | <MRL | 72.024 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  |  |  | 257 | 2022-07-23 03:44 | 212005-d2 | U | 0 | 0.883 | 448.566 | 0 | 0 | 1.172 | 1.36 | 0.835 | 2.703 | 2 | <MRL | <MRL | 897.132 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
173 | 2022-07-21 01:20 | WO-22270-112-A-1 | U | 0 | 1.418 | 2.563 | 72.186 | 0 | 0.638 | 0.855 | 1.181 | 3.652 | 1 | <MRL | <MRL | <MRL | 72.186 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 258 | 2022-07-23 04:20 | 212006-d2 | U | 0 | 0.83 | 490.756 | 1.676 | 0 | 1.077 | 0.561 | 2.218 | 3.409 | 2 | <MRL | <MRL | 981.512 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
174 | 2022-07-21 01:56 | WO-22270-113-A-1 | U | 0 | 1.485 | 1.551 | 94.178 | 0 | 0.999 | 0.372 | 2.63 | 3.505 | 1 | <MRL | <MRL | <MRL | 94.178 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 259 | 2022-07-23 04:56 | 212007-d2 | U | 0 | 1.345 | 652.269 | 1.572 | 0 | 1.491 | 0.645 | 0.901 | 2.969 | 2 | <MRL | <MRL | 1304.538 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
175 | 2022-07-21 02:32 | WO-22270-114-A-1 | U | 0 | 1.335 | 0 | 326.568 | 0 | 1.482 | 0.57 | 2.402 | 4.205 | 1 | <MRL | <MRL | <MRL | 326.568 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  |  |  | 260 | 2022-07-23 05:32 | 212008-d2 | U | 0 | 1.375 | 2324.998 | 3.302 | 0 | 1.675 | 2.072 | 1.104 | 2.746 | 2 | <MRL | <MRL | 4649.996 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
176 | 2022-07-21 03:08 | WO-22270-115-A-1 | U | 0 | 1.218 | 3.524 | 222.047 | 0 | 2.163 | 2.753 | 1.171 | 4.431 | 1 | <MRL | <MRL | <MRL | 222.047 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 261 | 2022-07-23 06:08 | 212009-d2 | U | 0 | 0.928 | 1553.208 | 1.554 | 0 | 0.839 | 0.439 | 1.139 | 2.174 | 2 | <MRL | <MRL | 3106.416 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
177 | 2022-07-21 03:44 | WO-22270-116-A-1 | U | 0 | 2.36 | 0 | 239.793 | 0 | 1.595 | 0.482 | 0.987 | 3.651 | 1 | <MRL | <MRL | <MRL | 239.793 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 262 | 2022-07-23 06:44 | 212010-d2 | U | 0 | 3.32 | 2159.757 | 3.578 | 0 | 0.633 | 0.586 | 1.438 | 2.838 | 2 | <MRL | <MRL | 4319.514 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
178 | 2022-07-21 04:20 | WO-22270-117-A-1 | U | 0 | 1.803 | 2.873 | 2.136 | 0 | 1.126 | 0.507 | 0.964 | 27.663 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 27.663 | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 263 | 2022-07-23 07:20 | 212011-d2 | U | 0 | 0.98 | 14.533 | 0 | 0 | 0.592 | 1.46 | 0.977 | 1.97 | 2 | <MRL | <MRL | 29.066 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
179 | 2022-07-21 04:56 | WO-22270-118-A-1 | U | 0 | 1.197 | 0 | 2.136 | 0 | 0.666 | 0.423 | 0.913 | 2.864 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 264 | 2022-07-23 07:56 | 212012-d2 | U | 0 | 1.21 | 11.442 | 0 | 0 | 1.784 | 0.663 | 2.858 | 2.336 | 2 | <MRL | <MRL | 22.884 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
180 | 2022-07-21 05:32 | WO-22270-119-A-1 | U | 0 | 1.762 | 0 | 1.872 | 0 | 1.779 | 0.726 | 2.894 | 4.091 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 265 | 2022-07-23 08:32 | 212013-d2 | U | 0 | 1.082 | 9.067 | 0 | 0 | 0.588 | 0.545 | 0.941 | 3.117 | 2 | <MRL | <MRL | 18.134 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
181 | 2022-07-21 06:08 | WO-22270-120-A-1 | U | 0 | 1.069 | 0 | 1.811 | 0 | 1.868 | 0.958 | 1.033 | 4.189 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  | IS4 |  |  |  |  | 266 | 2022-07-23 09:08 | 212014-d2 | U | 0 | 1.39 | 7.807 | 1.518 | 0 | 1.33 | 0.622 | 2.126 | 3.526 | 2 | <MRL | <MRL | 15.614 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
182 | 2022-07-21 06:44 | AB-220715-01 LFB-H 7 | Z | 172.757 | 171.667 | 189.674 | 192.398 | 156.023 | 192.64 | 195.667 | 158.717 | 190.922 | 1 | <MRL | 171.667 | 189.674 | 192.398 | 156.023 | 192.64 | 195.667 | 158.717 | 190.922 | No TrueV! | [12 empty] | 267 | 2022-07-23 09:44 | 210453-d2-reprep | U | 0 | 1.237 | 0 | 0 | 0 | 7.981 | 0.769 | 0.812 | 567.441 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | 15.962 | <MRL | <MRL | 1134.882 | [11 empty] | Out of holding time, electively reanalyzed
183 | 2022-07-21 07:20 | AB-220715-01 LRB 13_INJ001 | K | 0 | 0.708 | 0.574 | 1.13 | 0 | 0.579 | 0.447 | 1.31 | 0.522 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 274 | 2022-07-23 13:56 | 212015-d2 | U | 0 | 1.563 | 15.258 | 1.517 | 0 | 2.224 | 1.717 | 1.208 | 5.86 | 2 | <MRL | <MRL | 30.516 | <MRL | <MRL | <MRL | <MRL | <MRL | 11.72
184 | 2022-07-21 07:56 | AB-220715-01 LRB 13 | K | 0 | 0.858 | 0.337 | 0.956 | 0 | 0.599 | 0.354 | 1.057 | 0.374 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | [11 empty] | 275 | 2022-07-23 14:32 | 212016-d2 | U | 0 | 1.902 | 13.876 | 0 | 0 | 1.034 | 0.692 | 1.079 | 4.081 | 2 | <MRL | <MRL | 27.752 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
185 | 2022-07-21 08:32 | AB-220715-01 LD 6 | D | 0 | 1.889 | 0 | 1.984 | 0 | 1.785 | 0.814 | 0.93 | 3.868 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 276 | 2022-07-23 15:08 | 212017-d2 | U | 0 | 0.822 | 16.358 | 0 | 0 | 1.842 | 0.573 | 1.153 | 3.847 | 2 | <MRL | <MRL | 32.716 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
186 | 2022-07-21 09:08 | AB-220715-01 CCC-H 7 | Q | 369.407 | 335.708 | 385.941 | 375.749 | 303.514 | 387.221 | 384.409 | 317.825 | 372.392 | 1 | 369.407 | 335.708 | 385.941 | 375.749 | 303.514 | 387.221 | 384.409 | 317.825 | 372.392 | UL! |  |  |  |  | IS3 | [7 empty] | 277 | 2022-07-23 15:44 | 212018-d2 | U | 0 | 1.233 | 35.552 | 0 | 0 | 1.077 | 1.819 | 0.882 | 3.932 | 2 | <MRL | <MRL | 71.104 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
187 | 2022-07-21 09:44 | AB-220715-01 LFSM-H 6 | S | 0 | 178.901 | 202.898 | 189.642 | 145.053 | 200.076 | 195.386 | 146.784 | 201.023 | 1 | <MRL | 178.901 | 202.898 | 189.642 | 145.053 | 200.076 | 195.386 | 146.784 | 201.023 | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  |  |  |  | 278 | 2022-07-23 16:20 | 212019-d2 | U | 0 | 1.791 | 45.211 | 4.181 | 0 | 0.856 | 0.582 | 0.861 | 3.544 | 2 | <MRL | <MRL | 90.422 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
188 | 2022-07-21 10:20 | WO-22270-121-A-1 | U | 0 | 1.007 | 1.054 | 2.834 | 0 | 0 | 0.406 | 1.38 | 3.691 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 279 | 2022-07-23 16:56 | 210439-d2-reprep | U | 0 | 1.689 | 0 | 0 | 0 | 32.495 | 1.284 | 2.502 | 1038.324 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | 64.99 | <MRL | <MRL | 2076.648 | [11 empty] | Out of holding time, electively reanalyzed
189 | 2022-07-21 10:56 | WO-22270-122-A-1 | U | 0 | 2.476 | 0 | 4.153 | 0 | 1.193 | 0.714 | 0.971 | 3.841 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  | 280 | 2022-07-23 17:32 | 212020-d2 | U | 0 | 1.268 | 49.661 | 2.227 | 0 | 1.14 | 0.7 | 3.391 | 3.091 | 2 | <MRL | <MRL | 99.322 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
190 | 2022-07-21 11:32 | WO-22270-123-A-1 | U | 0 | 0.914 | 0 | 3.574 | 0 | 0.544 | 0.454 | 2.967 | 4.068 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 281 | 2022-07-23 18:08 | 212021-d2 | U | 0 | 1.112 | 190.082 | 1.457 | 0 | 1.973 | 0.67 | 1.227 | 4.366 | 2 | <MRL | <MRL | 380.164 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
191 | 2022-07-21 12:08 | WO-22270-124-A-1 | U | 0 | 1.17 | 0 | 8.903 | 0 | 3.316 | 0.583 | 1.115 | 4.517 | 1 | <MRL | <MRL | <MRL | 8.903 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 282 | 2022-07-23 18:44 | 212022-d2 | U | 0 | 1.364 | 166.615 | 1.564 | 0 | 4.744 | 0.578 | 0.929 | 7.36 | 2 | <MRL | <MRL | 333.23 | <MRL | <MRL | <MRL | <MRL | <MRL | 14.72
192 | 2022-07-21 12:44 | WO-22270-125-A-1 | U | 0 | 2.165 | 0 | 8.16 | 0 | 1.426 | 0.925 | 0.877 | 4.939 | 1 | <MRL | <MRL | <MRL | 8.16 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 283 | 2022-07-23 19:20 | 212023-d2 | U | 0 | 1.217 | 163.398 | 1.988 | 0 | 1.948 | 1.001 | 1.014 | 6.961 | 2 | <MRL | <MRL | 326.796 | <MRL | <MRL | <MRL | <MRL | <MRL | 13.922
193 | 2022-07-21 13:20 | WO-22270-126-A-1 | U | 0 | 2.014 | 0 | 8.852 | 0 | 2.024 | 0.878 | 0.952 | 3.937 | 1 | <MRL | <MRL | <MRL | 8.852 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  |  |  | 284 | 2022-07-23 19:56 | 212024-d2 | U | 0 | 1.053 | 157.924 | 1.455 | 0 | 1.469 | 1.21 | 0.88 | 7.471 | 2 | <MRL | <MRL | 315.848 | <MRL | <MRL | <MRL | <MRL | <MRL | 14.942
194 | 2022-07-21 13:56 | WO-22270-127-A-1 | U | 0 | 1.381 | 0 | 28.051 | 0 | 2.717 | 0.484 | 0.903 | 3.877 | 1 | <MRL | <MRL | <MRL | 28.051 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  |  |  | 287 | 2022-07-23 21:44 | 212030-d2 | U | 0 | 0.888 | 1495.148 | 2.292 | 0 | 1.261 | 0.658 | 2.261 | 7.764 | 2 | <MRL | <MRL | 2990.296 | <MRL | <MRL | <MRL | <MRL | <MRL | 15.528
195 | 2022-07-21 14:32 | WO-22270-128-A-1 | U | 0 | 2.627 | 0 | 33.692 | 0 | 1.561 | 0.52 | 1.558 | 3.491 | 1 | <MRL | <MRL | <MRL | 33.692 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  |  |  | 288 | 2022-07-23 22:20 | 212025-d2 | U | 0 | 0.892 | 590.481 | 1.491 | 0 | 1.588 | 1.144 | 2.233 | 7.844 | 2 | <MRL | <MRL | 1180.962 | <MRL | <MRL | <MRL | <MRL | <MRL | 15.688
196 | 2022-07-21 15:08 | WO-22270-129-A-1 | U | 0 | 0.913 | 1.124 | 38.125 | 0 | 2.706 | 0.837 | 0.979 | 4.154 | 1 | <MRL | <MRL | <MRL | 38.125 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 289 | 2022-07-23 22:56 | 212031-d2 | U | 0 | 1.04 | 1403.188 | 2.479 | 0 | 1.949 | 1.063 | 1.726 | 3.32 | 2 | <MRL | <MRL | 2806.376 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
197 | 2022-07-21 15:44 | WO-22270-130-A-1 | U | 0 | 1.542 | 0 | 39.499 | 0 | 21.233 | 1.533 | 0.829 | 5.842 | 1 | <MRL | <MRL | <MRL | 39.499 | <MRL | 21.233 | <MRL | <MRL | 5.842 | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS3 | IS4 |  |  |  |  | 290 | 2022-07-23 23:32 | 212026-d2 | U | 0 | 0.761 | 714.487 | 1.811 | 0 | 1.947 | 0.578 | 2.008 | 7.309 | 2 | <MRL | <MRL | 1428.974 | <MRL | <MRL | <MRL | <MRL | <MRL | 14.618
198 | 2022-07-21 16:20 | AB-220715-01 LRB 14_INJ001 | K | 0 | 0.835 | 0.484 | 1.109 | 0 | 0.587 | 0.321 | 1.409 | 0.356 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | [12 empty] | 291 | 2022-07-24 00:08 | 212032-d2 | U | 0 | 3.573 | 11190.616 | 3333.538 | 0 | 0.887 | 0.556 | 2.524 | 4.209 | 2 | <MRL | <MRL | 22381.232 | 6667.076 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! |  |  | UL | [6 empty] | R2
199 | 2022-07-21 16:56 | AB-220715-01 LRB 14 | K | 0 | 0.797 | 0.65 | 0 | 0 | 0.56 | 0.339 | 1.439 | 0.373 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 |  |  | IS4 |  |  | IS4 |  |  |  |  | 292 | 2022-07-24 00:44 | 212027-d2 | U | 0 | 0.826 | 629.943 | 1.486 | 0 | 0.916 | 0.646 | 1.053 | 8.219 | 2 | <MRL | <MRL | 1259.886 | <MRL | <MRL | <MRL | <MRL | <MRL | 16.438
200 | 2022-07-21 17:32 | WO-22270-131-A-1 | U | 0 | 1.28 | 0 | 142.268 | 0 | 1.455 | 0.796 | 0.743 | 3.517 | 1 | <MRL | <MRL | <MRL | 142.268 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  |  |  | 293 | 2022-07-24 01:20 | 212033-d2 | U | 0 | 2.963 | 12321.404 | 3521.747 | 0 | 1.283 | 0.678 | 2.154 | 5.232 | 2 | <MRL | <MRL | 24642.808 | 7043.494 | <MRL | <MRL | <MRL | <MRL | 10.464 | UL! |  |  | ul | [6 empty] | R2
201 | 2022-07-21 18:08 | WO-22270-132-A-1 | U | 0 | 1.683 | 0 | 112.938 | 0 | 3.691 | 0.725 | 0 | 3.637 | 1 | <MRL | <MRL | <MRL | 112.938 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 294 | 2022-07-24 01:56 | 210415-d2-reprep | U | 0 | 1.409 | 0 | 1.719 | 0 | 27.337 | 0.873 | 0.907 | 29.291 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | 54.674 | <MRL | <MRL | 58.582
202 | 2022-07-21 18:44 | WO-22270-133-A-1 | U | 0 | 1.435 | 0 | 83.818 | 0 | 2.816 | 2.182 | 3.056 | 4.59 | 1 | <MRL | <MRL | <MRL | 83.818 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  |  |  | 295 | 2022-07-24 02:32 | 212034-d2 | U | 0 | 2.631 | 11615.642 | 3344.616 | 0 | 1.155 | 1.058 | 1.4 | 3.552 | 2 | <MRL | <MRL | 23231.284 | 6689.232 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! |  |  | UL | [6 empty] | R2
203 | 2022-07-21 19:20 | WO-22270-134-A-1 | U | 0 | 2.327 | 0 | 2231.973 | 0 | 2.168 | 3.39 | 1.377 | 4.942 | 1 | <MRL | <MRL | <MRL | 2231.973 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  |  |  | 296 | 2022-07-24 03:08 | 212028-d2 | U | 0 | 1.068 | 1279.045 | 1.773 | 0 | 1.515 | 0.623 | 0.834 | 5.876 | 2 | <MRL | <MRL | 2558.09 | <MRL | <MRL | <MRL | <MRL | <MRL | 11.752
204 | 2022-07-21 19:56 | WO-22270-135-A-1 | U | 0 | 1.205 | 0 | 222.327 | 0 | 0.598 | 0.704 | 2.956 | 3.738 | 1 | <MRL | <MRL | <MRL | 222.327 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  |  |  | 297 | 2022-07-24 03:44 | 212029-d2 | U | 0 | 0.896 | 911.43 | 2.016 | 0 | 1.486 | 0.519 | 2.281 | 4.423 | 2 | <MRL | <MRL | 1822.86 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL
205 | 2022-07-21 20:32 | WO-22270-136-A-1 | U | 0 | 1.742 | 3.41 | 51.868 | 0 | 1.873 | 0.774 | 0.913 | 4.151 | 1 | <MRL | <MRL | <MRL | 51.868 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | [6 empty] | 302 | 2022-07-24 06:44 | 210405-d2-reprep | U | 0 | 2.387 | 37.659 | 9.482 | 0 | 2.417 | 1.101 | 1.251 | 38.635 | 2 | <MRL | <MRL | 75.318 | 18.964 | <MRL | <MRL | <MRL | <MRL | 77.27 | [11 empty] | Out of holding time, electively reanalyzed
206 | 2022-07-21 21:08 | WO-22270-137-A-1 | U | 0 | 1.387 | 11.283 | 1.438 | 0 | 1.092 | 0.447 | 0.983 | 2.975 | 1 | <MRL | <MRL | 11.283 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  | 304 | 2022-07-24 07:56 | 212035-d2 | U | 0 | 2.236 | 11785.487 | 3175.149 | 0 | 1.51 | 1.074 | 1.709 | 3.658 | 2 | <MRL | <MRL | 23570.974 | 6350.298 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! |  |  | UL | [6 empty] | R2
207 | 2022-07-21 21:44 | WO-22270-138-A-1 | U | 0 | 1.31 | 16.461 | 0 | 0 | 0.526 | 0.365 | 0.896 | 2.294 | 1 | <MRL | <MRL | 16.461 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  | 305 | 2022-07-24 08:32 | 212036-d2 | U | 0 | 2.206 | 12663.908 | 3534.884 | 0 | 0.69 | 0.541 | 1.114 | 2.071 | 2 | <MRL | <MRL | 25327.816 | 7069.768 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! |  |  | UL | [6 empty] | R2
208 | 2022-07-21 22:20 | WO-22270-139-A-1 | U | 0 | 1.429 | 15.189 | 0 | 0 | 0.633 | 0.538 | 0.876 | 4.91 | 1 | <MRL | <MRL | 15.189 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  | 306 | 2022-07-24 09:08 | 212037-d2 | U | 0 | 2.256 | 12265.672 | 3405.809 | 0 | 0.655 | 0.56 | 1.007 | 3.444 | 2 | <MRL | <MRL | 24531.344 | 6811.618 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! |  |  | UL | [6 empty] | R2
209 | 2022-07-21 22:56 | WO-22270-140-A-1 | U | 0 | 1.545 | 10.723 | 0 | 0 | 0.612 | 0.56 | 0.971 | 3.954 | 1 | <MRL | <MRL | 10.723 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 308 | 2022-07-24 10:20 | 212038-d2 | U | 0 | 2.686 | 11125.024 | 3483.05 | 0 | 3.116 | 0.626 | 1.336 | 1.971 | 2 | <MRL | <MRL | 22250.048 | 6966.1 | <MRL | <MRL | <MRL | <MRL | <MRL | UL!
210 | 2022-07-21 23:32 | AB-220715-01 LFB-H 8 | Z | 184.848 | 169.13 | 182.476 | 190.233 | 153.04 | 200.802 | 207.329 | 153.851 | 183.032 | 1 | <MRL | 169.13 | 182.476 | 190.233 | 153.04 | 200.802 | 207.329 | 153.851 | 183.032 | No TrueV! |  | IS3 |  |  | IS3 | [7 empty] | 309 | 2022-07-24 10:56 | 212039-d2 | U | 0 | 2.122 | 11343.019 | 3455.628 | 0 | 0.526 | 0.737 | 1.674 | 2.288 | 2 | <MRL | <MRL | 22686.038 | 6911.256 | <MRL | <MRL | <MRL | <MRL | <MRL | UL!
211 | 2022-07-22 00:08 | AB-220715-01 LRB 15_INJ001 | K | 0 | 1.064 | 0.909 | 1.636 | 1.4 | 1.407 | 1.026 | 1.582 | 1.466 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! |  | IS4 | [10 empty] | 310 | 2022-07-24 11:32 | 212040-d2 | U | 0 | 2.199 | 11381.093 | 3284.437 | 0 | 1.254 | 0.995 | 3.293 | 4.095 | 2 | <MRL | <MRL | 22762.186 | 6568.874 | <MRL | <MRL | <MRL | <MRL | <MRL | UL!
212 | 2022-07-22 00:44 | AB-220715-01 LRB 15 | K | 0 | 0.813 | 0.951 | 1.008 | 0 | 0.881 | 0.374 | 1.271 | 0.494 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | [10 empty] | 311 | 2022-07-24 12:08 | 212041-d2 | U | 0 | 2.467 | 10832.086 | 3592.031 | 0 | 2.265 | 1.019 | 1.242 | 5.134 | 2 | <MRL | <MRL | 21664.172 | 7184.062 | <MRL | <MRL | <MRL | <MRL | 10.268 | UL!
213 | 2022-07-22 01:20 | AB-220715-01 LD 7 | D | 0 | 1.639 | 7.487 | 0 | 0 | 1.55 | 0.653 | 1.203 | 4.949 | 1 | <MRL | <MRL | 7.487 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  | 312 | 2022-07-24 12:44 | 212042-d2 | U | 0 | 2.764 | 7278.211 | 2847.722 | 0 | 1.277 | 0.571 | 1.031 | 2.486 | 2 | <MRL | <MRL | 14556.422 | 5695.444 | <MRL | <MRL | <MRL | <MRL | <MRL
214 | 2022-07-22 01:56 | AB-220715-01 CCC-H 8 | Q | 380.715 | 336.433 | 385.036 | 378.446 | 316.872 | 374.878 | 379.099 | 325.053 | 375.361 | 1 | 380.715 | 336.433 | 385.036 | 378.446 | 316.872 | 374.878 | 379.099 | 325.053 | 375.361 | UL! |  | IS3 |  |  | IS3 |  |  | IS3 |  |  |  |  | 313 | 2022-07-24 13:20 | 212043-d2 | U | 0 | 2.036 | 7487.394 | 2860.632 | 0 | 2.189 | 0.812 | 6.395 | 6.272 | 2 | <MRL | <MRL | 14974.788 | 5721.264 | <MRL | <MRL | <MRL | 12.79 | 12.544
215 | 2022-07-22 02:32 | AB-220715-01 LFSM-H 7 | S | 0 | 176.03 | 220.478 | 188.25 | 153.039 | 196.063 | 196.726 | 150.528 | 201.943 | 1 | <MRL | 176.03 | 220.478 | 188.25 | 153.039 | 196.063 | 196.726 | 150.528 | 201.943 | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  |  |  |  | 314 | 2022-07-24 13:56 | 212044-d2 | U | 0 | 2.025 | 7969.456 | 2869.4 | 0 | 1.151 | 0.526 | 1.962 | 6.031 | 2 | <MRL | <MRL | 15938.912 | 5738.8 | <MRL | <MRL | <MRL | <MRL | 12.062
216 | 2022-07-22 03:08 | WO-22270-141-A-1 | U | 0 | 1.527 | 34.124 | 1.641 | 0 | 1.229 | 0.808 | 1.175 | 3.221 | 1 | <MRL | <MRL | 34.124 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 317 | 2022-07-24 15:44 | 212045-d2 | U | 0 | 2.3 | 7781.393 | 2748.222 | 0 | 0.993 | 0.51 | 1.99 | 3.187 | 2 | <MRL | <MRL | 15562.786 | 5496.444 | <MRL | <MRL | <MRL | <MRL | <MRL
217 | 2022-07-22 03:44 | WO-22270-142-A-1 | U | 0 | 1.463 | 33.507 | 0 | 0 | 1.132 | 0.486 | 1.051 | 2.606 | 1 | <MRL | <MRL | 33.507 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 318 | 2022-07-24 16:20 | 212046-d2 | U | 0 | 3.041 | 7105.418 | 2676.981 | 0 | 1.958 | 0.59 | 5.603 | 4.309 | 2 | <MRL | <MRL | 14210.836 | 5353.962 | <MRL | <MRL | <MRL | 11.206 | <MRL
218 | 2022-07-22 04:20 | WO-22270-143-A-1 | U | 0 | 2.015 | 26.498 | 0 | 0 | 0.601 | 0.608 | 0.984 | 3.584 | 1 | <MRL | <MRL | 26.498 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  | 319 | 2022-07-24 16:56 | 212047-d2 | U | 0 | 2.359 | 7173.014 | 2778.734 | 0 | 1.614 | 0.677 | 2.514 | 4.681 | 2 | <MRL | <MRL | 14346.028 | 5557.468 | <MRL | <MRL | <MRL | <MRL | <MRL
219 | 2022-07-22 04:56 | WO-22270-144-A-1 | U | 0 | 1.164 | 102.136 | 0 | 0 | 1.771 | 0.83 | 1.097 | 3.23 | 1 | <MRL | <MRL | 102.136 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 |  | IS4 |  |  |  |  | 320 | 2022-07-24 17:32 | 212048-d2 | U | 0 | 2.749 | 12554.794 | 3346.883 | 0 | 1.669 | 1.638 | 1.131 | 14.656 | 2 | <MRL | <MRL | 25109.588 | 6693.766 | <MRL | <MRL | <MRL | <MRL | 29.312 | UL!
220 | 2022-07-22 05:32 | WO-22270-145-A-1 | U | 0 | 0.689 | 88.027 | 0 | 0 | 1.335 | 0.525 | 0.844 | 2.649 | 1 | <MRL | <MRL | 88.027 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 321 | 2022-07-24 18:08 | 212049-d2 | U | 0 | 4.025 | 12151.767 | 3235.902 | 0 | 1.903 | 0.805 | 0.993 | 2.926 | 2 | <MRL | <MRL | 24303.534 | 6471.804 | <MRL | <MRL | <MRL | <MRL | <MRL | UL!
221 | 2022-07-22 06:08 | WO-22270-146-A-1 | U | 0 | 0.882 | 92.579 | 0 | 0 | 1.95 | 0.591 | 0.965 | 2.223 | 1 | <MRL | <MRL | 92.579 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 322 | 2022-07-24 18:44 | 212050-d2 | U | 0 | 3.068 | 7184.12 | 3348.427 | 0 | 2.498 | 1.048 | 0.951 | 8.092 | 2 | <MRL | <MRL | 14368.24 | 6696.854 | <MRL | <MRL | <MRL | <MRL | 16.184
222 | 2022-07-22 06:44 | WO-22270-147-A-1 | U | 0 | 1.916 | 324.339 | 0 | 0 | 2.685 | 0.489 | 1.445 | 3.744 | 1 | <MRL | <MRL | 324.339 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  | 323 | 2022-07-24 19:20 | 212051-d2 | U | 0 | 2.209 | 7307.135 | 3450.62 | 0 | 2.127 | 0.919 | 2.409 | 4.849 | 2 | <MRL | <MRL | 14614.27 | 6901.24 | <MRL | <MRL | <MRL | <MRL | <MRL
223 | 2022-07-22 07:20 | WO-22270-148-A-1 | U | 0 | 1.139 | 304.539 | 2.76 | 0 | 4.589 | 0.533 | 1.57 | 2.35 | 1 | <MRL | <MRL | 304.539 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 |  | IS4 |  |  |  |  | 324 | 2022-07-24 19:56 | 212052-d4 | U | 0 | 1.585 | 2491.878 | 1141.632 | 0 | 0.719 | 0.649 | 1.886 | 2.803 | 4 | <MRL | <MRL | 9967.512 | 4566.528 | <MRL | <MRL | <MRL | <MRL | <MRL
224 | 2022-07-22 07:56 | WO-22270-149-A-1 | U | 0 | 0.908 | 312.784 | 2.979 | 0 | 1.944 | 0.671 | 0.928 | 4.147 | 1 | <MRL | <MRL | 312.784 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 |  | IS4 |  |  |  |  | 325 | 2022-07-24 20:32 | 212053-d4 | U | 0 | 1.36 | 2772.695 | 1133.005 | 0 | 1.084 | 0.554 | 1.051 | 1.586 | 4 | <MRL | <MRL | 11090.78 | 4532.02 | <MRL | <MRL | <MRL | <MRL | <MRL
225 | 2022-07-22 08:32 | WO-22270-150-A-1 | U | 0 | 1.28 | 267.728 | 0 | 0 | 0.552 | 0.73 | 0.854 | 3.079 | 1 | <MRL | <MRL | 267.728 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 326 | 2022-07-24 21:08 | 212054-d4 | U | 0 | 1.044 | 2530.797 | 1166.087 | 0 | 0.644 | 0.508 | 1.058 | 2.469 | 4 | <MRL | <MRL | 10123.188 | 4664.348 | <MRL | <MRL | <MRL | <MRL | <MRL
226 | 2022-07-22 09:08 | AB-220715-01 LRB 16_INJ001 | K | 0 | 0.682 | 2.056 | 0 | 0 | 0.858 | 0.56 | 1.009 | 0.43 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d |  | IS4 | [7 empty] | 333 | 2022-07-25 01:20 | 212055-d4 | U | 0 | 1.511 | 2681.881 | 1085.099 | 0 | 0.693 | 0.525 | 1.66 | 1.787 | 4 | <MRL | <MRL | 10727.524 | 4340.396 | <MRL | <MRL | <MRL | <MRL | <MRL
227 | 2022-07-22 09:44 | AB-220715-01 LRB 16 | K | 0 | 0.881 | 1.823 | 0 | 0 | 0.865 | 0.378 | 0.899 | 0.429 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d |  | IS4 | [7 empty] | 334 | 2022-07-25 01:56 | 212056-d4 | U | 0 | 1.5 | 2794.788 | 1067.661 | 0 | 2.007 | 0.698 | 1.203 | 1.519 | 4 | <MRL | <MRL | 11179.152 | 4270.644 | <MRL | <MRL | <MRL | <MRL | <MRL
228 | 2022-07-22 10:20 | WO-22270-151-A-1 | U | 0 | 0.918 | 993.029 | 1.571 | 0 | 0.623 | 1.754 | 0.829 | 2.406 | 1 | <MRL | <MRL | 993.029 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 335 | 2022-07-25 02:32 | 212057-d4 | U | 0 | 0.92 | 2436.532 | 1051.989 | 0 | 1.524 | 0.721 | 2.83 | 1.186 | 4 | <MRL | <MRL | 9746.128 | 4207.956 | <MRL | <MRL | <MRL | <MRL | <MRL
229 | 2022-07-22 10:56 | WO-22270-152-A-1 | U | 0 | 1.351 | 1050.632 | 1.672 | 0 | 0.75 | 0.517 | 2.257 | 2.648 | 1 | <MRL | <MRL | 1050.632 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  |  |  |  | 336 | 2022-07-25 03:08 | 212058-d4 | U | 0 | 1.157 | 2567.789 | 1132.739 | 0 | 1.089 | 0.741 | 0.94 | 2.39 | 4 | <MRL | <MRL | 10271.156 | 4530.956 | <MRL | <MRL | <MRL | <MRL | <MRL
230 | 2022-07-22 11:32 | WO-22270-153-A-1 | U | 0 | 1.915 | 958.71 | 2.461 | 0 | 1.145 | 0.645 | 1.34 | 3.295 | 1 | <MRL | <MRL | 958.71 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  |  | 337 | 2022-07-25 03:44 | 212059-d4 | U | 0 | 0.964 | 2686.931 | 1066.892 | 0 | 1.694 | 0.6 | 0.968 | 2.002 | 4 | <MRL | <MRL | 10747.724 | 4267.568 | <MRL | <MRL | <MRL | <MRL | <MRL
231 | 2022-07-22 12:08 | WO-22270-154-A-1 | U | 0 | 0.893 | 2069.312 | 2.409 | 0 | 0.773 | 5.502 | 1.161 | 4.213 | 1 | <MRL | <MRL | 2069.312 | <MRL | <MRL | <MRL | 5.502 | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS3 |  |  |  |  | 338 | 2022-07-25 04:20 | 212060-d4 | U | 0 | 1.551 | 2806.729 | 1116.773 | 0 | 1.358 | 0.686 | 2.501 | 1.745 | 4 | <MRL | <MRL | 11226.916 | 4467.092 | <MRL | <MRL | <MRL | <MRL | <MRL
232 | 2022-07-22 12:44 | WO-22270-155-A-1 | U | 0 | 0.924 | 3164.557 | 2.974 | 0 | 0.664 | 0.599 | 0.936 | 3.675 | 1 | <MRL | <MRL | 3164.557 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4
233 | 2022-07-22 13:20 | WO-22270-156-A-1 | U | 0 | 1.656 | 2732.931 | 2.804 | 0 | 1.519 | 0.575 | 1.285 | 3.488 | 1 | <MRL | <MRL | 2732.931 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4
234 | 2022-07-22 13:56 | WO-22270-157-A-1 | U | 0 | 1.469 | 17.397 | 3.054 | 0 | 1.923 | 4.66 | 1.133 | 3.856 | 1 | <MRL | <MRL | 17.397 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4
235 | 2022-07-22 14:32 | WO-22270-158-A-1 | U | 0 | 1.107 | 7.677 | 1.143 | 0 | 1.129 | 1.068 | 2.294 | 2.164 | 1 | <MRL | <MRL | 7.677 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4
236 | 2022-07-22 15:08 | WO-22270-159-A-1 | U | 0 | 1.589 | 8.152 | 0 | 0 | 1.534 | 0.726 | 1.054 | 3.551 | 1 | <MRL | <MRL | 8.152 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS3 | IS4 | IS4
237 | 2022-07-22 15:44 | WO-22270-160-A-1 | U | 0 | 1.25 | 3.988 | 0 | 0 | 0.554 | 0.636 | 2.993 | 3.234 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | d | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
238 | 2022-07-22 16:20 | AB-220715-01 LFB-H 9 | Z | 225.666 | 160.929 | 180.037 | 195.058 | 144.71 | 199.266 | 209.15 | 147.606 | 201.143 | 1 | 225.666 | 160.929 | 180.037 | 195.058 | 144.71 | 199.266 | 209.15 | 147.606 | 201.143 | No TrueV! | IS3 | IS3 |  |  | IS3 |  |  | IS3
239 | 2022-07-22 16:56 | AB-220715-01 LRB 17_INJ001 | K | 0 | 1.121 | 0.812 | 1.817 | 1.497 | 1.693 | 1.235 | 1.555 | 1.538 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 |  |  | IS4
240 | 2022-07-22 17:32 | AB-220715-01 LRB 17 | K | 0 | 0.641 | 3.442 | 1.134 | 0 | 0.893 | 0.388 | 1.002 | 0.466 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4
241 | 2022-07-22 18:08 | AB-220715-01 LD 8 | D | 0 | 1.145 | 8.326 | 2.782 | 0 | 3.612 | 0.66 | 1.327 | 3.377 | 1 | <MRL | <MRL | 8.326 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
242 | 2022-07-22 18:44 | AB-220715-01 CCC-H 9 | Q | 361.131 | 341.763 | 369.546 | 391.441 | 304.674 | 386.79 | 391.626 | 311.608 | 363.614 | 1 | 361.131 | 341.763 | 369.546 | 391.441 | 304.674 | 386.79 | 391.626 | 311.608 | 363.614 | UL! |  | IS3 |  |  | IS3 |  |  | IS3
243 | 2022-07-22 19:20 | AB-220715-01 LFSM-H 8 | S | 0 | 178.667 | 214.18 | 184.277 | 144.849 | 202.262 | 199.472 | 151.908 | 187.394 | 1 | <MRL | 178.667 | 214.18 | 184.277 | 144.849 | 202.262 | 199.472 | 151.908 | 187.394 | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
244 | 2022-07-22 19:56 | WO-22270-161-A-1 | U | 0 | 1.741 | 12.151 | 1.612 | 0 | 0.627 | 1.19 | 1.224 | 2.606 | 1 | <MRL | <MRL | 12.151 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
245 | 2022-07-22 20:32 | WO-22270-162-A-1 | U | 0 | 1.197 | 15.364 | 0 | 0 | 0.958 | 0.54 | 1.038 | 3.833 | 1 | <MRL | <MRL | 15.364 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
246 | 2022-07-22 21:08 | WO-22270-163-A-1 | U | 0 | 1.477 | 14.745 | 0 | 0 | 1.607 | 0.599 | 1.276 | 2.784 | 1 | <MRL | <MRL | 14.745 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
247 | 2022-07-22 21:44 | WO-22270-164-A-1 | U | 0 | 1.355 | 41.544 | 1.282 | 0 | 1.813 | 1.221 | 2.595 | 3.054 | 1 | <MRL | <MRL | 41.544 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
248 | 2022-07-22 22:20 | WO-22270-165-A-1 | U | 0 | 1.596 | 41.894 | 0 | 0 | 0.662 | 0.495 | 0.961 | 3.212 | 1 | <MRL | <MRL | 41.894 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
249 | 2022-07-22 22:56 | WO-22270-166-A-1 | U | 0 | 1.124 | 43.64 | 0 | 0 | 0.702 | 0.432 | 1.05 | 3.827 | 1 | <MRL | <MRL | 43.64 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
250 | 2022-07-22 23:32 | WO-22270-167-A-1 | U | 0 | 1.721 | 148.683 | 0 | 0 | 1.69 | 1.157 | 1.072 | 5.225 | 1 | <MRL | <MRL | 148.683 | <MRL | <MRL | <MRL | <MRL | <MRL | 5.225 | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
251 | 2022-07-23 00:08 | WO-22270-168-A-1 | U | 0 | 1.091 | 139.298 | 0 | 0 | 1.087 | 0.534 | 2.366 | 2.786 | 1 | <MRL | <MRL | 139.298 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
252 | 2022-07-23 00:44 | WO-22270-169-A-1 | U | 0 | 1.6 | 211.612 | 1.952 | 0 | 1.623 | 0.82 | 1.398 | 3.554 | 1 | <MRL | <MRL | 211.612 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
253 | 2022-07-23 01:20 | WO-22270-170-A-1 | U | 0 | 1.209 | 223.747 | 0 | 0 | 0.549 | 0.459 | 2.394 | 2.996 | 1 | <MRL | <MRL | 223.747 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
254 | 2022-07-23 01:56 | AB-220715-01 LRB 18_INJ001 | K | 0 | 0.741 | 2.299 | 1.219 | 2.417 | 0.906 | 0.361 | 0.953 | 0.499 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4
255 | 2022-07-23 02:32 | WO-22209-62-A-1B | U | 0 | 1.04 | 0 | 0 | 0 | 3.26 | 0.453 | 0.892 | 114.95 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 114.95 | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  | Out of holding time, electively reanalyzed
256 | 2022-07-23 03:08 | AB-220715-01 LRB 18 | K | 0 | 0.778 | 1.268 | 1.199 | 0 | 0.945 | 0.363 | 0.933 | 0.455 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 |  |  | IS4
257 | 2022-07-23 03:44 | WO-22270-171-A-1 | U | 0 | 0.883 | 448.566 | 0 | 0 | 1.172 | 1.36 | 0.835 | 2.703 | 1 | <MRL | <MRL | 448.566 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
258 | 2022-07-23 04:20 | WO-22270-172-A-1 | U | 0 | 0.83 | 490.756 | 1.676 | 0 | 1.077 | 0.561 | 2.218 | 3.409 | 1 | <MRL | <MRL | 490.756 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
259 | 2022-07-23 04:56 | WO-22270-173-A-1 | U | 0 | 1.345 | 652.269 | 1.572 | 0 | 1.491 | 0.645 | 0.901 | 2.969 | 1 | <MRL | <MRL | 652.269 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
260 | 2022-07-23 05:32 | WO-22270-174-A-1 | U | 0 | 1.375 | 2324.998 | 3.302 | 0 | 1.675 | 2.072 | 1.104 | 2.746 | 1 | <MRL | <MRL | 2324.998 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
261 | 2022-07-23 06:08 | WO-22270-175-A-1 | U | 0 | 0.928 | 1553.208 | 1.554 | 0 | 0.839 | 0.439 | 1.139 | 2.174 | 1 | <MRL | <MRL | 1553.208 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
262 | 2022-07-23 06:44 | WO-22270-176-A-1 | U | 0 | 3.32 | 2159.757 | 3.578 | 0 | 0.633 | 0.586 | 1.438 | 2.838 | 1 | <MRL | <MRL | 2159.757 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
263 | 2022-07-23 07:20 | WO-22270-177-A-1 | U | 0 | 0.98 | 14.533 | 0 | 0 | 0.592 | 1.46 | 0.977 | 1.97 | 1 | <MRL | <MRL | 14.533 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 |  | IS4
264 | 2022-07-23 07:56 | WO-22270-178-A-1 | U | 0 | 1.21 | 11.442 | 0 | 0 | 1.784 | 0.663 | 2.858 | 2.336 | 1 | <MRL | <MRL | 11.442 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
265 | 2022-07-23 08:32 | WO-22270-179-A-1 | U | 0 | 1.082 | 9.067 | 0 | 0 | 0.588 | 0.545 | 0.941 | 3.117 | 1 | <MRL | <MRL | 9.067 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
266 | 2022-07-23 09:08 | WO-22270-180-A-1 | U | 0 | 1.39 | 7.807 | 1.518 | 0 | 1.33 | 0.622 | 2.126 | 3.526 | 1 | <MRL | <MRL | 7.807 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
267 | 2022-07-23 09:44 | WO-22209-64-A-1B | U | 0 | 1.237 | 0 | 0 | 0 | 7.981 | 0.769 | 0.812 | 567.441 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | 7.981 | <MRL | <MRL | 567.441 | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 |  | Out of holding time, electively reanalyzed
268 | 2022-07-23 10:20 | AB-220715-01 LFB-H 10 | Z | 148.473 | 165.749 | 181.985 | 193.018 | 154.851 | 200.351 | 196.216 | 164.645 | 192.246 | 1 | <MRL | 165.749 | 181.985 | 193.018 | 154.851 | 200.351 | 196.216 | 164.645 | 192.246 | No TrueV! |  | IS3 |  |  | IS3 |  |  | IS3
269 | 2022-07-23 10:56 | AB-220715-01 LRB 19_INJ001 | K | 0 | 1.233 | 2.874 | 1.801 | 2.675 | 1.349 | 0.997 | 1.411 | 1.205 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! |  | IS4
270 | 2022-07-23 11:32 | AB-220715-01 LRB 19 | K | 0 | 0.991 | 2.143 | 1.098 | 2.013 | 0.961 | 0.459 | 1.059 | 0.481 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 |  |  | IS4
271 | 2022-07-23 12:08 | AB-220715-01 LD 9 | D | 0 | 1.746 | 7.444 | 0 | 0 | 0.625 | 0.708 | 2.715 | 2.987 | 1 | <MRL | <MRL | 7.444 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
272 | 2022-07-23 12:44 | AB-220715-01 CCC-H 10 | Q | 368.792 | 339.369 | 386.392 | 388.865 | 312.752 | 398.474 | 383.256 | 345.56 | 383.207 | 1 | 368.792 | 339.369 | 386.392 | 388.865 | 312.752 | 398.474 | 383.256 | 345.56 | 383.207 | UL! |  | IS3 |  |  | IS3 | IS3 |  | IS3
273 | 2022-07-23 13:20 | AB-220715-01 LFSM-H 9 | S | 0 | 169.264 | 197.643 | 200.41 | 146.015 | 195.88 | 197.609 | 153.973 | 191.971 | 1 | <MRL | 169.264 | 197.643 | 200.41 | 146.015 | 195.88 | 197.609 | 153.973 | 191.971 | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
274 | 2022-07-23 13:56 | WO-22270-181-A-1 | U | 0 | 1.563 | 15.258 | 1.517 | 0 | 2.224 | 1.717 | 1.208 | 5.86 | 1 | <MRL | <MRL | 15.258 | <MRL | <MRL | <MRL | <MRL | <MRL | 5.86 | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
275 | 2022-07-23 14:32 | WO-22270-182-A-1 | U | 0 | 1.902 | 13.876 | 0 | 0 | 1.034 | 0.692 | 1.079 | 4.081 | 1 | <MRL | <MRL | 13.876 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
276 | 2022-07-23 15:08 | WO-22270-183-A-1 | U | 0 | 0.822 | 16.358 | 0 | 0 | 1.842 | 0.573 | 1.153 | 3.847 | 1 | <MRL | <MRL | 16.358 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
277 | 2022-07-23 15:44 | WO-22270-184-A-1 | U | 0 | 1.233 | 35.552 | 0 | 0 | 1.077 | 1.819 | 0.882 | 3.932 | 1 | <MRL | <MRL | 35.552 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
278 | 2022-07-23 16:20 | WO-22270-185-A-1 | U | 0 | 1.791 | 45.211 | 4.181 | 0 | 0.856 | 0.582 | 0.861 | 3.544 | 1 | <MRL | <MRL | 45.211 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4
279 | 2022-07-23 16:56 | WO-22209-50-A-1B | U | 0 | 1.689 | 0 | 0 | 0 | 32.495 | 1.284 | 2.502 | 1038.324 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | 32.495 | <MRL | <MRL | 1038.324 | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 |  | Out of holding time, electively reanalyzed
280 | 2022-07-23 17:32 | WO-22270-186-A-1 | U | 0 | 1.268 | 49.661 | 2.227 | 0 | 1.14 | 0.7 | 3.391 | 3.091 | 1 | <MRL | <MRL | 49.661 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
281 | 2022-07-23 18:08 | WO-22270-187-A-1 | U | 0 | 1.112 | 190.082 | 1.457 | 0 | 1.973 | 0.67 | 1.227 | 4.366 | 1 | <MRL | <MRL | 190.082 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4
282 | 2022-07-23 18:44 | WO-22270-188-A-1 | U | 0 | 1.364 | 166.615 | 1.564 | 0 | 4.744 | 0.578 | 0.929 | 7.36 | 1 | <MRL | <MRL | 166.615 | <MRL | <MRL | <MRL | <MRL | <MRL | 7.36 | UL! | IS3 | IS4 | IS4 | IS3 | IS4
283 | 2022-07-23 19:20 | WO-22270-189-A-1 | U | 0 | 1.217 | 163.398 | 1.988 | 0 | 1.948 | 1.001 | 1.014 | 6.961 | 1 | <MRL | <MRL | 163.398 | <MRL | <MRL | <MRL | <MRL | <MRL | 6.961 | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4
284 | 2022-07-23 19:56 | WO-22270-190-A-1 | U | 0 | 1.053 | 157.924 | 1.455 | 0 | 1.469 | 1.21 | 0.88 | 7.471 | 1 | <MRL | <MRL | 157.924 | <MRL | <MRL | <MRL | <MRL | <MRL | 7.471 | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 |  | IS4
285 | 2022-07-23 20:32 | AB-220715-01 LRB 20_INJ001 | K | 0 | 1.062 | 2.254 | 0 | 0 | 0.643 | 0.395 | 1.871 | 0.413 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! |  | IS4
286 | 2022-07-23 21:08 | AB-220715-01 LRB 20 | K | 0 | 0.784 | 1.802 | 0 | 2.185 | 0.678 | 0.65 | 0.947 | 0.406 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 |  |  | IS4 |  |  | IS4
287 | 2022-07-23 21:44 | WO-22270-196-A-1 | U | 0 | 0.888 | 1495.148 | 2.292 | 0 | 1.261 | 0.658 | 2.261 | 7.764 | 1 | <MRL | <MRL | 1495.148 | <MRL | <MRL | <MRL | <MRL | <MRL | 7.764 | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4
288 | 2022-07-23 22:20 | WO-22270-191-A-1 | U | 0 | 0.892 | 590.481 | 1.491 | 0 | 1.588 | 1.144 | 2.233 | 7.844 | 1 | <MRL | <MRL | 590.481 | <MRL | <MRL | <MRL | <MRL | <MRL | 7.844 | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4
289 | 2022-07-23 22:56 | WO-22270-197-A-1 | U | 0 | 1.04 | 1403.188 | 2.479 | 0 | 1.949 | 1.063 | 1.726 | 3.32 | 1 | <MRL | <MRL | 1403.188 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 |  | IS4
290 | 2022-07-23 23:32 | WO-22270-192-A-1 | U | 0 | 0.761 | 714.487 | 1.811 | 0 | 1.947 | 0.578 | 2.008 | 7.309 | 1 | <MRL | <MRL | 714.487 | <MRL | <MRL | <MRL | <MRL | <MRL | 7.309 | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4
291 | 2022-07-24 00:08 | WO-22270-198-A-1 | U | 0 | 3.573 | 11190.616 | 3333.538 | 0 | 0.887 | 0.556 | 2.524 | 4.209 | 1 | <MRL | <MRL | 11190.616 | 3333.538 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | UL | IS3 | IS3 | IS4 |  | IS4 |  | R2
292 | 2022-07-24 00:44 | WO-22270-193-A-1 | U | 0 | 0.826 | 629.943 | 1.486 | 0 | 0.916 | 0.646 | 1.053 | 8.219 | 1 | <MRL | <MRL | 629.943 | <MRL | <MRL | <MRL | <MRL | <MRL | 8.219 | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 |  | IS4
293 | 2022-07-24 01:20 | WO-22270-199-A-1 | U | 0 | 2.963 | 12321.404 | 3521.747 | 0 | 1.283 | 0.678 | 2.154 | 5.232 | 1 | <MRL | <MRL | 12321.404 | 3521.747 | <MRL | <MRL | <MRL | <MRL | 5.232 | UL! | IS3 | IS4 | UL | IS3 | IS3 | IS4 |  | IS4 |  | R2
294 | 2022-07-24 01:56 | WO-22209-26-A-1B | U | 0 | 1.409 | 0 | 1.719 | 0 | 27.337 | 0.873 | 0.907 | 29.291 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | 27.337 | <MRL | <MRL | 29.291 | UL! | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4
295 | 2022-07-24 02:32 | WO-22270-200-A-1 | U | 0 | 2.631 | 11615.642 | 3344.616 | 0 | 1.155 | 1.058 | 1.4 | 3.552 | 1 | <MRL | <MRL | 11615.642 | 3344.616 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | UL | IS3 | IS3 | IS4 | IS4 | IS4 |  | R2
296 | 2022-07-24 03:08 | WO-22270-194-A-1 | U | 0 | 1.068 | 1279.045 | 1.773 | 0 | 1.515 | 0.623 | 0.834 | 5.876 | 1 | <MRL | <MRL | 1279.045 | <MRL | <MRL | <MRL | <MRL | <MRL | 5.876 | UL! | IS3 | IS4 | IS4 | IS3 | IS4
297 | 2022-07-24 03:44 | WO-22270-195-A-1 | U | 0 | 0.896 | 911.43 | 2.016 | 0 | 1.486 | 0.519 | 2.281 | 4.423 | 1 | <MRL | <MRL | 911.43 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 |  | IS4
298 | 2022-07-24 04:20 | AB-220715-01 LFB-H 11 | Z | 187.116 | 167.433 | 231.801 | 198.179 | 151.857 | 190.277 | 199.559 | 160.42 | 192.084 | 1 | <MRL | 167.433 | 231.801 | 198.179 | 151.857 | 190.277 | 199.559 | 160.42 | 192.084 | No TrueV! |  | IS3 |  |  | IS3 |  |  | IS3 | IS3
299 | 2022-07-24 04:56 | AB-220715-01 LRB 21 | K | 0 | 0.881 | 1.799 | 1.323 | 0 | 0.894 | 0.413 | 2.293 | 0.455 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 |  |  | IS4
300 | 2022-07-24 05:32 | AB-220715-01 LD 10 | D | 0 | 1.452 | 5472.988 | 2201.609 | 0 | 1.37 | 0.723 | 2.393 | 2.957 | 1 | <MRL | <MRL | 5472.988 | 2201.609 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
301 | 2022-07-24 06:08 | AB-220715-01 CCC-H 11 | Q | 343.684 | 351.242 | 422.132 | 372.456 | 319.013 | 389.571 | 397.692 | 330.388 | 407.738 | 1 | 343.684 | 351.242 | 422.132 | 372.456 | 319.013 | 389.571 | 397.692 | 330.388 | 407.738 | UL! |  | IS3 |  |  | IS3 |  |  | IS3
302 | 2022-07-24 06:44 | WO-22209-16-A-1B | U | 0 | 2.387 | 37.659 | 9.482 | 0 | 2.417 | 1.101 | 1.251 | 38.635 | 1 | <MRL | <MRL | 37.659 | 9.482 | <MRL | <MRL | <MRL | <MRL | 38.635 | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 |  |  | Out of holding time, electively reanalyzed
303 | 2022-07-24 07:20 | AB-220715-01 LFSM-H 10 | S | 0 | 167.184 | 5413.034 | 2486.335 | 149.14 | 188.316 | 197.927 | 147.507 | 187.19 | 1 | <MRL | 167.184 | 5413.034 | 2486.335 | 149.14 | 188.316 | 197.927 | 147.507 | 187.19 | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
304 | 2022-07-24 07:56 | WO-22270-201-A-1 | U | 0 | 2.236 | 11785.487 | 3175.149 | 0 | 1.51 | 1.074 | 1.709 | 3.658 | 1 | <MRL | <MRL | 11785.487 | 3175.149 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | UL | IS3 | IS3 | IS4 | IS4 | IS4 |  | R2
305 | 2022-07-24 08:32 | WO-22270-202-A-1 | U | 0 | 2.206 | 12663.908 | 3534.884 | 0 | 0.69 | 0.541 | 1.114 | 2.071 | 1 | <MRL | <MRL | 12663.908 | 3534.884 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | UL | IS3 | IS3 | IS4 |  | IS4 |  | R2
306 | 2022-07-24 09:08 | WO-22270-203-A-1 | U | 0 | 2.256 | 12265.672 | 3405.809 | 0 | 0.655 | 0.56 | 1.007 | 3.444 | 1 | <MRL | <MRL | 12265.672 | 3405.809 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | UL | IS3 | IS3 | IS4 |  | IS4 |  | R2
307 | 2022-07-24 09:44 | AB-220715-01 LRB 21_INJ001 | K | 0 | 1.046 | 2.938 | 2.423 | 0 | 1.379 | 0.881 | 2.29 | 0.736 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! |  | IS4 | B
308 | 2022-07-24 10:20 | WO-22270-204-A-1 | U | 0 | 2.686 | 11125.024 | 3483.05 | 0 | 3.116 | 0.626 | 1.336 | 1.971 | 1 | <MRL | <MRL | 11125.024 | 3483.05 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4
309 | 2022-07-24 10:56 | WO-22270-205-A-1 | U | 0 | 2.122 | 11343.019 | 3455.628 | 0 | 0.526 | 0.737 | 1.674 | 2.288 | 1 | <MRL | <MRL | 11343.019 | 3455.628 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4
310 | 2022-07-24 11:32 | WO-22270-206-A-1 | U | 0 | 2.199 | 11381.093 | 3284.437 | 0 | 1.254 | 0.995 | 3.293 | 4.095 | 1 | <MRL | <MRL | 11381.093 | 3284.437 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4
311 | 2022-07-24 12:08 | WO-22270-207-A-1 | U | 0 | 2.467 | 10832.086 | 3592.031 | 0 | 2.265 | 1.019 | 1.242 | 5.134 | 1 | <MRL | <MRL | 10832.086 | 3592.031 | <MRL | <MRL | <MRL | <MRL | 5.134 | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 | IS4
312 | 2022-07-24 12:44 | WO-22270-208-A-1 | U | 0 | 2.764 | 7278.211 | 2847.722 | 0 | 1.277 | 0.571 | 1.031 | 2.486 | 1 | <MRL | <MRL | 7278.211 | 2847.722 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4
313 | 2022-07-24 13:20 | WO-22270-209-A-1 | U | 0 | 2.036 | 7487.394 | 2860.632 | 0 | 2.189 | 0.812 | 6.395 | 6.272 | 1 | <MRL | <MRL | 7487.394 | 2860.632 | <MRL | <MRL | <MRL | 6.395 | 6.272 | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
314 | 2022-07-24 13:56 | WO-22270-210-A-1 | U | 0 | 2.025 | 7969.456 | 2869.4 | 0 | 1.151 | 0.526 | 1.962 | 6.031 | 1 | <MRL | <MRL | 7969.456 | 2869.4 | <MRL | <MRL | <MRL | <MRL | 6.031 | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
315 | 2022-07-24 14:32 | AB-220715-01 LRB 22_INJ001 | K | 0 | 1.076 | 22.399 | 8.449 | 0 | 0.813 | 0.737 | 1.786 | 0.462 | 1 | <MRL | <MRL | 22.399 | 8.449 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! |  | IS4 |  |  | IS3
316 | 2022-07-24 15:08 | AB-220715-01 LRB 22 | K | 0 | 0.799 | 6.275 | 3.437 | 0 | 1.067 | 0.701 | 2.425 | 0.496 | 1 | <MRL | <MRL | 6.275 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 |  | IS4 |  |  | IS4
317 | 2022-07-24 15:44 | WO-22270-211-A-1 | U | 0 | 2.3 | 7781.393 | 2748.222 | 0 | 0.993 | 0.51 | 1.99 | 3.187 | 1 | <MRL | <MRL | 7781.393 | 2748.222 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
318 | 2022-07-24 16:20 | WO-22270-212-A-1 | U | 0 | 3.041 | 7105.418 | 2676.981 | 0 | 1.958 | 0.59 | 5.603 | 4.309 | 1 | <MRL | <MRL | 7105.418 | 2676.981 | <MRL | <MRL | <MRL | 5.603 | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
319 | 2022-07-24 16:56 | WO-22270-213-A-1 | U | 0 | 2.359 | 7173.014 | 2778.734 | 0 | 1.614 | 0.677 | 2.514 | 4.681 | 1 | <MRL | <MRL | 7173.014 | 2778.734 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
320 | 2022-07-24 17:32 | WO-22270-214-A-1 | U | 0 | 2.749 | 12554.794 | 3346.883 | 0 | 1.669 | 1.638 | 1.131 | 14.656 | 1 | <MRL | <MRL | 12554.794 | 3346.883 | <MRL | <MRL | <MRL | <MRL | 14.656 | UL! | IS3 | IS4 | IS4 | IS3 | IS3
321 | 2022-07-24 18:08 | WO-22270-215-A-1 | U | 0 | 4.025 | 12151.767 | 3235.902 | 0 | 1.903 | 0.805 | 0.993 | 2.926 | 1 | <MRL | <MRL | 12151.767 | 3235.902 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
322 | 2022-07-24 18:44 | WO-22270-216-A-1 | U | 0 | 3.068 | 7184.12 | 3348.427 | 0 | 2.498 | 1.048 | 0.951 | 8.092 | 1 | <MRL | <MRL | 7184.12 | 3348.427 | <MRL | <MRL | <MRL | <MRL | 8.092 | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
323 | 2022-07-24 19:20 | WO-22270-217-A-1 | U | 0 | 2.209 | 7307.135 | 3450.62 | 0 | 2.127 | 0.919 | 2.409 | 4.849 | 1 | <MRL | <MRL | 7307.135 | 3450.62 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
324 | 2022-07-24 19:56 | WO-22270-218-A-1 | U | 0 | 1.585 | 2491.878 | 1141.632 | 0 | 0.719 | 0.649 | 1.886 | 2.803 | 1 | <MRL | <MRL | 2491.878 | 1141.632 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
325 | 2022-07-24 20:32 | WO-22270-219-A-1 | U | 0 | 1.36 | 2772.695 | 1133.005 | 0 | 1.084 | 0.554 | 1.051 | 1.586 | 1 | <MRL | <MRL | 2772.695 | 1133.005 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
326 | 2022-07-24 21:08 | WO-22270-220-A-1 | U | 0 | 1.044 | 2530.797 | 1166.087 | 0 | 0.644 | 0.508 | 1.058 | 2.469 | 1 | <MRL | <MRL | 2530.797 | 1166.087 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
327 | 2022-07-24 21:44 | AB-220715-01 LFB-H 12 | Z | 144.503 | 165.072 | 207.993 | 182.251 | 145.325 | 186.411 | 192.991 | 152.715 | 184.98 | 1 | <MRL | 165.072 | 207.993 | 182.251 | 145.325 | 186.411 | 192.991 | 152.715 | 184.98 | No TrueV! |  | IS3 |  |  | IS3 |  |  | IS3
328 | 2022-07-24 22:20 | AB-220715-01 LRB 23_INJ001 | K | 0 | 1.162 | 1.639 | 2.016 | 1.448 | 1.26 | 1.078 | 1.597 | 1.255 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! |  | IS4 | b |  | IS4
329 | 2022-07-24 22:56 | AB-220715-01 LRB 23 | K | 0 | 0.726 | 6.134 | 2.845 | 0 | 0.779 | 0.425 | 2.333 | 0.472 | 1 | <MRL | <MRL | 6.134 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | b |  | IS4 |  |  | IS4
330 | 2022-07-24 23:32 | AB-220715-01 LD 11 | D | 0 | 1.08 | 2421.487 | 1091.661 | 4.165 | 1.364 | 0.555 | 0.907 | 1.575 | 1 | <MRL | <MRL | 2421.487 | 1091.661 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
331 | 2022-07-25 00:08 | AB-220715-01 CCC-H 12 | Q | 337.382 | 330.585 | 418.533 | 384.909 | 310.509 | 386.724 | 390.787 | 319.683 | 381.938 | 1 | 337.382 | 330.585 | 418.533 | 384.909 | 310.509 | 386.724 | 390.787 | 319.683 | 381.938 | UL! |  | IS3 |  |  | IS3 |  |  | IS3
332 | 2022-07-25 00:44 | AB-220715-01 LFSM-H 11 | S | 0 | 165.797 | 2977.339 | 1328.5 | 144.818 | 194.494 | 194.899 | 151.821 | 194.326 | 1 | <MRL | 165.797 | 2977.339 | 1328.5 | 144.818 | 194.494 | 194.899 | 151.821 | 194.326 | No TrueV! | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
333 | 2022-07-25 01:20 | WO-22270-221-A-1 | U | 0 | 1.511 | 2681.881 | 1085.099 | 0 | 0.693 | 0.525 | 1.66 | 1.787 | 1 | <MRL | <MRL | 2681.881 | 1085.099 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
334 | 2022-07-25 01:56 | WO-22270-222-A-1 | U | 0 | 1.5 | 2794.788 | 1067.661 | 0 | 2.007 | 0.698 | 1.203 | 1.519 | 1 | <MRL | <MRL | 2794.788 | 1067.661 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
335 | 2022-07-25 02:32 | WO-22270-223-A-1 | U | 0 | 0.92 | 2436.532 | 1051.989 | 0 | 1.524 | 0.721 | 2.83 | 1.186 | 1 | <MRL | <MRL | 2436.532 | 1051.989 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
336 | 2022-07-25 03:08 | WO-22270-224-A-1 | U | 0 | 1.157 | 2567.789 | 1132.739 | 0 | 1.089 | 0.741 | 0.94 | 2.39 | 1 | <MRL | <MRL | 2567.789 | 1132.739 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
337 | 2022-07-25 03:44 | WO-22270-225-A-1 | U | 0 | 0.964 | 2686.931 | 1066.892 | 0 | 1.694 | 0.6 | 0.968 | 2.002 | 1 | <MRL | <MRL | 2686.931 | 1066.892 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
338 | 2022-07-25 04:20 | WO-22270-226-A-1 | U | 0 | 1.551 | 2806.729 | 1116.773 | 0 | 1.358 | 0.686 | 2.501 | 1.745 | 1 | <MRL | <MRL | 2806.729 | 1116.773 | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
339 | 2022-07-25 04:56 | AB-220715-01 LRB 24_INJ001 | K | 0 | 0.766 | 8.854 | 4.418 | 2.309 | 0.64 | 0.348 | 0.875 | 0.375 | 1 | <MRL | <MRL | 8.854 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! |  | IS4 | b | b
340 | 2022-07-25 05:32 | AB-220715-01 LRB 24 | K | 0 | 0.849 | 4.17 | 2.162 | 2.368 | 0.668 | 0.373 | 2.338 | 0.408 | 1 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | UL! | IS3 | IS4 | b | IS4 | IS4 |  |  | IS4
341 | 2022-07-25 06:08 | AB-220715-01 CCC-H 13 | Q | 332.365 | 341.506 | 376.546 | 374.334 | 303.614 | 388.839 | 386.732 | 323.323 | 387.108 | 1 | 332.365 | 341.506 | 376.546 | 374.334 | 303.614 | 388.839 | 386.732 | 323.323 | 387.108 | UL! |  | IS3 |  |  | IS3 |  |  | IS3
342 | 2022-07-25 06:44 | AB-220715-01 QCS 2 | Q | 161.011 | 171.1 | 199.036 | 193.156 | 158.561 | 196.615 | 209.797 | 162.451 | 191.063 | 1 | <MRL | 171.1 | 199.036 | 193.156 | 158.561 | 196.615 | 209.797 | 162.451 | 191.063 | UL! | [7 empty] | IS3
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## Internal Standard Recoveries
| MPFBA | MPFBS | MPHFPO-DA | MPFHxA | MPFHxS | MPFOA | MPFNA | MPFOS | MPFDA | MPFBA | MPFBS | MPHFPO-DA | MPFHxA | MPFHxS | MPFOA | MPFNA | MPFOS | MPFDA | MPFBA | MPFBS | MPHFPO-DA | MPFHxA | MPFHxS | MPFOA | MPFNA | MPFOS | MPFDA
AB-220715-01 Cal Std 1 1 | 96.56% | 131.1% | 131.7% | 130.9% | 141.5% | 143.5% | 134.6% | 126.5% | 128.4% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 Cal Std 2 1 | 55.79% | 100.4% | 93.3% | 97.73% | 109.5% | 109.2% | 110.1% | 108% | 107.8% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 Cal Std 3 1 | 96.91% | 115.4% | 111.5% | 116.6% | 129.5% | 127% | 124.8% | 121.5% | 125.4% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 Cal Std 4 1 | 116.6% | 111.5% | 118.8% | 114.3% | 123.9% | 124.5% | 127.3% | 124% | 129.6% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 Cal Std 5 1 | 97.85% | 111.7% | 108.1% | 109.2% | 115.5% | 117.6% | 117.5% | 115.4% | 117.8% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 Cal Std 6 1 | 100% | 99.5% | 99.95% | 100% | 110% | 108% | 111.6% | 109.9% | 112.9% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 Cal Std 7 1 | 94.79% | 92.98% | 91.55% | 88.96% | 88.59% | 86.62% | 97.74% | 100% | 100% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 Cal Std 8 1 | 107.3% | 104.3% | 113.4% | 110.2% | 100% | 100% | 100% | 95.88% | 93.24% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 Cal Std 9 1 | 107.2% | 100% | 108.7% | 107.1% | 89.13% | 84.81% | 84.65% | 84% | 77.63% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 Cal Std 10 1 | 114.3% | 87.84% | 100% | 89.99% | 62.96% | 60.9% | 64.31% | 61.24% | 54.38% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 Cal Std 11 1 | 102.9% | 66.75% | 87.36% | 72.41% | 42.41% | 42.47% | 44.1% | 42.77% | 36.34% | FALSE | FALSE | FALSE | FALSE | IS3 | IS3 | IS3 | IS3 | IS3 |  |  |  |  | IS3 | IS3 | IS3 | IS3 | IS3
AB-220715-01 Cal Std 12 1 | 109.4% | 50.74% | 65.05% | 49.96% | 26.38% | 26.86% | 28.81% | 29.03% | 23.15% | FALSE | FALSE | FALSE | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  |  |  | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
AB-220715-01 Cal Std 13 1 | 85.67% | 31.19% | 39.86% | 28.76% | 14.72% | 13.01% | 14.99% | 17.13% | 11.66% | FALSE | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
AB-220715-01 LFB-H 1 | 115.1% | 110% | 118.7% | 113.6% | 99.09% | 100.1% | 100.3% | 90.46% | 92.61% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 1_INJ001 | 89.1% | 117.7% | 121.3% | 121.5% | 119.8% | 125.1% | 119.8% | 112.4% | 121.3% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 1 | 47.91% | 77.72% | 78.59% | 82.12% | 96.98% | 100.4% | 99.11% | 95.02% | 100.6% | IS3 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | IS3
AB-220715-01 CCC-H 1 | 99.03% | 94.03% | 110.6% | 102% | 81.44% | 81.09% | 81.54% | 76.29% | 75.03% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 IBC 1 | 108.5% | 106.7% | 119% | 113.2% | 97.71% | 98.65% | 101.1% | 91.26% | 101.9% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 QCS 1 | 104.5% | 105.8% | 116.2% | 108.8% | 101.4% | 100.6% | 101.8% | 90.78% | 94.74% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
WO-22270-01-A-1 | 13.04% | 16.86% | 18.55% | 20.54% | 34.49% | 37.47% | 53.16% | 67.4% | 67.27% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-02-A-1 | 14.62% | 20.69% | 22.78% | 24.95% | 39.09% | 43.44% | 60.75% | 78.13% | 76.37% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-03-A-1 | 16.77% | 20.84% | 22.39% | 25.32% | 38.8% | 43.38% | 62.76% | 78.51% | 77.56% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-04-A-1 | 12.9% | 18.7% | 21.49% | 22.22% | 36.9% | 40.32% | 57.05% | 73.67% | 71.81% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-05-A-1 | 14.31% | 22.46% | 24.63% | 26.44% | 40.08% | 42.88% | 63.88% | 78.83% | 77.81% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-06-A-1 | 17.31% | 18.99% | 20.18% | 22.51% | 37.22% | 41.1% | 60.21% | 73.12% | 76.04% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-07-A-1 | 13.66% | 19.81% | 23.45% | 24.6% | 37.88% | 41.58% | 60.42% | 73.36% | 76% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-08-A-1 | 9.625% | 15.29% | 17.58% | 18.25% | 32.67% | 36.08% | 52.58% | 64.99% | 66.96% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-09-A-1 | 14.16% | 22.03% | 25.05% | 26.71% | 38.52% | 43.55% | 64.64% | 78.33% | 81.2% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-10-A-1 | 11.95% | 21.07% | 23.27% | 24.23% | 38.57% | 42.58% | 62.6% | 76.97% | 79.57% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 LRB 2_INJ001 | 87.09% | 104.8% | 109.4% | 110.8% | 113.5% | 118.7% | 116.7% | 109.7% | 117.7% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 2 | 99.71% | 105.2% | 111.8% | 110.4% | 112.7% | 119.3% | 117.1% | 105.7% | 116.2% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
WO-22270-11-A-1 | 12.46% | 21.26% | 23.61% | 23.88% | 36.96% | 41.07% | 61.78% | 74.14% | 79.08% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-12-A-1 | 22.7% | 19.58% | 21.58% | 23.84% | 35.82% | 39.46% | 59.61% | 71.99% | 74.8% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-13-A-1 | 18.29% | 14.29% | 16.56% | 17.53% | 30.63% | 35.13% | 51.61% | 62.99% | 66.56% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-14-A-1 | 14.48% | 19.61% | 23.79% | 24.02% | 36.37% | 42.09% | 61.75% | 72.14% | 78.88% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-15-A-1 | 11.38% | 18.83% | 21.37% | 22.92% | 35.93% | 40.19% | 60.18% | 74.27% | 74.59% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-16-A-1 | 14.3% | 22.13% | 25.57% | 27.5% | 38.26% | 43.58% | 64.65% | 76.81% | 82.75% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-17-A-1 | 13.83% | 18.57% | 19.39% | 22.29% | 33.23% | 38.99% | 55.68% | 66.28% | 73.33% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-18-A-1 | 11.99% | 18.3% | 19.55% | 21.73% | 32.78% | 37.62% | 56.87% | 66.51% | 74.52% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-19-A-1 | 11.97% | 18.86% | 19.97% | 23.04% | 33.33% | 38.72% | 58.59% | 68.63% | 76.77% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-20-A-1 | 14.77% | 16.42% | 20.63% | 22.52% | 30.67% | 38.5% | 57.62% | 61.31% | 72.85% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 LFB-H 2 | 102.6% | 96.03% | 108.2% | 103.8% | 92.08% | 98.56% | 99.56% | 84.31% | 91.11% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 3_INJ001 | 116.1% | 103.2% | 113.4% | 117.1% | 111.8% | 126.6% | 123.6% | 108.2% | 123.2% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 3 | 58.18% | 102.8% | 114.1% | 113.4% | 106.4% | 114.9% | 112.8% | 99.4% | 110.8% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LD 1 | 14.76% | 20.08% | 24.41% | 25.33% | 35.55% | 40.08% | 60.77% | 70.93% | 77.61% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 CCC-H 2 | 79.54% | 81.89% | 100.9% | 97.55% | 73.21% | 81.09% | 84.2% | 71.33% | 71.96% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LFSM-H 1 | 18.44% | 18.92% | 22.55% | 22.53% | 31.19% | 36.13% | 52.21% | 60.45% | 64.03% | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | FALSE | FALSE | FALSE | IS3 | IS4 | IS3 | IS3 | IS3 | IS3
WO-22270-21-A-1 | 10.39% | 17.06% | 20.38% | 20.76% | 31.53% | 38.01% | 54.51% | 63.44% | 70.01% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-22-A-1 | 15.87% | 20.37% | 23.51% | 24.83% | 36.16% | 41.38% | 62.36% | 70.08% | 78.45% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-23-A-1 | 15.92% | 19.99% | 22.8% | 24.59% | 36.12% | 41.41% | 60.06% | 68.21% | 76.28% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-24-A-1 | 12.89% | 17.93% | 20.69% | 22.15% | 32.83% | 38.32% | 57.65% | 64.02% | 72.1% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-25-A-1 | 13.91% | 17.39% | 22.15% | 24.19% | 35.18% | 40.67% | 60.66% | 69.34% | 77.88% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-26-A-1 | 11.08% | 14.22% | 17.44% | 18.46% | 30.52% | 35.37% | 53.79% | 60.6% | 68.55% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-27-A-1 | 16.91% | 18.58% | 24.19% | 24.77% | 33.11% | 40.68% | 61.37% | 66.35% | 77.94% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-28-A-1 | 11.64% | 15.57% | 21.29% | 22.59% | 32.98% | 39.17% | 59.49% | 67.77% | 74.29% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-29-A-1 | 14.35% | 19.3% | 24.61% | 26.8% | 36.39% | 44.05% | 67.99% | 74.26% | 80.35% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-30-A-1 | 12.77% | 17.89% | 22.92% | 24.39% | 34.9% | 42.11% | 63.29% | 71.5% | 74.71% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 LRB 4_INJ001 | 109.2% | 89.58% | 106.8% | 109.1% | 106.6% | 120.8% | 120.1% | 104.7% | 118.4% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 4 | 63.64% | 92.03% | 113.5% | 108.7% | 103.3% | 114.1% | 113.3% | 94% | 112% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
WO-22270-31-A-1 | 13.26% | 18.72% | 21.01% | 25.7% | 35.26% | 40.54% | 61.23% | 69.61% | 79.78% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-32-A-1 | 7.495% | 7.385% | 9.877% | 10.42% | 21.74% | 25.22% | 38.6% | 42.2% | 49.41% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-33-A-1 | 11.89% | 12.31% | 14.49% | 17.94% | 28.87% | 34.55% | 51% | 56.92% | 66.14% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-34-A-1 | 7.789% | 11.08% | 14.82% | 16.11% | 26.28% | 31.67% | 48.6% | 53.07% | 61.25% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-35-A-1 | 11.88% | 15.45% | 21.81% | 23.02% | 32.14% | 39.26% | 58.46% | 66.72% | 74.44% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-36-A-1 | 14.56% | 23.47% | 28.89% | 32.44% | 39.7% | 46.84% | 70.94% | 79.23% | 88.13% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-37-A-1 | 11.27% | 14.12% | 20.2% | 20.69% | 29.9% | 35.63% | 55.44% | 61.13% | 70.28% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-38-A-1 | 11.18% | 14.63% | 21.7% | 21.29% | 29.98% | 37.1% | 55.22% | 63.7% | 75.11% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-39-A-1 | 12.33% | 14.25% | 20.34% | 20.89% | 29.43% | 35.38% | 54.54% | 59.55% | 70.12% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-40-A-1 | 12.53% | 13.84% | 19.64% | 20.35% | 29.54% | 36.32% | 54.06% | 60.3% | 70.34% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 LFB-H 3 | 92.96% | 76.5% | 97.23% | 100.4% | 83.98% | 93.37% | 98.29% | 80.28% | 89.84% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 5_INJ001 | 109.5% | 81.27% | 112.2% | 111.8% | 104.3% | 120.6% | 120% | 101.5% | 119.8% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 5 | 66.08% | 81.45% | 101.4% | 105.2% | 98.09% | 110.4% | 107.8% | 91.19% | 107.4% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LD 2 | 11.84% | 14.36% | 18.57% | 20.23% | 29.72% | 35.66% | 52.69% | 58.26% | 70.78% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 CCC-H 3 | 88.29% | 69.04% | 95.68% | 92.69% | 69.16% | 78.45% | 82.97% | 65.92% | 72.38% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LFSM-H 2 | 13.06% | 14.72% | 22.21% | 22.36% | 28.58% | 33.34% | 47.12% | 49.69% | 57.59% | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | FALSE | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
WO-22270-41-A-1 | 14.09% | 13.19% | 18.14% | 19.11% | 28.25% | 35.67% | 55.62% | 58.43% | 69.33% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-42-A-1 | 11.52% | 16.04% | 19.86% | 23.76% | 31.49% | 38.12% | 57.15% | 63.14% | 73.97% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-43-A-1 | 11.9% | 14.85% | 20.07% | 21.72% | 30.77% | 35.53% | 53.2% | 59.45% | 70.18% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-44-A-1 | 16.88% | 16.41% | 22.09% | 24.04% | 32.95% | 38.6% | 59.9% | 63.4% | 77.02% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-45-A-1 | 13.25% | 18.03% | 23.63% | 25.45% | 35.2% | 41.86% | 62.86% | 68.41% | 77.02% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-52-A-1 | 7.788% | 7.589% | 11.11% | 11.16% | 22.55% | 28.71% | 43.62% | 45.21% | 54.43% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-46-A-1 | 12.86% | 16.53% | 23.59% | 23.9% | 32.87% | 39.79% | 59.82% | 63.9% | 78.55% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-47-A-1 | 19.06% | 13.72% | 19.62% | 20.02% | 30.77% | 36.94% | 58.23% | 61.5% | 71.82% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-53-A-1 | 14.45% | 17.53% | 24.91% | 26.32% | 34.75% | 44.34% | 70.8% | 74.17% | 89.43% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-48-A-1 | 11.02% | 15.93% | 22.36% | 24.35% | 33.18% | 40.71% | 62.37% | 67.1% | 79.56% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-54-A-1 | 18.56% | 26.14% | 35.4% | 38.3% | 54.17% | 69.36% | 107.3% | 104.1% | 118.1% | IS3 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4
WO-22270-49-A-1 | 12.17% | 13.7% | 18.05% | 19.01% | 30.61% | 37.11% | 55.32% | 58.75% | 72.56% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-55-A-1 | 12.61% | 10.1% | 14.56% | 13.44% | 20.82% | 25.29% | 32.74% | 28.36% | 36.09% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-50-A-1 | 15.03% | 17.04% | 22.99% | 25.28% | 31.8% | 38.71% | 59.66% | 61.95% | 75.36% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-56-A-1 | 13.24% | 16.7% | 21.7% | 23.49% | 31.08% | 38.26% | 59.24% | 63.32% | 76.99% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-57-A-1 | 19.44% | 21.94% | 28.15% | 31.3% | 42.74% | 55.41% | 90.88% | 86.82% | 106.5% | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4
AB-220715-01 LRB 6_INJ001 | 108.7% | 81.03% | 104.3% | 102.7% | 100% | 116.2% | 115.1% | 94.69% | 113.9% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 6 | 59.05% | 82.02% | 105.4% | 104.7% | 95.16% | 112.3% | 113.8% | 93.09% | 107.9% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
WO-22270-58-A-1 | 6.308% | 7.103% | 10.48% | 11.64% | 20.16% | 27.56% | 45.3% | 46.09% | 57% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-51-A-1 | 12.4% | 19.65% | 25.04% | 25.24% | 35.03% | 42.9% | 67.21% | 71.44% | 85.71% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-59-A-1 | 13.27% | 15.81% | 22.73% | 23.44% | 29.33% | 38.25% | 61.56% | 62.86% | 74.95% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-60-A-1 | 11.2% | 12.74% | 17.11% | 19.9% | 29.4% | 37.17% | 59.14% | 60.54% | 76% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 LFB-H 4 | 95.34% | 72.02% | 99.92% | 100.9% | 79.17% | 92.78% | 96.2% | 74.9% | 87.74% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 7_INJ001 | 106.7% | 80.61% | 109.7% | 107% | 92.84% | 109.9% | 105.7% | 92.31% | 108% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 7 | 48.99% | 73.12% | 95.18% | 97.41% | 78.16% | 89.78% | 84.81% | 74.33% | 85% | IS3 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | IS3
AB-220715-01 LD 3 | 14.33% | 15.64% | 20.82% | 22.08% | 27.71% | 33.94% | 54.37% | 60.3% | 70.41% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 CCC-H 4 | 67.61% | 60.16% | 84.79% | 79.08% | 58.02% | 63.08% | 64.12% | 59.11% | 59.62% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LFSM-H 3 | 12.66% | 15.03% | 22.02% | 24.13% | 31.94% | 35.31% | 54.15% | 63.37% | 64% | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | FALSE | FALSE | FALSE | IS3 | IS4 | IS3 | IS3 | IS3 | IS3
WO-22270-61-A-1 | 12.81% | 13.93% | 21.06% | 19.99% | 29.79% | 35.59% | 55.03% | 65.01% | 76.21% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-62-A-1 | 11.78% | 12.49% | 19.4% | 17.73% | 26.76% | 33.35% | 44.04% | 50.85% | 68.22% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-63-A-1 | 9.378% | 8.547% | 12.38% | 13.49% | 21.37% | 26.5% | 39.16% | 43.52% | 53.75% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-64-A-1 | 8.423% | 7.461% | 10.8% | 11.28% | 19.95% | 23.6% | 35.25% | 39.3% | 49.17% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-65-A-1 | 16.12% | 18.39% | 26.02% | 28.23% | 35.26% | 43.98% | 66.9% | 75.44% | 86.39% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-66-A-1 | 14.37% | 13.71% | 20.54% | 22.01% | 25.72% | 32.92% | 49.54% | 55.76% | 67.39% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-67-A-1 | 10.6% | 11.59% | 16.35% | 17.17% | 24.4% | 28.75% | 44.08% | 48.12% | 61.33% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3
WO-22270-68-A-1 | 10.67% | 9.618% | 14.37% | 15.66% | 22.8% | 27.05% | 41.53% | 47.22% | 58.06% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-69-A-1 | 6.649% | 4.929% | 7.696% | 8.594% | 15.68% | 19.64% | 29.54% | 32.35% | 40.45% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-70-A-1 | 8.246% | 8.858% | 13.09% | 13.35% | 21.38% | 26.43% | 38.81% | 44.09% | 54.79% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 LRB 8_INJ001 | 124.9% | 78.86% | 119% | 112% | 91.58% | 108.6% | 102.1% | 90.21% | 101.2% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 8 | 55.52% | 71.4% | 95.08% | 95.7% | 79.04% | 89.27% | 87.05% | 77.1% | 82.57% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
WO-22270-71-A-1 | 14.1% | 12.28% | 19.43% | 19.24% | 26.15% | 32.28% | 49.09% | 54.31% | 66.64% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-72-A-1 | 11.99% | 13.5% | 19.24% | 19.64% | 25.68% | 31.68% | 48.76% | 53.46% | 65.75% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-73-A-1 | 9.659% | 10.88% | 15.41% | 16.1% | 23.36% | 29.21% | 44.32% | 47.83% | 60.26% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-74-A-1 | 11.25% | 10.73% | 15.4% | 16.23% | 22.72% | 27.59% | 43.89% | 47.13% | 60.19% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-75-A-1 | 13.37% | 16.56% | 22.18% | 23.95% | 28.23% | 34.61% | 51.8% | 56.46% | 69.27% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-76-A-1 | 7.179% | 5.368% | 7.738% | 8.387% | 15.74% | 18.9% | 28.89% | 28.92% | 37.86% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-77-A-1 | 14.83% | 16.87% | 24.53% | 24.23% | 27.53% | 36.15% | 54.36% | 57.96% | 73.35% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-78-A-1 | 17.58% | 18.68% | 26.82% | 26.4% | 29.46% | 36.5% | 56.56% | 58.25% | 75.69% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-79-A-1 | 18.08% | 17.85% | 25.47% | 25.93% | 28.73% | 36.2% | 54.91% | 60.74% | 74.53% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-80-A-1 | 17.88% | 20.63% | 28.28% | 27.83% | 29.78% | 36.64% | 57.59% | 60.39% | 76.69% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 LFB-H 5 | 79.85% | 61.85% | 84.29% | 79.65% | 66.34% | 69.52% | 68.81% | 60.94% | 64.84% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 9_INJ001 | 110.6% | 73.32% | 96.18% | 94.69% | 85.53% | 95.22% | 91.75% | 78.57% | 95.67% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 9 | 76.83% | 70.79% | 94.66% | 92.17% | 75.07% | 87.97% | 84.81% | 72.01% | 84.38% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LD 4 | 13.91% | 15.63% | 22.61% | 22.75% | 25.54% | 33.74% | 50.76% | 54.4% | 67.12% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 CCC-H 5 | 98.43% | 63.58% | 91.86% | 85.49% | 56.46% | 65.35% | 63.53% | 55.41% | 59.32% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LFSM-H 4 | 19.25% | 17.31% | 25.47% | 25.28% | 26.93% | 32.54% | 47.4% | 52.28% | 59.25% | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | FALSE | FALSE | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3
WO-22270-81-A-1 | 10.04% | 10.77% | 15.5% | 15.91% | 22.23% | 27.95% | 42.99% | 45.47% | 58.45% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-82-A-1 | 16.48% | 17.25% | 24.58% | 25.4% | 31.02% | 39.67% | 62.46% | 63.81% | 79.02% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-83-A-1 | 16.41% | 17.99% | 25.9% | 26.8% | 29.97% | 38.05% | 58.63% | 62.5% | 78.15% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-84-A-1 | 16.61% | 15.89% | 25.07% | 24.51% | 28.52% | 36.15% | 56.93% | 59.73% | 73.51% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-85-A-1 | 16.98% | 17.8% | 25.06% | 24.76% | 28.81% | 37.07% | 58.03% | 61.69% | 76.84% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-86-A-1 | 16.8% | 16.53% | 23.71% | 24.04% | 29.12% | 38.57% | 59.7% | 62.8% | 75.64% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-87-A-1 | 16.29% | 17.67% | 23.69% | 25.25% | 28.61% | 39.65% | 61.15% | 65% | 72.28% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-88-A-1 | 14.74% | 15.84% | 22.8% | 22.5% | 28.17% | 35.05% | 55.85% | 59.25% | 75% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-89-A-1 | 11.65% | 9.964% | 12.92% | 13.36% | 21.82% | 27.62% | 42.49% | 44.66% | 58.49% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4
WO-22270-90-A-1 | 11.45% | 11.4% | 16.43% | 16.66% | 22.3% | 28.87% | 38.57% | 40.74% | 58.39% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4
AB-220715-01 LRB 10_INJ001 | 81.02% | 65.22% | 80.63% | 87.64% | 75.82% | 89.31% | 80.66% | 64.18% | 84.8% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 10 | 29.24% | 44.11% | 57.76% | 58.9% | 53.85% | 66.12% | 67.07% | 58.09% | 66.11% | IS3 | IS4 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | IS3 | IS4
WO-22270-91-A-1 | 17.62% | 17.46% | 23.19% | 23.76% | 28.36% | 36.71% | 56.27% | 59.12% | 73.85% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-92-A-1 | 18.37% | 18.51% | 24.42% | 26.43% | 28.37% | 37.61% | 62.2% | 64.08% | 78.52% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-93-A-1 | 20.45% | 19.61% | 27.3% | 27.4% | 30.59% | 39.54% | 61.15% | 63.27% | 78.9% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-94-A-1 | 16.71% | 14.81% | 21.93% | 21.16% | 27.17% | 36.94% | 60.45% | 59.04% | 71.89% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-95-A-1 | 12.21% | 12.33% | 17.57% | 17.96% | 23.4% | 31.01% | 47.06% | 50.1% | 66.44% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4
WO-22270-96-A-1 | 20.42% | 18.48% | 26.68% | 25.83% | 28.42% | 36.93% | 57.42% | 59.99% | 76.13% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-97-A-1 | 19.38% | 18.38% | 25.43% | 26.78% | 28.23% | 36.32% | 57.18% | 59.57% | 76.02% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-98-A-1 | 17.35% | 18.82% | 24.87% | 25.55% | 28.5% | 35.72% | 56.64% | 56.59% | 73.89% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-99-A-1 | 9.912% | 5.832% | 8.423% | 9.114% | 15.67% | 20.16% | 29.71% | 26.47% | 40.05% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-100-A-1 | 10.99% | 10.75% | 14.88% | 16.03% | 21.47% | 27.87% | 42.54% | 42.68% | 57.13% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 LFB-H 6 | 63.1% | 51.63% | 72.69% | 71.08% | 57.97% | 67.46% | 72.07% | 59.33% | 69.13% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 11_INJ001 | 121.6% | 77% | 103.3% | 101.2% | 81.07% | 98.27% | 96.38% | 75.54% | 94.02% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 11 | 83.91% | 74.59% | 95.06% | 100.1% | 76.12% | 94.97% | 90.68% | 74.44% | 89.17% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LD 5 | 7.742% | 5.795% | 8.257% | 8.589% | 15.82% | 19.9% | 30.3% | 29.34% | 40.31% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 CCC-H 6 | 83.1% | 54.07% | 78.39% | 75.02% | 52.9% | 63.73% | 66.65% | 56.97% | 60.66% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LFSM-H 5 | 10.25% | 9.39% | 13.6% | 13.97% | 19.37% | 23.11% | 35.14% | 36.27% | 45.89% | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
WO-22270-101-A-1 | 20.37% | 18.79% | 25.96% | 27.78% | 31.88% | 40.72% | 66.01% | 66.76% | 84.72% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-102-A-1 | 18.93% | 18.09% | 24.28% | 26.08% | 27.48% | 36.59% | 56.1% | 57.35% | 76.29% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-103-A-1 | 18.06% | 22.63% | 29.87% | 32.55% | 43.31% | 58.9% | 90.8% | 84.93% | 105.2% | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-104-A-1 | 11.29% | 12.81% | 18.76% | 18.86% | 23.84% | 31.11% | 48.63% | 50.94% | 66.55% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4
WO-22270-105-A-1 | 11.63% | 13.1% | 18.03% | 18.88% | 25.35% | 33.27% | 51.57% | 50.57% | 66.15% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-106-A-1 | 14.83% | 14.83% | 22.75% | 24.04% | 26.32% | 37% | 58.36% | 56.75% | 73.53% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-107-A-1 | 19.85% | 19.38% | 26.37% | 27.56% | 29.5% | 39.38% | 61.44% | 60.51% | 79.17% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-108-A-1 | 17.26% | 16.11% | 23.83% | 25.06% | 28.05% | 38.42% | 61.93% | 58.62% | 70.98% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-109-A-1 | 14.33% | 15.45% | 22.99% | 22.68% | 26.27% | 35.28% | 55.78% | 54.04% | 73.89% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-110-A-1 | 14.42% | 14.63% | 21.05% | 22.76% | 24.28% | 34.66% | 53.78% | 54.99% | 73.01% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
AB-220715-01 LRB 12_INJ001 | 87.28% | 56.67% | 79.32% | 82.48% | 65.26% | 83.07% | 87.49% | 69.83% | 87.72% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 12 | 38.84% | 57.14% | 74.71% | 78.53% | 61.28% | 76.01% | 78.89% | 65.38% | 79.78% | IS3 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | IS3
WO-22270-111-A-1 | 14.15% | 13.96% | 19.79% | 19.98% | 23.08% | 31.41% | 48.56% | 48.45% | 66.14% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4
WO-22270-112-A-1 | 13.41% | 14.72% | 22.66% | 23.01% | 25.04% | 36.06% | 57.56% | 54.59% | 70.89% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-113-A-1 | 17.35% | 16.38% | 24.22% | 26.28% | 25.49% | 35.19% | 52.32% | 50.63% | 70.11% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-114-A-1 | 12.09% | 11.7% | 16.05% | 16.18% | 22.37% | 30.29% | 47.52% | 47.35% | 65.74% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4
WO-22270-115-A-1 | 16.08% | 16.27% | 23.72% | 24.14% | 26% | 38.37% | 59.13% | 57.03% | 67.7% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-116-A-1 | 17.34% | 17.81% | 22.94% | 24.25% | 26.81% | 36.62% | 61.36% | 60.36% | 75.79% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-117-A-1 | 15.16% | 15.67% | 21.18% | 22.98% | 24.05% | 32.64% | 49.25% | 47.76% | 58.49% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-118-A-1 | 15.11% | 15.12% | 20.64% | 22.07% | 25.1% | 33.31% | 54.69% | 51.7% | 71.32% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-119-A-1 | 10.65% | 11.43% | 15.74% | 16.72% | 22.02% | 27.41% | 43.47% | 40.37% | 59.6% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-120-A-1 | 15.13% | 14.41% | 19.57% | 22.23% | 23.74% | 32.21% | 51.01% | 49.23% | 63.84% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  | IS4
AB-220715-01 LFB-H 7 | 82.69% | 51.48% | 74.43% | 73.66% | 51.22% | 64.61% | 70.59% | 57.05% | 61.74% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 13_INJ001 | 81.89% | 59.25% | 81.76% | 81.82% | 62.8% | 80.04% | 82.52% | 66.91% | 81.34% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 13 | 28.78% | 52.67% | 67.58% | 73.55% | 53.83% | 64.88% | 66.1% | 51.01% | 63.69% | IS3 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | IS3
AB-220715-01 LD 6 | 10.94% | 10.62% | 15.53% | 16.75% | 19.78% | 27.3% | 42.95% | 39.97% | 59.4% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 CCC-H 7 | 90.52% | 52.45% | 78.26% | 77.81% | 48.34% | 60.5% | 64.94% | 52.77% | 54.28% | FALSE | FALSE | FALSE | FALSE | IS3 | FALSE | FALSE | FALSE | FALSE |  |  |  |  | IS3
AB-220715-01 LFSM-H 6 | 12.46% | 12.53% | 18.6% | 19.78% | 20.8% | 28.36% | 45.33% | 45.08% | 56.95% | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | FALSE | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
WO-22270-121-A-1 | 19.42% | 19.66% | 25.9% | 27.7% | 25.2% | 34.46% | 49.98% | 49.34% | 60.51% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-122-A-1 | 8.341% | 6.685% | 10.32% | 10.83% | 16.43% | 22.1% | 31.47% | 29.36% | 43.75% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-123-A-1 | 15.22% | 16.39% | 20.04% | 22.75% | 22.24% | 33.7% | 55.05% | 51.45% | 67.2% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-124-A-1 | 18.26% | 16.98% | 24.54% | 25.27% | 26.54% | 36.01% | 56.08% | 53.52% | 65.43% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-125-A-1 | 19.12% | 16.44% | 21% | 24.78% | 25.66% | 38.15% | 54.7% | 52.23% | 61.65% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-126-A-1 | 12.42% | 10.25% | 14.85% | 16.76% | 20.18% | 28.57% | 43.78% | 41.03% | 59.1% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4
WO-22270-127-A-1 | 15.98% | 12.74% | 19.29% | 19.86% | 22.39% | 30.49% | 42.56% | 38.66% | 55.27% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4
WO-22270-128-A-1 | 12.94% | 12.72% | 17.43% | 19.89% | 22.89% | 30.55% | 47.26% | 43.95% | 63.84% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4
WO-22270-129-A-1 | 15.32% | 16.39% | 21.31% | 23.67% | 23.74% | 34.64% | 55.93% | 53.57% | 74.22% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-130-A-1 | 11.1% | 7.986% | 13.89% | 13.23% | 17.86% | 27.5% | 43.74% | 37.99% | 55.6% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS3 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS3 | IS4
AB-220715-01 LRB 14_INJ001 | 77.5% | 52.89% | 72.85% | 75.88% | 56.18% | 74.89% | 79.16% | 59.67% | 78.75% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
AB-220715-01 LRB 14 | 27.33% | 38.38% | 50.42% | 52.86% | 44.22% | 58.63% | 63.77% | 47.47% | 61.85% | IS3 | IS4 | FALSE | FALSE | IS4 | FALSE | FALSE | IS4 | FALSE | IS3 | IS4 |  |  | IS4 |  |  | IS4
WO-22270-131-A-1 | 11.77% | 10.82% | 15.55% | 16.44% | 20.9% | 28.36% | 45.2% | 40.19% | 56% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4
WO-22270-132-A-1 | 17.14% | 15.75% | 22.61% | 24.32% | 24.83% | 37.31% | 59.54% | 53.26% | 69% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-133-A-1 | 11.28% | 10.52% | 14.54% | 17.07% | 20.46% | 29.02% | 47.19% | 43.97% | 60.39% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4
WO-22270-134-A-1 | 12.7% | 11.09% | 17.13% | 15.99% | 21.5% | 32.37% | 48.75% | 45.43% | 61.56% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4
WO-22270-135-A-1 | 10.24% | 8.684% | 13.53% | 13.35% | 18.78% | 26.86% | 44.61% | 38.87% | 57.85% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4
WO-22270-136-A-1 | 6.294% | 10.61% | 15.2% | 18.78% | 19.62% | 37% | 65.65% | 58.24% | 70.67% | IS3 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-137-A-1 | 10.54% | 11.9% | 18.16% | 20.93% | 18.12% | 31.22% | 46.09% | 34.9% | 47.74% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | d
WO-22270-138-A-1 | 6.098% | 8.836% | 14.76% | 16.53% | 15.82% | 27.75% | 46.62% | 37.75% | 42.78% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | d
WO-22270-139-A-1 | 5.309% | 4.679% | 8.257% | 10.26% | 13.3% | 20.67% | 32.53% | 27.49% | 44.96% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | d
WO-22270-140-A-1 | 9.934% | 9.155% | 14.6% | 15.74% | 17.61% | 27.17% | 41.5% | 36.03% | 54.96% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  | d
AB-220715-01 LFB-H 8 | 83.31% | 45.79% | 66% | 64.83% | 44.16% | 56.44% | 62.53% | 54.32% | 66.09% | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE | FALSE | FALSE | FALSE |  | IS3 |  |  | IS3
AB-220715-01 LRB 15_INJ001 | 52.9% | 45.86% | 63.37% | 73.72% | 58.06% | 79.34% | 79.35% | 63.65% | 82.56% | FALSE | IS4 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE |  | IS4
AB-220715-01 LRB 15 | 36.96% | 45.92% | 68.94% | 69.18% | 57.77% | 75.38% | 75.22% | 60.86% | 85.24% | IS3 | IS4 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | IS3 | IS4
AB-220715-01 LD 7 | 10.16% | 6.762% | 12.45% | 11.76% | 16.82% | 23.22% | 36.47% | 32.26% | 47.56% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 CCC-H 8 | 65.25% | 42.82% | 65% | 65.95% | 40.14% | 50.35% | 54.84% | 44.86% | 50.14% | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE |  | IS3 |  |  | IS3 |  |  | IS3
AB-220715-01 LFSM-H 7 | 13.72% | 10.77% | 16.49% | 19.74% | 18.78% | 27.56% | 40.23% | 38.04% | 50.5% | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | FALSE | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
WO-22270-141-A-1 | 13.68% | 12.75% | 19.53% | 21.68% | 21.76% | 31.4% | 48.73% | 46.18% | 63.2% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  | d
WO-22270-142-A-1 | 9.647% | 11% | 17.49% | 17.35% | 19.08% | 30.31% | 46.43% | 40.58% | 60.96% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  | d
WO-22270-143-A-1 | 8.576% | 5.42% | 8.759% | 9.442% | 13.76% | 21.41% | 32.36% | 29.22% | 44.92% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | d
WO-22270-144-A-1 | 16.56% | 12.7% | 19.36% | 21.8% | 21.78% | 34.36% | 51.09% | 44.52% | 65.06% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 |  | IS4 |  | d
WO-22270-145-A-1 | 17.18% | 12.53% | 21.14% | 20.43% | 20.78% | 30.88% | 48.63% | 43.03% | 67.67% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  | d
WO-22270-146-A-1 | 14.08% | 7.948% | 13.16% | 14.61% | 16.46% | 25.49% | 39.33% | 32.97% | 52.1% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  | d
WO-22270-147-A-1 | 7.916% | 4.174% | 7.433% | 7.812% | 12.89% | 20.38% | 33.71% | 28.03% | 37.01% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | d
WO-22270-148-A-1 | 15.35% | 14.45% | 21.43% | 24.6% | 21.62% | 36.05% | 54.06% | 44.62% | 68.04% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 |  | IS4 |  | d
WO-22270-149-A-1 | 14.85% | 14.85% | 19.94% | 23.4% | 22.58% | 34.44% | 50.5% | 45.23% | 67.08% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 |  | IS4 |  | d
WO-22270-150-A-1 | 14.09% | 10.82% | 17.73% | 18.67% | 18.99% | 30.39% | 44.49% | 39.88% | 59.71% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  | d
AB-220715-01 LRB 16_INJ001 | 49.58% | 36.06% | 52.74% | 58.4% | 44.25% | 65.94% | 71.38% | 54.87% | 79% | IS3 | IS4 | FALSE | FALSE | IS4 | FALSE | FALSE | FALSE | FALSE | IS3 | IS4 | d |  | IS4 |  |  |  |  | d
AB-220715-01 LRB 16 | 45.37% | 34.81% | 50.26% | 58.46% | 45.85% | 64.7% | 68.2% | 56.33% | 75.91% | IS3 | IS4 | FALSE | FALSE | IS4 | FALSE | FALSE | FALSE | FALSE | IS3 | IS4 | d |  | IS4 |  |  |  |  | d
WO-22270-151-A-1 | 9.512% | 10.13% | 15.68% | 17.94% | 19.05% | 27.8% | 42.05% | 37.55% | 56.9% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  | d
WO-22270-152-A-1 | 14.75% | 12.86% | 17.46% | 20.9% | 20.34% | 30.7% | 47.23% | 41.66% | 63.74% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  | d
WO-22270-153-A-1 | 6.431% | 3.827% | 6.047% | 7.196% | 11.77% | 18.66% | 31.98% | 28.29% | 40.62% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | d
WO-22270-154-A-1 | 12.18% | 11.39% | 18.04% | 20.11% | 20.04% | 30.24% | 46.32% | 40.8% | 60.36% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS3 | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS3 |  | d
WO-22270-155-A-1 | 15.23% | 18.62% | 25.47% | 34.03% | 42.71% | 67.88% | 87.73% | 66.8% | 94.67% | IS3 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | FALSE | IS3 | IS4 | d | IS3 | IS4 |  |  |  |  | d
WO-22270-156-A-1 | 10.46% | 9.897% | 13.16% | 17.31% | 18.12% | 28.11% | 46.19% | 38.17% | 61.77% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  | d
WO-22270-157-A-1 | 11.25% | 11.13% | 18.04% | 18.84% | 20.67% | 32.84% | 44.52% | 40.46% | 62.42% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  | d
WO-22270-158-A-1 | 11.12% | 10.85% | 16.26% | 19.4% | 18.87% | 29.55% | 47.6% | 38.87% | 57.71% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 | IS4 | IS4 |  | d
WO-22270-159-A-1 | 9.343% | 10.05% | 17.51% | 18.78% | 26.64% | 45.21% | 66.34% | 53.97% | 81.82% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | d | IS3 | IS4 | IS4 |  |  |  | d
WO-22270-160-A-1 | 9.397% | 6.297% | 10.3% | 12.02% | 13.92% | 21.51% | 33.06% | 27.71% | 40.67% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | d | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | d
AB-220715-01 LFB-H 9 | 37.98% | 38.25% | 57.82% | 58.44% | 37.99% | 52.04% | 56.05% | 44.73% | 55.23% | IS3 | IS3 | FALSE | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE | IS3 | IS3 |  |  | IS3 |  |  | IS3
AB-220715-01 LRB 17_INJ001 | 48.17% | 38.51% | 56.49% | 62.92% | 48.76% | 67.1% | 71.83% | 54.29% | 78.58% | IS3 | IS4 | FALSE | FALSE | IS4 | FALSE | FALSE | FALSE | FALSE | IS3 | IS4 |  |  | IS4
AB-220715-01 LRB 17 | 42.07% | 42.22% | 53.42% | 60.13% | 51.75% | 63.27% | 71.28% | 58.22% | 74.6% | IS3 | IS4 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | IS3 | IS4
AB-220715-01 LD 8 | 10.07% | 6.825% | 10.48% | 11.09% | 14.9% | 21.67% | 30.48% | 27.81% | 42.42% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 CCC-H 9 | 81.72% | 44.6% | 67.26% | 62.57% | 40.74% | 50.9% | 52.03% | 45.61% | 51.22% | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE |  | IS3 |  |  | IS3 |  |  | IS3
AB-220715-01 LFSM-H 8 | 12.38% | 8.513% | 13.33% | 15.23% | 17.57% | 23.74% | 36.77% | 34.83% | 47.32% | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
WO-22270-161-A-1 | 7.514% | 8.467% | 13.44% | 13.1% | 16.02% | 23.22% | 36.63% | 34.83% | 51.49% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-162-A-1 | 8.469% | 9.989% | 14.92% | 15.47% | 19.01% | 25.97% | 42.42% | 39.5% | 54.09% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-163-A-1 | 5.914% | 4.767% | 8.401% | 8.678% | 13.14% | 18.31% | 29.69% | 27.77% | 41.08% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-164-A-1 | 13.56% | 12.68% | 20.6% | 21.15% | 21.42% | 32.05% | 47.26% | 42.64% | 63.2% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-165-A-1 | 7.283% | 11.52% | 18.57% | 17.66% | 19.21% | 28.97% | 44.55% | 40.01% | 54.73% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-166-A-1 | 10.12% | 11.55% | 17.8% | 18.78% | 21.72% | 29.83% | 46.31% | 42.45% | 64.46% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-167-A-1 | 5.582% | 6.479% | 10.45% | 10.93% | 15.82% | 22.82% | 40.07% | 32.66% | 53.64% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-168-A-1 | 11.58% | 11.65% | 16.99% | 18.31% | 19.33% | 29.4% | 45.41% | 41.43% | 58.74% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-169-A-1 | 8.101% | 6.314% | 8.63% | 10.1% | 14.73% | 21.52% | 36.07% | 31.24% | 45.09% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-170-A-1 | 12.66% | 9.412% | 14.78% | 15.59% | 19.04% | 27.79% | 43.86% | 39.82% | 56.75% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
AB-220715-01 LRB 18_INJ001 | 46.98% | 42.2% | 58.63% | 60.9% | 53.78% | 69.57% | 74.68% | 59.55% | 77.89% | IS3 | IS4 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | IS3 | IS4
WO-22209-62-A-1B | 8.644% | 7.122% | 11.83% | 12.93% | 14.74% | 26.2% | 41.37% | 30.37% | 50.46% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
AB-220715-01 LRB 18 | 24.67% | 38.38% | 54.82% | 56.2% | 47.88% | 60.72% | 65.37% | 55.18% | 68.28% | IS3 | IS4 | FALSE | FALSE | IS4 | FALSE | FALSE | FALSE | FALSE | IS3 | IS4 |  |  | IS4
WO-22270-171-A-1 | 13.83% | 12.84% | 18.51% | 20.4% | 20.89% | 29.68% | 45.32% | 41.67% | 59.17% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-172-A-1 | 12.05% | 12.27% | 19.11% | 19.82% | 20.31% | 29.77% | 45.38% | 41.43% | 60.13% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-173-A-1 | 11.1% | 13.3% | 19.1% | 21.07% | 20.33% | 30.46% | 48.28% | 41.39% | 64.71% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-174-A-1 | 9.207% | 10.92% | 15.16% | 16.49% | 19.81% | 31.44% | 47.13% | 41.23% | 55.72% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-175-A-1 | 10.57% | 11.32% | 17.72% | 18.87% | 20.64% | 29.31% | 46.18% | 41.56% | 60.16% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-176-A-1 | 7.629% | 5.656% | 9.187% | 9.781% | 14.3% | 21.71% | 34.21% | 29.94% | 45.48% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-177-A-1 | 10.66% | 12.43% | 20.02% | 21.85% | 24.5% | 36.83% | 55.4% | 49.27% | 70.16% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 |  | IS4
WO-22270-178-A-1 | 4.801% | 5.431% | 8.732% | 9.901% | 12.46% | 18.12% | 30.59% | 28.52% | 43.71% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-179-A-1 | 11.5% | 8.032% | 12.23% | 13.05% | 17.05% | 24.11% | 37.72% | 34.17% | 51.61% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-180-A-1 | 10.71% | 10.07% | 16.17% | 17.25% | 19.37% | 28.33% | 46.38% | 39.88% | 57.74% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22209-64-A-1B | 10.52% | 6.555% | 11.38% | 12.49% | 14.45% | 25.93% | 40.97% | 30.55% | 43.95% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4
AB-220715-01 LFB-H 10 | 74.42% | 43.89% | 63.36% | 64.24% | 43.38% | 53.89% | 61.18% | 48.1% | 61.41% | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE |  | IS3 |  |  | IS3 |  |  | IS3
AB-220715-01 LRB 19_INJ001 | 52.01% | 45.74% | 62.67% | 66.15% | 54.44% | 76.49% | 81.41% | 62.51% | 86.92% | FALSE | IS4 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE |  | IS4
AB-220715-01 LRB 19 | 29.65% | 39.85% | 61.97% | 63.16% | 46.59% | 64.94% | 69.17% | 53.58% | 76.02% | IS3 | IS4 | FALSE | FALSE | IS4 | FALSE | FALSE | FALSE | FALSE | IS3 | IS4 |  |  | IS4
AB-220715-01 LD 9 | 10.83% | 7.669% | 12.29% | 13.58% | 16.6% | 25.56% | 39.72% | 36.16% | 51.67% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
AB-220715-01 CCC-H 10 | 77.32% | 43.38% | 61.07% | 61.5% | 38.03% | 49.74% | 55.66% | 41.24% | 51.62% | FALSE | IS3 | FALSE | FALSE | IS3 | IS3 | FALSE | IS3 | FALSE |  | IS3 |  |  | IS3 | IS3 |  | IS3
AB-220715-01 LFSM-H 9 | 6.967% | 7.659% | 12.3% | 12.34% | 15.18% | 21.24% | 33.63% | 30.05% | 41.54% | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
WO-22270-181-A-1 | 11.22% | 11.21% | 18.28% | 20.25% | 20.13% | 30.43% | 43.16% | 38.5% | 60.57% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-182-A-1 | 9.552% | 6.818% | 11.79% | 12.65% | 15.37% | 25.16% | 40.64% | 31.97% | 52.59% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-183-A-1 | 13.78% | 12.14% | 17.08% | 19.33% | 20.06% | 29.85% | 47.53% | 39.62% | 61.52% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-184-A-1 | 13.67% | 11.79% | 19.27% | 20.31% | 19.67% | 29.37% | 46.94% | 39.34% | 64.35% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-185-A-1 | 12.75% | 15.14% | 24.73% | 25.86% | 28.16% | 50.91% | 72.95% | 59.34% | 73.97% | IS3 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS3 | IS4
WO-22209-50-A-1B | 10.19% | 8.44% | 13.74% | 16.34% | 16.67% | 28.12% | 45.14% | 34.72% | 41.81% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4
WO-22270-186-A-1 | 8.66% | 6.084% | 9.335% | 10.37% | 13.61% | 21.12% | 33.95% | 28.75% | 43.17% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
WO-22270-187-A-1 | 32.27% | 21.63% | 33.62% | 36.68% | 25.77% | 40.04% | 64.61% | 52.85% | 76.3% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4
WO-22270-188-A-1 | 32.97% | 23.55% | 34.41% | 39.24% | 28.94% | 50.05% | 72.02% | 57.89% | 79.34% | IS3 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS3 | IS4
WO-22270-189-A-1 | 25.76% | 20.12% | 33.06% | 33.08% | 27.31% | 43.33% | 70.47% | 54.49% | 83.39% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4
WO-22270-190-A-1 | 16.63% | 15.18% | 24.66% | 24.31% | 21.09% | 33.55% | 53.85% | 46.61% | 72.16% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 |  | IS4
AB-220715-01 LRB 20_INJ001 | 73.33% | 43.35% | 69.18% | 70.39% | 51.86% | 73.81% | 77.71% | 56.77% | 82.11% | FALSE | IS4 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE |  | IS4
AB-220715-01 LRB 20 | 31.51% | 37.16% | 56.05% | 58.47% | 39.79% | 60.61% | 62.91% | 45.43% | 63.55% | IS3 | IS4 | FALSE | FALSE | IS4 | FALSE | FALSE | IS4 | FALSE | IS3 | IS4 |  |  | IS4 |  |  | IS4
WO-22270-196-A-1 | 32.76% | 20.99% | 33.49% | 36.16% | 24.46% | 41.22% | 65.39% | 52.01% | 81.79% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4
WO-22270-191-A-1 | 33.24% | 23.27% | 34.66% | 39.6% | 27.29% | 46.4% | 72.32% | 57.01% | 83.04% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4
WO-22270-197-A-1 | 15.85% | 12.26% | 19.57% | 21.77% | 20.23% | 32.67% | 51.75% | 41.59% | 66.97% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 |  | IS4
WO-22270-192-A-1 | 38.18% | 26.57% | 37.31% | 40.48% | 27.06% | 42.13% | 66.35% | 56.31% | 86.67% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4
WO-22270-198-A-1 | 18.41% | 14.86% | 18.4% | 20.13% | 22.37% | 35.29% | 55.06% | 44.09% | 71.75% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | UL | IS3 | IS3 | IS4 |  | IS4 |  | UL
WO-22270-193-A-1 | 25.96% | 18.34% | 29.24% | 30.36% | 22.8% | 36.47% | 59.08% | 48.74% | 71.79% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 |  | IS4
WO-22270-199-A-1 | 21.86% | 14.17% | 17.7% | 19.18% | 21.56% | 34.2% | 52.88% | 43.2% | 69.24% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | UL | IS3 | IS3 | IS4 |  | IS4 |  | ul
WO-22209-26-A-1B | 13.38% | 7.439% | 12.64% | 13.49% | 14.62% | 24.99% | 40.49% | 29.78% | 51.42% | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | FALSE | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4
WO-22270-200-A-1 | 15.57% | 13.28% | 16.92% | 19.02% | 20.58% | 33.73% | 49.7% | 40.78% | 63.34% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | UL | IS3 | IS3 | IS4 | IS4 | IS4 |  | UL
WO-22270-194-A-1 | 37.94% | 26.14% | 38.17% | 44.89% | 30.98% | 52.66% | 79.26% | 60.87% | 94.62% | IS3 | IS4 | IS4 | IS3 | IS4 | FALSE | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS3 | IS4
WO-22270-195-A-1 | 19.49% | 15.07% | 24.5% | 25.81% | 21.63% | 33.73% | 51.83% | 43.81% | 69.34% | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS4 | IS4 |  | IS4
AB-220715-01 LFB-H 11 | 93.32% | 48.25% | 73.51% | 75.26% | 46.03% | 64.89% | 64.13% | 44.6% | 38.74% | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE | FALSE | IS3 | IS3 |  | IS3 |  |  | IS3 |  |  | IS3 | IS3
AB-220715-01 LRB 21 | 35.07% | 40.13% | 66.17% | 66.45% | 45.21% | 67.06% | 71.98% | 50.39% | 73.33% | IS3 | IS4 | FALSE | FALSE | IS4 | FALSE | FALSE | FALSE | FALSE | IS3 | IS4 |  |  | IS4
AB-220715-01 LD 10 | 13.33% | 12.3% | 17.65% | 19.9% | 20.46% | 33.79% | 55.73% | 45.67% | 68.29% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
AB-220715-01 CCC-H 11 | 100.3% | 46.61% | 78.58% | 79.86% | 40.15% | 57.56% | 62.9% | 43.1% | 51.8% | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE |  | IS3 |  |  | IS3 |  |  | IS3
WO-22209-16-A-1B | 18.64% | 9.493% | 21.38% | 20.71% | 17.77% | 31.99% | 46.47% | 34.2% | 59.53% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4
AB-220715-01 LFSM-H 10 | 10.68% | 8.557% | 14.24% | 13.44% | 16.17% | 26.09% | 39.35% | 32.75% | 50.37% | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | FALSE | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
WO-22270-201-A-1 | 12.09% | 10.06% | 14.65% | 15.98% | 18.2% | 27.88% | 43.93% | 37.32% | 59.21% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | UL | IS3 | IS3 | IS4 | IS4 | IS4 |  | UL
WO-22270-202-A-1 | 16.15% | 9.644% | 13.52% | 14.4% | 20.06% | 30.87% | 55.03% | 43.13% | 65.55% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | UL | IS3 | IS3 | IS4 |  | IS4 |  | UL
WO-22270-203-A-1 | 17.69% | 13.97% | 16.95% | 20.25% | 22.32% | 32.56% | 56.48% | 42.15% | 64.65% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | UL | IS3 | IS3 | IS4 |  | IS4 |  | UL
AB-220715-01 LRB 21_INJ001 | 59.6% | 41.68% | 67.77% | 68.74% | 51.17% | 75.47% | 82.26% | 56.91% | 82.52% | FALSE | IS4 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE |  | IS4 | B | [6 empty] | b
WO-22270-204-A-1 | 12.99% | 6.313% | 10.98% | 9.855% | 14.04% | 23.11% | 38.5% | 32.45% | 52.52% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4
WO-22270-205-A-1 | 14.51% | 11.83% | 17.43% | 17.56% | 19.91% | 31.93% | 48.15% | 38.69% | 61.07% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4
WO-22270-206-A-1 | 9.805% | 7.842% | 14.03% | 13.66% | 15.41% | 27.57% | 41.29% | 31.24% | 54.67% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4
WO-22270-207-A-1 | 12.47% | 5.439% | 10.93% | 9.152% | 12.61% | 23.6% | 38.25% | 26.4% | 46.38% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 | IS4 | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 | IS4
WO-22270-208-A-1 | 11.58% | 10.66% | 17.39% | 16.83% | 17.7% | 30.89% | 48.55% | 35.81% | 62.19% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | IS4 | IS4
WO-22270-209-A-1 | 17.6% | 15.05% | 21.25% | 21.39% | 20.43% | 36.9% | 56.07% | 41.47% | 71.47% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
WO-22270-210-A-1 | 22.1% | 14.41% | 22.34% | 22.85% | 21.71% | 37.76% | 60.09% | 43.06% | 74.05% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
AB-220715-01 LRB 22_INJ001 | 70.44% | 41.75% | 76.83% | 78.09% | 49.53% | 79.19% | 89.41% | 56.13% | 89.45% | FALSE | IS4 | FALSE | FALSE | IS3 | FALSE | FALSE | FALSE | FALSE |  | IS4 |  |  | IS3
AB-220715-01 LRB 22 | 32.64% | 30% | 49.19% | 53.89% | 35.04% | 57.75% | 64.24% | 39.37% | 62.02% | IS3 | IS4 | IS4 | FALSE | IS4 | FALSE | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 |  | IS4 |  |  | IS4
WO-22270-211-A-1 | 15.38% | 13.6% | 19.58% | 21.84% | 20.75% | 36.12% | 56.96% | 42.96% | 71.67% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
WO-22270-212-A-1 | 18.72% | 10.85% | 18.34% | 19.18% | 18.79% | 32.42% | 53.79% | 40.26% | 67.85% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
WO-22270-213-A-1 | 22.77% | 13.7% | 23.47% | 22.81% | 21.52% | 36.81% | 62.23% | 46.33% | 73.41% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
WO-22270-214-A-1 | 33.89% | 19.73% | 24.41% | 29.41% | 25.91% | 50.18% | 76.33% | 54.57% | 84.31% | IS3 | IS4 | IS4 | IS3 | IS3 | FALSE | FALSE | FALSE | FALSE | IS3 | IS4 | IS4 | IS3 | IS3
WO-22270-215-A-1 | 11.89% | 8.707% | 15.54% | 16.83% | 18.25% | 33.93% | 52.41% | 38.31% | 63% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
WO-22270-216-A-1 | 18.68% | 12.68% | 23.46% | 22.09% | 20.72% | 32.38% | 55.56% | 42.8% | 69.62% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
WO-22270-217-A-1 | 15.53% | 11.21% | 21.24% | 19.93% | 19.57% | 34.39% | 51.29% | 37.09% | 65.79% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
WO-22270-218-A-1 | 20.51% | 14.62% | 28.23% | 26.42% | 22.65% | 38.31% | 53.64% | 42.84% | 70.62% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
WO-22270-219-A-1 | 23.12% | 17.57% | 29.91% | 31.28% | 24.15% | 43.31% | 63.3% | 43.25% | 80.44% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
WO-22270-220-A-1 | 17.68% | 13.9% | 24.68% | 25.22% | 20.82% | 36.74% | 56.1% | 40.09% | 67.9% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
AB-220715-01 LFB-H 12 | 80.06% | 40.55% | 72.04% | 79.31% | 44.73% | 70.13% | 75.19% | 46.84% | 70.21% | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE |  | IS3 |  |  | IS3 |  |  | IS3
AB-220715-01 LRB 23_INJ001 | 78.6% | 41.63% | 73.27% | 79.08% | 49.97% | 81.52% | 87.7% | 54.79% | 86.31% | FALSE | IS4 | FALSE | FALSE | IS4 | FALSE | FALSE | FALSE | FALSE |  | IS4 | b |  | IS4 |  |  |  |  | b
AB-220715-01 LRB 23 | 26.34% | 39.62% | 64.68% | 69.4% | 40.3% | 66.44% | 70.74% | 43.5% | 66.59% | IS3 | IS4 | FALSE | FALSE | IS4 | FALSE | FALSE | IS4 | FALSE | IS3 | IS4 | b |  | IS4 |  |  | IS4 |  | b
AB-220715-01 LD 11 | 21.56% | 16.83% | 29.78% | 30.74% | 23.2% | 42.01% | 61.32% | 44.84% | 76.21% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
AB-220715-01 CCC-H 12 | 82.81% | 40.58% | 73.17% | 75.58% | 38.8% | 60.43% | 66.9% | 40.88% | 58.31% | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE |  | IS3 |  |  | IS3 |  |  | IS3
AB-220715-01 LFSM-H 11 | 16.45% | 13.61% | 24.72% | 25.21% | 20.27% | 34.32% | 49.78% | 35.04% | 58.85% | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | FALSE | IS3 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
WO-22270-221-A-1 | 12.5% | 17.26% | 29.59% | 30.02% | 24.44% | 41.33% | 59.39% | 45.98% | 70.19% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
WO-22270-222-A-1 | 15.88% | 13.46% | 22.29% | 25.46% | 20.79% | 35.76% | 53.26% | 37.5% | 60.8% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
WO-22270-223-A-1 | 22.37% | 17.08% | 28.82% | 30.37% | 22.84% | 41.02% | 59.19% | 43.72% | 69.9% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
WO-22270-224-A-1 | 23.08% | 17.27% | 29.15% | 28.96% | 23.45% | 39.44% | 55.32% | 41.58% | 66.7% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
WO-22270-225-A-1 | 15.95% | 15.63% | 26.7% | 29.82% | 22.08% | 39.67% | 58.79% | 42.68% | 73.39% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
WO-22270-226-A-1 | 23.96% | 18.03% | 28.84% | 31.2% | 23.06% | 39.86% | 56.72% | 42.56% | 68.78% | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | FALSE | IS4 | FALSE | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 |  | IS4
AB-220715-01 LRB 24_INJ001 | 82.44% | 42.95% | 72.04% | 82.35% | 51.77% | 83.86% | 90.81% | 56.2% | 89.26% | FALSE | IS4 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE |  | IS4 | b | b | [5 empty] | b | b
AB-220715-01 LRB 24 | 20.02% | 24.58% | 43.73% | 48.48% | 37.74% | 61.93% | 67.32% | 42.54% | 66.72% | IS3 | IS4 | IS4 | IS4 | IS4 | FALSE | FALSE | IS4 | FALSE | IS3 | IS4 | b | IS4 | IS4 |  |  | IS4 |  | b
AB-220715-01 CCC-H 13 | 78.64% | 41.09% | 79% | 78.26% | 38.92% | 64.54% | 68.85% | 40.48% | 60.06% | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE | FALSE | IS3 | FALSE |  | IS3 |  |  | IS3 |  |  | IS3
AB-220715-01 QCS 2 | 99.15% | 52.82% | 98.74% | 98.94% | 50.64% | 80.89% | 83.15% | 49.2% | 80.74% | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | IS3 | FALSE | [7 empty] | IS3

## Chain of Custody
# | Time Stamp: | Sample ID: | Pruned Sample ID: | Sample I.D. | Lab ID | Preparation Date: | Match? | Holding Time:
302 | 2022-07-24 06:44 | 210405-d2-reprep | 210405 | 042222PFAS_IS141_04C | 210405 | 2022-04-25 | MATCH | 90.2809
294 | 2022-07-24 01:56 | 210415-d2-reprep | 210415 | 042022PFAS_IS141_13C | 210415 | 2022-04-25 | MATCH | 90.0809
279 | 2022-07-23 16:56 | 210439-d2-reprep | 210439 | 042022PFAS_IS141_20C DUP | 210439 | 2022-04-25 | MATCH | 89.7059
255 | 2022-07-23 02:32 | 210451-d2-reprep-shaken | 210451 | 042022PFAS_IS141_24B | 210451 | 2022-04-25 | MATCH | 89.1059
267 | 2022-07-23 09:44 | 210453-d2-reprep | 210453 | 042022PFAS_IS141_25A | 210453 | 2022-04-25 | MATCH | 89.4059
20 | 2022-07-17 05:32 | 211827-d2 | 211827 | 061522PFAS_IS144_01A | 211827 | 2022-06-15 | MATCH | 32.2309
21 | 2022-07-17 06:08 | 211828-d2 | 211828 | 061522PFAS_IS144_01B | 211828 | 2022-06-15 | MATCH | 32.2559
22 | 2022-07-17 06:44 | 211829-d2 | 211829 | 061522PFAS_IS144_01C | 211829 | 2022-06-15 | MATCH | 32.2809
23 | 2022-07-17 07:20 | 211830-d2 | 211830 | 061522PFAS_IS144_01AF | 211830 | 2022-06-15 | MATCH | 32.3059
24 | 2022-07-17 07:56 | 211831-d2 | 211831 | 061522PFAS_IS144_01BF | 211831 | 2022-06-15 | MATCH | 32.3309
25 | 2022-07-17 08:32 | 211832-d2 | 211832 | 061522PFAS_IS144_01CF | 211832 | 2022-06-15 | MATCH | 32.3559
26 | 2022-07-17 09:08 | 211833-d2 | 211833 | 061522PFAS_IS144_02A | 211833 | 2022-06-15 | MATCH | 32.3809
27 | 2022-07-17 09:44 | 211834-d2 | 211834 | 061522PFAS_IS144_02B | 211834 | 2022-06-15 | MATCH | 32.4059
28 | 2022-07-17 10:20 | 211835-d2 | 211835 | 061522PFAS_IS144_02C | 211835 | 2022-06-15 | MATCH | 32.4309
29 | 2022-07-17 10:56 | 211836-d2 | 211836 | 061522PFAS_IS144_02C DUP | 211836 | 2022-06-15 | MATCH | 32.4559
32 | 2022-07-17 12:44 | 211837-d2 | 211837 | 061522PFAS_IS144_03A | 211837 | 2022-06-15 | MATCH | 32.5309
33 | 2022-07-17 13:20 | 211838-d2 | 211838 | 061522PFAS_IS144_03B | 211838 | 2022-06-15 | MATCH | 32.5559
34 | 2022-07-17 13:56 | 211847-d2 | 211847 | 061522PFAS_IS144_03C | 211847 | 2022-06-15 | MATCH | 32.5809
35 | 2022-07-17 14:32 | 211848-d2 | 211848 | 061522PFAS_IS144_04A | 211848 | 2022-06-15 | MATCH | 32.6059
36 | 2022-07-17 15:08 | 211849-d2 | 211849 | 061522PFAS_IS144_04B | 211849 | 2022-06-15 | MATCH | 32.6309
37 | 2022-07-17 15:44 | 211850-d2 | 211850 | 061522PFAS_IS144_04C | 211850 | 2022-06-15 | MATCH | 32.6559
38 | 2022-07-17 16:20 | 211851-d2 | 211851 | 061322PFAS_IS144_11A | 211851 | 2022-06-15 | MATCH | 32.6809
39 | 2022-07-17 16:56 | 211852-d2 | 211852 | 061322PFAS_IS144_11B | 211852 | 2022-06-15 | MATCH | 32.7059
40 | 2022-07-17 17:32 | 211853-d2 | 211853 | 061322PFAS_IS144_11C | 211853 | 2022-06-15 | MATCH | 32.7309
41 | 2022-07-17 18:08 | 211854-d2 | 211854 | 061322PFAS_IS144_11C DUP | 211854 | 2022-06-15 | MATCH | 32.7559
48 | 2022-07-17 22:20 | 211855-d2 | 211855 | 061322PFAS_IS144_12A | 211855 | 2022-06-15 | MATCH | 32.9309
49 | 2022-07-17 22:56 | 211856-d2 | 211856 | 061322PFAS_IS144_12B | 211856 | 2022-06-15 | MATCH | 32.9559
50 | 2022-07-17 23:32 | 211857-d2 | 211857 | 061322PFAS_IS144_12C | 211857 | 2022-06-15 | MATCH | 32.9809
51 | 2022-07-18 00:08 | 211858-d2 | 211858 | 061322PFAS_IS144_13A | 211858 | 2022-06-15 | MATCH | 33.0059
52 | 2022-07-18 00:44 | 211859-d2 | 211859 | 061322PFAS_IS144_13B | 211859 | 2022-06-15 | MATCH | 33.0309
53 | 2022-07-18 01:20 | 211860-d2 | 211860 | 061322PFAS_IS144_13C | 211860 | 2022-06-15 | MATCH | 33.0559
54 | 2022-07-18 01:56 | 211861-d2 | 211861 | 061322PFAS_IS144_14A | 211861 | 2022-06-15 | MATCH | 33.0809
55 | 2022-07-18 02:32 | 211862-d2 | 211862 | 061322PFAS_IS144_14B | 211862 | 2022-06-15 | MATCH | 33.1059
56 | 2022-07-18 03:08 | 211863-d2 | 211863 | 061322PFAS_IS144_14C | 211863 | 2022-06-15 | MATCH | 33.1309
57 | 2022-07-18 03:44 | 211864-d2 | 211864 | 061322PFAS_IS144_14C DUP | 211864 | 2022-06-15 | MATCH | 33.1559
60 | 2022-07-18 05:32 | 211865-d2 | 211865 | 061322PFAS_IS144_15A | 211865 | 2022-06-15 | MATCH | 33.2309
61 | 2022-07-18 06:08 | 211866-d2 | 211866 | 061322PFAS_IS144_15B | 211866 | 2022-06-15 | MATCH | 33.2559
62 | 2022-07-18 06:44 | 211867-d2 | 211867 | 061322PFAS_IS144_15C | 211867 | 2022-06-15 | MATCH | 33.2809
63 | 2022-07-18 07:20 | 211868-d2 | 211868 | 061322PFAS_IS144_16A | 211868 | 2022-06-15 | MATCH | 33.3059
64 | 2022-07-18 07:56 | 211869-d2 | 211869 | 061322PFAS_IS144_16B | 211869 | 2022-06-15 | MATCH | 33.3309
65 | 2022-07-18 08:32 | 211870-d2 | 211870 | 061322PFAS_IS144_16C | 211870 | 2022-06-15 | MATCH | 33.3559
66 | 2022-07-18 09:08 | 211871-d2 | 211871 | 061322PFAS_IS144_17A | 211871 | 2022-06-15 | MATCH | 33.3809
67 | 2022-07-18 09:44 | 211872-d2 | 211872 | 061322PFAS_IS144_17B | 211872 | 2022-06-15 | MATCH | 33.4059
68 | 2022-07-18 10:20 | 211873-d2 | 211873 | 061322PFAS_IS144_17C | 211873 | 2022-06-15 | MATCH | 33.4309
69 | 2022-07-18 10:56 | 211874-d2 | 211874 | 061322PFAS_IS144_17C DUP | 211874 | 2022-06-15 | MATCH | 33.4559
76 | 2022-07-18 15:08 | 211875-d2 | 211875 | 061322PFAS_IS144_18A | 211875 | 2022-06-15 | MATCH | 33.6309
77 | 2022-07-18 15:44 | 211876-d2 | 211876 | 061322PFAS_IS144_18B | 211876 | 2022-06-15 | MATCH | 33.6559
78 | 2022-07-18 16:20 | 211877-d2 | 211877 | 061322PFAS_IS144_18C | 211877 | 2022-06-15 | MATCH | 33.6809
79 | 2022-07-18 16:56 | 211878-d2 | 211878 | 061322PFAS_IS144_19A | 211878 | 2022-06-15 | MATCH | 33.7059
80 | 2022-07-18 17:32 | 211879-d2 | 211879 | 061322PFAS_IS144_19B | 211879 | 2022-06-15 | MATCH | 33.7309
82 | 2022-07-18 18:44 | 211880-d2 | 211880 | 061322PFAS_IS144_19C | 211880 | 2022-06-15 | MATCH | 33.7809
83 | 2022-07-18 19:20 | 211881-d2 | 211881 | 061322PFAS_IS144_20A | 211881 | 2022-06-15 | MATCH | 33.8059
85 | 2022-07-18 20:32 | 211882-d2 | 211882 | 061322PFAS_IS144_20B | 211882 | 2022-06-15 | MATCH | 33.8559
87 | 2022-07-18 21:44 | 211883-d2 | 211883 | 061322PFAS_IS144_20C | 211883 | 2022-06-15 | MATCH | 33.9059
89 | 2022-07-18 22:56 | 211884-d2 | 211884 | 061322PFAS_IS144_20C DUP | 211884 | 2022-06-15 | MATCH | 33.9559
95 | 2022-07-19 02:32 | 211885-d2 | 211885 | 061322PFAS_IS144_21A | 211885 | 2022-06-15 | MATCH | 34.1059
81 | 2022-07-18 18:08 | 211886-d2 | 211886 | 061322PFAS_IS144_21B | 211886 | 2022-06-15 | MATCH | 33.7559
84 | 2022-07-18 19:56 | 211887-d2 | 211887 | 061322PFAS_IS144_21C | 211887 | 2022-06-15 | MATCH | 33.8309
86 | 2022-07-18 21:08 | 211888-d2 | 211888 | 061322PFAS_IS144_22A | 211888 | 2022-06-15 | MATCH | 33.8809
88 | 2022-07-18 22:20 | 211889-d2 | 211889 | 061322PFAS_IS144_22B | 211889 | 2022-06-15 | MATCH | 33.9309
90 | 2022-07-18 23:32 | 211890-d2 | 211890 | 061322PFAS_IS144_22C | 211890 | 2022-06-15 | MATCH | 33.9809
91 | 2022-07-19 00:08 | 211891-d2 | 211891 | 061322PFAS_IS144_23A | 211891 | 2022-06-15 | MATCH | 34.0059
94 | 2022-07-19 01:56 | 211892-d2 | 211892 | 061322PFAS_IS144_23B | 211892 | 2022-06-15 | MATCH | 34.0809
96 | 2022-07-19 03:08 | 211893-d2 | 211893 | 061322PFAS_IS144_23C | 211893 | 2022-06-15 | MATCH | 34.1309
97 | 2022-07-19 03:44 | 211894-d2 | 211894 | 061322PFAS_IS144_23C DUP | 211894 | 2022-06-15 | MATCH | 34.1559
104 | 2022-07-19 07:56 | 211895-d2 | 211895 | 061322PFAS_IS144_24A | 211895 | 2022-06-15 | MATCH | 34.3309
105 | 2022-07-19 08:32 | 211896-d2 | 211896 | 061322PFAS_IS144_24B | 211896 | 2022-06-15 | MATCH | 34.3559
106 | 2022-07-19 09:08 | 211897-d2 | 211897 | 061322PFAS_IS144_24C | 211897 | 2022-06-15 | MATCH | 34.3809
107 | 2022-07-19 09:44 | 211898-d2 | 211898 | 061322PFAS_IS144_25A | 211898 | 2022-06-15 | MATCH | 34.4059
108 | 2022-07-19 10:20 | 211899-d2 | 211899 | 061322PFAS_IS144_25B | 211899 | 2022-06-15 | MATCH | 34.4309
109 | 2022-07-19 10:56 | 211900-d2 | 211900 | 061322PFAS_IS144_25C | 211900 | 2022-06-15 | MATCH | 34.4559
110 | 2022-07-19 11:32 | 211901-d2 | 211901 | 061322PFAS_IS144_26A | 211901 | 2022-06-15 | MATCH | 34.4809
111 | 2022-07-19 12:08 | 211902-d2 | 211902 | 061322PFAS_IS144_26B | 211902 | 2022-06-15 | MATCH | 34.5059
112 | 2022-07-19 12:44 | 211903-d2 | 211903 | 061322PFAS_IS144_26C | 211903 | 2022-06-15 | MATCH | 34.5309
113 | 2022-07-19 13:20 | 211904-d2 | 211904 | 061322PFAS_IS144_26C DUP | 211904 | 2022-06-15 | MATCH | 34.5559
116 | 2022-07-19 15:08 | 211905-d2 | 211905 | 061322PFAS_IS144_27A | 211905 | 2022-06-15 | MATCH | 34.6309
117 | 2022-07-19 15:44 | 211906-d2 | 211906 | 061322PFAS_IS144_27B | 211906 | 2022-06-15 | MATCH | 34.6559
118 | 2022-07-19 16:20 | 211907-d2 | 211907 | 061322PFAS_IS144_27C | 211907 | 2022-06-15 | MATCH | 34.6809
119 | 2022-07-19 16:56 | 211908-d2 | 211908 | 061322PFAS_IS144_28A | 211908 | 2022-06-15 | MATCH | 34.7059
120 | 2022-07-19 17:32 | 211909-d2 | 211909 | 061322PFAS_IS144_28B | 211909 | 2022-06-15 | MATCH | 34.7309
121 | 2022-07-19 18:08 | 211910-d2 | 211910 | 061322PFAS_IS144_28C | 211910 | 2022-06-15 | MATCH | 34.7559
122 | 2022-07-19 18:44 | 211911-d2 | 211911 | 061422PFAS_IS144_29A | 211911 | 2022-06-15 | MATCH | 34.7809
123 | 2022-07-19 19:20 | 211912-d2 | 211912 | 061422PFAS_IS144_29B | 211912 | 2022-06-15 | MATCH | 34.8059
124 | 2022-07-19 19:56 | 211913-d2 | 211913 | 061422PFAS_IS144_29C | 211913 | 2022-06-15 | MATCH | 34.8309
125 | 2022-07-19 20:32 | 211914-d2 | 211914 | 061422PFAS_IS144_29C DUP | 211914 | 2022-06-15 | MATCH | 34.8559
132 | 2022-07-20 00:44 | 211915-d2 | 211915 | 061422PFAS_IS144_30A | 211915 | 2022-06-15 | MATCH | 35.0309
133 | 2022-07-20 01:20 | 211916-d2 | 211916 | 061422PFAS_IS144_30B | 211916 | 2022-06-15 | MATCH | 35.0559
134 | 2022-07-20 01:56 | 211917-d2 | 211917 | 061422PFAS_IS144_30C | 211917 | 2022-06-15 | MATCH | 35.0809
135 | 2022-07-20 02:32 | 211918-d2 | 211918 | 061422PFAS_IS144_31A | 211918 | 2022-06-15 | MATCH | 35.1059
136 | 2022-07-20 03:08 | 211919-d2 | 211919 | 061422PFAS_IS144_31B | 211919 | 2022-06-15 | MATCH | 35.1309
137 | 2022-07-20 03:44 | 211920-d2 | 211920 | 061422PFAS_IS144_31C | 211920 | 2022-06-15 | MATCH | 35.1559
138 | 2022-07-20 04:20 | 211921-d2 | 211921 | 061422PFAS_IS144_32A | 211921 | 2022-06-15 | MATCH | 35.1809
139 | 2022-07-20 04:56 | 211922-d2 | 211922 | 061422PFAS_IS144_32B | 211922 | 2022-06-15 | MATCH | 35.2059
140 | 2022-07-20 05:32 | 211923-d2 | 211923 | 061422PFAS_IS144_32C | 211923 | 2022-06-15 | MATCH | 35.2309
141 | 2022-07-20 06:08 | 211924-d2 | 211924 | 061422PFAS_IS144_32C DUP | 211924 | 2022-06-15 | MATCH | 35.2559
144 | 2022-07-20 07:56 | 211925-d2 | 211925 | 061422PFAS_IS144_33A | 211925 | 2022-06-15 | MATCH | 35.3309
145 | 2022-07-20 08:32 | 211926-d2 | 211926 | 061422PFAS_IS144_33B | 211926 | 2022-06-15 | MATCH | 35.3559
146 | 2022-07-20 09:08 | 211927-d2 | 211927 | 061422PFAS_IS144_33C | 211927 | 2022-06-15 | MATCH | 35.3809
147 | 2022-07-20 09:44 | 211928-d2 | 211928 | 061422PFAS_IS144_34A | 211928 | 2022-06-15 | MATCH | 35.4059
148 | 2022-07-20 10:20 | 211929-d2 | 211929 | 061422PFAS_IS144_34B | 211929 | 2022-06-15 | MATCH | 35.4309
149 | 2022-07-20 10:56 | 211930-d2 | 211930 | 061422PFAS_IS144_34C | 211930 | 2022-06-15 | MATCH | 35.4559
150 | 2022-07-20 11:32 | 211931-d2 | 211931 | 061422PFAS_IS144_35A | 211931 | 2022-06-15 | MATCH | 35.4809
151 | 2022-07-20 12:08 | 211932-d2 | 211932 | 061422PFAS_IS144_35B | 211932 | 2022-06-15 | MATCH | 35.5059
152 | 2022-07-20 12:44 | 211933-d2 | 211933 | 061422PFAS_IS144_35C | 211933 | 2022-06-15 | MATCH | 35.5309
153 | 2022-07-20 13:20 | 211934-d2 | 211934 | 061422PFAS_IS144_35C DUP | 211934 | 2022-06-15 | MATCH | 35.5559
160 | 2022-07-20 17:32 | 211935-d2 | 211935 | 061422PFAS_IS144_36A | 211935 | 2022-06-15 | MATCH | 35.7309
161 | 2022-07-20 18:08 | 211936-d2 | 211936 | 061422PFAS_IS144_36B | 211936 | 2022-06-15 | MATCH | 35.7559
162 | 2022-07-20 18:44 | 211937-d2 | 211937 | 061422PFAS_IS144_36C | 211937 | 2022-06-15 | MATCH | 35.7809
163 | 2022-07-20 19:20 | 211938-d2 | 211938 | 061422PFAS_IS144_37A | 211938 | 2022-06-15 | MATCH | 35.8059
164 | 2022-07-20 19:56 | 211939-d2 | 211939 | 061422PFAS_IS144_37B | 211939 | 2022-06-15 | MATCH | 35.8309
165 | 2022-07-20 20:32 | 211940-d2 | 211940 | 061422PFAS_IS144_37C | 211940 | 2022-06-15 | MATCH | 35.8559
166 | 2022-07-20 21:08 | 211941-d2 | 211941 | 061422PFAS_IS144_38A | 211941 | 2022-06-15 | MATCH | 35.8809
167 | 2022-07-20 21:44 | 211942-d2 | 211942 | 061422PFAS_IS144_38B | 211942 | 2022-06-15 | MATCH | 35.9059
168 | 2022-07-20 22:20 | 211943-d2 | 211943 | 061422PFAS_IS144_38C | 211943 | 2022-06-15 | MATCH | 35.9309
169 | 2022-07-20 22:56 | 211944-d2 | 211944 | 061422PFAS_IS144_38C DUP | 211944 | 2022-06-15 | MATCH | 35.9559
172 | 2022-07-21 00:44 | 211945-d2 | 211945 | 061422PFAS_IS144_39A | 211945 | 2022-06-15 | MATCH | 36.0309
173 | 2022-07-21 01:20 | 211946-d2 | 211946 | 061422PFAS_IS144_39B | 211946 | 2022-06-15 | MATCH | 36.0559
174 | 2022-07-21 01:56 | 211947-d2 | 211947 | 061422PFAS_IS144_39C | 211947 | 2022-06-15 | MATCH | 36.0809
175 | 2022-07-21 02:32 | 211948-d2 | 211948 | 061422PFAS_IS144_40A | 211948 | 2022-06-15 | MATCH | 36.1059
176 | 2022-07-21 03:08 | 211949-d2 | 211949 | 061422PFAS_IS144_40B | 211949 | 2022-06-15 | MATCH | 36.1309
177 | 2022-07-21 03:44 | 211950-d2 | 211950 | 061422PFAS_IS144_40C | 211950 | 2022-06-15 | MATCH | 36.1559
178 | 2022-07-21 04:20 | 211951-d2 | 211951 | 061422PFAS_IS144_41A | 211951 | 2022-06-15 | MATCH | 36.1809
179 | 2022-07-21 04:56 | 211952-d2 | 211952 | 061422PFAS_IS144_41B | 211952 | 2022-06-15 | MATCH | 36.2059
180 | 2022-07-21 05:32 | 211953-d2 | 211953 | 061422PFAS_IS144_41C | 211953 | 2022-06-15 | MATCH | 36.2309
181 | 2022-07-21 06:08 | 211954-d2 | 211954 | 061422PFAS_IS144_41C DUP | 211954 | 2022-06-15 | MATCH | 36.2559
188 | 2022-07-21 10:20 | 211955-d2 | 211955 | 061422PFAS_IS144_42A | 211955 | 2022-06-15 | MATCH | 36.4309
189 | 2022-07-21 10:56 | 211956-d2 | 211956 | 061422PFAS_IS144_42B | 211956 | 2022-06-15 | MATCH | 36.4559
190 | 2022-07-21 11:32 | 211957-d2 | 211957 | 061422PFAS_IS144_42C | 211957 | 2022-06-15 | MATCH | 36.4809
191 | 2022-07-21 12:08 | 211958-d2 | 211958 | 061422PFAS_IS144_43A | 211958 | 2022-06-15 | MATCH | 36.5059
192 | 2022-07-21 12:44 | 211959-d2 | 211959 | 061422PFAS_IS144_43B | 211959 | 2022-06-15 | MATCH | 36.5309
193 | 2022-07-21 13:20 | 211960-d2 | 211960 | 061422PFAS_IS144_43C | 211960 | 2022-06-15 | MATCH | 36.5559
194 | 2022-07-21 13:56 | 211961-d2 | 211961 | 061422PFAS_IS144_44A | 211961 | 2022-06-15 | MATCH | 36.5809
195 | 2022-07-21 14:32 | 211962-d2 | 211962 | 061422PFAS_IS144_44B | 211962 | 2022-06-15 | MATCH | 36.6059
196 | 2022-07-21 15:08 | 211963-d2 | 211963 | 061422PFAS_IS144_44C | 211963 | 2022-06-15 | MATCH | 36.6309
197 | 2022-07-21 15:44 | 211964-d2 | 211964 | 061422PFAS_IS144_44C DUP | 211964 | 2022-06-15 | MATCH | 36.6559
200 | 2022-07-21 17:32 | 211965-d2 | 211965 | 061422PFAS_IS144_45A | 211965 | 2022-06-15 | MATCH | 36.7309
201 | 2022-07-21 18:08 | 211966-d2 | 211966 | 061422PFAS_IS144_45B | 211966 | 2022-06-15 | MATCH | 36.7559
202 | 2022-07-21 18:44 | 211967-d2 | 211967 | 061422PFAS_IS144_45C | 211967 | 2022-06-15 | MATCH | 36.7809
203 | 2022-07-21 19:20 | 211968-d2 | 211968 | 061422PFAS_IS144_46A | 211968 | 2022-06-15 | MATCH | 36.8059
204 | 2022-07-21 19:56 | 211969-d2 | 211969 | 061422PFAS_IS144_46B | 211969 | 2022-06-15 | MATCH | 36.8309
205 | 2022-07-21 20:32 | 211970-d2 | 211970 | 061422PFAS_IS144_46C | 211970 | 2022-06-15 | MATCH | 36.8559
206 | 2022-07-21 21:08 | 211971-d2 | 211971 | 061522PFAS_IS144_47A | 211971 | 2022-06-15 | MATCH | 36.8809
207 | 2022-07-21 21:44 | 211972-d2 | 211972 | 061522PFAS_IS144_47B | 211972 | 2022-06-15 | MATCH | 36.9059
208 | 2022-07-21 22:20 | 211973-d2 | 211973 | 061522PFAS_IS144_47C | 211973 | 2022-06-15 | MATCH | 36.9309
209 | 2022-07-21 22:56 | 211974-d2 | 211974 | 061522PFAS_IS144_47C DUP | 211974 | 2022-06-15 | MATCH | 36.9559
216 | 2022-07-22 03:08 | 211975-d2 | 211975 | 061522PFAS_IS144_48A | 211975 | 2022-06-15 | MATCH | 37.1309
217 | 2022-07-22 03:44 | 211976-d2 | 211976 | 061522PFAS_IS144_48B | 211976 | 2022-06-15 | MATCH | 37.1559
218 | 2022-07-22 04:20 | 211977-d2 | 211977 | 061522PFAS_IS144_48C | 211977 | 2022-06-15 | MATCH | 37.1809
219 | 2022-07-22 04:56 | 211978-d2 | 211978 | 061522PFAS_IS144_49A | 211978 | 2022-06-15 | MATCH | 37.2059
220 | 2022-07-22 05:32 | 211979-d2 | 211979 | 061522PFAS_IS144_49B | 211979 | 2022-06-15 | MATCH | 37.2309
221 | 2022-07-22 06:08 | 211980-d2 | 211980 | 061522PFAS_IS144_49C | 211980 | 2022-06-15 | MATCH | 37.2559
222 | 2022-07-22 06:44 | 211981-d2 | 211981 | 061522PFAS_IS144_50A | 211981 | 2022-06-15 | MATCH | 37.2809
223 | 2022-07-22 07:20 | 211982-d2 | 211982 | 061522PFAS_IS144_50B | 211982 | 2022-06-15 | MATCH | 37.3059
224 | 2022-07-22 07:56 | 211983-d2 | 211983 | 061522PFAS_IS144_50C | 211983 | 2022-06-15 | MATCH | 37.3309
225 | 2022-07-22 08:32 | 211984-d2 | 211984 | 061522PFAS_IS144_50C DUP | 211984 | 2022-06-15 | MATCH | 37.3559
228 | 2022-07-22 10:20 | 211985-d2 | 211985 | 061522PFAS_IS144_51A | 211985 | 2022-06-15 | MATCH | 37.4309
229 | 2022-07-22 10:56 | 211986-d2 | 211986 | 061522PFAS_IS144_51B | 211986 | 2022-06-15 | MATCH | 37.4559
230 | 2022-07-22 11:32 | 211987-d2 | 211987 | 061522PFAS_IS144_51C | 211987 | 2022-06-15 | MATCH | 37.4809
231 | 2022-07-22 12:08 | 211988-d2 | 211988 | 061522PFAS_IS144_52A | 211988 | 2022-06-15 | MATCH | 37.5059
232 | 2022-07-22 12:44 | 211989-d2 | 211989 | 061522PFAS_IS144_52B | 211989 | 2022-06-15 | MATCH | 37.5309
233 | 2022-07-22 13:20 | 211990-d2 | 211990 | 061522PFAS_IS144_52C | 211990 | 2022-06-15 | MATCH | 37.5559
234 | 2022-07-22 13:56 | 211991-d2 | 211991 | 061522PFAS_IS144_53A | 211991 | 2022-06-15 | MATCH | 37.5809
235 | 2022-07-22 14:32 | 211992-d2 | 211992 | 061522PFAS_IS144_53B | 211992 | 2022-06-15 | MATCH | 37.6059
236 | 2022-07-22 15:08 | 211993-d2 | 211993 | 061522PFAS_IS144_53C | 211993 | 2022-06-15 | MATCH | 37.6309
237 | 2022-07-22 15:44 | 211994-d2 | 211994 | 061522PFAS_IS144_53C DUP | 211994 | 2022-06-15 | MATCH | 37.6559
244 | 2022-07-22 19:56 | 211995-d2 | 211995 | 061522PFAS_IS144_54A | 211995 | 2022-06-15 | MATCH | 37.8309
245 | 2022-07-22 20:32 | 211996-d2 | 211996 | 061522PFAS_IS144_54B | 211996 | 2022-06-15 | MATCH | 37.8559
246 | 2022-07-22 21:08 | 211997-d2 | 211997 | 061522PFAS_IS144_54C | 211997 | 2022-06-15 | MATCH | 37.8809
247 | 2022-07-22 21:44 | 211998-d2 | 211998 | 061522PFAS_IS144_55A | 211998 | 2022-06-15 | MATCH | 37.9059
248 | 2022-07-22 22:20 | 211999-d2 | 211999 | 061522PFAS_IS144_55B | 211999 | 2022-06-15 | MATCH | 37.9309
249 | 2022-07-22 22:56 | 212000-d2 | 212000 | 061522PFAS_IS144_55C | 212000 | 2022-06-15 | MATCH | 37.9559
250 | 2022-07-22 23:32 | 212001-d2 | 212001 | 061522PFAS_IS144_56A | 212001 | 2022-06-15 | MATCH | 37.9809
251 | 2022-07-23 00:08 | 212002-d2 | 212002 | 061522PFAS_IS144_56B | 212002 | 2022-06-15 | MATCH | 38.0059
252 | 2022-07-23 00:44 | 212003-d2 | 212003 | 061522PFAS_IS144_56C | 212003 | 2022-06-15 | MATCH | 38.0309
253 | 2022-07-23 01:20 | 212004-d2 | 212004 | 061522PFAS_IS144_56C DUP | 212004 | 2022-06-15 | MATCH | 38.0559
257 | 2022-07-23 03:44 | 212005-d2 | 212005 | 061522PFAS_IS144_57A | 212005 | 2022-06-15 | MATCH | 38.1559
258 | 2022-07-23 04:20 | 212006-d2 | 212006 | 061522PFAS_IS144_57B | 212006 | 2022-06-15 | MATCH | 38.1809
259 | 2022-07-23 04:56 | 212007-d2 | 212007 | 061522PFAS_IS144_57C | 212007 | 2022-06-15 | MATCH | 38.2059
260 | 2022-07-23 05:32 | 212008-d2 | 212008 | 061522PFAS_IS144_58A | 212008 | 2022-06-15 | MATCH | 38.2309
261 | 2022-07-23 06:08 | 212009-d2 | 212009 | 061522PFAS_IS144_58B | 212009 | 2022-06-15 | MATCH | 38.2559
262 | 2022-07-23 06:44 | 212010-d2 | 212010 | 061522PFAS_IS144_58C | 212010 | 2022-06-15 | MATCH | 38.2809
263 | 2022-07-23 07:20 | 212011-d2 | 212011 | 061522PFAS_IS144_59A | 212011 | 2022-06-15 | MATCH | 38.3059
264 | 2022-07-23 07:56 | 212012-d2 | 212012 | 061522PFAS_IS144_59B | 212012 | 2022-06-15 | MATCH | 38.3309
265 | 2022-07-23 08:32 | 212013-d2 | 212013 | 061522PFAS_IS144_59C | 212013 | 2022-06-15 | MATCH | 38.3559
266 | 2022-07-23 09:08 | 212014-d2 | 212014 | 061522PFAS_IS144_59C DUP | 212014 | 2022-06-15 | MATCH | 38.3809
274 | 2022-07-23 13:56 | 212015-d2 | 212015 | 061522PFAS_IS144_60A | 212015 | 2022-06-15 | MATCH | 38.5809
275 | 2022-07-23 14:32 | 212016-d2 | 212016 | 061522PFAS_IS144_60B | 212016 | 2022-06-15 | MATCH | 38.6059
276 | 2022-07-23 15:08 | 212017-d2 | 212017 | 061522PFAS_IS144_60C | 212017 | 2022-06-15 | MATCH | 38.6309
277 | 2022-07-23 15:44 | 212018-d2 | 212018 | 061522PFAS_IS144_61A | 212018 | 2022-06-15 | MATCH | 38.6559
278 | 2022-07-23 16:20 | 212019-d2 | 212019 | 061522PFAS_IS144_61B | 212019 | 2022-06-15 | MATCH | 38.6809
280 | 2022-07-23 17:32 | 212020-d2 | 212020 | 061522PFAS_IS144_61C | 212020 | 2022-06-15 | MATCH | 38.7309
281 | 2022-07-23 18:08 | 212021-d2 | 212021 | 061522PFAS_IS144_62A | 212021 | 2022-06-15 | MATCH | 38.7559
282 | 2022-07-23 18:44 | 212022-d2 | 212022 | 061522PFAS_IS144_62B | 212022 | 2022-06-15 | MATCH | 38.7809
283 | 2022-07-23 19:20 | 212023-d2 | 212023 | 061522PFAS_IS144_62C | 212023 | 2022-06-15 | MATCH | 38.8059
284 | 2022-07-23 19:56 | 212024-d2 | 212024 | 061522PFAS_IS144_62C DUP | 212024 | 2022-06-15 | MATCH | 38.8309
288 | 2022-07-23 22:20 | 212025-d2 | 212025 | 061522PFAS_IS144_63A | 212025 | 2022-06-15 | MATCH | 38.9309
290 | 2022-07-23 23:32 | 212026-d2 | 212026 | 061522PFAS_IS144_63B | 212026 | 2022-06-15 | MATCH | 38.9809
292 | 2022-07-24 00:44 | 212027-d2 | 212027 | 061522PFAS_IS144_63C | 212027 | 2022-06-15 | MATCH | 39.0309
296 | 2022-07-24 03:08 | 212028-d2 | 212028 | 061522PFAS_IS144_64A | 212028 | 2022-06-15 | MATCH | 39.1309
297 | 2022-07-24 03:44 | 212029-d2 | 212029 | 061522PFAS_IS144_64B | 212029 | 2022-06-15 | MATCH | 39.1559
287 | 2022-07-23 21:44 | 212030-d2 | 212030 | 061522PFAS_IS144_64C | 212030 | 2022-06-15 | MATCH | 38.9059
289 | 2022-07-23 22:56 | 212031-d2 | 212031 | 061522PFAS_IS144_64C DUP | 212031 | 2022-06-15 | MATCH | 38.9559
291 | 2022-07-24 00:08 | 212032-d2 | 212032 | 060622PFAS_IS144_05A | 212032 | 2022-06-15 | MATCH | 39.0059
293 | 2022-07-24 01:20 | 212033-d2 | 212033 | 060622PFAS_IS144_05B | 212033 | 2022-06-15 | MATCH | 39.0559
295 | 2022-07-24 02:32 | 212034-d2 | 212034 | 060622PFAS_IS144_05C | 212034 | 2022-06-15 | MATCH | 39.1059
304 | 2022-07-24 07:56 | 212035-d2 | 212035 | 060622PFAS_IS144_05C DUP | 212035 | 2022-06-15 | MATCH | 39.3309
305 | 2022-07-24 08:32 | 212036-d2 | 212036 | 060622PFAS_IS144_05AF | 212036 | 2022-06-15 | MATCH | 39.3559
306 | 2022-07-24 09:08 | 212037-d2 | 212037 | 060622PFAS_IS144_05BF | 212037 | 2022-06-15 | MATCH | 39.3809
308 | 2022-07-24 10:20 | 212038-d2 | 212038 | 060622PFAS_IS144_05CF | 212038 | 2022-06-15 | MATCH | 39.4309
309 | 2022-07-24 10:56 | 212039-d2 | 212039 | 061522PFAS_IS144_06A | 212039 | 2022-06-15 | MATCH | 39.4559
310 | 2022-07-24 11:32 | 212040-d2 | 212040 | 061522PFAS_IS144_06B | 212040 | 2022-06-15 | MATCH | 39.4809
311 | 2022-07-24 12:08 | 212041-d2 | 212041 | 061522PFAS_IS144_06C | 212041 | 2022-06-15 | MATCH | 39.5059
312 | 2022-07-24 12:44 | 212042-d2 | 212042 | 060622PFAS_IS144_07A | 212042 | 2022-06-15 | MATCH | 39.5309
313 | 2022-07-24 13:20 | 212043-d2 | 212043 | 060622PFAS_IS144_07B | 212043 | 2022-06-15 | MATCH | 39.5559
314 | 2022-07-24 13:56 | 212044-d2 | 212044 | 060622PFAS_IS144_07C | 212044 | 2022-06-15 | MATCH | 39.5809
317 | 2022-07-24 15:44 | 212045-d2 | 212045 | 060622PFAS_IS144_07AF | 212045 | 2022-06-15 | MATCH | 39.6559
318 | 2022-07-24 16:20 | 212046-d2 | 212046 | 060622PFAS_IS144_07BF | 212046 | 2022-06-15 | MATCH | 39.6809
319 | 2022-07-24 16:56 | 212047-d2 | 212047 | 060622PFAS_IS144_07CF | 212047 | 2022-06-15 | MATCH | 39.7059
320 | 2022-07-24 17:32 | 212048-d2 | 212048 | 061522PFAS_IS144_08A | 212048 | 2022-06-15 | MATCH | 39.7309
321 | 2022-07-24 18:08 | 212049-d2 | 212049 | 061522PFAS_IS144_08B | 212049 | 2022-06-15 | MATCH | 39.7559
322 | 2022-07-24 18:44 | 212050-d2 | 212050 | 061522PFAS_IS144_08C | 212050 | 2022-06-15 | MATCH | 39.7809
323 | 2022-07-24 19:20 | 212051-d2 | 212051 | 061522PFAS_IS144_08C DUP | 212051 | 2022-06-15 | MATCH | 39.8059
324 | 2022-07-24 19:56 | 212052-d4 | 212052 | 060622PFAS_IS144_09A | 212052 | 2022-06-15 | MATCH | 39.8309
325 | 2022-07-24 20:32 | 212053-d4 | 212053 | 060622PFAS_IS144_09B | 212053 | 2022-06-15 | MATCH | 39.8559
326 | 2022-07-24 21:08 | 212054-d4 | 212054 | 060622PFAS_IS144_09C | 212054 | 2022-06-15 | MATCH | 39.8809
333 | 2022-07-25 01:20 | 212055-d4 | 212055 | 060622PFAS_IS144_09AF | 212055 | 2022-06-15 | MATCH | 40.0559
334 | 2022-07-25 01:56 | 212056-d4 | 212056 | 060622PFAS_IS144_09BF | 212056 | 2022-06-15 | MATCH | 40.0809
335 | 2022-07-25 02:32 | 212057-d4 | 212057 | 060622PFAS_IS144_09CF | 212057 | 2022-06-15 | MATCH | 40.1059
336 | 2022-07-25 03:08 | 212058-d4 | 212058 | 061522PFAS_IS144_10A | 212058 | 2022-06-15 | MATCH | 40.1309
337 | 2022-07-25 03:44 | 212059-d4 | 212059 | 061522PFAS_IS144_10B | 212059 | 2022-06-15 | MATCH | 40.1559
338 | 2022-07-25 04:20 | 212060-d4 | 212060 | 061522PFAS_IS144_10C | 212060 | 2022-06-15 | MATCH | 40.1809

## Sheet1_1
Batch Folder Path: | D:\TraceFinderData\Projects\PFAS Online SPE\2022-07-15\2022-07-15-AB2
Batch Name: | 2022-07-15-AB2
Last Acquisition Date: | 2022-07-26 04:21
User Name: | Quantiva
Sample Count: | 342
Sample Amount - N/F
Data File | 11-Cl-PF3OUDS | 4-2:FTS | 6-2:FTS | 8-2:FTS | 9-Cl-PF3ONS | FBSA | FHxA | HFPO-DA | L-PFBS | L-PFHpS | L-PFHxS | L-PFOS | NaDONA | N-Et-FOSA | N-Me-FOSA | PFBA | PFDA | PFDoA | PFDS | PFHpA | PFHxA | PFHxDA | PFNA | PFNS | PFOA | PFODA | PFOSA | PFPeA | PFPeS | PFTeDA | PFTrDA | PFUdA | All Flags | Sample Type
AB-220715-01 Cal Std 1 1 | N/A | 1.818 | 9.435 | 352.285 | 1.048 | -25.205 | 4.359 | N/F | 1.235 | 1.124 | 1.65 | 1.566 | N/F | N/F | 1.04 | N/F | 1.209 | N/F | N/F | 1.262 | N/F | N/A | 1.19 | 0.951 | 1.244 | N/A | 1.179 | N/F | N/F | N/F | N/F | N/A |  | Cal Std
AB-220715-01 Cal Std 2 1-inj2 | N/A | 1.734 | 18.276 | -345.458 | 2.017 | 65.003 | 4.311 | 2.756 | 1.809 | 2.157 | 2.259 | 2.219 | N/F | N/A | 2.076 | N/F | 2.178 | N/F | N/F | 2.237 | 2.915 | N/A | 2.12 | 2.059 | 2.18 | N/A | 2.12 | N/F | N/F | N/A | N/F | N/A |  | Cal Std
AB-220715-01 Cal Std 3 1-inj2 | N/A | 3.2 | -40.277 | -529.529 | 3.942 | -12.433 | 4.266 | 4.106 | 3.71 | 3.579 | 3.685 | 3.587 | N/F | N/F | 3.849 | N/F | 4.036 | N/F | N/F | 4.048 | 4.587 | N/A | 3.898 | 3.712 | 3.99 | N/A | 3.716 | 3.263 | N/F | N/A | N/F | N/A |  | Cal Std
AB-220715-01 Cal Std 4 1 | N/A | 8.952 | -16.464 | -322.404 | 10.076 | 20.114 | 4.354 | 7.465 | 8.981 | 9.519 | 7.953 | 8.162 | N/F | N/A | 9.418 | N/F | 9.592 | N/F | N/A | 9.603 | 9.617 | N/A | 9.647 | 8.899 | 9.57 | N/A | 9.444 | 9.819 | N/F | N/A | N/F | N/A |  | Cal Std
AB-220715-01 Cal Std 5 1 | N/A | 18.959 | -40.753 | -114.05 | 19.101 | -26.361 | 4.303 | 20.512 | 15.471 | 19.107 | 14.817 | 15.556 | N/F | N/A | 19.672 | N/F | 19.349 | N/F | N/A | 18.541 | 17.849 | N/A | 19.29 | 18.904 | 18.982 | N/A | 19.008 | 20.334 | N/F | N/F | N/F | N/A |  | Cal Std
AB-220715-01 Cal Std 6 1 | N/A | 37.3 | -38.037 | -166.428 | 39.009 | -5.741 | 4.374 | 39.383 | 32.123 | 38.467 | 29.069 | 31.304 | N/F | N/F | 39.567 | N/F | 38.803 | N/F | N/F | 37.884 | 35.309 | N/A | 38.945 | 38.728 | 38.825 | N/A | 38.797 | 37.267 | N/F | N/F | N/F | N/A |  | Cal Std
AB-220715-01 Cal Std 7 1 | N/F | 87.826 | 179.625 | -102.989 | 100.642 | 224.831 | N/F | 88.145 | 78.091 | 101.464 | 72.445 | 76.777 | 187.562 | N/F | 98.961 | N/F | 96.261 | N/F | N/F | 94.675 | 90.964 | N/A | 96.207 | 101.248 | 96.896 | N/A | 98.559 | 98.822 | N/F | N/F | N/F | N/A |  | Cal Std
AB-220715-01 Cal Std 8 1 | N/A | 182.894 | 411.664 | -12.086 | 203.297 | 115.919 | 4.689 | 187.477 | 165.757 | 185.591 | 158.86 | 164.702 | N/F | N/F | 203.68 | 192.335 | 195.086 | N/F | N/F | 195.745 | 189.927 | N/A | 197.798 | 197.698 | 195 | N/A | 200.636 | 226.938 | N/F | N/F | N/F | N/A |  | Cal Std
AB-220715-01 Cal Std 9 1 | N/F | 400.125 | 753.755 | 1343.928 | 396.449 | 675.453 | N/F | 413.164 | 345.579 | 343.07 | 319.1 | 341.327 | N/F | N/F | 413.521 | 410.126 | 385.574 | N/F | N/A | 388.162 | 381.334 | N/A | 403.989 | 393.469 | 385.596 | N/A | 401.989 | 455.957 | N/F | N/F | N/A | N/A |  | Cal Std
AB-220715-01 Cal Std 10 1 | N/F | 930.676 | 706.139 | 1761.782 | 966.546 | 745.421 | N/F | 983.792 | 830.668 | 809.341 | 810.873 | 841.7 | 558.702 | N/F | 974.347 | 930.971 | 924.724 | N/F | N/A | 921.421 | 867.467 | N/A | 919.292 | 933.907 | 924.669 | N/A | 964.723 | 1000.928 | N/F | N/F | N/F | N/A |  | Cal Std
AB-220715-01 Cal Std 11 1 | N/F | 2009.026 | 1279.872 | 4274.39 | 2030.168 | N/F | 11.732 | 1959.323 | 1747.461 | 1709.43 | 1770.213 | 1829.534 | 1305.379 | N/F | 2008.033 | 2075.212 | 1916.836 | N/F | N/A | 1874.033 | 1844.46 | N/A | 1936.193 | 1864.811 | 1893.058 | N/A | 1984.146 | 2126.978 | N/F | N/F | N/F | N/A |  | Cal Std
AB-220715-01 Cal Std 12 1 | N/F | 4015.57 | 3942.167 | 3984.467 | 4160.47 | N/F | 41.168 | 4112.801 | 3719.585 | 3910.973 | 4004.568 | 3973.032 | 2721 | N/F | 4161.904 | 4301.969 | 3952.891 | N/F | N/A | 3972.222 | 4155.832 | N/A | 4063.315 | 3932.077 | 3945.156 | N/A | 4079.025 | 3641.29 | N/F | N/F | N/F | N/A |  | Cal Std
AB-220715-01 Cal Std 13 1 | N/F | 10078.922 | 10611.598 | 7653.093 | 9844.236 | N/F | 16193.444 | 9957.078 | 8816.874 | 9807.66 | 9042.279 | 9202.601 | 12327.357 | N/F | 9840.932 | 9689.387 | 10230.461 | N/F | N/A | 10257.165 | 10175.739 | N/A | 10085.115 | 9603.318 | 10261.834 | N/A | 9973.66 | 10235.505 | N/F | N/F | N/F | N/A |  | Cal Std
AB-220715-01 LFB-H 1 | N/F | 189.043 | -130.954 | 1514.823 | 218.691 | N/F | 2352.962 | 199.413 | 167.941 | 185.56 | 154.586 | 168.032 | 400.189 | N/F | 203.567 | 184.824 | 196.5 | N/F | N/F | 191.394 | 198.905 | N/A | 197.021 | 195.76 | 195.518 | N/A | 205.593 | 229.72 | N/F | N/F | N/F | N/A |  | Unknown
AB-220715-01 LRB 1_INJ001 | N/F | 1.922 | -117.186 | 698.465 | 1.097 | N/F | 12.054 | N/F | N/F | 0.867 | N/F | 1.258 | N/F | N/F | 1.2 | N/F | 0.984 | N/F | N/F | 0.81 | 1.749 | N/A | 0.917 | 0.762 | 1.349 | N/A | 1.682 | -0.523 | N/F | N/F | N/F | N/A | D,Q | Unknown
AB-220715-01 LRB 1 | N/F | 3.077 | -73.294 | 957.11 | 0.953 | N/F | 8.461 | N/F | N/F | 1.256 | N/F | 1.494 | 188.165 | N/F | 1.111 | N/F | 1.116 | N/F | N/F | 0.943 | 3.07 | N/A | 0.967 | 0.782 | 1.317 | N/A | 1.382 | 2.295 | N/F | N/F | N/F | N/A | D,Q | Unknown
AB-220715-01 CCC-H 1 | N/F | 394.539 | N/F | 2955.157 | 428.954 | N/F | 4591.05 | 402.071 | 352.033 | 364.636 | 325.628 | 344.494 | 641.003 | N/F | 413.658 | 402.94 | 398.521 | N/F | N/A | 405.085 | 385.132 | N/A | 401.704 | 391.166 | 391.959 | N/A | 409.662 | 494.343 | N/F | N/A | N/F | N/A |  | Unknown
AB-220715-01 IBC 1 | N/F | 182.915 | 86.836 | 145.893 | 212.813 | N/F | 29.709 | 197.324 | 173.785 | 185.06 | 163.267 | 171.927 | 406.554 | N/F | 2.943 | 191.996 | 195.534 | N/F | N/F | 202.175 | 192.575 | N/A | 202.398 | 3.904 | 196.228 | N/A | 2.735 | 234.057 | N/F | N/F | N/F | N/A |  | Unknown
AB-220715-01 QCS 1 | N/F | 1.868 | 606.745 | 170.048 | 3.173 | N/F | N/F | 153.347 | 176.792 | 180.09 | 161.888 | 172.792 | 189.522 | N/F | 1.243 | 180.3 | 191.245 | N/F | N/F | 196.7 | 187.09 | N/A | 198.336 | 199.617 | 193.982 | N/A | 1.77 | 243.092 | N/F | N/F | N/F | N/A |  | Unknown
WO-22270-01-A-1 | N/F | 24.292 | 85.372 | 177.742 | 0.409 | 6175.874 | N/F | N/F | N/F | 1.417 | N/F | 1.107 | N/F | N/F | 0.963 | N/F | 4.515 | N/F | N/A | N/F | N/F | N/A | 1.284 | 0.343 | 2.455 | N/A | 0.639 | N/F | N/F | N/F | N/F | N/A | C | Unknown
WO-22270-02-A-1 | N/F | 13.878 | 51.058 | -146.244 | 0.801 | 5948.936 | N/F | N/F | N/F | 0.979 | N/F | 1.166 | N/F | N/F | 0.762 | N/F | 4.114 | N/F | N/A | N/F | N/F | N/A | 0.801 | 0.191 | 2.629 | N/A | 0.531 | N/F | N/F | N/F | N/F | N/A | C | Unknown
WO-22270-03-A-1 | N/F | 33.292 | -332.436 | -202.581 | 0.706 | 3782.3 | N/F | N/F | N/F | 0.649 | N/F | 0.982 | N/F | N/F | 0.799 | N/F | 3.879 | N/F | N/A | N/F | N/F | N/A | 0.734 | 0.2 | 1.389 | N/A | 0.504 | N/F | N/F | N/F | N/F | N/A |  | Unknown
WO-22270-04-A-1 | N/F | 24.578 | 193.196 | -82.625 | 0.323 | 4588.594 | 6.732 | N/F | N/F | N/F | N/F | 1.082 | N/F | N/F | 0.848 | N/F | 3.959 | N/F | N/A | N/F | N/F | N/A | 0.854 | 0.254 | 1.569 | N/A | 0.453 | N/F | N/F | N/A | N/F | N/A |  | Unknown
WO-22270-05-A-1 | N/F | 28.831 | -222.137 | -251.473 | 0.341 | 4983.912 | N/F | N/F | N/F | N/F | N/F | 1.161 | N/F | N/F | 0.771 | N/F | 4.702 | N/F | N/A | 1.496 | N/F | N/A | 0.677 | 0.982 | 3.116 | N/A | 0.414 | N/F | N/F | N/F | N/F | N/A |  | Unknown
WO-22270-06-A-1 | N/F | 36.226 | -107.551 | -283.067 | 0.535 | 3286.822 | N/F | 3.727 | N/F | 0.676 | N/F | 1.226 | N/F | N/F | 0.863 | N/F | 3.878 | N/F | N/A | N/F | N/F | N/A | 1.073 | 0.106 | 3.167 | N/A | 0.422 | N/F | N/F | N/A | N/A | N/A |  | Unknown
WO-22270-07-A-1 | N/F | 21.669 | -92.364 | 98.181 | 0.253 | 2959.834 | 7.197 | N/F | N/F | N/F | N/F | 0.834 | N/F | N/F | 0.84 | N/F | 3.862 | N/F | N/A | N/F | N/F | N/A | 0.509 | 0.118 | 2.314 | N/A | 0.383 | N/F | N/F | N/A | N/F | N/A |  | Unknown
WO-22270-08-A-1 | N/F | 24.927 | -48.896 | -286.629 | 0.513 | 7371.95 | 6.516 | N/F | N/F | 0.849 | N/F | 0.938 | N/F | N/F | 0.649 | N/F | 4.404 | N/F | N/A | N/F | N/F | N/A | 0.848 | 0.911 | 2.862 | N/A | 0.411 | N/F | N/F | N/A | N/F | N/A | C | Unknown
WO-22270-09-A-1 | N/F | 26.592 | -139.748 | -324.447 | 0.287 | 1554.936 | 4.981 | N/F | N/F | 0.657 | N/F | 2.294 | N/F | N/F | 0.735 | N/F | 5.182 | N/F | N/A | N/F | N/F | N/A | 1.009 | 0.595 | 1.798 | N/A | 0.437 | N/F | N/F | N/A | N/F | N/A |  | Unknown
WO-22270-10-A-1 | N/F | 27.52 | -171.328 | 55.63 | 0.221 | 2215.248 | 5.815 | N/F | N/F | 0.65 | N/F | 0.933 | N/F | N/A | 0.734 | N/F | 6.309 | N/F | N/A | N/F | N/F | N/A | 0.804 | 0.56 | 2.843 | N/A | 0.354 | N/F | N/F | N/F | N/F | N/A |  | Unknown
AB-220715-01 LRB 2_INJ001 | N/F | 3.785 | -245.748 | -68.982 | 0.379 | 42.315 | 4.974 | N/F | N/F | 0.341 | 1.014 | 0.918 | N/F | N/A | 0.917 | N/F | 0.443 | N/F | N/A | N/F | N/F | N/A | 0.708 | 0.203 | 1.595 | N/A | 0.342 | -0.603 | N/F | N/A | N/F | N/A | D,Q | Unknown
AB-220715-01 LRB 2 | N/F | 4.965 | -84.895 | -481.747 | 0.387 | 62.421 | 4.839 | N/F | N/F | 0.338 | N/F | 1.055 | N/F | N/F | 0.915 | N/F | 0.367 | N/F | N/A | N/F | N/F | N/A | 0.802 | 0.047 | 1.568 | N/A | 0.398 | -0.336 | N/F | N/A | N/F | N/A | D,Q | Unknown
WO-22270-11-A-1 | N/F | 15.981 | -320.785 | -106.147 | 0.41 | 430.438 | N/F | N/F | N/F | 1.135 | N/F | 2.047 | 187.783 | N/A | 0.851 | N/F | 4.526 | N/F | N/A | N/F | N/F | N/A | 0.482 | 0.546 | 1.094 | N/A | 0.366 | N/F | N/F | N/F | N/A | N/A |  | Unknown
WO-22270-12-A-1 | N/F | 28.86 | -155.394 | -173.09 | 0.203 | 1147.317 | 6.071 | N/F | N/F | 0.537 | N/F | 0.921 | N/F | N/A | 0.77 | N/F | 3.046 | N/F | N/A | 1.008 | N/F | N/A | 0.455 | 0.057 | 2.504 | N/A | 0.369 | N/F | N/F | N/F | N/F | N/A |  | Unknown
WO-22270-13-A-1 | N/F | 22.779 | -80.718 | -23.135 | 0.36 | 1149.207 | 5.747 | 1.644 | N/F | 0.623 | 0.942 | 0.867 | N/F | N/A | 0.975 | N/F | 4.196 | N/F | N/A | N/F | N/F | N/A | 0.516 | 0.481 | N/F | N/A | 0.425 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-14-A-1 | N/F | 14.976 | -181.535 | -466.736 | 0.184 | 2286.665 | 5.506 | N/F | N/F | 1.446 | N/F | 0.847 | N/F | N/A | 1.071 | N/F | 4.113 | N/A | N/A | N/F | N/F | N/A | 0.372 | 0.448 | N/F | N/A | 0.367 | N/F | N/F | N/A | N/F | N/A |  | Unknown
WO-22270-15-A-1 | N/F | 24.808 | -120.171 | -380.016 | 0.226 | 1863.968 | 5.736 | N/F | N/F | 0.82 | N/F | 1.973 | 187.636 | N/A | 1.086 | N/F | 2.456 | N/F | N/A | N/F | N/F | N/A | 0.761 | 0.5 | 2.884 | N/A | 0.374 | N/F | N/F | N/F | N/F | N/A |  | Unknown
WO-22270-16-A-1 | N/F | 31.508 | 5.894 | -300.746 | 0.2 | 2588.437 | 6.466 | N/F | N/F | 0.757 | N/F | 0.965 | N/F | N/A | 0.934 | N/F | 4.631 | N/A | N/A | N/F | N/F | N/A | 0.681 | 0.554 | 2.347 | N/A | 0.356 | N/F | N/F | N/F | N/F | N/A |  | Unknown
WO-22270-17-A-1 | N/F | 28.338 | -180.179 | -310.139 | 0.257 | 281.051 | 5.48 | N/F | N/F | 0.768 | N/F | 0.832 | 187.841 | N/A | 0.946 | N/F | 4.525 | N/F | N/A | 0.974 | N/F | N/A | 0.916 | 0.04 | 3.796 | N/A | 0.338 | N/F | N/F | N/F | N/F | N/A |  | Unknown
WO-22270-18-A-1 | N/F | 31.56 | 47.573 | -136.151 | 0.203 | 589.695 | 5.553 | N/F | N/F | 1.356 | 1.431 | 0.83 | 187.979 | N/F | 1.221 | N/F | 1.851 | N/F | N/A | N/F | N/F | N/A | 0.394 | 0.528 | 1.384 | N/A | 0.354 | N/F | N/F | N/A | N/F | N/A |  | Unknown
WO-22270-19-A-1 | N/F | 25.166 | -106.945 | -465.721 | 0.216 | 2527.997 | 5.513 | 1.644 | N/F | 1.189 | N/F | 0.866 | N/F | N/A | 0.93 | N/F | 3.432 | N/F | N/A | N/F | N/F | N/A | 0.482 | 0.539 | 1.812 | N/A | 0.351 | N/F | N/F | N/F | N/A | N/A |  | Unknown
WO-22270-20-A-1 | N/F | 53.096 | -359.243 | -466.001 | 0.214 | 208.349 | 4.773 | 2.717 | N/F | 1.284 | N/F | 0.99 | 187.828 | N/A | 0.88 | N/F | 5.105 | N/F | N/A | N/F | N/F | N/A | 0.447 | 0.487 | 2.423 | N/A | 0.382 | N/F | N/A | N/F | N/F | N/A |  | Unknown
AB-220715-01 LFB-H 2 | N/A | 184.906 | 113630.304 | 979.726 | 205.592 | 2546.837 | 2525.067 | 194.921 | 161.04 | 182.148 | 152.172 | 160.442 | 389.669 | N/A | 205.831 | 159.874 | 193.606 | N/F | N/A | 190.051 | 187.047 | N/A | 192.039 | 183.021 | 190.583 | N/A | 199.766 | 217.366 | N/F | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 3_INJ001 | N/F | 6.882 | 7.627 | 251.555 | 0.409 | 263.867 | 7.266 | N/F | N/F | 0.45 | 1.095 | 1.016 | N/F | N/A | 0.916 | N/F | 0.521 | N/F | N/A | N/F | 1.04 | N/A | 0.487 | 0.238 | 1.137 | N/A | 0.747 | -1.623 | N/F | N/A | N/F | N/A | D,Q | Unknown
AB-220715-01 LRB 3 | N/F | 3.868 | -114.955 | 27.184 | 0.229 | 914.719 | 5.356 | N/F | N/F | 0.584 | 0.969 | 0.842 | N/F | N/A | 0.987 | N/F | 0.4 | N/F | N/A | N/F | 0.977 | N/A | 0.453 | 0.407 | 1.107 | N/A | 0.393 | -0.875 | N/A | N/A | N/F | N/A | D,Q | Unknown
AB-220715-01 LD 1 | N/A | 37.29 | 20.936 | -120.81 | 0.204 | 1966.004 | 5.821 | 2.029 | N/F | 0.683 | 1.804 | 0.834 | N/F | N/A | 0.997 | N/F | 3.428 | N/F | N/A | 0.935 | N/F | N/A | 0.621 | 0.216 | 2.867 | N/A | 0.39 | N/F | N/A | N/A | N/F | N/A |  | Unknown
AB-220715-01 CCC-H 2 | N/F | 387.163 | 397810.785 | 35616.145 | 388.636 | 4429248.499 | 4596.8 | 399.928 | 347.787 | 348.057 | 320.885 | 336.485 | 621.876 | N/A | 401.9 | 364.718 | 386.616 | N/F | N/A | 382.625 | 380.048 | N/A | 387.617 | 350.75 | 378.815 | N/A | 388.012 | 542.884 | N/A | N/A | N/F | N/A |  | Unknown
AB-220715-01 LFSM-H 1 | N/F | 200.539 | 202856.498 | 45270.74 | 199.423 | 2276612.707 | 4148.379 | 188.888 | 164.433 | 228.183 | 151.916 | 149.476 | 365.052 | N/A | 200.828 | N/F | 196.881 | N/A | N/A | 194.52 | 194.271 | N/A | 194.736 | 215.661 | 194.121 | N/A | 203.344 | 142.788 | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-21 | N/F | 22.425 | 6.685 | 65.907 | 0.473 | 6553.62 | 11.657 | 2.659 | N/F | 1.599 | N/F | 1.043 | 188.036 | N/A | 0.863 | N/F | 3.786 | N/F | N/A | 1.496 | N/F | N/A | 0.654 | 0.388 | 4.374 | N/A | 0.79 | N/F | N/F | N/F | N/A | N/A | C | Unknown
WO-22270-22 | N/F | 32.433 | -70.052 | -379.965 | 0.284 | 1030.003 | 6.405 | N/F | N/F | 1.279 | N/F | 0.808 | N/F | N/A | 0.842 | N/F | 3.208 | N/F | N/A | N/F | N/F | N/A | 0.608 | 0.207 | 1.838 | N/A | 0.55 | N/F | N/F | N/F | N/F | N/A |  | Unknown
WO-22270-23 | N/F | 38.47 | -210.781 | 740.972 | 0.24 | 1422.215 | 7.404 | N/F | N/F | 1.168 | N/F | 0.896 | N/F | N/A | 0.95 | N/F | 4.061 | N/F | N/A | 0.833 | 2.005 | N/A | 0.641 | 0.508 | 3.327 | N/A | 0.393 | N/F | N/F | N/F | N/F | N/A |  | Unknown
WO-22270-24 | N/F | 30.713 | 111.542 | -484.943 | 0.228 | 957.291 | 6.573 | N/F | N/F | 1.376 | N/F | 1.016 | N/F | N/A | 0.921 | N/F | 3.347 | N/A | N/A | N/F | N/F | N/A | 0.582 | 0.295 | 2.263 | N/A | 0.356 | N/F | N/F | N/A | N/F | N/A |  | Unknown
WO-22270-25 | N/F | 22.896 | -281.866 | -239.709 | 0.217 | 656.068 | 6.098 | N/F | N/F | 0.7 | N/F | 1.775 | 187.758 | N/A | 0.765 | N/F | 4.386 | N/F | N/A | N/F | N/F | N/A | 0.552 | 0.22 | N/F | N/A | 0.404 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-26 | N/F | 37.312 | 132.811 | -316.779 | N/F | 1561.491 | 5.905 | N/F | N/F | 0.866 | N/F | 2.491 | N/F | N/A | 1.161 | N/F | 3.774 | N/A | N/A | 0.998 | N/F | N/A | 0.619 | 0.558 | 2.693 | N/A | 0.337 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-27 | N/F | 23.13 | 2.289 | -407.11 | 0.191 | 1001.232 | 4.838 | N/F | N/F | 1.137 | N/F | 0.946 | 187.897 | N/A | 1.037 | N/F | 3.312 | N/A | N/A | 0.948 | N/F | N/A | 0.47 | 0.307 | 1.492 | N/A | 0.385 | N/F | N/F | N/A | N/F | N/A |  | Unknown
WO-22270-28 | N/F | 24.68 | -121.78 | -432.129 | 0.18 | 265 | 5.098 | N/F | N/F | 0.697 | N/F | 0.892 | N/F | N/A | 0.904 | N/F | 4.549 | N/A | N/A | N/F | N/F | N/A | 0.471 | 0.034 | N/F | N/A | 0.324 | N/F | N/F | N/A | N/A | N/A |  | Unknown
WO-22270-29 | N/F | 27.328 | -147.154 | 93.347 | 0.186 | 2261.44 | 5.124 | N/F | N/F | 0.668 | N/F | 0.826 | N/F | N/A | 0.64 | N/F | 4.642 | N/A | N/A | N/F | N/F | N/A | 0.539 | 0.263 | 0.799 | N/A | 0.35 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-30 | N/F | 25.218 | -186.109 | 12.111 | 0.177 | 198.397 | 4.866 | 2.817 | N/F | 0.926 | N/F | 1.847 | N/F | N/A | 0.762 | N/F | 4.086 | N/F | N/A | 1.255 | 1.854 | N/A | 0.836 | 0.24 | 2.664 | N/A | 0.313 | N/F | N/A | N/A | N/F | N/A |  | Unknown
AB-220715-01 LRB 4_INJ001 | N/F | 5.692 | 22.562 | -121 | 0.171 | -16.951 | 4.367 | N/F | N/F | 0.348 | 0.968 | 1.507 | N/F | N/A | 1.156 | N/F | 0.328 | N/F | N/A | N/F | 1.03 | N/A | 0.597 | 0.239 | 1.15 | N/A | 0.286 | -2.028 | N/A | N/A | N/F | N/A | D,Q | Unknown
AB-220715-01 LRB 4 | N/F | 7.094 | -130.591 | -115.158 | 0.178 | 20.383 | 4.614 | N/F | N/F | 0.405 | 0.951 | 0.839 | N/F | N/A | 1.018 | N/F | 0.33 | N/F | N/F | N/F | 1.05 | N/A | 0.398 | 0.146 | 1.046 | N/A | 0.284 | -0.494 | N/F | N/F | N/F | N/A | D,Q | Unknown
WO-22270-31 | N/F | 38.18 | 157.496 | 13.156 | 0.179 | N/F | 4.846 | 0.624 | N/F | 1.16 | N/F | 0.864 | 187.756 | N/A | 0.831 | N/F | 2.993 | N/A | N/A | N/F | N/F | N/A | 1.049 | 0.261 | 1.614 | N/A | 0.373 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-32 | N/F | 38.234 | 40.312 | N/F | 0.22 | 107.206 | 4.777 | N/F | N/F | 0.528 | N/F | 2.278 | 187.979 | N/A | 0.911 | N/F | 3.971 | N/F | N/A | N/F | N/F | N/A | 0.671 | 0.328 | 1.812 | N/A | 0.345 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-33 | N/A | 28.871 | -19.527 | 245.309 | 0.182 | 557.936 | 5.647 | 0.752 | N/F | 0.664 | N/F | 0.899 | 187.795 | N/A | 0.892 | N/F | 3.255 | N/A | N/A | N/F | N/F | N/A | 0.42 | 0.19 | 2.319 | N/A | 0.331 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-34 | N/F | 49.009 | 483.692 | 20.303 | 0.512 | N/F | 4.704 | N/F | N/F | 2.173 | N/F | 1.188 | 188.404 | N/A | 0.868 | N/F | 4.461 | N/A | N/A | 1.774 | N/F | N/A | 0.998 | 0.873 | 3.094 | N/A | 0.358 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-35 | N/F | 25.252 | -140.842 | 672.749 | 0.178 | 94.903 | 4.293 | N/F | N/F | 0.97 | N/F | 1.696 | 188.321 | N/A | 0.897 | N/F | 2.965 | N/A | N/A | 1.682 | N/F | N/A | 1.591 | 0.313 | 0.962 | N/A | 0.331 | -1.244 | N/A | N/A | N/F | N/A | D,Q | Unknown
WO-22270-36 | N/F | 16.799 | -241.736 | -206.944 | 0.176 | 582.572 | 4.784 | 0.078 | N/F | 0.609 | N/F | 1.719 | N/F | N/A | 0.795 | N/F | 3.607 | N/F | N/A | N/F | N/F | N/A | 0.517 | 0.259 | 1.501 | N/A | 0.335 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-37 | N/F | 39.245 | 56.791 | 124.41 | 0.18 | N/F | 4.343 | N/F | 0.915 | 0.687 | N/F | 1.024 | 187.741 | N/A | 0.849 | N/F | 3.229 | N/A | N/A | N/F | N/F | N/A | 0.447 | 0.183 | 1.387 | N/A | 0.308 | N/F | N/A | N/A | N/F | N/A | D,Q | Unknown
WO-22270-38 | N/F | 39.077 | -11.461 | -242.322 | 0.173 | 73.51 | 4.53 | N/F | 1.299 | 0.749 | N/F | 0.958 | 187.744 | N/A | 0.899 | N/F | 2.328 | N/A | N/A | N/F | N/F | N/A | 0.48 | 0.376 | 1.752 | N/A | 0.349 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-39 | N/A | 45.013 | -159.732 | 170.022 | 0.183 | 140.387 | 5.026 | N/F | N/F | 0.773 | N/F | 2.016 | 187.866 | N/A | 0.908 | N/F | 2.578 | N/F | N/A | N/F | N/F | N/A | 0.429 | 0.211 | 1.962 | N/A | 0.346 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-40 | N/F | 42.996 | -338.47 | -239.222 | 0.17 | 421.879 | 4.508 | 1.034 | N/F | 0.742 | N/F | 0.952 | 187.789 | N/A | 0.862 | N/F | 3.462 | N/F | N/A | N/F | N/F | N/A | 1.428 | 0.104 | 0.761 | N/A | 0.33 | N/F | N/A | N/A | N/F | N/A |  | Unknown
AB-220715-01 LFB-H 3 | N/A | 191.433 | 204953.821 | 1065183.17 | 198.893 | 2534851.855 | 2425.241 | 196.235 | 168.028 | 177.805 | 153.757 | 164.912 | 396.715 | N/A | 204.158 | 157.99 | 198.894 | N/F | N/A | 194.119 | 190.785 | N/A | 199.696 | 173.364 | 199.713 | N/A | 203.969 | 205.399 | N/A | N/A | N/F | N/A |  | Unknown
AB-220715-01 LRB 5_INJ001 | N/F | 7.584 | -281.959 | 949.865 | 0.402 | 701.571 | 6.195 | -0.174 | N/F | 0.458 | 1.63 | 0.988 | 187.563 | N/A | 1.024 | N/F | 0.496 | N/F | N/A | N/F | 0.976 | N/A | 0.45 | 0.224 | 0.766 | N/A | 0.681 | -1.056 | N/A | N/A | N/F | N/A | D,Q | Unknown
AB-220715-01 LRB 5 | N/F | 8.422 | 130.084 | 79.684 | 0.244 | 584.304 | 4.815 | -0.157 | 0.683 | 0.348 | 1.69 | 1.616 | N/F | N/A | 1.069 | N/F | 0.431 | N/F | N/A | N/F | 1.012 | N/A | 0.41 | 0.144 | 0.861 | N/A | 0.392 | 0.369 | N/A | N/F | N/F | N/A | D,Q | Unknown
AB-220715-01 LD 2 | N/F | 45.096 | -20.249 | 335.456 | 0.193 | 560.067 | 6.321 | N/F | 1.048 | 0.866 | 2.332 | 1.012 | 188.049 | N/A | 0.873 | N/F | 2.589 | N/F | N/A | 1.392 | N/F | N/A | 1.059 | 0.308 | 4.203 | N/A | 0.334 | N/F | N/A | N/A | N/F | N/A |  | Unknown
AB-220715-01 CCC-H 3 | N/A | 388.814 | 399990.35 | 2271882.462 | 389.626 | 5436115.528 | 4664.145 | 397.811 | 340.157 | 346.509 | 317.243 | 332.226 | 616.069 | N/A | 398.999 | 361.667 | 389.63 | N/A | N/A | 390.952 | 382.755 | N/A | 391.511 | 328.05 | 391.053 | N/A | 405.276 | 409.892 | N/A | N/A | N/F | N/A |  | Unknown
AB-220715-01 LFSM-H 2 | N/F | 202.915 | 199860.162 | 1136033.921 | 209.084 | 3807941.254 | 3781.142 | 184.578 | 168.055 | 209.163 | 150.282 | 152.35 | 372.471 | N/A | 205.796 | N/F | 196.473 | N/F | N/A | 198.245 | 183.486 | N/A | 196.189 | 205.737 | 198.046 | N/A | 198.227 | 161.408 | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-41 | N/F | 36.121 | -79.897 | 843.726 | 0.243 | 1065.53 | 7.473 | N/F | N/F | 0.758 | N/F | 0.97 | N/F | N/A | 0.884 | N/F | 2.578 | N/F | N/A | N/F | N/F | N/A | 0.66 | 0.34 | 0.959 | N/A | 0.412 | N/F | N/A | N/F | N/A | N/A |  | Unknown
WO-22270-42 | N/F | 29.638 | -151.906 | 53.387 | 0.214 | 1780.241 | 5.602 | 2.162 | 1.077 | 0.842 | N/F | 0.841 | 188.125 | N/A | 0.877 | N/F | 1.633 | N/A | N/A | N/F | N/F | N/A | 0.428 | 0.434 | 1.255 | N/A | 0.378 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-43 | N/A | 48.714 | -32.864 | -302.054 | 0.204 | 1765.114 | 5.076 | N/F | N/F | 0.681 | N/F | 0.917 | 187.736 | N/A | 0.792 | N/F | 2.633 | N/F | N/A | N/F | N/F | N/A | 0.663 | 0.425 | 4.642 | N/A | 0.357 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-44 | N/F | 31.123 | 487.277 | 444.975 | 0.199 | 672.882 | 4.694 | 1.368 | N/F | 0.792 | N/F | 0.928 | 188.324 | N/A | 0.911 | N/F | 1.905 | N/A | N/A | N/F | N/F | N/A | 0.442 | 0.325 | 1.306 | N/A | 0.303 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-45 | N/A | 33.222 | -210.097 | 379.669 | 0.176 | 1818.866 | 4.687 | 3.345 | 1.289 | 0.595 | N/F | 0.858 | 187.705 | N/A | 0.762 | N/F | 1.785 | N/F | N/A | 3.221 | N/F | N/A | 1.441 | 0.212 | N/F | N/A | 0.359 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-52 | N/A | 59.001 | 94.326 | 484.088 | 0.189 | 216.345 | 4.697 | 3.526 | 1.25 | 0.597 | N/F | 2.063 | 188.149 | N/A | 1.078 | N/F | 2.789 | N/A | N/A | N/F | N/F | N/A | 0.646 | 0.284 | 1.301 | N/A | 0.334 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-46 | N/F | 31.053 | 253.48 | 44.055 | 0.183 | N/F | 4.479 | 2.103 | 1.039 | 0.679 | N/F | 1.511 | 187.71 | N/A | 0.843 | N/F | 3.148 | N/F | N/A | N/F | N/F | N/A | 0.561 | 0.255 | 1.241 | N/A | 0.35 | N/F | N/A | N/A | N/F | N/A | D,Q | Unknown
WO-22270-47 | N/F | 32.404 | 52.566 | 434.154 | 0.186 | 610.917 | 4.535 | 0.861 | N/F | 0.8 | N/F | 0.922 | 187.967 | N/A | 0.825 | N/F | 1.894 | N/A | N/A | N/F | N/F | N/A | 0.441 | 0.241 | 1.16 | N/A | 0.323 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-53 | N/A | 37.429 | 139.613 | -116.772 | 0.176 | N/F | 4.363 | 1.472 | 1.507 | 0.769 | N/F | 0.826 | 187.683 | N/A | 0.94 | N/F | 3.553 | N/F | N/A | N/F | N/F | N/A | 0.685 | 0.189 | 1.634 | N/A | 0.37 | N/F | N/A | N/A | N/F | N/A | D,Q | Unknown
WO-22270-48 | N/F | 29.045 | 109.329 | 137.623 | 0.181 | 453.766 | 5.949 | 1.281 | N/F | 0.682 | N/F | 1.93 | 187.782 | N/A | 0.915 | N/F | 3.977 | N/A | N/A | N/F | N/F | N/A | 0.378 | 0.218 | 0.488 | N/A | 0.324 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-54 | N/A | 20.422 | -137.626 | 327.654 | 0.172 | N/F | 4.419 | 3.071 | N/F | 0.602 | N/F | 1.31 | 187.582 | N/A | 0.893 | N/F | 3.216 | N/F | N/A | N/F | 1.304 | N/A | 0.375 | 0.155 | 0.828 | N/A | 0.316 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-49 | N/A | 32.428 | 28.722 | -73.564 | 0.172 | 1899.826 | 4.755 | N/F | N/F | 0.677 | N/F | 0.988 | 187.898 | N/A | 0.803 | N/F | 4.442 | N/A | N/A | N/F | N/F | N/A | 0.497 | 0.263 | 0.776 | N/A | 0.366 | 17.052 | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-55 | N/A | 37.057 | -7.986 | 38.086 | 0.21 | 220.186 | N/F | 2.916 | 1.289 | 0.726 | N/F | 3.016 | 188.052 | N/A | 0.837 | N/F | 2.961 | N/A | N/A | N/F | N/F | N/A | 3.128 | 0.453 | N/F | N/A | 0.404 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-50 | N/F | 31.937 | 300.571 | -79.736 | 0.179 | 157.73 | 4.388 | N/F | 0.742 | 0.711 | N/F | 0.949 | 187.973 | N/A | 0.951 | N/F | 4.275 | N/F | N/A | N/F | N/F | N/A | 1.212 | 0.365 | 0.833 | N/A | 0.328 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-56 | N/A | 42.16 | -76.376 | -201.312 | 0.177 | 1464.114 | 4.694 | N/F | 1.308 | 0.534 | N/F | 0.891 | 187.762 | N/A | 0.859 | N/F | 4.414 | N/F | N/A | N/F | N/F | N/A | 0.793 | 0.117 | 1.763 | N/A | 0.366 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-57 | N/F | 20.638 | -208.802 | -85.049 | 0.168 | 190.818 | 4.445 | 2.437 | 1.272 | 0.728 | N/F | 1.085 | 187.988 | N/A | 0.722 | N/F | 4.114 | N/A | N/A | N/F | N/F | N/A | 0.685 | 0.05 | 0.912 | N/A | 0.34 | N/F | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 6_INJ001 | N/F | 8.845 | -55.064 | 91.123 | 0.182 | 4.099 | 4.463 | 0.438 | N/F | 0.286 | 1.676 | 1.236 | 187.637 | N/A | 1.069 | N/F | 0.36 | N/F | N/F | N/F | 0.966 | N/A | 0.346 | 0.121 | 0.485 | N/A | 0.325 | -2.567 | N/A | N/A | N/F | N/A | D,Q | Unknown
AB-220715-01 LRB 6 | N/A | 7.103 | -46.72 | -203.463 | 0.182 | N/F | 4.335 | 1.195 | 0.557 | 0.301 | 1.456 | 0.934 | 187.569 | N/A | 1.084 | N/F | 0.365 | N/A | N/F | N/F | 1.007 | N/A | 0.342 | 0.135 | 0.535 | N/A | 0.329 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-58 | N/A | 60.784 | 439.04 | 170.062 | 0.175 | N/F | N/F | N/F | 2.421 | 0.965 | N/F | 1.217 | 187.901 | N/A | 0.856 | N/F | 3.988 | N/A | N/A | N/F | N/F | N/A | 0.383 | 0.348 | 1.42 | N/A | 0.34 | N/F | N/A | N/A | N/F | N/A | D,Q | Unknown
WO-22270-51 | N/F | 25.845 | 658.879 | 489.993 | 0.173 | N/F | 4.602 | N/F | N/F | 0.761 | N/F | 1.713 | 187.606 | N/A | 0.761 | N/F | 2.866 | N/F | N/A | N/F | N/F | N/A | 0.944 | 0.186 | 4.097 | N/A | 0.316 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-59 | N/A | 38.107 | 32.296 | -100.643 | 0.171 | 34.57 | 4.641 | N/F | 1.069 | 0.537 | N/F | 1.563 | 187.745 | N/A | 0.972 | N/F | 3.171 | N/F | N/A | N/F | N/F | N/A | 0.473 | 0.127 | 0.765 | N/A | 0.349 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-60 | N/F | 40.536 | 9.291 | 702.691 | 0.174 | 260.992 | 4.547 | N/F | 2.623 | 0.69 | N/F | 0.999 | 187.794 | N/A | 0.912 | N/F | 3.518 | N/A | N/A | N/F | N/F | N/A | 1.577 | 0.091 | 3.544 | N/A | 0.34 | N/F | N/A | N/A | N/F | N/A |  | Unknown
AB-220715-01 LFB-H 4 | N/A | 181.393 | 205949.66 | 1215643.239 | 204.863 | 2564720.119 | 2282.495 | 197.68 | 175.872 | 182.263 | 157.175 | 165.163 | 401.16 | N/A | 209.15 | 162.744 | 198.772 | N/A | N/A | 196.702 | 199.009 | N/A | 204.168 | 164.372 | 199.897 | N/A | 203.969 | 200.013 | N/A | N/A | N/F | N/A |  | Unknown
AB-220715-01 LRB 7_INJ001 | N/F | 8.831 | 96.429 | 1035.383 | 0.398 | 859.46 | 7.397 | -0.005 | 0.669 | 0.499 | 1.998 | 1.365 | 187.729 | N/A | 1.318 | N/F | 0.506 | N/F | N/A | N/F | 1.05 | N/A | 0.567 | 0.236 | 1.489 | N/A | 0.715 | -2.539 | N/A | N/F | N/F | N/A | D,Q | Unknown
AB-220715-01 LRB 7 | N/F | 9.454 | -45.321 | 589.02 | 0.231 | 481.512 | 4.788 | 0.5 | N/F | 0.381 | 2.109 | 0.94 | 187.633 | N/F | 1.398 | N/F | 0.397 | N/F | N/A | N/F | 0.992 | N/A | 0.776 | 0.176 | 1.449 | N/A | 0.397 | N/F | N/A | N/F | N/F | N/A |  | Unknown
AB-220715-01 LD 3 | N/F | 41.583 | 548.163 | 628.968 | 0.204 | 539.1 | 5.759 | N/F | 1.254 | 1.038 | N/F | 0.955 | 187.727 | N/A | 0.871 | N/F | 3.264 | N/F | N/A | 1.952 | N/F | N/A | 0.487 | 0.373 | 1.999 | N/A | 0.344 | 5.741 | N/A | N/F | N/F | N/A |  | Unknown
AB-220715-01 CCC-H 4 | N/A | 348.664 | 380225.664 | 2358966.15 | 391.157 | 6954813.794 | 5338.629 | 388.596 | 332.142 | 353.136 | 321.727 | 331.024 | 614.304 | N/A | 402.846 | 336.451 | 382.399 | N/F | N/A | 384.014 | 377.704 | N/A | 386.028 | 362.543 | 387.419 | N/A | 404.275 | 472.636 | N/A | N/F | N/F | N/A |  | Unknown
AB-220715-01 LFSM-H 3 | N/F | 181.278 | 207815.95 | 1162763.306 | 196.814 | 4876860.321 | 4987.276 | 191.02 | 176.858 | 235.491 | 151.448 | 154.815 | 363.347 | N/A | 203.658 | N/F | 200.285 | N/F | N/A | 196.918 | 182.049 | N/A | 198.045 | 211.178 | 197.926 | N/A | 206.493 | 187.576 | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-61 | N/F | 40.546 | 76.828 | -169.15 | 0.26 | 1770.586 | 6.746 | N/F | 1.818 | 1.026 | N/F | 2.014 | 187.654 | N/A | 0.84 | N/F | 3.348 | N/F | N/A | N/F | N/F | N/A | 0.482 | 0.37 | 0.83 | N/A | 0.423 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-62 | N/F | 40.311 | -94.139 | -139.346 | 0.226 | 496.242 | 5.911 | 1.249 | 1.636 | 0.981 | N/F | 2.345 | 187.819 | N/A | 0.712 | N/F | 3.215 | N/F | N/A | N/F | N/F | N/A | 0.586 | 0.099 | 1.122 | N/A | 0.334 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-63 | N/F | 51.491 | -196.216 | 241.773 | 0.209 | 240.577 | 6.226 | N/F | 2.378 | 1.4 | N/F | 1.061 | 188.027 | N/A | 0.752 | N/F | 3.361 | N/F | N/A | N/F | N/F | N/A | 0.534 | 0.363 | 1.792 | N/A | 0.38 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-64 | N/F | 65.189 | 558.335 | N/F | 0.199 | 819.809 | 5.129 | 2.491 | 2.725 | 1.01 | N/F | 0.884 | 187.843 | N/A | 0.94 | N/F | 3.534 | N/F | N/A | N/F | N/F | N/A | 3.413 | 0.591 | N/F | N/A | 0.376 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-65 | N/F | 32.06 | 1042.275 | 483.673 | 0.186 | 291.884 | 4.407 | N/F | 0.788 | 0.562 | N/F | 1.601 | 187.928 | N/A | 0.924 | N/F | 3.516 | N/F | N/A | N/F | 1.145 | N/A | 0.555 | 0.195 | 0.741 | N/A | 0.335 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-66 | N/A | 43.678 | 534.79 | 16.302 | 0.193 | N/F | 4.632 | 3.031 | 3.244 | 0.777 | N/F | 2.321 | 187.655 | N/A | 0.814 | N/F | 3.2 | N/F | N/A | N/F | N/F | N/A | 0.722 | 0.011 | 1.099 | N/A | 0.308 | N/F | N/A | N/A | N/F | N/A | D,Q | Unknown
WO-22270-67 | N/F | 47.877 | 490.498 | -232.519 | 0.188 | 350.882 | 4.846 | 2.759 | 2.024 | 0.644 | N/F | 2.084 | 187.85 | N/A | 0.909 | N/F | 4.157 | N/F | N/A | N/F | N/F | N/A | 20.342 | 0.369 | 2.636 | N/A | 0.363 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-68 | N/A | 41.071 | 1114.668 | N/F | 0.182 | N/F | 4.7 | 2.365 | 3.843 | 0.874 | N/F | 2.14 | 187.807 | N/A | 1.157 | N/F | 3.686 | N/F | N/A | N/F | N/F | N/A | 0.457 | 0.253 | 1.27 | N/A | 0.383 | N/F | N/A | N/F | N/F | N/A | D,Q | Unknown
WO-22270-69 | N/F | 90.629 | 384.049 | 1850.524 | 0.192 | 171.812 | 5.382 | N/F | 3.852 | 0.802 | N/F | 2.992 | 187.693 | N/A | 1.174 | N/F | 4.846 | N/F | N/A | N/F | N/F | N/A | 0.518 | 0.655 | 0.793 | N/A | 0.412 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-70 | N/A | 73.601 | 962.263 | -384.289 | 0.175 | N/F | 4.381 | 1.632 | 1.922 | 0.735 | N/F | 1.318 | 187.983 | N/A | 0.962 | N/F | 4.055 | N/F | N/A | 1.749 | N/F | N/A | 1.354 | 0.155 | 3.853 | N/A | 0.366 | N/F | N/A | N/F | N/F | N/A | D,Q | Unknown
AB-220715-01 LRB 8_INJ001 | N/A | 9.116 | -264.025 | -300.723 | 0.17 | -0.296 | 4.321 | -0.188 | 0.733 | 0.324 | 2.078 | 0.885 | 187.552 | N/A | 0.907 | N/F | 0.33 | N/F | N/A | N/F | 0.985 | N/A | 0.692 | 0.152 | 1.451 | N/A | 0.316 | -1.484 | N/A | N/A | N/F | N/A | D,Q | Unknown
AB-220715-01 LRB 8 | N/F | 9.295 | -210.421 | 29.482 | 0.201 | 25.296 | 4.508 | 0.346 | 0.678 | 0.359 | 2.103 | 1.442 | 187.603 | N/F | 0.877 | N/F | 0.371 | N/F | N/A | N/F | 0.981 | N/A | 0.846 | 0.207 | 1.359 | N/A | 0.335 | -0.242 | N/A | N/F | N/F | N/A | D,Q | Unknown
WO-22270-71 | N/A | 35.578 | 140.711 | -272.815 | 0.184 | 738.168 | 5.023 | 2.742 | 1.464 | 0.85 | N/F | 0.896 | 187.673 | N/A | 0.712 | N/F | 8.605 | N/F | N/A | N/F | N/F | N/A | 5.465 | 0.188 | 2.626 | N/A | 0.322 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-72 | N/A | 44.232 | 190.339 | 176.56 | 0.177 | 636.91 | 4.338 | 1.019 | 1.316 | 0.652 | N/F | 0.981 | 187.778 | N/A | 0.836 | N/F | 3.996 | N/F | N/A | N/F | N/F | N/A | 0.393 | 0.147 | 1.363 | N/A | 0.305 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-73 | N/A | 48.91 | -64.362 | -459.024 | 0.17 | 269.895 | N/F | 3.521 | 1.954 | 0.816 | N/F | 0.921 | 187.732 | N/A | 1.086 | N/F | 3.362 | N/F | N/A | N/F | N/F | N/A | 0.46 | 0.18 | 0.576 | N/A | 0.312 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-74 | N/A | 52.99 | 856.849 | -72.531 | 0.173 | 419.554 | 4.292 | 2.192 | 1.892 | 0.727 | N/F | 2.049 | 187.763 | N/A | 0.88 | N/F | 6.787 | N/F | N/A | 1.274 | N/F | N/A | 2.505 | 0.23 | 2.129 | N/A | 0.332 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-75 | N/F | 39.104 | 201.744 | -251.403 | 0.179 | 210.064 | 4.407 | 2.943 | 1.166 | 0.885 | N/F | 1.729 | 187.918 | N/A | 0.966 | N/F | 3.346 | N/F | N/A | 2.022 | N/F | N/A | 1.793 | 0.281 | 4.871 | N/A | 0.336 | N/F | N/A | N/F | N/A | N/A |  | Unknown
WO-22270-76 | N/F | 103.752 | 497.725 | N/F | 0.179 | 307.607 | 4.292 | 3.26 | 2.297 | 1.211 | N/F | 3.292 | 188.311 | N/A | 1.122 | N/F | 4.155 | N/F | N/A | N/F | N/F | N/A | 0.522 | 0.298 | 1.654 | N/A | 0.319 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-77 | N/A | 35.759 | 1.786 | -106.899 | 0.175 | 248.929 | 4.574 | 0.62 | 1.209 | 1.108 | N/F | 2.161 | 187.938 | N/A | 0.582 | N/F | 19.796 | N/F | N/A | 1.767 | 1.755 | N/A | 0.445 | 0.251 | 0.829 | N/A | 0.318 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-78 | N/F | 35.054 | 507.47 | 509.707 | 0.169 | N/F | N/F | 2.357 | 0.789 | 0.783 | N/F | 1.883 | 188.028 | N/A | 0.815 | N/F | 3.982 | N/F | N/A | 1.746 | 1.926 | N/A | 0.575 | 0.158 | 0.94 | N/A | 0.304 | N/F | N/A | N/F | N/F | N/A | D,Q | Unknown
WO-22270-79 | N/A | 51.327 | 469.012 | -236.571 | 0.172 | 123.785 | N/F | 1.421 | 1.09 | 0.796 | N/F | 1.036 | 187.933 | N/A | 0.863 | N/F | 3.306 | N/F | N/A | 0.995 | 1.998 | N/A | 0.403 | 0.072 | 1.095 | N/A | 0.318 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-80 | N/F | 44.766 | 229.273 | 357.704 | 0.164 | N/F | 4.443 | 2.245 | 1.521 | 0.677 | N/F | 0.867 | 187.781 | N/A | 0.85 | N/F | 3.667 | N/F | N/A | 1.982 | 2.097 | N/A | 0.983 | 0.204 | 1.857 | N/A | 0.296 | N/F | N/A | N/F | N/F | N/A | D,Q | Unknown
AB-220715-01 LFB-H 5 | N/A | 188.365 | 202840.503 | 1234515.59 | 198.619 | 3055789.999 | 2799.149 | 211.05 | 174.405 | 168.726 | 155.005 | 164.39 | 417.344 | N/F | 204.941 | 190.537 | 199.853 | N/F | N/A | 197.117 | 205.523 | N/A | 206.05 | 176.772 | 201.716 | N/A | 203.13 | 232.48 | N/A | N/F | N/A | N/A |  | Unknown
AB-220715-01 LRB 9_INJ001 | N/F | 8.502 | 60.389 | 1267.874 | 0.399 | 641.222 | 7.26 | 0.111 | 0.709 | 0.406 | 2.124 | 1.658 | 187.684 | N/A | 1.081 | N/F | 0.497 | N/F | N/A | N/F | 1.066 | N/A | 0.986 | 0.186 | 1.724 | N/A | 0.698 | -1.639 | N/A | N/F | N/F | N/A | D,Q | Unknown
AB-220715-01 LRB 9 | N/F | 7.767 | 156.527 | 24.489 | 0.216 | 309.813 | 4.813 | 0.494 | 0.742 | 0.341 | 2.23 | 0.933 | 187.609 | N/A | 1.129 | N/F | 0.364 | N/F | N/A | 0.52 | 1.009 | N/A | 1.115 | 0.228 | 1.482 | N/A | 0.364 | -1.159 | N/A | N/F | N/F | N/A | D,Q | Unknown
AB-220715-01 LD 4 | N/A | 40.77 | 286.021 | -1.409 | 0.186 | 2852.671 | 6.915 | 3.027 | 2.589 | 0.733 | N/F | 2.064 | 187.912 | N/F | 0.839 | N/F | 3.69 | N/F | N/A | 2.485 | 2.379 | N/A | 1.152 | 0.082 | 3.311 | N/A | 0.333 | N/F | N/A | N/F | N/A | N/A |  | Unknown
AB-220715-01 CCC-H 5 | N/A | 356.422 | 388334.722 | 2343649.659 | 390.247 | 4984878.821 | 4802.419 | 389.461 | 338.354 | 350.896 | 315.397 | 332.343 | 604.509 | N/A | 395.002 | 356.628 | 386.242 | N/F | N/A | 376.721 | 384.448 | N/A | 395.37 | 338.568 | 376.568 | N/A | 403.856 | 386.002 | N/A | N/F | N/A | N/A |  | Unknown
AB-220715-01 LFSM-H 4 | N/F | 189.439 | 202997.037 | 1126695.618 | 201.674 | 3303844.392 | 3928.87 | 190.142 | 170.383 | 232.842 | 149.526 | 151.141 | 362.469 | N/A | 206.914 | N/F | 201.625 | N/F | N/A | 194.19 | 193.605 | N/A | 204.594 | 202.219 | 196.17 | N/A | 203.272 | 148.778 | N/A | N/F | N/A | N/A |  | Unknown
WO-22270-81 | N/F | 52.993 | 642.221 | 78.951 | 0.265 | 555.646 | 8.044 | 3.201 | 1.934 | 0.985 | N/F | 3.031 | 188.613 | N/A | 1.113 | N/F | 3.721 | N/F | N/A | 1.954 | 3.35 | N/A | 0.552 | 0.218 | 1.584 | N/A | 0.404 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-82 | N/F | 41.309 | 1672.951 | 1063.712 | 0.205 | N/F | 6.641 | 2.934 | 1.181 | 0.923 | N/F | 0.86 | 187.815 | N/A | 0.714 | N/F | 3.553 | N/F | N/A | 3.331 | 3.299 | N/A | 0.474 | 0.263 | 2.325 | N/A | 0.357 | N/F | N/A | N/F | N/F | N/A | D,Q | Unknown
WO-22270-83 | N/F | 48.205 | 1269.882 | 239.84 | 0.195 | N/F | 5.685 | 2.189 | 1.222 | 0.894 | N/F | 1.866 | 187.94 | N/A | 0.769 | N/F | 3.516 | N/F | N/A | N/F | 3.92 | N/A | 0.568 | 0.344 | 0.903 | N/A | 0.337 | N/F | N/A | N/F | N/F | N/A | D,Q | Unknown
WO-22270-84 | N/F | 37.294 | 112.183 | 481.908 | 0.191 | N/F | 4.693 | N/F | 2.623 | 0.787 | N/F | 0.926 | 187.662 | N/F | 0.915 | N/F | 2.515 | N/F | N/A | N/F | 10.43 | N/A | 0.444 | 0.271 | 1.798 | N/A | 0.378 | N/F | N/A | N/F | N/F | N/A | D,Q | Unknown
WO-22270-85 | N/F | 37.171 | 90.487 | 244.487 | 0.185 | 33.24 | 5.408 | 1.331 | 1.233 | 1.466 | N/F | 2.173 | 187.764 | N/A | 0.852 | N/F | 3.296 | N/F | N/A | 2.354 | 8.256 | N/A | 0.342 | 0.262 | 1.847 | N/A | 0.349 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-86 | N/A | 41.693 | 635.704 | 53.279 | 0.184 | N/F | 4.755 | 1.607 | 1.149 | 1 | N/F | 0.937 | 187.897 | N/A | 0.729 | N/F | 3.299 | N/F | N/A | 3.729 | 8.317 | N/A | 0.451 | 0.303 | 3.486 | N/A | 0.299 | N/F | N/A | N/F | N/F | N/A | D,Q | Unknown
WO-22270-87 | N/A | 39.548 | -111.161 | -463.869 | 0.177 | 239.196 | 5.035 | N/F | N/F | 0.846 | N/F | 1.704 | 187.954 | N/A | 0.647 | N/F | 3.777 | N/F | N/A | 2.319 | 29.138 | N/A | 0.522 | 0.254 | 0.651 | N/A | 0.351 | N/F | N/A | N/F | N/A | N/A |  | Unknown
WO-22270-88 | N/F | 55.023 | 920.01 | -135.586 | 0.174 | 83.849 | 4.619 | 2.168 | 1.205 | 0.764 | N/F | 2.057 | 188.26 | N/A | 0.884 | N/F | 4.311 | N/F | N/A | N/F | 29.571 | N/A | 0.458 | 0.22 | 0.822 | N/A | 0.362 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-89 | N/A | 68.22 | 1232.885 | 19.6 | 0.169 | 402.393 | 4.281 | 1.877 | 1.45 | 1.078 | N/F | 2.282 | 187.746 | N/A | 0.896 | N/F | 3.765 | N/F | N/A | 3.42 | 34.603 | N/A | 0.764 | 0.105 | 1.466 | N/A | 0.329 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-90 | N/A | 46.72 | 769.791 | 277.473 | 0.172 | N/F | 4.656 | 1.717 | 1.248 | 0.748 | N/F | 0.996 | 187.886 | N/A | 0.978 | N/F | 3.436 | N/F | N/A | 2.13 | 34.536 | N/A | N/F | 0.221 | N/F | N/A | 0.354 | N/F | N/A | N/F | N/F | N/A | D,Q | Unknown
AB-220715-01 LRB 10_INJ001 | N/F | 10.169 | 78.436 | 467.545 | 0.175 | N/F | 4.356 | 0.364 | 0.855 | 0.33 | N/F | 0.955 | 187.712 | N/A | 0.945 | N/F | 0.39 | N/F | N/F | N/F | N/F | N/A | 1.144 | 0.173 | 1.382 | N/A | 0.316 | -0.968 | N/A | N/F | N/F | N/A | D,Q | Unknown
AB-220715-01 LRB 10 | N/A | 17.776 | 203.248 | -481.869 | 0.177 | 91.877 | 4.334 | 0.581 | 0.656 | 0.382 | N/F | 1.082 | 187.804 | N/F | 1.182 | N/F | 0.442 | N/F | N/F | N/F | 1.06 | N/A | 1.02 | 0.278 | 1.482 | N/A | 0.305 | 1.344 | N/A | N/F | N/F | N/A | D,Q | Unknown
WO-22270-91 | N/F | 30.101 | -151.174 | 41.871 | 0.178 | N/F | 5.171 | N/F | 0.964 | 0.79 | N/F | 1.001 | 187.829 | N/A | 0.865 | N/F | 2.934 | N/F | N/A | 1.424 | 89.122 | N/A | 0.341 | 0.183 | 2.673 | N/A | 0.305 | N/F | N/A | N/F | N/F | N/A | D,Q | Unknown
WO-22270-92 | N/A | 58.989 | 413.207 | -252.14 | 0.175 | 226.358 | 4.403 | 1.92 | 1.274 | 0.775 | N/F | 1.027 | 187.98 | N/A | 0.716 | N/F | 3.82 | N/F | N/A | 3.867 | 118.451 | N/A | 0.488 | 0.184 | 1.678 | N/A | 0.314 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-93 | N/F | 35.932 | 363.93 | 300.949 | 0.174 | 2129.979 | 4.325 | 1.009 | 2.193 | 0.707 | N/F | 1.807 | 187.78 | N/A | 0.83 | N/F | 4.301 | N/F | N/A | 3.048 | 95.675 | N/A | 0.505 | 0.036 | 2.634 | N/A | 0.33 | N/F | N/A | N/F | N/A | N/A |  | Unknown
WO-22270-94 | N/F | 39.687 | 2083.552 | -172.736 | 0.171 | N/F | 4.406 | 1.973 | 1.385 | 0.72 | N/F | 0.892 | 187.733 | N/A | 0.644 | N/F | 3.73 | N/F | N/A | 3.31 | 513.964 | N/A | 0.553 | 0.014 | 4.213 | N/A | 0.324 | -1.63 | N/A | N/A | N/F | N/A | D,Q | Unknown
WO-22270-95 | N/A | 48.156 | 132.114 | -385.437 | 0.171 | 56.814 | 4.569 | 1.934 | 1.772 | 1.084 | N/F | 0.933 | 187.851 | N/F | 0.954 | N/F | 3.59 | N/F | N/A | 1.419 | 296.758 | N/A | 0.443 | 0.018 | 1.156 | N/A | 0.298 | N/F | N/A | N/F | N/F | N/A |  | Unknown
WO-22270-96 | N/F | 30.667 | 115.319 | 355.413 | 0.172 | 69.572 | 4.278 | N/F | 0.84 | 0.912 | N/F | 1.081 | 187.768 | N/A | 0.941 | N/F | 3.898 | N/F | N/A | 1.253 | 429.411 | N/A | 0.395 | 0.337 | 2.135 | N/A | 0.34 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-97 | N/F | 34.976 | 1095.1 | -298.073 | 0.172 | N/F | 4.85 | 1.411 | 0.982 | 0.805 | N/F | 0.924 | 187.836 | N/A | 0.823 | N/F | 3.657 | N/F | N/A | N/F | 2.168 | N/A | 0.561 | 0.151 | 0.986 | N/A | 0.306 | N/F | N/A | N/A | N/F | N/A | D,Q | Unknown
WO-22270-98 | N/F | 35.134 | 1017.672 | 656.634 | 0.172 | N/F | 4.272 | 3.012 | 0.913 | 0.602 | N/F | 0.86 | N/F | N/A | 0.931 | N/F | 4.501 | N/F | N/A | N/F | 1.918 | N/A | 0.427 | 0.235 | 0.611 | N/A | 0.313 | N/F | N/A | N/A | N/F | N/A | D,Q | Unknown
WO-22270-99 | N/F | 74.339 | 2248.231 | 1226.484 | 0.171 | N/F | 6.182 | 2.588 | 1.609 | 0.856 | N/F | 2.825 | 187.718 | N/A | 0.984 | N/F | 4.353 | N/F | N/A | N/F | N/F | N/A | 1.034 | 0.217 | N/F | N/A | 0.304 | N/F | N/F | N/A | N/F | N/A | D,Q | Unknown
WO-22270-100 | N/F | 64.394 | -56.538 | -403.909 | 0.17 | N/F | N/F | N/F | 2.049 | 0.622 | N/F | 0.868 | 188.207 | N/F | 0.921 | N/F | 4.313 | N/F | N/A | 1.066 | 1.921 | N/A | 0.447 | 0.224 | 1.687 | N/A | 0.3 | N/F | N/A | N/F | N/F | N/A | D,Q | Unknown
AB-220715-01 LFB-H 6 | N/A | 181.741 | 196443.003 | 1149571.427 | 191.557 | 3292088.252 | 2844.398 | 182.753 | 166.991 | 181.226 | 148.72 | 161.946 | 384.475 | N/A | 187.724 | 156.313 | 188.731 | N/F | N/A | 188.302 | 191.319 | N/A | 188.34 | 167.236 | 191.836 | N/A | 201.151 | 226.894 | N/A | N/A | N/F | N/A |  | Unknown
AB-220715-01 LRB 11_INJ001 | N/F | 8.472 | 317.579 | 207.284 | 0.355 | 461.488 | 6.762 | 0.603 | 0.684 | 0.41 | N/F | 2.04 | 187.903 | N/A | 1.102 | N/F | 0.503 | N/F | N/A | N/F | N/F | N/A | 0.798 | 0.141 | 1.569 | N/A | 0.62 | -2.486 | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LRB 11 | N/A | 10.077 | 165.446 | 342.618 | 0.208 | N/F | 4.591 | 0.293 | 0.596 | 0.385 | N/F | 1.677 | 187.683 | N/A | 1.098 | N/F | 0.383 | N/F | N/A | N/F | 0.998 | N/A | 0.811 | 0.045 | 1.468 | N/A | 0.409 | -2.011 | N/A | N/A | N/F | N/A | D,Q | Unknown
AB-220715-01 LD 5 | N/F | 97.484 | 1542.264 | -323.222 | 0.193 | 156.442 | 5.354 | 1.745 | 2.34 | 0.802 | N/F | 3.174 | 187.951 | N/A | 0.988 | N/F | 4.628 | N/F | N/A | N/F | 2.824 | N/A | 1.161 | 0.37 | 3.298 | N/A | 0.336 | N/F | N/F | N/A | N/A | N/A |  | Unknown
AB-220715-01 CCC-H 6 | N/F | 350.486 | 385600.225 | 2262197.668 | 367.431 | 5424472.255 | 4953.977 | 375.438 | 350.691 | 359.107 | 302.006 | 313.329 | 595.579 | N/A | 403.343 | 377.628 | 378.911 | N/A | N/A | 378.394 | 368.549 | N/A | 377.013 | 332.589 | 372.428 | N/A | 384.168 | 421.07 | N/A | N/A | N/F | N/A |  | Unknown
AB-220715-01 LFSM-H 5 | N/F | 201.428 | 200130.033 | 1142518.349 | 208.381 | 3197608.075 | 4363.9 | 184.913 | 171.217 | 204.164 | 144.456 | 149.42 | 365.498 | N/A | 205.917 | N/F | 198.493 | N/F | N/A | 192.215 | 184.057 | N/A | 197.638 | 196.268 | 196.1 | N/A | 202.464 | 153.964 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-101 | N/F | 33.731 | 1059.816 | -503.82 | 0.239 | 1008.589 | 8.518 | 3.681 | 0.982 | 0.53 | N/F | 1.145 | 187.717 | N/A | 0.972 | N/F | 4.04 | N/F | N/A | 0.779 | 2.944 | N/A | 0.434 | 0.381 | 0.731 | N/A | 0.407 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-102 | N/A | 24.256 | 1326.149 | -413.123 | 0.205 | 680.382 | 4.615 | 1.641 | 1.622 | 0.456 | N/F | 2.04 | 187.754 | N/A | 1.146 | N/F | 4.044 | N/F | N/A | N/F | 2.814 | N/A | 0.503 | 0.372 | 1.723 | N/A | 0.369 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-103 | N/F | 60.937 | -39.804 | -518.56 | 0.192 | 124.878 | 5.063 | 1.6 | 1.09 | 0.578 | N/F | 1.683 | 187.861 | N/A | 0.778 | N/F | 4.974 | N/F | N/A | N/F | 3.613 | N/A | 0.843 | 0.211 | 1.333 | N/A | 0.315 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-104 | N/F | 42.933 | 1198.628 | -448.666 | 0.207 | 317.913 | 5.368 | 1.636 | 1.437 | 0.719 | N/F | 1.001 | 187.697 | N/F | 0.939 | N/F | 3.97 | N/F | N/A | N/F | 8.259 | N/A | 0.365 | 0.216 | 0.625 | N/A | 0.375 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-105 | N/F | 93.769 | 19.379 | -319.987 | 0.184 | 110.89 | 4.965 | 1.889 | 1.197 | 0.688 | N/F | 1.069 | 187.671 | N/A | 0.593 | N/F | 4.763 | N/F | N/A | N/F | 7.025 | N/A | 0.782 | 0.222 | 0.608 | N/A | 0.361 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-106 | N/F | 67.49 | 3127.405 | 744.399 | 0.212 | 371.567 | 5.199 | N/F | 0.963 | 0.71 | N/F | 2.235 | 187.958 | N/A | 0.65 | N/F | 4.019 | N/F | N/A | 1.448 | 7.749 | N/A | 0.619 | 0.388 | 0.523 | N/A | 0.365 | 4.342 | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-107 | N/A | 33.765 | 169.512 | -418.587 | 0.174 | 25.086 | 4.403 | 2.217 | 1.928 | 0.596 | N/F | 1.906 | 187.928 | N/F | 0.678 | N/F | 3.585 | N/F | N/A | 1.513 | 26.028 | N/A | 0.499 | 0.253 | 0.535 | N/A | 0.352 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-108 | N/A | 52.107 | 626.182 | 591.075 | 0.17 | 59.577 | 4.5 | 3.237 | 1.173 | 0.584 | N/F | 1.045 | N/F | N/A | 0.59 | N/F | 3.661 | N/F | N/A | 1.854 | 24.938 | N/A | 0.489 | 0.095 | 0.557 | N/A | 0.357 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-109 | N/A | 42.539 | 349.224 | -344.628 | 0.175 | 307.754 | 4.903 | 1.92 | 1.642 | 0.578 | N/F | 0.961 | 187.735 | N/A | 0.69 | N/F | 4.195 | N/F | N/A | N/F | 36.818 | N/A | 1.347 | 0.271 | 0.617 | N/A | 0.322 | N/F | N/A | N/A | N/F | N/A |  | Unknown
WO-22270-110 | N/F | 46.505 | 646.813 | -492.082 | 0.169 | 584.479 | 4.524 | 2.75 | 0.734 | 0.697 | N/F | 0.98 | 187.836 | N/A | 0.973 | N/F | 3.898 | N/F | N/A | N/F | 35.931 | N/A | 1.286 | 0.176 | 1.446 | N/A | 0.301 | N/F | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 12_INJ001 | N/A | 16.116 | 1.775 | 453.877 | 0.175 | 116.664 | 4.28 | 0.578 | 0.705 | 0.443 | N/F | 1.732 | 187.671 | N/A | 1.094 | N/F | 0.332 | N/F | N/A | N/F | N/F | N/A | 0.335 | 0.258 | 0.51 | N/A | 0.346 | N/F | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 12 | N/F | 16.842 | -140.753 | 171.064 | 0.18 | 1060.489 | 4.318 | 0.319 | 0.946 | 0.42 | N/F | 1.105 | 187.657 | N/A | 1.079 | N/F | 0.346 | N/F | N/A | N/F | N/F | N/A | 0.321 | 0.211 | 0.517 | N/A | 0.315 | -1.605 | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-111 | N/A | 42.733 | 203.767 | -459.801 | 0.185 | N/F | 4.502 | 1.185 | 1.267 | 0.755 | N/F | 0.886 | 187.818 | N/A | 0.911 | N/F | 3.531 | N/F | N/A | N/F | 72.024 | N/A | 0.505 | 0.188 | 0.756 | N/A | 0.298 | N/F | N/A | N/A | N/F | N/A | D,Q | Unknown
WO-22270-112 | N/A | 58.537 | 1266.19 | -248.505 | 0.168 | N/F | 4.417 | 2.563 | 1.418 | 0.598 | N/F | 1.181 | 187.96 | N/A | 0.682 | N/F | 3.652 | N/F | N/A | N/F | 72.186 | N/A | 0.855 | 0.243 | 0.638 | N/A | 0.315 | N/F | N/A | N/A | N/F | N/A | D,Q | Unknown
WO-22270-113 | N/A | 36.305 | 730.178 | N/F | 0.201 | N/F | 4.282 | 1.551 | 1.485 | 0.621 | N/F | 2.63 | 187.671 | N/F | 0.875 | N/F | 3.505 | N/F | N/A | 0.633 | 94.178 | N/A | 0.372 | 0.313 | 0.999 | N/A | 0.305 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-114 | N/A | 57.432 | 808.673 | -229.463 | 0.186 | N/F | 4.303 | N/F | 1.335 | 0.655 | N/F | 2.402 | 188.007 | N/A | 0.841 | N/F | 4.205 | N/F | N/A | 1.115 | 326.568 | N/A | 0.57 | 0.257 | 1.482 | N/A | 0.313 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-115 | N/F | 43.75 | 1783.148 | -110.945 | 0.175 | N/F | 4.424 | 3.524 | 1.218 | 0.736 | N/F | 1.171 | 188.204 | N/A | 0.7 | N/F | 4.431 | N/F | N/A | N/F | 222.047 | N/A | 2.753 | 0.144 | 2.163 | N/A | 0.315 | N/F | N/A | N/A | N/F | N/A | D,Q | Unknown
WO-22270-116 | N/A | 67.219 | 3334.089 | 68.538 | 0.185 | 335.165 | 4.367 | N/F | 2.36 | 0.911 | N/F | 0.987 | 187.883 | N/F | 0.592 | N/F | 3.651 | N/F | N/A | N/F | 239.793 | N/A | 0.482 | 0.256 | 1.595 | N/A | 0.38 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-117 | N/F | 31.489 | 893.006 | -483.667 | 0.178 | N/F | N/F | 2.873 | 1.803 | 0.843 | N/F | 0.964 | 187.986 | N/A | 0.989 | N/F | 27.663 | N/F | N/A | 0.705 | 2.136 | N/A | 0.507 | 0.194 | 1.126 | N/A | 0.371 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-118 | N/A | 47.377 | 2470.661 | -403.327 | 0.172 | 455.012 | 4.465 | N/F | 1.197 | 0.874 | N/F | 0.913 | 187.763 | N/F | 0.597 | N/F | 2.864 | N/F | N/A | N/F | 2.136 | N/A | 0.423 | 0.298 | 0.666 | N/A | 0.322 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-119 | N/A | 71.148 | 581.211 | 542.148 | 0.213 | 139.33 | 4.473 | N/F | 1.762 | 1.048 | N/F | 2.894 | 187.708 | N/F | 0.845 | N/F | 4.091 | N/F | N/A | N/F | 1.872 | N/A | 0.726 | 0.204 | 1.779 | N/A | 0.362 | 11.111 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-120 | N/F | 83.8 | 321.139 | -327.12 | 0.167 | 455.36 | N/F | N/F | 1.069 | 0.711 | N/F | 1.033 | 187.728 | N/A | 0.884 | N/F | 4.189 | N/F | N/A | N/F | 1.811 | N/A | 0.958 | 0.322 | 1.868 | N/A | 0.311 | N/F | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LFB-H 7 | N/A | 204.781 | 198140.72 | 1057135.503 | 165.166 | 2447239.965 | 2508.736 | 189.674 | 171.667 | 189.285 | 156.023 | 158.717 | 383.497 | N/A | 206.892 | 172.757 | 190.922 | N/F | N/A | 191.718 | 192.398 | N/A | 195.667 | 152.165 | 192.64 | N/A | 200.014 | 191.329 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 13_INJ001 | N/F | 12.2 | -234.346 | 273.809 | 0.345 | 1370.912 | 6.716 | 0.574 | 0.708 | 0.572 | N/F | 1.31 | 187.698 | N/A | 1.272 | N/F | 0.522 | N/F | N/A | N/F | 1.13 | N/A | 0.447 | 0.413 | 0.579 | N/A | 0.665 | N/F | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 13 | N/F | 17.745 | 132.549 | N/F | 0.218 | 1747.628 | 5.293 | 0.337 | 0.858 | 0.329 | N/F | 1.057 | 187.644 | N/A | 1.183 | N/F | 0.374 | N/F | N/A | N/F | 0.956 | N/A | 0.354 | 0.267 | 0.599 | N/A | 0.465 | -2.027 | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LD 6 | N/F | 54.984 | 2628.138 | 1682.051 | 0.218 | 278.044 | 5.22 | N/F | 1.889 | 0.628 | N/F | 0.93 | 187.859 | N/A | 0.785 | N/F | 3.868 | N/F | N/A | 2.465 | 1.984 | N/A | 0.814 | 0.468 | 1.785 | N/A | 0.364 | N/F | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 CCC-H 7 | N/F | 359.39 | 389323.119 | 2156754.132 | 349.245 | 4615790.133 | 4460.001 | 385.941 | 335.708 | 364.811 | 303.514 | 317.825 | 588.131 | N/A | 405.387 | 369.407 | 372.392 | N/F | N/A | 377.006 | 375.749 | N/A | 384.409 | 298.732 | 387.221 | N/A | 383.41 | 384.022 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LFSM-H 6 | N/F | 186.26 | 204489.742 | 1159280.825 | 194.374 | 3757884.147 | 3600.105 | 202.898 | 178.901 | 221.409 | 145.053 | 146.784 | 370.14 | N/A | 202.432 | N/F | 201.023 | N/F | N/A | 197.115 | 189.642 | N/A | 195.386 | 186.132 | 200.076 | N/A | 200.886 | 191.886 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-121 | N/F | 46.695 | 388.276 | -219.126 | 0.228 | 307.211 | 6.823 | 1.054 | 1.007 | 0.852 | N/F | 1.38 | 187.79 | N/A | 0.887 | N/F | 3.691 | N/F | N/A | N/F | 2.834 | N/A | 0.406 | 0.44 | N/F | N/A | 0.426 | -1.099 | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-122 | N/F | 88.084 | 2073.15 | -459.482 | 0.208 | 174.02 | 9.192 | N/F | 2.476 | 0.818 | N/F | 0.971 | 187.813 | N/A | 1.098 | N/F | 3.841 | N/F | N/A | N/F | 4.153 | N/A | 0.714 | 0.706 | 1.193 | N/A | 0.364 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-123 | N/F | 46.07 | 3104.927 | 31.206 | 0.192 | N/F | 5.338 | N/F | 0.914 | 1.248 | N/F | 2.967 | 187.818 | N/A | 0.632 | N/F | 4.068 | N/F | N/A | N/F | 3.574 | N/A | 0.454 | 0.383 | 0.544 | N/A | 0.411 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-124 | N/F | 58.248 | 2032.92 | -216.765 | 0.184 | N/F | 5.106 | N/F | 1.17 | 0.612 | N/F | 1.115 | 187.719 | N/F | 0.612 | N/F | 4.517 | N/F | N/A | N/F | 8.903 | N/A | 0.583 | 0.417 | 3.316 | N/A | 0.33 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-125 | N/A | 61.89 | 4734.78 | -74.305 | 0.233 | 88.921 | 4.529 | N/F | 2.165 | 0.891 | N/F | 0.877 | 188.119 | N/A | 0.564 | N/F | 4.939 | N/F | N/A | N/F | 8.16 | N/A | 0.925 | 0.512 | 1.426 | N/A | 0.396 | 0.078 | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-126 | N/F | 44.648 | 3852.046 | -400.763 | 0.191 | N/F | 4.402 | N/F | 2.014 | 1.035 | N/F | 0.952 | 188.795 | N/A | 0.597 | N/F | 3.937 | N/F | N/A | N/F | 8.852 | N/A | 0.878 | 0.378 | 2.024 | N/A | 0.35 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-127 | N/F | 30.429 | 1416.199 | 1192.418 | 0.176 | N/F | 5.269 | N/F | 1.381 | 0.63 | N/F | 0.903 | 187.978 | N/A | 1.067 | N/F | 3.877 | N/F | N/A | N/F | 28.051 | N/A | 0.484 | 0.49 | 2.717 | N/A | 0.322 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-128 | N/F | 41.731 | 1489.561 | -254.157 | 0.182 | N/F | 4.334 | N/F | 2.627 | 0.862 | N/F | 1.558 | 188.016 | N/A | 0.83 | N/F | 3.491 | N/F | N/A | N/F | 33.692 | N/A | 0.52 | 0.306 | 1.561 | N/A | 0.336 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-129 | N/F | 49.417 | 2775.292 | -400.293 | 0.172 | N/F | 4.462 | 1.124 | 0.913 | 0.466 | N/F | 0.979 | 188.075 | N/A | 0.817 | N/F | 4.154 | N/F | N/A | 3.087 | 38.125 | N/A | 0.837 | 0.025 | 2.706 | N/A | 0.32 | N/F | N/A | N/A | N/F | N/A | D,Q | Unknown
WO-22270-130 | N/F | 141.321 | 3730.449 | 201.143 | N/F | N/F | N/F | N/F | 1.542 | 0.742 | N/F | 0.829 | 187.988 | N/A | 0.677 | N/F | 5.842 | N/F | N/A | 2.975 | 39.499 | N/A | 1.533 | 0.3 | 21.233 | N/A | 0.333 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LRB 14_INJ001 | N/F | 14.765 | 1040.067 | -183.837 | 0.168 | 42.386 | 4.34 | 0.484 | 0.835 | 0.445 | N/F | 1.409 | 187.638 | N/A | 1.058 | N/F | 0.356 | N/F | N/A | 0.486 | 1.109 | N/A | 0.321 | 0.184 | 0.587 | N/A | 0.324 | N/F | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 14 | N/A | 19.671 | 1331.876 | -156.099 | 0.197 | N/F | 4.342 | 0.65 | 0.797 | 0.511 | N/F | 1.439 | 187.706 | N/A | 1.258 | N/F | 0.373 | N/F | N/A | N/F | N/F | N/A | 0.339 | 0.231 | 0.56 | N/A | 0.337 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-131 | N/A | 319.097 | 1734.18 | 204.249 | 0.172 | N/F | 4.494 | N/F | 1.28 | 0.781 | N/F | 0.743 | 188.045 | N/A | 0.864 | N/F | 3.517 | N/F | N/A | 1.2 | 142.268 | N/A | 0.796 | 0.298 | 1.455 | N/A | 0.336 | -1.639 | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-132 | N/F | 60.479 | 2786.299 | 534.822 | 0.17 | N/F | N/F | N/F | 1.683 | 0.556 | N/F | N/F | 187.859 | N/A | 0.573 | N/F | 3.637 | N/F | N/A | N/F | 112.938 | N/A | 0.725 | 0.036 | 3.691 | N/A | 0.294 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-133 | N/F | 58.792 | 2230.904 | 571.219 | 0.169 | N/F | 4.358 | N/F | 1.435 | 1.299 | N/F | 3.056 | 188.927 | N/A | 0.773 | N/F | 4.59 | N/F | N/A | 1.669 | 83.818 | N/A | 2.182 | 0.032 | 2.816 | N/A | 0.311 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-134 | N/A | 137.963 | 3885.937 | 1275.706 | 0.166 | 433.112 | 5.014 | N/F | 2.327 | 1.098 | N/F | 1.377 | 188.853 | N/F | 0.54 | N/F | 4.942 | N/F | N/A | 5.25 | 2231.973 | N/A | 3.39 | 0.259 | 2.168 | N/A | 0.36 | 17.895 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-135 | N/F | 93.675 | 1434.034 | 81.724 | 0.176 | N/F | 5.045 | N/F | 1.205 | 0.905 | N/F | 2.956 | 188.61 | N/A | 0.614 | N/F | 3.738 | N/F | N/A | 1.671 | 222.327 | N/A | 0.704 | 0.397 | 0.598 | N/A | 0.314 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-136 | N/A | 1380.136 | 7192.438 | 232.137 | 0.189 | N/F | 4.61 | 3.41 | 1.742 | 0.611 | N/F | 0.913 | 188.042 | N/A | 0.468 | N/F | 4.151 | N/A | N/A | 3.756 | 51.868 | N/A | 0.774 | 0.054 | 1.873 | N/A | 0.301 | 27.799 | N/F | N/A | N/A | N/A |  | Unknown
WO-22270-137 | N/A | 244.19 | 551.055 | -215.286 | 0.186 | N/F | N/F | 11.283 | 1.387 | 0.948 | N/F | 0.983 | 187.795 | N/A | 0.449 | N/F | 2.975 | N/A | N/A | N/F | 1.438 | N/A | 0.447 | 0.015 | 1.092 | N/A | 0.328 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-138 | N/A | 722.708 | 1782.187 | 92.015 | 0.168 | N/F | N/F | 16.461 | 1.31 | 0.972 | N/F | 0.896 | 188.95 | N/A | 0.185 | N/F | 2.294 | N/F | N/A | 2.151 | N/F | N/A | 0.365 | 0.139 | 0.526 | N/A | 0.328 | 23.02 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-139 | N/A | 1405.259 | 1938.607 | -290.602 | 0.168 | N/F | 4.337 | 15.189 | 1.429 | 1.625 | N/F | 0.876 | 188.102 | N/A | 0.808 | N/F | 4.91 | N/A | N/A | N/F | N/F | N/A | 0.538 | 0.008 | 0.633 | N/A | 0.296 | 21.595 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-140 | N/A | 849.187 | 3952.986 | 20.313 | 0.168 | 505.27 | N/F | 10.723 | 1.545 | 0.744 | N/F | 0.971 | 188.247 | N/A | 0.827 | N/F | 3.954 | N/A | N/A | N/F | N/F | N/A | 0.56 | 0.063 | 0.612 | N/A | 0.276 | N/F | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LFB-H 8 | N/A | 180.431 | 192046.504 | 1076773.877 | 183.957 | 2047300.17 | 2290.255 | 182.476 | 169.13 | 199.841 | 153.04 | 153.851 | 386.009 | N/A | 200.563 | 184.848 | 183.032 | N/A | N/A | 189.842 | 190.233 | N/A | 207.329 | 159.763 | 200.802 | N/A | 200.042 | 167.174 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 15_INJ001 | N/A | 88.971 | 381.149 | 5342.364 | 0.959 | 8907.323 | 12.588 | 0.909 | 1.064 | 0.907 | 1.4 | 1.582 | 188.028 | N/A | 1.146 | N/F | 1.466 | N/A | N/A | 1.447 | 1.636 | N/A | 1.026 | 0.767 | 1.407 | N/A | 1.591 | N/F | N/A | N/A | N/A | N/A | C | Unknown
AB-220715-01 LRB 15 | N/A | 99.359 | 681.894 | 467.245 | 0.25 | N/F | 4.737 | 0.951 | 0.813 | 0.365 | N/F | 1.271 | 187.941 | N/A | 0.92 | N/F | 0.494 | N/A | N/A | N/F | 1.008 | N/A | 0.374 | 0.07 | 0.881 | N/A | 0.592 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LD 7 | N/A | 1153.376 | 616.662 | 18.357 | 0.204 | 146.822 | 5.082 | 7.487 | 1.639 | 0.603 | N/F | 1.203 | 188.821 | N/A | 0.704 | N/F | 4.949 | N/A | N/A | N/F | N/F | N/A | 0.653 | 0.169 | 1.55 | N/A | 0.407 | 46.344 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 CCC-H 8 | N/A | 357.065 | 400397.555 | 2350663.417 | 389.103 | 5106299.443 | 3966.62 | 385.036 | 336.433 | 377.342 | 316.872 | 325.053 | 574.823 | N/A | 401.32 | 380.715 | 375.361 | N/A | N/A | 390.578 | 378.446 | N/A | 379.099 | 342.018 | 374.878 | N/A | 395.998 | 415.555 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LFSM-H 7 | N/A | 872.937 | 207451.099 | 1065725.275 | 191.747 | 2848347.37 | 3079.549 | 220.478 | 176.03 | 236.177 | 153.039 | 150.528 | 364.538 | N/A | 201.623 | N/F | 201.943 | N/A | N/A | 195.526 | 188.25 | N/A | 196.726 | 176.145 | 196.063 | N/A | 204.965 | 156.285 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-141 | N/A | 432.865 | 1434.497 | 3331.26 | 0.55 | 7084.771 | 11.738 | 34.124 | 1.527 | 1.058 | N/F | 1.175 | 188.02 | N/A | 0.992 | N/F | 3.221 | N/A | N/A | 1.068 | 1.641 | N/A | 0.808 | 0.847 | 1.229 | N/A | 0.838 | 68.821 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-142 | N/A | 449.446 | 1714.475 | 239.509 | 0.321 | N/F | 5.797 | 33.507 | 1.463 | 0.945 | N/F | 1.051 | 187.917 | N/A | 0.865 | N/F | 2.606 | N/A | N/A | N/F | N/F | N/A | 0.486 | 0.41 | 1.132 | N/A | 0.492 | -1.685 | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-143 | N/A | 1319.375 | 1620.926 | -97.186 | 0.357 | N/F | 6.583 | 26.498 | 2.015 | 1.735 | N/F | 0.984 | 188.245 | N/A | 0.786 | N/F | 3.584 | N/A | N/A | N/F | N/F | N/A | 0.608 | 0.733 | 0.601 | N/A | 0.437 | 4.384 | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-144 | N/A | 752.207 | 3690.267 | -196.4 | 0.194 | 163.597 | 5.183 | 102.136 | 1.164 | 0.743 | N/F | 1.097 | 189.092 | N/A | 0.747 | N/F | 3.23 | N/A | N/A | 4.989 | N/F | N/A | 0.83 | 0.504 | 1.771 | N/A | 0.377 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-145 | N/A | 909.389 | 1765.158 | -255.025 | 0.226 | 81.577 | 4.453 | 88.027 | 0.689 | 0.617 | N/F | 0.844 | 188.287 | N/A | 0.809 | N/F | 2.649 | N/A | N/A | 2.422 | N/F | N/A | 0.525 | 0.393 | 1.335 | N/A | 0.352 | 54.026 | N/F | N/A | N/A | N/A |  | Unknown
WO-22270-146 | N/A | 537.5 | 339.421 | -457.809 | 0.194 | N/F | 4.289 | 92.579 | 0.882 | 0.943 | N/F | 0.965 | 189.194 | N/A | 0.883 | N/F | 2.223 | N/A | N/A | N/F | N/F | N/A | 0.591 | 0.494 | 1.95 | N/A | 0.339 | 80.268 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-147 | N/A | 1107.263 | 1385.91 | -28.236 | 0.187 | N/F | N/F | 324.339 | 1.916 | 0.863 | N/F | 1.445 | 189.12 | N/A | 0.57 | N/F | 3.744 | N/A | N/A | N/F | N/F | N/A | 0.489 | 0.757 | 2.685 | N/A | 0.41 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-148 | N/A | 977.58 | 1846.577 | 737.936 | 0.176 | N/F | 4.377 | 304.539 | 1.139 | 0.872 | N/F | 1.57 | 188.053 | N/A | 0.631 | N/F | 2.35 | N/A | N/A | 3.04 | 2.76 | N/A | 0.533 | 0.307 | 4.589 | N/A | 0.352 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-149 | N/A | 624.84 | 1688.909 | -125.136 | 0.179 | N/F | 4.476 | 312.784 | 0.908 | 0.747 | N/F | 0.928 | 188.119 | N/F | 0.96 | N/F | 4.147 | N/A | N/A | N/F | 2.979 | N/A | 0.671 | 0.253 | 1.944 | N/A | 0.313 | 46.838 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-150 | N/A | 644.971 | 359.833 | -17.626 | 0.176 | N/F | 4.301 | 267.728 | 1.28 | 0.919 | N/F | 0.854 | 188.123 | N/A | 0.974 | N/F | 3.079 | N/A | N/A | N/F | N/F | N/A | 0.73 | 0.253 | 0.552 | N/A | 0.382 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LRB 16_INJ001 | N/A | 159.939 | 196.658 | 1089.619 | 0.192 | 1.891 | 4.362 | 2.056 | 0.682 | 0.373 | N/F | 1.009 | 187.928 | N/A | 0.884 | N/F | 0.43 | N/A | N/A | N/F | N/F | N/A | 0.56 | 0.085 | 0.858 | N/A | 0.36 | -0.437 | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LRB 16 | N/A | 133.202 | 535.015 | 1735.266 | 0.175 | N/F | 4.336 | 1.823 | 0.881 | 0.292 | N/F | 0.899 | 187.647 | N/A | 0.901 | N/F | 0.429 | N/A | N/A | N/F | N/F | N/A | 0.378 | 0.101 | 0.865 | N/A | 0.314 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-151 | N/A | 508.489 | 927.013 | -181.719 | 0.195 | N/F | N/F | 993.029 | 0.918 | 0.587 | N/F | 0.829 | 187.726 | N/A | 0.887 | N/F | 2.406 | N/A | N/A | N/F | 1.571 | N/A | 1.754 | 0.22 | 0.623 | N/A | 0.323 | N/F | N/F | N/A | N/A | N/A | D,Q | Unknown
WO-22270-152 | N/A | 497.604 | 2413.122 | 813.222 | 0.176 | N/F | 4.416 | 1050.632 | 1.351 | 0.585 | N/F | 2.257 | 187.693 | N/A | 0.879 | N/F | 2.648 | N/A | N/A | N/F | 1.672 | N/A | 0.517 | 0.2 | 0.75 | N/A | 0.352 | N/F | N/F | N/A | N/A | N/A | D,Q | Unknown
WO-22270-153 | N/A | 859.471 | 746.388 | -451.643 | 0.19 | N/F | N/F | 958.71 | 1.915 | 0.69 | N/F | 1.34 | 187.772 | N/A | 0.4 | N/F | 3.295 | N/A | N/A | N/F | 2.461 | N/A | 0.645 | 0.221 | 1.145 | N/A | 0.337 | N/F | N/F | N/A | N/A | N/A | D,Q | Unknown
WO-22270-154 | N/A | 401.255 | 27.113 | 225.824 | 0.179 | N/F | 4.634 | 2069.312 | 0.893 | 0.451 | N/F | 1.161 | 187.703 | N/A | 0.885 | N/F | 4.213 | N/A | N/A | N/F | 2.409 | N/A | 5.502 | 0.058 | 0.773 | N/A | 0.318 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-155 | N/A | 333.591 | 604.733 | -147.316 | 0.17 | N/F | 4.405 | 3164.557 | 0.924 | 0.464 | N/F | 0.936 | 187.619 | N/A | 0.648 | N/F | 3.675 | N/A | N/A | N/F | 2.974 | N/A | 0.599 | 0.068 | 0.664 | N/A | 0.314 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-156 | N/A | 793.249 | 1015.968 | -511.051 | 0.225 | N/F | N/F | 2732.931 | 1.656 | 0.574 | N/F | 1.285 | 188.53 | N/A | 0.544 | N/F | 3.488 | N/A | N/A | N/F | 2.804 | N/A | 0.575 | 0.223 | 1.519 | N/A | 0.332 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-157 | N/A | 348.709 | 102.602 | -347.872 | 0.166 | N/F | N/F | 17.397 | 1.469 | 0.6 | N/F | 1.133 | 187.771 | N/A | 0.886 | N/F | 3.856 | N/A | N/A | 2.832 | 3.054 | N/A | 4.66 | 0.122 | 1.923 | N/A | 0.342 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-158 | N/A | 797.264 | 1109.268 | 252.499 | 0.178 | N/F | N/F | 7.677 | 1.107 | 0.994 | N/F | 2.294 | 188.131 | N/A | 0.61 | N/F | 2.164 | N/A | N/A | N/F | 1.143 | N/A | 1.068 | 0.065 | 1.129 | N/A | 0.301 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-159 | N/A | 674.126 | 829.057 | 356.948 | 0.177 | 122.794 | 4.395 | 8.152 | 1.589 | 0.622 | N/F | 1.054 | 187.663 | N/A | 0.608 | N/F | 3.551 | N/A | N/A | N/F | N/F | N/A | 0.726 | 0.067 | 1.534 | N/A | 0.307 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-160 | N/A | 699.334 | 3558.733 | -352.482 | 0.189 | N/F | N/F | 3.988 | 1.25 | 0.894 | N/F | 2.993 | 187.698 | N/A | 0.525 | N/F | 3.234 | N/A | N/A | N/F | N/F | N/A | 0.636 | 0.235 | 0.554 | N/A | 0.302 | N/F | N/F | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LFB-H 9 | N/A | 179.273 | 184173.045 | 1193753.672 | 181.484 | 3562575.448 | 2000.378 | 180.037 | 160.929 | 200.056 | 144.71 | 147.606 | 385.193 | N/A | 199.5 | 225.666 | 201.143 | N/A | N/A | 199.053 | 195.058 | N/A | 209.15 | 151.798 | 199.266 | N/A | 197.382 | 302.704 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 17_INJ001 | N/A | 189.576 | 860.666 | 7786.963 | 1.198 | 7931.958 | 15.366 | 0.812 | 1.121 | 1.196 | 1.497 | 1.555 | 188.161 | N/A | 1.533 | N/F | 1.538 | N/A | N/A | 1.256 | 1.817 | N/A | 1.235 | 1.184 | 1.693 | N/A | 1.839 | 18.523 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 17 | N/A | 192.221 | 324.303 | 1625.266 | 0.227 | 206.856 | 4.743 | 3.442 | 0.641 | 0.382 | N/F | 1.002 | 187.939 | N/A | 1.023 | N/F | 0.466 | N/A | N/A | N/F | 1.134 | N/A | 0.388 | 0.102 | 0.893 | N/A | 0.432 | N/F | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LD 8 | N/A | 603.311 | 1097.445 | 505.6 | 0.215 | 172.239 | 5.395 | 8.326 | 1.145 | 0.718 | N/F | 1.327 | 188.188 | N/A | 0.834 | N/F | 3.377 | N/A | N/A | N/F | 2.782 | N/A | 0.66 | 0.198 | 3.612 | N/A | 0.426 | 25.975 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 CCC-H 9 | N/A | 370.536 | 401819.365 | 2284987.744 | 376.49 | 4169724.864 | 4286.54 | 369.546 | 341.763 | 373.974 | 304.674 | 311.608 | 600.951 | N/A | 397.55 | 361.131 | 363.614 | N/A | N/A | 392.862 | 391.441 | N/A | 391.626 | 324.939 | 386.79 | N/A | 400.534 | 373.29 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LFSM-H 8 | N/A | 545.767 | 193390.292 | 1078186.147 | 199.404 | 2273336.683 | 3525.523 | 214.18 | 178.667 | 216.133 | 144.849 | 151.908 | 365.028 | N/A | 198.021 | N/F | 187.394 | N/A | N/A | 193.345 | 184.277 | N/A | 199.472 | 168.893 | 202.262 | N/A | 205.813 | 126.079 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-161 | N/A | 374.117 | 1237.971 | 629.047 | 0.386 | 482.122 | 9.39 | 12.151 | 1.741 | 0.943 | N/F | 1.224 | 188.191 | N/A | 0.991 | N/F | 2.606 | N/A | N/A | N/F | 1.612 | N/A | 1.19 | 0.727 | 0.627 | N/A | 0.586 | 9.032 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-162 | N/A | 530.721 | 683.792 | 241.509 | 0.276 | N/F | 6.067 | 15.364 | 1.197 | 0.672 | N/F | 1.038 | 187.782 | N/A | 1.125 | N/F | 3.833 | N/A | N/A | 2.793 | N/F | N/A | 0.54 | 0.442 | 0.958 | N/A | 0.41 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-163 | N/A | 608.499 | -141.249 | 776.718 | 0.228 | N/F | 6.085 | 14.745 | 1.477 | 1.035 | N/F | 1.276 | 187.807 | N/A | 1.065 | N/F | 2.784 | N/A | N/A | N/F | N/F | N/A | 0.599 | 0.587 | 1.607 | N/A | 0.488 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-164 | N/A | 477.573 | 702.703 | -527.035 | 0.199 | N/F | 5.163 | 41.544 | 1.355 | 0.496 | N/F | 2.595 | 187.674 | N/A | 1.033 | N/F | 3.054 | N/A | N/A | N/F | 1.282 | N/A | 1.221 | 0.297 | 1.813 | N/A | 0.38 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-165 | N/A | 456.255 | 260.59 | -217.347 | 0.207 | N/F | 4.69 | 41.894 | 1.596 | 0.811 | N/F | 0.961 | 188.16 | N/A | 0.736 | N/F | 3.212 | N/A | N/A | N/F | N/F | N/A | 0.495 | 0.399 | 0.662 | N/A | 0.434 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-166 | N/A | 468.243 | 2022.892 | -171.061 | 0.177 | N/F | 4.774 | 43.64 | 1.124 | 0.617 | N/F | 1.05 | 187.76 | N/A | 0.863 | N/F | 3.827 | N/A | N/A | N/F | N/F | N/A | 0.432 | 0.292 | 0.702 | N/A | 0.361 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-167 | N/A | 926.237 | 1857.922 | 260.741 | 0.186 | N/F | N/F | 148.683 | 1.721 | 0.739 | N/F | 1.072 | 188.278 | N/A | 0.624 | N/F | 5.225 | N/A | N/A | N/F | N/F | N/A | 1.157 | 0.185 | 1.69 | N/A | 0.319 | 68.971 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-168 | N/A | 737.249 | 978.444 | -190.548 | 0.191 | N/F | 4.384 | 139.298 | 1.091 | 0.65 | N/F | 2.366 | 188.308 | N/A | 0.73 | N/F | 2.786 | N/A | N/A | 3.339 | N/F | N/A | 0.534 | 0.247 | 1.087 | N/A | 0.353 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-169 | N/A | 1388.371 | 4127.374 | 119.983 | 0.193 | 265.354 | 4.49 | 211.612 | 1.6 | 0.818 | N/F | 1.398 | 188.907 | N/A | 0.789 | N/F | 3.554 | N/A | N/A | N/F | 1.952 | N/A | 0.82 | 0.445 | 1.623 | N/A | 0.394 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-170 | N/A | 691.204 | 2073.372 | -161.839 | 0.18 | N/F | 4.516 | 223.747 | 1.209 | 0.737 | N/F | 2.394 | 188.543 | N/A | 0.582 | N/F | 2.996 | N/A | N/A | N/F | N/F | N/A | 0.459 | 0.174 | 0.549 | N/A | 0.352 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LRB 18_INJ001 | N/A | 114.063 | 93.2 | 168.151 | 0.176 | N/F | 4.473 | 2.299 | 0.741 | 0.333 | 2.417 | 0.953 | 187.594 | N/A | 1.014 | N/F | 0.499 | N/A | N/A | N/F | 1.219 | N/A | 0.361 | 0.135 | 0.906 | N/A | 0.303 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
210451-d2-reprep-shaken | N/A | 1110.491 | 2013.411 | 613.096 | 0.185 | N/F | 4.846 | N/F | 1.04 | 1.202 | N/F | 0.892 | 188.334 | N/A | 0.895 | N/F | 114.95 | N/A | N/A | N/F | N/F | N/A | 0.453 | 0.216 | 3.26 | N/A | 0.296 | N/F | N/F | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LRB 18 | N/A | 111.461 | -59.955 | 1106.857 | 0.192 | N/F | 4.314 | 1.268 | 0.778 | 0.381 | N/F | 0.933 | 188.018 | N/A | 1.142 | N/F | 0.455 | N/A | N/A | N/F | 1.199 | N/A | 0.363 | 0.089 | 0.945 | N/A | 0.332 | 12.141 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-171 | N/A | 363.353 | 129.772 | -388.098 | 0.17 | 657.429 | 4.75 | 448.566 | 0.883 | 0.618 | N/F | 0.835 | 187.728 | N/A | 1.015 | N/F | 2.703 | N/A | N/A | N/F | N/F | N/A | 1.36 | 0.196 | 1.172 | N/A | 0.315 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-172 | N/A | 388.823 | 842.025 | 598.295 | 0.172 | N/F | 4.437 | 490.756 | 0.83 | 0.609 | N/F | 2.218 | 187.709 | N/A | 0.977 | N/F | 3.409 | N/A | N/A | N/F | 1.676 | N/A | 0.561 | 0.164 | 1.077 | N/A | 0.307 | 2.457 | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-173 | N/A | 555.787 | 1393.404 | 238.921 | 0.184 | 150.209 | 4.72 | 652.269 | 1.345 | 0.7 | N/F | 0.901 | 187.941 | N/A | 0.804 | N/F | 2.969 | N/A | N/A | N/F | 1.572 | N/A | 0.645 | 0.195 | 1.491 | N/A | 0.328 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-174 | N/A | 740.05 | 1611.789 | -318.849 | 0.174 | N/F | 4.334 | 2324.998 | 1.375 | 0.569 | N/F | 1.104 | 187.935 | N/A | 0.676 | N/F | 2.746 | N/A | N/A | 3.038 | 3.302 | N/A | 2.072 | 0.088 | 1.675 | N/A | 0.318 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-175 | N/A | 338.075 | 355.31 | -364.433 | 0.175 | N/F | N/F | 1553.208 | 0.928 | 0.703 | N/F | 1.139 | 187.774 | N/A | 0.901 | N/F | 2.174 | N/A | N/A | N/F | 1.554 | N/A | 0.439 | 0.139 | 0.839 | N/A | 0.291 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-176 | N/A | 873.16 | 1346.751 | 626.725 | 0.174 | N/F | N/F | 2159.757 | 3.32 | 0.703 | N/F | 1.438 | 188.086 | N/A | 0.911 | N/F | 2.838 | N/A | N/A | N/F | 3.578 | N/A | 0.586 | 0.226 | 0.633 | N/A | 0.32 | 11.94 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-177 | N/A | 380.104 | 588.062 | -470.427 | 0.175 | N/F | 4.969 | 14.533 | 0.98 | 0.554 | N/F | 0.977 | 187.79 | N/A | 0.786 | N/F | 1.97 | N/A | N/A | N/F | N/F | N/A | 1.46 | 0.105 | 0.592 | N/A | 0.332 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-178 | N/A | 712.816 | 747.72 | -397.303 | 0.175 | N/F | N/F | 11.442 | 1.21 | 0.733 | N/F | 2.858 | 187.987 | N/F | 1.206 | N/F | 2.336 | N/A | N/A | N/F | N/F | N/A | 0.663 | 0.153 | 1.784 | N/A | 0.323 | 135.129 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-179 | N/A | 498.837 | 246.649 | -344.837 | 0.191 | N/F | N/F | 9.067 | 1.082 | 0.897 | N/F | 0.941 | 187.734 | N/A | 0.862 | N/F | 3.117 | N/A | N/A | N/F | N/F | N/A | 0.545 | 0.064 | 0.588 | N/A | 0.326 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-180 | N/A | 471.512 | 503.558 | 80.265 | 0.178 | N/F | 4.363 | 7.807 | 1.39 | 0.577 | N/F | 2.126 | 187.797 | N/A | 0.858 | N/F | 3.526 | N/A | N/A | N/F | 1.518 | N/A | 0.622 | 0.113 | 1.33 | N/A | 0.301 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
210453-d2-reprep | N/A | 497.715 | 249.565 | 476.014 | 0.195 | N/F | 4.428 | N/F | 1.237 | 1.769 | N/F | 0.812 | 187.997 | N/A | 0.728 | N/F | 567.441 | N/A | N/A | N/F | N/F | N/A | 0.769 | 0.043 | 7.981 | N/A | 0.326 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LFB-H 10 | N/A | 183.313 | 192175.585 | 1136760.867 | 190.493 | 2233134.065 | 2259.241 | 181.985 | 165.749 | 195.644 | 154.851 | 164.645 | 397.271 | N/A | 197.667 | 148.473 | 192.246 | N/A | N/A | 193.519 | 193.018 | N/A | 196.216 | 148.837 | 200.351 | N/A | 200.62 | 189.03 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 19_INJ001 | N/A | 92.578 | 448.323 | 4643.06 | 0.879 | 9282.624 | 14.475 | 2.874 | 1.233 | 0.889 | 2.675 | 1.411 | 188.01 | N/A | 1.206 | N/F | 1.205 | N/A | N/A | 1.117 | 1.801 | N/A | 0.997 | 0.552 | 1.349 | N/A | 1.321 | 8.994 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 19 | N/A | 146.981 | 184.569 | -292.191 | 0.243 | 1660.244 | 5.742 | 2.143 | 0.991 | 0.589 | 2.013 | 1.059 | 187.579 | N/A | 0.956 | N/F | 0.481 | N/A | N/A | N/F | 1.098 | N/A | 0.459 | 0.267 | 0.961 | N/A | 0.439 | 25.896 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LD 9 | N/A | 480.119 | 1791.344 | N/F | 0.194 | N/F | 5.206 | 7.444 | 1.746 | 0.798 | N/F | 2.715 | 187.866 | N/A | 0.629 | N/F | 2.987 | N/A | N/A | N/F | N/F | N/A | 0.708 | 0.3 | 0.625 | N/A | 0.427 | 39.821 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 CCC-H 10 | N/A | 385.09 | 396357.497 | 2134411.167 | 382.755 | 4321233.184 | 4306.494 | 386.392 | 339.369 | 373.044 | 312.752 | 345.56 | 614.869 | N/A | 405.738 | 368.792 | 383.207 | N/A | N/A | 394.703 | 388.865 | N/A | 383.256 | 298.051 | 398.474 | N/A | 402.651 | 384.608 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LFSM-H 9 | N/A | 632.759 | 186689.026 | 1051632.937 | 193.501 | 3677637.883 | 4115.085 | 197.643 | 169.264 | 213.425 | 146.015 | 153.973 | 362.349 | N/A | 208.652 | N/F | 191.971 | N/A | N/A | 192.767 | 200.41 | N/A | 197.609 | 168.227 | 195.88 | N/A | 197.637 | 180.703 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-181 | N/A | 295.477 | 412.958 | 2691.353 | 0.71 | 8203.047 | 14.155 | 15.258 | 1.563 | 1.08 | N/F | 1.208 | 188.12 | N/A | 0.966 | N/F | 5.86 | N/A | N/A | 2.08 | 1.517 | N/A | 1.717 | 0.586 | 2.224 | N/A | 0.841 | N/F | N/A | N/A | N/A | N/A | C | Unknown
WO-22270-182 | N/A | 903.516 | 2511.945 | 457.954 | 0.269 | 499.822 | 7.563 | 13.876 | 1.902 | 0.83 | N/F | 1.079 | 187.786 | N/A | 0.879 | N/F | 4.081 | N/A | N/A | N/F | N/F | N/A | 0.692 | 0.569 | 1.034 | N/A | 0.479 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-183 | N/A | 305.831 | 888.062 | 55.856 | 0.234 | 67.282 | 4.935 | 16.358 | 0.822 | 0.759 | N/F | 1.153 | 187.968 | N/A | 0.864 | N/F | 3.847 | N/A | N/A | N/F | N/F | N/A | 0.573 | 0.345 | 1.842 | N/A | 0.472 | 12.809 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-184 | N/A | 425.218 | 4601.483 | 124.736 | 0.199 | N/F | 5.179 | 35.552 | 1.233 | 0.676 | N/F | 0.882 | 187.909 | N/A | 0.645 | N/F | 3.932 | N/A | N/A | N/F | N/F | N/A | 1.819 | 0.334 | 1.077 | N/A | 0.399 | 35.396 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-185 | N/A | 551.563 | 903.84 | 314.46 | 0.184 | N/F | 4.92 | 45.211 | 1.791 | 0.698 | N/F | 0.861 | 188.492 | N/A | 0.518 | N/F | 3.544 | N/A | N/A | 1.62 | 4.181 | N/A | 0.582 | 0.303 | 0.856 | N/A | 0.354 | 81.177 | N/A | N/A | N/A | N/A |  | Unknown
210439-d2-reprep | N/A | 507.148 | 761.16 | -222.136 | 0.182 | N/F | 4.659 | N/F | 1.689 | 0.746 | N/F | 2.502 | 188.551 | N/A | 0.781 | N/F | 1038.324 | N/A | N/A | N/F | N/F | N/A | 1.284 | 0.092 | 32.495 | N/A | 0.318 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-186 | N/A | 942.291 | 370.595 | 705.64 | 0.19 | N/F | 4.644 | 49.661 | 1.268 | 0.689 | N/F | 3.391 | 187.894 | N/A | 0.732 | N/F | 3.091 | N/A | N/A | N/F | 2.227 | N/A | 0.7 | 0.442 | 1.14 | N/A | 0.483 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-187 | N/A | 264.215 | 1699.485 | -34.082 | 0.184 | N/F | 4.328 | 190.082 | 1.112 | 0.721 | N/F | 1.227 | 187.651 | N/A | 0.732 | N/F | 4.366 | N/A | N/A | N/F | 1.457 | N/A | 0.67 | 0.223 | 1.973 | N/A | 0.353 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-188 | N/A | 368.475 | 1408.27 | -90.525 | 0.211 | 473.302 | 4.93 | 166.615 | 1.364 | 0.714 | N/F | 0.929 | 187.935 | N/A | 0.632 | N/F | 7.36 | N/A | N/A | 1.105 | 1.564 | N/A | 0.578 | 0.255 | 4.744 | N/A | 0.383 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-189 | N/A | 716.649 | 4049.28 | -225.788 | 0.181 | 509.394 | 4.848 | 163.398 | 1.217 | 0.695 | N/F | 1.014 | 187.647 | N/A | 0.549 | N/F | 6.961 | N/A | N/A | 2.777 | 1.988 | N/A | 1.001 | 0.189 | 1.948 | N/A | 0.306 | 2.342 | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-190 | N/A | 586.393 | 870.088 | 150.108 | 0.182 | N/F | N/F | 157.924 | 1.053 | 0.873 | N/F | 0.88 | 188.33 | N/A | 0.738 | N/F | 7.471 | N/A | N/A | N/F | 1.455 | N/A | 1.21 | 0.147 | 1.469 | N/A | 0.289 | 3.279 | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LRB 20_INJ001 | N/A | 105.621 | 190.516 | 439.517 | 0.171 | 311.764 | 4.68 | 2.254 | 1.062 | 0.338 | N/F | 1.871 | 187.677 | N/A | 0.943 | N/F | 0.413 | N/A | N/A | N/F | N/F | N/A | 0.395 | 0.215 | 0.643 | N/A | 0.343 | -1.08 | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LRB 20 | N/A | 122.89 | 719.619 | -451.517 | 0.178 | 435.444 | N/F | 1.802 | 0.784 | 0.361 | 2.185 | 0.947 | 187.61 | N/A | 1.019 | N/F | 0.406 | N/A | N/A | N/F | N/F | N/A | 0.65 | 0.224 | 0.678 | N/A | 0.307 | 0.728 | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-196 | N/A | 245.233 | 1155.914 | 39.231 | 0.173 | 140.209 | 4.328 | 1495.148 | 0.888 | 0.719 | N/F | 2.261 | 187.869 | N/A | 0.841 | N/F | 7.764 | N/A | N/A | N/F | 2.292 | N/A | 0.658 | 0.132 | 1.261 | N/A | 0.299 | N/F | N/F | N/A | N/A | N/A |  | Unknown
WO-22270-191 | N/A | 342.918 | 2497.322 | -288.367 | 0.171 | N/F | 4.28 | 590.481 | 0.892 | 0.636 | N/F | 2.233 | 187.985 | N/A | 0.697 | N/F | 7.844 | N/A | N/A | N/F | 1.491 | N/A | 1.144 | 0.121 | 1.588 | N/A | 0.309 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-197 | N/A | 480.823 | -40.592 | -350.025 | 0.175 | N/F | N/F | 1403.188 | 1.04 | 0.603 | N/F | 1.726 | 187.977 | N/A | 0.806 | N/F | 3.32 | N/A | N/A | N/F | 2.479 | N/A | 1.063 | 0.148 | 1.949 | N/A | 0.343 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-192 | N/A | 260.774 | 1358.077 | 491.542 | 0.179 | N/F | N/F | 714.487 | 0.761 | 0.671 | N/F | 2.008 | 187.851 | N/A | 0.711 | N/F | 7.309 | N/A | N/A | 1.041 | 1.811 | N/A | 0.578 | 0.131 | 1.947 | N/A | 0.326 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-198 | N/A | 379.117 | 1293.838 | -395.564 | 0.173 | N/F | 4.272 | 11190.616 | 3.573 | 0.562 | N/F | 2.524 | 187.739 | N/A | 1.082 | N/F | 4.209 | N/A | N/A | 0.957 | 3333.538 | N/A | 0.556 | 0.111 | 0.887 | N/A | 0.327 | 4.961 | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-193 | N/A | 334.998 | 354.494 | 100.364 | 0.185 | N/F | N/F | 629.943 | 0.826 | 0.573 | N/F | 1.053 | 187.674 | N/A | 0.66 | N/F | 8.219 | N/A | N/A | N/F | 1.486 | N/A | 0.646 | 0.088 | 0.916 | N/A | 0.321 | N/F | N/F | N/A | N/A | N/A | D,Q | Unknown
WO-22270-199 | N/A | 592.626 | 1474.833 | 83.177 | 0.172 | N/F | N/F | 12321.404 | 2.963 | 0.733 | N/F | 2.154 | 187.77 | N/A | 0.828 | N/F | 5.232 | N/A | N/A | N/F | 3521.747 | N/A | 0.678 | 0.135 | 1.283 | N/A | 0.304 | -1.357 | N/A | N/A | N/A | N/A | D,Q | Unknown
210415-d2-reprep | N/A | 945.963 | 819.894 | -237.992 | N/F | N/F | N/F | N/F | 1.409 | 1.116 | N/F | 0.907 | 187.862 | N/A | 0.807 | N/F | 29.291 | N/A | N/A | N/F | 1.719 | N/A | 0.873 | 0.243 | 27.337 | N/A | 0.329 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-200 | N/A | 660.136 | 2288.562 | 13.805 | 0.178 | N/F | 4.291 | 11615.642 | 2.631 | 0.562 | N/F | 1.4 | 187.651 | N/A | 0.902 | N/F | 3.552 | N/A | N/A | N/F | 3344.616 | N/A | 1.058 | 0.104 | 1.155 | N/A | 0.326 | N/F | N/F | N/A | N/A | N/A | D,Q | Unknown
WO-22270-194 | N/A | 195.729 | 1946.31 | 4566.076 | 0.178 | N/F | 4.458 | 1279.045 | 1.068 | 0.678 | N/F | 0.834 | 187.632 | N/A | 0.722 | N/F | 5.876 | N/A | N/A | 0.627 | 1.773 | N/A | 0.623 | 0.048 | 1.515 | N/A | 0.295 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-195 | N/A | 271.791 | 620.599 | -299.107 | 0.172 | 21.34 | 4.526 | 911.43 | 0.896 | 0.622 | N/F | 2.281 | 187.748 | N/A | 0.928 | N/F | 4.423 | N/A | N/A | N/F | 2.016 | N/A | 0.519 | 0.114 | 1.486 | N/A | 0.305 | N/F | N/F | N/A | N/A | N/A |  | Unknown
AB-220715-01 LFB-H 11 | N/A | 180.599 | 203287.697 | 1079076.758 | 180.881 | 1913705.613 | 2001.294 | 231.801 | 167.433 | 176.526 | 151.857 | 160.42 | 392.271 | N/A | 204.886 | 187.116 | 192.084 | N/A | N/A | 195.694 | 198.179 | N/A | 199.559 | 143.564 | 190.277 | N/A | 203.133 | 180.657 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 21 | N/A | 134.845 | -144.747 | -176.01 | 0.247 | 1452.803 | 5.844 | 1.799 | 0.881 | 0.4 | N/F | 2.293 | 187.912 | N/A | 0.833 | N/F | 0.455 | N/A | N/A | N/F | 1.323 | N/A | 0.413 | 0.178 | 0.894 | N/A | 0.447 | 11.457 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LD 10 | N/A | 724.351 | 2453.882 | 275.408 | 0.205 | 70.681 | 5.027 | 5472.988 | 1.452 | 0.657 | N/F | 2.393 | 187.695 | N/A | 0.573 | N/F | 2.957 | N/A | N/A | N/F | 2201.609 | N/A | 0.723 | 0.153 | 1.37 | N/A | 0.356 | N/F | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 CCC-H 11 | N/A | 360.377 | 411309.975 | 2209183.059 | 370.046 | 3728072.273 | 3494.515 | 422.132 | 351.242 | 357.395 | 319.013 | 330.388 | 615.979 | N/A | 416.753 | 343.684 | 407.738 | N/A | N/A | 383.878 | 372.456 | N/A | 397.692 | 278.725 | 389.571 | N/A | 396.039 | 358.14 | N/A | N/A | N/A | N/A |  | Unknown
210405-d2-reprep | N/A | 1284.279 | 3668.587 | 1764.813 | 0.232 | N/F | 4.619 | 37.659 | 2.387 | 1.022 | N/F | 1.251 | 188.845 | N/A | 0.88 | N/F | 38.635 | N/A | N/A | N/F | 9.482 | N/A | 1.101 | 0.178 | 2.417 | N/A | 0.459 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LFSM-H 10 | N/A | 1642.322 | 187709.896 | 1081933.816 | 194.187 | 2597560.901 | 3679.328 | 5413.034 | 167.184 | 216.672 | 149.14 | 147.507 | 351.761 | N/A | 197.066 | N/F | 187.19 | N/A | N/A | 187.935 | 2486.335 | N/A | 197.927 | 150.219 | 188.316 | N/A | 199.607 | 153.272 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-201 | N/A | 886.991 | 2818.686 | 2055.531 | 0.422 | 356.689 | 9.555 | 11785.487 | 2.236 | 1.119 | N/F | 1.709 | 188.017 | N/A | 0.838 | N/F | 3.658 | N/A | N/A | N/F | 3175.149 | N/A | 1.074 | 0.607 | 1.51 | N/A | 0.767 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-202 | N/A | 395.964 | 1768.526 | 189.722 | 0.257 | 75.292 | 5.169 | 12663.908 | 2.206 | 0.821 | N/F | 1.114 | 188.514 | N/A | 0.869 | N/F | 2.071 | N/A | N/A | N/F | 3534.884 | N/A | 0.541 | 0.345 | 0.69 | N/A | 0.383 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-203 | N/A | 342.351 | 1350.103 | 264.293 | 0.213 | N/F | 5.828 | 12265.672 | 2.256 | 0.677 | N/F | 1.007 | 187.666 | N/A | 0.818 | N/F | 3.444 | N/A | N/A | N/F | 3405.809 | N/A | 0.56 | 0.231 | 0.655 | N/A | 0.394 | 15.565 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 21_INJ001 | N/A | 106.422 | 435.685 | 3460.962 | 0.821 | 8843.147 | 12.058 | 2.938 | 1.046 | 0.919 | N/F | 2.29 | 188.236 | N/A | 1.136 | N/F | 0.736 | N/A | N/A | 1.119 | 2.423 | N/A | 0.881 | 0.282 | 1.379 | N/A | 0.957 | 3.968 | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-204 | N/A | 856.475 | 1298.825 | -35.948 | 0.203 | N/F | 4.913 | 11125.024 | 2.686 | 0.718 | N/F | 1.336 | 188.017 | N/A | 0.797 | N/F | 1.971 | N/A | N/A | N/F | 3483.05 | N/A | 0.626 | 0.216 | 3.116 | N/A | 0.392 | 30.705 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-205 | N/A | 632.934 | 3274.024 | -87.21 | 0.18 | N/F | 4.354 | 11343.019 | 2.122 | 0.71 | N/F | 1.674 | 188.175 | N/A | 0.626 | N/F | 2.288 | N/A | N/A | N/F | 3455.628 | N/A | 0.737 | 0.352 | 0.526 | N/A | 0.346 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-206 | N/A | 719.857 | 4357.568 | 405.929 | 0.179 | N/F | 4.486 | 11381.093 | 2.199 | 0.941 | N/F | 3.293 | 188.811 | N/A | 0.876 | N/F | 4.095 | N/A | N/A | N/F | 3284.437 | N/A | 0.995 | 0.292 | 1.254 | N/A | 0.351 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-207 | N/A | 796.364 | 1447.769 | -161.357 | 0.185 | N/F | N/F | 10832.086 | 2.467 | 0.652 | N/F | 1.242 | 188.359 | N/A | 0.768 | N/F | 5.134 | N/A | N/A | N/F | 3592.031 | N/A | 1.019 | 0.316 | 2.265 | N/A | 0.312 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-208 | N/A | 405.594 | 1736.863 | N/F | 0.184 | N/F | N/F | 7278.211 | 2.764 | 0.781 | N/F | 1.031 | 188.354 | N/A | 1.023 | N/F | 2.486 | N/A | N/A | N/F | 2847.722 | N/A | 0.571 | 0.21 | 1.277 | N/A | 0.333 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-209 | N/A | 374.189 | 622.034 | 188.453 | 0.199 | 133.364 | 4.296 | 7487.394 | 2.036 | 0.816 | N/F | 6.395 | 188.347 | N/A | 0.794 | N/F | 6.272 | N/A | N/A | N/F | 2860.632 | N/A | 0.812 | 0.099 | 2.189 | N/A | 0.312 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-210 | N/A | 326.197 | 744.226 | 371.848 | 0.175 | N/F | N/F | 7969.456 | 2.025 | 0.593 | N/F | 1.962 | 187.818 | N/A | 0.609 | N/F | 6.031 | N/A | N/A | N/F | 2869.4 | N/A | 0.526 | 0.131 | 1.151 | N/A | 0.312 | N/F | N/F | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LRB 22_INJ001 | N/A | 100.5 | 122.105 | -135.885 | 0.183 | 126.266 | 4.268 | 22.399 | 1.076 | 0.345 | N/F | 1.786 | 187.615 | N/A | 0.884 | N/F | 0.462 | N/A | N/A | N/F | 8.449 | N/A | 0.737 | 0.085 | 0.813 | N/A | 0.311 | -0.51 | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LRB 22 | N/A | 113.172 | 309.636 | 568.275 | 0.205 | N/F | 4.39 | 6.275 | 0.799 | 0.441 | N/F | 2.425 | 187.634 | N/A | 0.899 | N/F | 0.496 | N/A | N/A | 1.177 | 3.437 | N/A | 0.701 | 0.152 | 1.067 | N/A | 0.338 | 0.009 | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-211 | N/A | 208.129 | 417.085 | -215.488 | 0.175 | 57.754 | 4.326 | 7781.393 | 2.3 | 0.545 | N/F | 1.99 | 187.851 | N/A | 0.823 | N/F | 3.187 | N/A | N/A | N/F | 2748.222 | N/A | 0.51 | 0.055 | 0.993 | N/A | 0.373 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-212 | N/A | 262.063 | -15.73 | -295.261 | 0.177 | N/F | 4.431 | 7105.418 | 3.041 | 0.755 | N/F | 5.603 | 188.206 | N/A | 0.926 | N/F | 4.309 | N/A | N/A | N/F | 2676.981 | N/A | 0.59 | 0.108 | 1.958 | N/A | 0.312 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-213 | N/A | 200.2 | 1102.797 | -263 | 0.175 | N/F | N/F | 7173.014 | 2.359 | 0.615 | N/F | 2.514 | 187.737 | N/A | 0.813 | N/F | 4.681 | N/A | N/A | N/F | 2778.734 | N/A | 0.677 | 0.09 | 1.614 | N/A | 0.296 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-214 | N/A | 385.503 | 2133.271 | 89.812 | 0.164 | 68.858 | N/F | 12554.794 | 2.749 | 0.677 | N/F | 1.131 | 188.393 | N/A | 0.489 | N/F | 14.656 | N/A | N/A | N/F | 3346.883 | N/A | 1.638 | 0.046 | 1.669 | N/A | 0.343 | N/F | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-215 | N/A | 969.521 | 2632.211 | 200.106 | 0.167 | N/F | 4.357 | 12151.767 | 4.025 | 0.661 | N/F | 0.993 | 189.071 | N/A | 0.527 | N/F | 2.926 | N/A | N/A | 1.526 | 3235.902 | N/A | 0.805 | 0.034 | 1.903 | N/A | 0.34 | 27.392 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-216 | N/A | 380.126 | 506.498 | 92.728 | 0.175 | N/F | N/F | 7184.12 | 3.068 | 0.528 | N/F | 0.951 | 188.512 | N/A | 0.902 | N/F | 8.092 | N/A | N/A | N/F | 3348.427 | N/A | 1.048 | 0.072 | 2.498 | N/A | 0.292 | 33.383 | N/F | N/A | N/A | N/A |  | Unknown
WO-22270-217 | N/A | 348.426 | 263.928 | 916.392 | 0.176 | N/F | N/F | 7307.135 | 2.209 | 0.754 | N/F | 2.409 | N/F | N/A | 0.834 | N/F | 4.849 | N/A | N/A | N/F | 3450.62 | N/A | 0.919 | 0.117 | 2.127 | N/A | 0.292 | 9.253 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-218 | N/A | 228.452 | 265.157 | 233.356 | 0.168 | N/F | 4.291 | 2491.878 | 1.585 | 0.584 | N/F | 1.886 | 187.632 | N/A | 0.949 | N/F | 2.803 | N/A | N/A | N/F | 1141.632 | N/A | 0.649 | 0.058 | 0.719 | N/A | 0.363 | N/F | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-219 | N/A | 409.526 | 2373.125 | 373.959 | 0.17 | 286.527 | N/F | 2772.695 | 1.36 | 0.388 | N/F | 1.051 | 187.822 | N/A | 0.655 | N/F | 1.586 | N/A | N/A | N/F | 1133.005 | N/A | 0.554 | 0.045 | 1.084 | N/A | 0.315 | N/F | N/F | N/A | N/A | N/A |  | Unknown
WO-22270-220 | N/A | 472.426 | 988.494 | -230.532 | 0.168 | N/F | N/F | 2530.797 | 1.044 | 0.612 | N/F | 1.058 | 188.204 | N/A | 0.831 | N/F | 2.469 | N/A | N/A | N/F | 1166.087 | N/A | 0.508 | 0.052 | 0.644 | N/A | 0.306 | 17.684 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LFB-H 12 | N/A | 169.776 | 193948.501 | 1075322.799 | 182.539 | 1827906.594 | 1835.054 | 207.993 | 165.072 | 172.498 | 145.325 | 152.715 | 381.404 | N/A | 197.609 | 144.503 | 184.98 | N/A | N/A | 190.744 | 182.251 | N/A | 192.991 | 125.822 | 186.411 | N/A | 192.954 | 180.173 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 23_INJ001 | N/A | 96.558 | 113.309 | 6826.043 | 0.917 | 7421.516 | 12.544 | 1.639 | 1.162 | 0.967 | 1.448 | 1.597 | 188.309 | N/A | 1.086 | N/F | 1.255 | N/A | N/A | 1.25 | 2.016 | N/A | 1.078 | 0.504 | 1.26 | N/A | 1.566 | -0.91 | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LRB 23 | N/A | 93.238 | -31.758 | N/F | 0.271 | 2951.116 | 5.326 | 6.134 | 0.726 | 0.347 | N/F | 2.333 | 187.589 | N/A | 1.155 | N/F | 0.472 | N/A | N/A | 1.286 | 2.845 | N/A | 0.425 | 0.224 | 0.779 | N/A | 0.543 | 1.188 | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LD 11 | N/A | 205.586 | 1037.342 | -173.164 | 0.186 | 160.869 | 5.301 | 2421.487 | 1.08 | 0.656 | 4.165 | 0.907 | 187.723 | N/A | 0.736 | N/F | 1.575 | N/A | N/A | 1.654 | 1091.661 | N/A | 0.555 | 0.161 | 1.364 | N/A | 0.422 | N/F | N/F | N/A | N/A | N/A |  | Unknown
AB-220715-01 CCC-H 12 | N/A | 387.532 | 368562.575 | 2167249.335 | 363.821 | 3658441.908 | 3538.283 | 418.533 | 330.585 | 347.972 | 310.509 | 319.683 | 615.692 | N/A | 410.172 | 337.382 | 381.938 | N/A | N/A | 382.884 | 384.909 | N/A | 390.787 | 242.254 | 386.724 | N/A | 392.201 | 374.953 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LFSM-H 11 | N/A | 446.832 | 189558.942 | 1120806.553 | 199.866 | 3254113.065 | 2898.003 | 2977.339 | 165.797 | 216.358 | 144.818 | 151.821 | 367.706 | N/A | 201.771 | N/F | 194.326 | N/A | N/A | 194.698 | 1328.5 | N/A | 194.899 | 141.519 | 194.494 | N/A | 211.115 | 211.572 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-221 | N/A | 177.795 | 1144.229 | 605.924 | 0.247 | 753.691 | 5.817 | 2681.881 | 1.511 | 0.588 | N/F | 1.66 | 187.952 | N/A | 0.898 | N/F | 1.787 | N/A | N/A | N/F | 1085.099 | N/A | 0.525 | 0.109 | 0.693 | N/A | 0.579 | 4.277 | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-222 | N/A | 261.145 | 568.837 | 1030.072 | 0.227 | 102.121 | 5.862 | 2794.788 | 1.5 | 0.739 | N/F | 1.203 | 187.81 | N/A | 0.89 | N/F | 1.519 | N/A | N/A | N/F | 1067.661 | N/A | 0.698 | 0.168 | 2.007 | N/A | 0.466 | 2.197 | N/A | N/A | N/A | N/A | D,Q | Unknown
WO-22270-223 | N/A | 157.285 | 760.486 | 75.86 | 0.209 | 482.191 | 4.712 | 2436.532 | 0.92 | 0.6 | N/F | 2.83 | N/F | N/A | 0.883 | N/F | 1.186 | N/A | N/A | N/F | 1051.989 | N/A | 0.721 | 0.316 | 1.524 | N/A | 0.405 | 26.471 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-224 | N/A | 205.76 | -95.656 | -210.894 | 0.183 | N/F | 5.082 | 2567.789 | 1.157 | 0.579 | N/F | 0.94 | 188.189 | N/A | 0.929 | N/F | 2.39 | N/A | N/A | N/F | 1132.739 | N/A | 0.741 | 0.235 | 1.089 | N/A | 0.403 | 10.66 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-225 | N/A | 190.438 | 492.339 | -40.619 | 0.193 | N/F | 4.38 | 2686.931 | 0.964 | 0.575 | N/F | 0.968 | 188.426 | N/A | 0.949 | N/F | 2.002 | N/A | N/A | N/F | 1066.892 | N/A | 0.6 | 0.198 | 1.694 | N/A | 0.348 | 28.228 | N/A | N/A | N/A | N/A |  | Unknown
WO-22270-226 | N/A | 422.537 | 465.862 | 446.896 | 0.177 | N/F | 4.497 | 2806.729 | 1.551 | 0.433 | N/F | 2.501 | 187.79 | N/A | 0.743 | N/F | 1.745 | N/A | N/A | N/F | 1116.773 | N/A | 0.686 | 0.162 | 1.358 | N/A | 0.342 | 24.547 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 LRB 24_INJ001 | N/A | 73.102 | -7.809 | N/F | 0.179 | 167.865 | 4.315 | 8.854 | 0.766 | 0.493 | 2.309 | 0.875 | 187.65 | N/A | 0.991 | N/F | 0.375 | N/A | N/A | N/F | 4.418 | N/A | 0.348 | 0.135 | 0.64 | N/A | 0.329 | 3.423 | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 LRB 24 | N/A | 105.884 | -52.849 | N/F | 0.182 | 1755.915 | 4.396 | 4.17 | 0.849 | 0.323 | 2.368 | 2.338 | 187.764 | N/A | 1.382 | N/F | 0.408 | N/A | N/A | N/F | 2.162 | N/A | 0.373 | 0.196 | 0.668 | N/A | 0.343 | 3.801 | N/A | N/A | N/A | N/A | D,Q | Unknown
AB-220715-01 CCC-H 13 | N/A | 355.433 | 396434.361 | 2226765.266 | 384.528 | 4053095.978 | 3487.488 | 376.546 | 341.506 | 359.734 | 303.614 | 323.323 | 601.413 | N/A | 400.31 | 332.365 | 387.108 | N/A | N/A | 380.562 | 374.334 | N/A | 386.732 | 246.296 | 388.839 | N/A | 396.766 | 412.81 | N/A | N/A | N/A | N/A |  | Unknown
AB-220715-01 QCS 2 | N/A | 194.411 | 209001.638 | 1163656.531 | 197.955 | 31773.153 | 33.662 | 199.036 | 171.1 | 178.018 | 158.561 | 162.451 | 413.278 | N/A | 4.614 | 161.011 | 191.063 | N/A | N/A | 201.953 | 193.156 | N/A | 209.797 | 2.703 | 196.615 | N/A | 4.913 | 229.537 | N/A | N/A | N/A | N/A |  | Unknown

## Sheet2
Batch Folder Path: | D:\TraceFinderData\Projects\PFAS Online SPE\2022-07-15\2022-07-15-AB2
Batch Name: | 2022-07-15-AB2
Last Acquisition Date: | 2022-07-26 04:21
User Name: | Quantiva
Sample Count: | 342
Sample Amount - Zero
Data File | 11-Cl-PF3OUDS | 4-2:FTS | 6-2:FTS | 8-2:FTS | 9-Cl-PF3ONS | FBSA | FHxA | HFPO-DA | L-PFBS | L-PFHpS | L-PFHxS | L-PFOS | NaDONA | N-Et-FOSA | N-Me-FOSA | PFBA | PFDA | PFDoA | PFDS | PFHpA | PFHxA | PFHxDA | PFNA | PFNS | PFOA | PFODA | PFOSA | PFPeA | PFPeS | PFTeDA | PFTrDA | PFUdA | All Flags | Sample Type | Sample Acquisition Date
AB-220715-01 Cal Std 1 1 | 0 | 1.818 | 9.435 | 352.285 | 1.048 | -25.205 | 4.359 | 0 | 1.235 | 1.124 | 1.65 | 1.566 | 0 | 0 | 1.04 | 0 | 1.209 | 0 | 0 | 1.262 | 0 | 0 | 1.19 | 0.951 | 1.244 | 0 | 1.179 | 0 | 0 | 0 | 0 | 0 |  | Cal Std | 2022-07-16 18:08
AB-220715-01 Cal Std 2 1 | 0 | 1.734 | 18.276 | -345.458 | 2.017 | 65.003 | 4.311 | 2.756 | 1.809 | 2.157 | 2.259 | 2.219 | 0 | 0 | 2.076 | 0 | 2.178 | 0 | 0 | 2.237 | 2.915 | 0 | 2.12 | 2.059 | 2.18 | 0 | 2.12 | 0 | 0 | 0 | 0 | 0 |  | Cal Std | 2022-07-16 18:44
AB-220715-01 Cal Std 3 1 | 0 | 3.2 | -40.277 | -529.529 | 3.942 | -12.433 | 4.266 | 4.106 | 3.71 | 3.579 | 3.685 | 3.587 | 0 | 0 | 3.849 | 0 | 4.036 | 0 | 0 | 4.048 | 4.587 | 0 | 3.898 | 3.712 | 3.99 | 0 | 3.716 | 3.263 | 0 | 0 | 0 | 0 |  | Cal Std | 2022-07-16 19:20
AB-220715-01 Cal Std 4 1 | 0 | 8.952 | -16.464 | -322.404 | 10.076 | 20.114 | 4.354 | 7.465 | 8.981 | 9.519 | 7.953 | 8.162 | 0 | 0 | 9.418 | 0 | 9.592 | 0 | 0 | 9.603 | 9.617 | 0 | 9.647 | 8.899 | 9.57 | 0 | 9.444 | 9.819 | 0 | 0 | 0 | 0 |  | Cal Std | 2022-07-16 19:56
AB-220715-01 Cal Std 5 1 | 0 | 18.959 | -40.753 | -114.05 | 19.101 | -26.361 | 4.303 | 20.512 | 15.471 | 19.107 | 14.817 | 15.556 | 0 | 0 | 19.672 | 0 | 19.349 | 0 | 0 | 18.541 | 17.849 | 0 | 19.29 | 18.904 | 18.982 | 0 | 19.008 | 20.334 | 0 | 0 | 0 | 0 |  | Cal Std | 2022-07-16 20:32
AB-220715-01 Cal Std 6 1 | 0 | 37.3 | -38.037 | -166.428 | 39.009 | -5.741 | 4.374 | 39.383 | 32.123 | 38.467 | 29.069 | 31.304 | 0 | 0 | 39.567 | 0 | 38.803 | 0 | 0 | 37.884 | 35.309 | 0 | 38.945 | 38.728 | 38.825 | 0 | 38.797 | 37.267 | 0 | 0 | 0 | 0 |  | Cal Std | 2022-07-16 21:08
AB-220715-01 Cal Std 7 1 | 0 | 87.826 | 179.625 | -102.989 | 100.642 | 224.831 | 0 | 88.145 | 78.091 | 101.464 | 72.445 | 76.777 | 187.562 | 0 | 98.961 | 0 | 96.261 | 0 | 0 | 94.675 | 90.964 | 0 | 96.207 | 101.248 | 96.896 | 0 | 98.559 | 98.822 | 0 | 0 | 0 | 0 |  | Cal Std | 2022-07-16 21:44
AB-220715-01 Cal Std 8 1 | 0 | 182.894 | 411.664 | -12.086 | 203.297 | 115.919 | 4.689 | 187.477 | 165.757 | 185.591 | 158.86 | 164.702 | 0 | 0 | 203.68 | 192.335 | 195.086 | 0 | 0 | 195.745 | 189.927 | 0 | 197.798 | 197.698 | 195 | 0 | 200.636 | 226.938 | 0 | 0 | 0 | 0 |  | Cal Std | 2022-07-16 22:20
AB-220715-01 Cal Std 9 1 | 0 | 400.125 | 753.755 | 1343.928 | 396.449 | 675.453 | 0 | 413.164 | 345.579 | 343.07 | 319.1 | 341.327 | 0 | 0 | 413.521 | 410.126 | 385.574 | 0 | 0 | 388.162 | 381.334 | 0 | 403.989 | 393.469 | 385.596 | 0 | 401.989 | 455.957 | 0 | 0 | 0 | 0 |  | Cal Std | 2022-07-16 22:56
AB-220715-01 Cal Std 10 1 | 0 | 930.676 | 706.139 | 1761.782 | 966.546 | 745.421 | 0 | 983.792 | 830.668 | 809.341 | 810.873 | 841.7 | 558.702 | 0 | 974.347 | 930.971 | 924.724 | 0 | 0 | 921.421 | 867.467 | 0 | 919.292 | 933.907 | 924.669 | 0 | 964.723 | 1000.928 | 0 | 0 | 0 | 0 |  | Cal Std | 2022-07-16 23:32
AB-220715-01 Cal Std 11 1 | 0 | 2009.026 | 1279.872 | 4274.39 | 2030.168 | 0 | 11.732 | 1959.323 | 1747.461 | 1709.43 | 1770.213 | 1829.534 | 1305.379 | 0 | 2008.033 | 2075.212 | 1916.836 | 0 | 0 | 1874.033 | 1844.46 | 0 | 1936.193 | 1864.811 | 1893.058 | 0 | 1984.146 | 2126.978 | 0 | 0 | 0 | 0 |  | Cal Std | 2022-07-17 00:08
AB-220715-01 Cal Std 12 1 | 0 | 4015.57 | 3942.167 | 3984.467 | 4160.47 | 0 | 41.168 | 4112.801 | 3719.585 | 3910.973 | 4004.568 | 3973.032 | 2721 | 0 | 4161.904 | 4301.969 | 3952.891 | 0 | 0 | 3972.222 | 4155.832 | 0 | 4063.315 | 3932.077 | 3945.156 | 0 | 4079.025 | 3641.29 | 0 | 0 | 0 | 0 |  | Cal Std | 2022-07-17 00:44
AB-220715-01 Cal Std 13 1 | 0 | 10078.922 | 10611.598 | 7653.093 | 9844.236 | 0 | 16193.444 | 9957.078 | 8816.874 | 9807.66 | 9042.279 | 9202.601 | 12327.357 | 0 | 9840.932 | 9689.387 | 10230.461 | 0 | 0 | 10257.165 | 10175.739 | 0 | 10085.115 | 9603.318 | 10261.834 | 0 | 9973.66 | 10235.505 | 0 | 0 | 0 | 0 |  | Cal Std | 2022-07-17 01:20
AB-220715-01 LFB-H 1 | 0 | 189.043 | -130.954 | 1514.823 | 218.691 | 0 | 2352.962 | 199.413 | 167.941 | 185.56 | 154.586 | 168.032 | 400.189 | 0 | 203.567 | 184.824 | 196.5 | 0 | 0 | 191.394 | 198.905 | 0 | 197.021 | 195.76 | 195.518 | 0 | 205.593 | 229.72 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 01:56
AB-220715-01 LRB 1_INJ001 | 0 | 1.922 | -117.186 | 698.465 | 1.097 | 0 | 12.054 | 0 | 0 | 0.867 | 0 | 1.258 | 0 | 0 | 1.2 | 0 | 0.984 | 0 | 0 | 0.81 | 1.749 | 0 | 0.917 | 0.762 | 1.349 | 0 | 1.682 | -0.523 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-17 02:32
AB-220715-01 LRB 1 | 0 | 3.077 | -73.294 | 957.11 | 0.953 | 0 | 8.461 | 0 | 0 | 1.256 | 0 | 1.494 | 188.165 | 0 | 1.111 | 0 | 1.116 | 0 | 0 | 0.943 | 3.07 | 0 | 0.967 | 0.782 | 1.317 | 0 | 1.382 | 2.295 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-17 03:08
AB-220715-01 CCC-H 1 | 0 | 394.539 | 0 | 2955.157 | 428.954 | 0 | 4591.05 | 402.071 | 352.033 | 364.636 | 325.628 | 344.494 | 641.003 | 0 | 413.658 | 402.94 | 398.521 | 0 | 0 | 405.085 | 385.132 | 0 | 401.704 | 391.166 | 391.959 | 0 | 409.662 | 494.343 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 03:44
AB-220715-01 IBC 1 | 0 | 182.915 | 86.836 | 145.893 | 212.813 | 0 | 29.709 | 197.324 | 173.785 | 185.06 | 163.267 | 171.927 | 406.554 | 0 | 2.943 | 191.996 | 195.534 | 0 | 0 | 202.175 | 192.575 | 0 | 202.398 | 3.904 | 196.228 | 0 | 2.735 | 234.057 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 04:20
AB-220715-01 QCS 1 | 0 | 1.868 | 606.745 | 170.048 | 3.173 | 0 | 0 | 153.347 | 176.792 | 180.09 | 161.888 | 172.792 | 189.522 | 0 | 1.243 | 180.3 | 191.245 | 0 | 0 | 196.7 | 187.09 | 0 | 198.336 | 199.617 | 193.982 | 0 | 1.77 | 243.092 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 04:56
WO-22270-01-A-1 | 0 | 24.292 | 85.372 | 177.742 | 0.409 | 6175.874 | 0 | 0 | 0 | 1.417 | 0 | 1.107 | 0 | 0 | 0.963 | 0 | 4.515 | 0 | 0 | 0 | 0 | 0 | 1.284 | 0.343 | 2.455 | 0 | 0.639 | 0 | 0 | 0 | 0 | 0 | C | Unknown | 2022-07-17 05:32
WO-22270-02-A-1 | 0 | 13.878 | 51.058 | -146.244 | 0.801 | 5948.936 | 0 | 0 | 0 | 0.979 | 0 | 1.166 | 0 | 0 | 0.762 | 0 | 4.114 | 0 | 0 | 0 | 0 | 0 | 0.801 | 0.191 | 2.629 | 0 | 0.531 | 0 | 0 | 0 | 0 | 0 | C | Unknown | 2022-07-17 06:08
WO-22270-03-A-1 | 0 | 33.292 | -332.436 | -202.581 | 0.706 | 3782.3 | 0 | 0 | 0 | 0.649 | 0 | 0.982 | 0 | 0 | 0.799 | 0 | 3.879 | 0 | 0 | 0 | 0 | 0 | 0.734 | 0.2 | 1.389 | 0 | 0.504 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 06:44
WO-22270-04-A-1 | 0 | 24.578 | 193.196 | -82.625 | 0.323 | 4588.594 | 6.732 | 0 | 0 | 0 | 0 | 1.082 | 0 | 0 | 0.848 | 0 | 3.959 | 0 | 0 | 0 | 0 | 0 | 0.854 | 0.254 | 1.569 | 0 | 0.453 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 07:20
WO-22270-05-A-1 | 0 | 28.831 | -222.137 | -251.473 | 0.341 | 4983.912 | 0 | 0 | 0 | 0 | 0 | 1.161 | 0 | 0 | 0.771 | 0 | 4.702 | 0 | 0 | 1.496 | 0 | 0 | 0.677 | 0.982 | 3.116 | 0 | 0.414 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 07:56
WO-22270-06-A-1 | 0 | 36.226 | -107.551 | -283.067 | 0.535 | 3286.822 | 0 | 3.727 | 0 | 0.676 | 0 | 1.226 | 0 | 0 | 0.863 | 0 | 3.878 | 0 | 0 | 0 | 0 | 0 | 1.073 | 0.106 | 3.167 | 0 | 0.422 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 08:32
WO-22270-07-A-1 | 0 | 21.669 | -92.364 | 98.181 | 0.253 | 2959.834 | 7.197 | 0 | 0 | 0 | 0 | 0.834 | 0 | 0 | 0.84 | 0 | 3.862 | 0 | 0 | 0 | 0 | 0 | 0.509 | 0.118 | 2.314 | 0 | 0.383 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 09:08
WO-22270-08-A-1 | 0 | 24.927 | -48.896 | -286.629 | 0.513 | 7371.95 | 6.516 | 0 | 0 | 0.849 | 0 | 0.938 | 0 | 0 | 0.649 | 0 | 4.404 | 0 | 0 | 0 | 0 | 0 | 0.848 | 0.911 | 2.862 | 0 | 0.411 | 0 | 0 | 0 | 0 | 0 | C | Unknown | 2022-07-17 09:44
WO-22270-09-A-1 | 0 | 26.592 | -139.748 | -324.447 | 0.287 | 1554.936 | 4.981 | 0 | 0 | 0.657 | 0 | 2.294 | 0 | 0 | 0.735 | 0 | 5.182 | 0 | 0 | 0 | 0 | 0 | 1.009 | 0.595 | 1.798 | 0 | 0.437 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 10:20
WO-22270-10-A-1 | 0 | 27.52 | -171.328 | 55.63 | 0.221 | 2215.248 | 5.815 | 0 | 0 | 0.65 | 0 | 0.933 | 0 | 0 | 0.734 | 0 | 6.309 | 0 | 0 | 0 | 0 | 0 | 0.804 | 0.56 | 2.843 | 0 | 0.354 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 10:56
AB-220715-01 LRB 2_INJ001 | 0 | 3.785 | -245.748 | -68.982 | 0.379 | 42.315 | 4.974 | 0 | 0 | 0.341 | 1.014 | 0.918 | 0 | 0 | 0.917 | 0 | 0.443 | 0 | 0 | 0 | 0 | 0 | 0.708 | 0.203 | 1.595 | 0 | 0.342 | -0.603 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-17 11:32
AB-220715-01 LRB 2 | 0 | 4.965 | -84.895 | -481.747 | 0.387 | 62.421 | 4.839 | 0 | 0 | 0.338 | 0 | 1.055 | 0 | 0 | 0.915 | 0 | 0.367 | 0 | 0 | 0 | 0 | 0 | 0.802 | 0.047 | 1.568 | 0 | 0.398 | -0.336 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-17 12:08
WO-22270-11-A-1 | 0 | 15.981 | -320.785 | -106.147 | 0.41 | 430.438 | 0 | 0 | 0 | 1.135 | 0 | 2.047 | 187.783 | 0 | 0.851 | 0 | 4.526 | 0 | 0 | 0 | 0 | 0 | 0.482 | 0.546 | 1.094 | 0 | 0.366 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 12:44
WO-22270-12-A-1 | 0 | 28.86 | -155.394 | -173.09 | 0.203 | 1147.317 | 6.071 | 0 | 0 | 0.537 | 0 | 0.921 | 0 | 0 | 0.77 | 0 | 3.046 | 0 | 0 | 1.008 | 0 | 0 | 0.455 | 0.057 | 2.504 | 0 | 0.369 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 13:20
WO-22270-13-A-1 | 0 | 22.779 | -80.718 | -23.135 | 0.36 | 1149.207 | 5.747 | 1.644 | 0 | 0.623 | 0.942 | 0.867 | 0 | 0 | 0.975 | 0 | 4.196 | 0 | 0 | 0 | 0 | 0 | 0.516 | 0.481 | 0 | 0 | 0.425 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 13:56
WO-22270-14-A-1 | 0 | 14.976 | -181.535 | -466.736 | 0.184 | 2286.665 | 5.506 | 0 | 0 | 1.446 | 0 | 0.847 | 0 | 0 | 1.071 | 0 | 4.113 | 0 | 0 | 0 | 0 | 0 | 0.372 | 0.448 | 0 | 0 | 0.367 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 14:32
WO-22270-15-A-1 | 0 | 24.808 | -120.171 | -380.016 | 0.226 | 1863.968 | 5.736 | 0 | 0 | 0.82 | 0 | 1.973 | 187.636 | 0 | 1.086 | 0 | 2.456 | 0 | 0 | 0 | 0 | 0 | 0.761 | 0.5 | 2.884 | 0 | 0.374 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 15:08
WO-22270-16-A-1 | 0 | 31.508 | 5.894 | -300.746 | 0.2 | 2588.437 | 6.466 | 0 | 0 | 0.757 | 0 | 0.965 | 0 | 0 | 0.934 | 0 | 4.631 | 0 | 0 | 0 | 0 | 0 | 0.681 | 0.554 | 2.347 | 0 | 0.356 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 15:44
WO-22270-17-A-1 | 0 | 28.338 | -180.179 | -310.139 | 0.257 | 281.051 | 5.48 | 0 | 0 | 0.768 | 0 | 0.832 | 187.841 | 0 | 0.946 | 0 | 4.525 | 0 | 0 | 0.974 | 0 | 0 | 0.916 | 0.04 | 3.796 | 0 | 0.338 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 16:20
WO-22270-18-A-1 | 0 | 31.56 | 47.573 | -136.151 | 0.203 | 589.695 | 5.553 | 0 | 0 | 1.356 | 1.431 | 0.83 | 187.979 | 0 | 1.221 | 0 | 1.851 | 0 | 0 | 0 | 0 | 0 | 0.394 | 0.528 | 1.384 | 0 | 0.354 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 16:56
WO-22270-19-A-1 | 0 | 25.166 | -106.945 | -465.721 | 0.216 | 2527.997 | 5.513 | 1.644 | 0 | 1.189 | 0 | 0.866 | 0 | 0 | 0.93 | 0 | 3.432 | 0 | 0 | 0 | 0 | 0 | 0.482 | 0.539 | 1.812 | 0 | 0.351 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 17:32
WO-22270-20-A-1 | 0 | 53.096 | -359.243 | -466.001 | 0.214 | 208.349 | 4.773 | 2.717 | 0 | 1.284 | 0 | 0.99 | 187.828 | 0 | 0.88 | 0 | 5.105 | 0 | 0 | 0 | 0 | 0 | 0.447 | 0.487 | 2.423 | 0 | 0.382 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 18:08
AB-220715-01 LFB-H 2 | 0 | 184.906 | 113630.304 | 979.726 | 205.592 | 2546.837 | 2525.067 | 194.921 | 161.04 | 182.148 | 152.172 | 160.442 | 389.669 | 0 | 205.831 | 159.874 | 193.606 | 0 | 0 | 190.051 | 187.047 | 0 | 192.039 | 183.021 | 190.583 | 0 | 199.766 | 217.366 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 18:44
AB-220715-01 LRB 3_INJ001 | 0 | 6.882 | 7.627 | 251.555 | 0.409 | 263.867 | 7.266 | 0 | 0 | 0.45 | 1.095 | 1.016 | 0 | 0 | 0.916 | 0 | 0.521 | 0 | 0 | 0 | 1.04 | 0 | 0.487 | 0.238 | 1.137 | 0 | 0.747 | -1.623 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-17 19:20
AB-220715-01 LRB 3 | 0 | 3.868 | -114.955 | 27.184 | 0.229 | 914.719 | 5.356 | 0 | 0 | 0.584 | 0.969 | 0.842 | 0 | 0 | 0.987 | 0 | 0.4 | 0 | 0 | 0 | 0.977 | 0 | 0.453 | 0.407 | 1.107 | 0 | 0.393 | -0.875 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-17 19:56
AB-220715-01 LD 1 | 0 | 37.29 | 20.936 | -120.81 | 0.204 | 1966.004 | 5.821 | 2.029 | 0 | 0.683 | 1.804 | 0.834 | 0 | 0 | 0.997 | 0 | 3.428 | 0 | 0 | 0.935 | 0 | 0 | 0.621 | 0.216 | 2.867 | 0 | 0.39 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 20:32
AB-220715-01 CCC-H 2 | 0 | 387.163 | 397810.785 | 35616.145 | 388.636 | 4429248.499 | 4596.8 | 399.928 | 347.787 | 348.057 | 320.885 | 336.485 | 621.876 | 0 | 401.9 | 364.718 | 386.616 | 0 | 0 | 382.625 | 380.048 | 0 | 387.617 | 350.75 | 378.815 | 0 | 388.012 | 542.884 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 21:08
AB-220715-01 LFSM-H 1 | 0 | 200.539 | 202856.498 | 45270.74 | 199.423 | 2276612.707 | 4148.379 | 188.888 | 164.433 | 228.183 | 151.916 | 149.476 | 365.052 | 0 | 200.828 | 0 | 196.881 | 0 | 0 | 194.52 | 194.271 | 0 | 194.736 | 215.661 | 194.121 | 0 | 203.344 | 142.788 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 21:44
WO-22270-21-A-1 | 0 | 22.425 | 6.685 | 65.907 | 0.473 | 6553.62 | 11.657 | 2.659 | 0 | 1.599 | 0 | 1.043 | 188.036 | 0 | 0.863 | 0 | 3.786 | 0 | 0 | 1.496 | 0 | 0 | 0.654 | 0.388 | 4.374 | 0 | 0.79 | 0 | 0 | 0 | 0 | 0 | C | Unknown | 2022-07-17 22:20
WO-22270-22-A-1 | 0 | 32.433 | -70.052 | -379.965 | 0.284 | 1030.003 | 6.405 | 0 | 0 | 1.279 | 0 | 0.808 | 0 | 0 | 0.842 | 0 | 3.208 | 0 | 0 | 0 | 0 | 0 | 0.608 | 0.207 | 1.838 | 0 | 0.55 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 22:56
WO-22270-23-A-1 | 0 | 38.47 | -210.781 | 740.972 | 0.24 | 1422.215 | 7.404 | 0 | 0 | 1.168 | 0 | 0.896 | 0 | 0 | 0.95 | 0 | 4.061 | 0 | 0 | 0.833 | 2.005 | 0 | 0.641 | 0.508 | 3.327 | 0 | 0.393 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-17 23:32
WO-22270-24-A-1 | 0 | 30.713 | 111.542 | -484.943 | 0.228 | 957.291 | 6.573 | 0 | 0 | 1.376 | 0 | 1.016 | 0 | 0 | 0.921 | 0 | 3.347 | 0 | 0 | 0 | 0 | 0 | 0.582 | 0.295 | 2.263 | 0 | 0.356 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 00:08
WO-22270-25-A-1 | 0 | 22.896 | -281.866 | -239.709 | 0.217 | 656.068 | 6.098 | 0 | 0 | 0.7 | 0 | 1.775 | 187.758 | 0 | 0.765 | 0 | 4.386 | 0 | 0 | 0 | 0 | 0 | 0.552 | 0.22 | 0 | 0 | 0.404 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 00:44
WO-22270-26-A-1 | 0 | 37.312 | 132.811 | -316.779 | 0 | 1561.491 | 5.905 | 0 | 0 | 0.866 | 0 | 2.491 | 0 | 0 | 1.161 | 0 | 3.774 | 0 | 0 | 0.998 | 0 | 0 | 0.619 | 0.558 | 2.693 | 0 | 0.337 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 01:20
WO-22270-27-A-1 | 0 | 23.13 | 2.289 | -407.11 | 0.191 | 1001.232 | 4.838 | 0 | 0 | 1.137 | 0 | 0.946 | 187.897 | 0 | 1.037 | 0 | 3.312 | 0 | 0 | 0.948 | 0 | 0 | 0.47 | 0.307 | 1.492 | 0 | 0.385 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 01:56
WO-22270-28-A-1 | 0 | 24.68 | -121.78 | -432.129 | 0.18 | 265 | 5.098 | 0 | 0 | 0.697 | 0 | 0.892 | 0 | 0 | 0.904 | 0 | 4.549 | 0 | 0 | 0 | 0 | 0 | 0.471 | 0.034 | 0 | 0 | 0.324 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 02:32
WO-22270-29-A-1 | 0 | 27.328 | -147.154 | 93.347 | 0.186 | 2261.44 | 5.124 | 0 | 0 | 0.668 | 0 | 0.826 | 0 | 0 | 0.64 | 0 | 4.642 | 0 | 0 | 0 | 0 | 0 | 0.539 | 0.263 | 0.799 | 0 | 0.35 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 03:08
WO-22270-30-A-1 | 0 | 25.218 | -186.109 | 12.111 | 0.177 | 198.397 | 4.866 | 2.817 | 0 | 0.926 | 0 | 1.847 | 0 | 0 | 0.762 | 0 | 4.086 | 0 | 0 | 1.255 | 1.854 | 0 | 0.836 | 0.24 | 2.664 | 0 | 0.313 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 03:44
AB-220715-01 LRB 4_INJ001 | 0 | 5.692 | 22.562 | -121 | 0.171 | -16.951 | 4.367 | 0 | 0 | 0.348 | 0.968 | 1.507 | 0 | 0 | 1.156 | 0 | 0.328 | 0 | 0 | 0 | 1.03 | 0 | 0.597 | 0.239 | 1.15 | 0 | 0.286 | -2.028 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-18 04:20
AB-220715-01 LRB 4 | 0 | 7.094 | -130.591 | -115.158 | 0.178 | 20.383 | 4.614 | 0 | 0 | 0.405 | 0.951 | 0.839 | 0 | 0 | 1.018 | 0 | 0.33 | 0 | 0 | 0 | 1.05 | 0 | 0.398 | 0.146 | 1.046 | 0 | 0.284 | -0.494 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-18 04:56
WO-22270-31-A-1 | 0 | 38.18 | 157.496 | 13.156 | 0.179 | 0 | 4.846 | 0.624 | 0 | 1.16 | 0 | 0.864 | 187.756 | 0 | 0.831 | 0 | 2.993 | 0 | 0 | 0 | 0 | 0 | 1.049 | 0.261 | 1.614 | 0 | 0.373 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 05:32
WO-22270-32-A-1 | 0 | 38.234 | 40.312 | 0 | 0.22 | 107.206 | 4.777 | 0 | 0 | 0.528 | 0 | 2.278 | 187.979 | 0 | 0.911 | 0 | 3.971 | 0 | 0 | 0 | 0 | 0 | 0.671 | 0.328 | 1.812 | 0 | 0.345 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 06:08
WO-22270-33-A-1 | 0 | 28.871 | -19.527 | 245.309 | 0.182 | 557.936 | 5.647 | 0.752 | 0 | 0.664 | 0 | 0.899 | 187.795 | 0 | 0.892 | 0 | 3.255 | 0 | 0 | 0 | 0 | 0 | 0.42 | 0.19 | 2.319 | 0 | 0.331 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 06:44
WO-22270-34-A-1 | 0 | 49.009 | 483.692 | 20.303 | 0.512 | 0 | 4.704 | 0 | 0 | 2.173 | 0 | 1.188 | 188.404 | 0 | 0.868 | 0 | 4.461 | 0 | 0 | 1.774 | 0 | 0 | 0.998 | 0.873 | 3.094 | 0 | 0.358 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 07:20
WO-22270-35-A-1 | 0 | 25.252 | -140.842 | 672.749 | 0.178 | 94.903 | 4.293 | 0 | 0 | 0.97 | 0 | 1.696 | 188.321 | 0 | 0.897 | 0 | 2.965 | 0 | 0 | 1.682 | 0 | 0 | 1.591 | 0.313 | 0.962 | 0 | 0.331 | -1.244 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-18 07:56
WO-22270-36-A-1 | 0 | 16.799 | -241.736 | -206.944 | 0.176 | 582.572 | 4.784 | 0.078 | 0 | 0.609 | 0 | 1.719 | 0 | 0 | 0.795 | 0 | 3.607 | 0 | 0 | 0 | 0 | 0 | 0.517 | 0.259 | 1.501 | 0 | 0.335 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 08:32
WO-22270-37-A-1 | 0 | 39.245 | 56.791 | 124.41 | 0.18 | 0 | 4.343 | 0 | 0.915 | 0.687 | 0 | 1.024 | 187.741 | 0 | 0.849 | 0 | 3.229 | 0 | 0 | 0 | 0 | 0 | 0.447 | 0.183 | 1.387 | 0 | 0.308 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-18 09:08
WO-22270-38-A-1 | 0 | 39.077 | -11.461 | -242.322 | 0.173 | 73.51 | 4.53 | 0 | 1.299 | 0.749 | 0 | 0.958 | 187.744 | 0 | 0.899 | 0 | 2.328 | 0 | 0 | 0 | 0 | 0 | 0.48 | 0.376 | 1.752 | 0 | 0.349 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 09:44
WO-22270-39-A-1 | 0 | 45.013 | -159.732 | 170.022 | 0.183 | 140.387 | 5.026 | 0 | 0 | 0.773 | 0 | 2.016 | 187.866 | 0 | 0.908 | 0 | 2.578 | 0 | 0 | 0 | 0 | 0 | 0.429 | 0.211 | 1.962 | 0 | 0.346 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 10:20
WO-22270-40-A-1 | 0 | 42.996 | -338.47 | -239.222 | 0.17 | 421.879 | 4.508 | 1.034 | 0 | 0.742 | 0 | 0.952 | 187.789 | 0 | 0.862 | 0 | 3.462 | 0 | 0 | 0 | 0 | 0 | 1.428 | 0.104 | 0.761 | 0 | 0.33 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 10:56
AB-220715-01 LFB-H 3 | 0 | 191.433 | 204953.821 | 1065183.17 | 198.893 | 2534851.855 | 2425.241 | 196.235 | 168.028 | 177.805 | 153.757 | 164.912 | 396.715 | 0 | 204.158 | 157.99 | 198.894 | 0 | 0 | 194.119 | 190.785 | 0 | 199.696 | 173.364 | 199.713 | 0 | 203.969 | 205.399 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 11:32
AB-220715-01 LRB 5_INJ001 | 0 | 7.584 | -281.959 | 949.865 | 0.402 | 701.571 | 6.195 | -0.174 | 0 | 0.458 | 1.63 | 0.988 | 187.563 | 0 | 1.024 | 0 | 0.496 | 0 | 0 | 0 | 0.976 | 0 | 0.45 | 0.224 | 0.766 | 0 | 0.681 | -1.056 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-18 12:08
AB-220715-01 LRB 5 | 0 | 8.422 | 130.084 | 79.684 | 0.244 | 584.304 | 4.815 | -0.157 | 0.683 | 0.348 | 1.69 | 1.616 | 0 | 0 | 1.069 | 0 | 0.431 | 0 | 0 | 0 | 1.012 | 0 | 0.41 | 0.144 | 0.861 | 0 | 0.392 | 0.369 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-18 12:44
AB-220715-01 LD 2 | 0 | 45.096 | -20.249 | 335.456 | 0.193 | 560.067 | 6.321 | 0 | 1.048 | 0.866 | 2.332 | 1.012 | 188.049 | 0 | 0.873 | 0 | 2.589 | 0 | 0 | 1.392 | 0 | 0 | 1.059 | 0.308 | 4.203 | 0 | 0.334 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 13:20
AB-220715-01 CCC-H 3 | 0 | 388.814 | 399990.35 | 2271882.462 | 389.626 | 5436115.528 | 4664.145 | 397.811 | 340.157 | 346.509 | 317.243 | 332.226 | 616.069 | 0 | 398.999 | 361.667 | 389.63 | 0 | 0 | 390.952 | 382.755 | 0 | 391.511 | 328.05 | 391.053 | 0 | 405.276 | 409.892 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 13:56
AB-220715-01 LFSM-H 2 | 0 | 202.915 | 199860.162 | 1136033.921 | 209.084 | 3807941.254 | 3781.142 | 184.578 | 168.055 | 209.163 | 150.282 | 152.35 | 372.471 | 0 | 205.796 | 0 | 196.473 | 0 | 0 | 198.245 | 183.486 | 0 | 196.189 | 205.737 | 198.046 | 0 | 198.227 | 161.408 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 14:32
WO-22270-41-A-1 | 0 | 36.121 | -79.897 | 843.726 | 0.243 | 1065.53 | 7.473 | 0 | 0 | 0.758 | 0 | 0.97 | 0 | 0 | 0.884 | 0 | 2.578 | 0 | 0 | 0 | 0 | 0 | 0.66 | 0.34 | 0.959 | 0 | 0.412 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 15:08
WO-22270-42-A-1 | 0 | 29.638 | -151.906 | 53.387 | 0.214 | 1780.241 | 5.602 | 2.162 | 1.077 | 0.842 | 0 | 0.841 | 188.125 | 0 | 0.877 | 0 | 1.633 | 0 | 0 | 0 | 0 | 0 | 0.428 | 0.434 | 1.255 | 0 | 0.378 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 15:44
WO-22270-43-A-1 | 0 | 48.714 | -32.864 | -302.054 | 0.204 | 1765.114 | 5.076 | 0 | 0 | 0.681 | 0 | 0.917 | 187.736 | 0 | 0.792 | 0 | 2.633 | 0 | 0 | 0 | 0 | 0 | 0.663 | 0.425 | 4.642 | 0 | 0.357 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 16:20
WO-22270-44-A-1 | 0 | 31.123 | 487.277 | 444.975 | 0.199 | 672.882 | 4.694 | 1.368 | 0 | 0.792 | 0 | 0.928 | 188.324 | 0 | 0.911 | 0 | 1.905 | 0 | 0 | 0 | 0 | 0 | 0.442 | 0.325 | 1.306 | 0 | 0.303 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 16:56
WO-22270-45-A-1 | 0 | 33.222 | -210.097 | 379.669 | 0.176 | 1818.866 | 4.687 | 3.345 | 1.289 | 0.595 | 0 | 0.858 | 187.705 | 0 | 0.762 | 0 | 1.785 | 0 | 0 | 3.221 | 0 | 0 | 1.441 | 0.212 | 0 | 0 | 0.359 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 17:32
WO-22270-52-A-1 | 0 | 59.001 | 94.326 | 484.088 | 0.189 | 216.345 | 4.697 | 3.526 | 1.25 | 0.597 | 0 | 2.063 | 188.149 | 0 | 1.078 | 0 | 2.789 | 0 | 0 | 0 | 0 | 0 | 0.646 | 0.284 | 1.301 | 0 | 0.334 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 18:08
WO-22270-46-A-1 | 0 | 31.053 | 253.48 | 44.055 | 0.183 | 0 | 4.479 | 2.103 | 1.039 | 0.679 | 0 | 1.511 | 187.71 | 0 | 0.843 | 0 | 3.148 | 0 | 0 | 0 | 0 | 0 | 0.561 | 0.255 | 1.241 | 0 | 0.35 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-18 18:44
WO-22270-47-A-1 | 0 | 32.404 | 52.566 | 434.154 | 0.186 | 610.917 | 4.535 | 0.861 | 0 | 0.8 | 0 | 0.922 | 187.967 | 0 | 0.825 | 0 | 1.894 | 0 | 0 | 0 | 0 | 0 | 0.441 | 0.241 | 1.16 | 0 | 0.323 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 19:20
WO-22270-53-A-1 | 0 | 37.429 | 139.613 | -116.772 | 0.176 | 0 | 4.363 | 1.472 | 1.507 | 0.769 | 0 | 0.826 | 187.683 | 0 | 0.94 | 0 | 3.553 | 0 | 0 | 0 | 0 | 0 | 0.685 | 0.189 | 1.634 | 0 | 0.37 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-18 19:56
WO-22270-48-A-1 | 0 | 29.045 | 109.329 | 137.623 | 0.181 | 453.766 | 5.949 | 1.281 | 0 | 0.682 | 0 | 1.93 | 187.782 | 0 | 0.915 | 0 | 3.977 | 0 | 0 | 0 | 0 | 0 | 0.378 | 0.218 | 0.488 | 0 | 0.324 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 20:32
WO-22270-54-A-1 | 0 | 20.422 | -137.626 | 327.654 | 0.172 | 0 | 4.419 | 3.071 | 0 | 0.602 | 0 | 1.31 | 187.582 | 0 | 0.893 | 0 | 3.216 | 0 | 0 | 0 | 1.304 | 0 | 0.375 | 0.155 | 0.828 | 0 | 0.316 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 21:08
WO-22270-49-A-1 | 0 | 32.428 | 28.722 | -73.564 | 0.172 | 1899.826 | 4.755 | 0 | 0 | 0.677 | 0 | 0.988 | 187.898 | 0 | 0.803 | 0 | 4.442 | 0 | 0 | 0 | 0 | 0 | 0.497 | 0.263 | 0.776 | 0 | 0.366 | 17.052 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 21:44
WO-22270-55-A-1 | 0 | 37.057 | -7.986 | 38.086 | 0.21 | 220.186 | 0 | 2.916 | 1.289 | 0.726 | 0 | 3.016 | 188.052 | 0 | 0.837 | 0 | 2.961 | 0 | 0 | 0 | 0 | 0 | 3.128 | 0.453 | 0 | 0 | 0.404 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 22:20
WO-22270-50-A-1 | 0 | 31.937 | 300.571 | -79.736 | 0.179 | 157.73 | 4.388 | 0 | 0.742 | 0.711 | 0 | 0.949 | 187.973 | 0 | 0.951 | 0 | 4.275 | 0 | 0 | 0 | 0 | 0 | 1.212 | 0.365 | 0.833 | 0 | 0.328 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 22:56
WO-22270-56-A-1 | 0 | 42.16 | -76.376 | -201.312 | 0.177 | 1464.114 | 4.694 | 0 | 1.308 | 0.534 | 0 | 0.891 | 187.762 | 0 | 0.859 | 0 | 4.414 | 0 | 0 | 0 | 0 | 0 | 0.793 | 0.117 | 1.763 | 0 | 0.366 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-18 23:32
WO-22270-57-A-1 | 0 | 20.638 | -208.802 | -85.049 | 0.168 | 190.818 | 4.445 | 2.437 | 1.272 | 0.728 | 0 | 1.085 | 187.988 | 0 | 0.722 | 0 | 4.114 | 0 | 0 | 0 | 0 | 0 | 0.685 | 0.05 | 0.912 | 0 | 0.34 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 00:08
AB-220715-01 LRB 6_INJ001 | 0 | 8.845 | -55.064 | 91.123 | 0.182 | 4.099 | 4.463 | 0.438 | 0 | 0.286 | 1.676 | 1.236 | 187.637 | 0 | 1.069 | 0 | 0.36 | 0 | 0 | 0 | 0.966 | 0 | 0.346 | 0.121 | 0.485 | 0 | 0.325 | -2.567 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-19 00:44
AB-220715-01 LRB 6 | 0 | 7.103 | -46.72 | -203.463 | 0.182 | 0 | 4.335 | 1.195 | 0.557 | 0.301 | 1.456 | 0.934 | 187.569 | 0 | 1.084 | 0 | 0.365 | 0 | 0 | 0 | 1.007 | 0 | 0.342 | 0.135 | 0.535 | 0 | 0.329 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-19 01:20
WO-22270-58-A-1 | 0 | 60.784 | 439.04 | 170.062 | 0.175 | 0 | 0 | 0 | 2.421 | 0.965 | 0 | 1.217 | 187.901 | 0 | 0.856 | 0 | 3.988 | 0 | 0 | 0 | 0 | 0 | 0.383 | 0.348 | 1.42 | 0 | 0.34 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-19 01:56
WO-22270-51-A-1 | 0 | 25.845 | 658.879 | 489.993 | 0.173 | 0 | 4.602 | 0 | 0 | 0.761 | 0 | 1.713 | 187.606 | 0 | 0.761 | 0 | 2.866 | 0 | 0 | 0 | 0 | 0 | 0.944 | 0.186 | 4.097 | 0 | 0.316 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 02:32
WO-22270-59-A-1 | 0 | 38.107 | 32.296 | -100.643 | 0.171 | 34.57 | 4.641 | 0 | 1.069 | 0.537 | 0 | 1.563 | 187.745 | 0 | 0.972 | 0 | 3.171 | 0 | 0 | 0 | 0 | 0 | 0.473 | 0.127 | 0.765 | 0 | 0.349 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 03:08
WO-22270-60-A-1 | 0 | 40.536 | 9.291 | 702.691 | 0.174 | 260.992 | 4.547 | 0 | 2.623 | 0.69 | 0 | 0.999 | 187.794 | 0 | 0.912 | 0 | 3.518 | 0 | 0 | 0 | 0 | 0 | 1.577 | 0.091 | 3.544 | 0 | 0.34 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 03:44
AB-220715-01 LFB-H 4 | 0 | 181.393 | 205949.66 | 1215643.239 | 204.863 | 2564720.119 | 2282.495 | 197.68 | 175.872 | 182.263 | 157.175 | 165.163 | 401.16 | 0 | 209.15 | 162.744 | 198.772 | 0 | 0 | 196.702 | 199.009 | 0 | 204.168 | 164.372 | 199.897 | 0 | 203.969 | 200.013 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 04:20
AB-220715-01 LRB 7_INJ001 | 0 | 8.831 | 96.429 | 1035.383 | 0.398 | 859.46 | 7.397 | -0.005 | 0.669 | 0.499 | 1.998 | 1.365 | 187.729 | 0 | 1.318 | 0 | 0.506 | 0 | 0 | 0 | 1.05 | 0 | 0.567 | 0.236 | 1.489 | 0 | 0.715 | -2.539 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-19 04:56
AB-220715-01 LRB 7 | 0 | 9.454 | -45.321 | 589.02 | 0.231 | 481.512 | 4.788 | 0.5 | 0 | 0.381 | 2.109 | 0.94 | 187.633 | 0 | 1.398 | 0 | 0.397 | 0 | 0 | 0 | 0.992 | 0 | 0.776 | 0.176 | 1.449 | 0 | 0.397 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 05:32
AB-220715-01 LD 3 | 0 | 41.583 | 548.163 | 628.968 | 0.204 | 539.1 | 5.759 | 0 | 1.254 | 1.038 | 0 | 0.955 | 187.727 | 0 | 0.871 | 0 | 3.264 | 0 | 0 | 1.952 | 0 | 0 | 0.487 | 0.373 | 1.999 | 0 | 0.344 | 5.741 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 06:08
AB-220715-01 CCC-H 4 | 0 | 348.664 | 380225.664 | 2358966.15 | 391.157 | 6954813.794 | 5338.629 | 388.596 | 332.142 | 353.136 | 321.727 | 331.024 | 614.304 | 0 | 402.846 | 336.451 | 382.399 | 0 | 0 | 384.014 | 377.704 | 0 | 386.028 | 362.543 | 387.419 | 0 | 404.275 | 472.636 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 06:44
AB-220715-01 LFSM-H 3 | 0 | 181.278 | 207815.95 | 1162763.306 | 196.814 | 4876860.321 | 4987.276 | 191.02 | 176.858 | 235.491 | 151.448 | 154.815 | 363.347 | 0 | 203.658 | 0 | 200.285 | 0 | 0 | 196.918 | 182.049 | 0 | 198.045 | 211.178 | 197.926 | 0 | 206.493 | 187.576 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 07:20
WO-22270-61-A-1 | 0 | 40.546 | 76.828 | -169.15 | 0.26 | 1770.586 | 6.746 | 0 | 1.818 | 1.026 | 0 | 2.014 | 187.654 | 0 | 0.84 | 0 | 3.348 | 0 | 0 | 0 | 0 | 0 | 0.482 | 0.37 | 0.83 | 0 | 0.423 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 07:56
WO-22270-62-A-1 | 0 | 40.311 | -94.139 | -139.346 | 0.226 | 496.242 | 5.911 | 1.249 | 1.636 | 0.981 | 0 | 2.345 | 187.819 | 0 | 0.712 | 0 | 3.215 | 0 | 0 | 0 | 0 | 0 | 0.586 | 0.099 | 1.122 | 0 | 0.334 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 08:32
WO-22270-63-A-1 | 0 | 51.491 | -196.216 | 241.773 | 0.209 | 240.577 | 6.226 | 0 | 2.378 | 1.4 | 0 | 1.061 | 188.027 | 0 | 0.752 | 0 | 3.361 | 0 | 0 | 0 | 0 | 0 | 0.534 | 0.363 | 1.792 | 0 | 0.38 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 09:08
WO-22270-64-A-1 | 0 | 65.189 | 558.335 | 0 | 0.199 | 819.809 | 5.129 | 2.491 | 2.725 | 1.01 | 0 | 0.884 | 187.843 | 0 | 0.94 | 0 | 3.534 | 0 | 0 | 0 | 0 | 0 | 3.413 | 0.591 | 0 | 0 | 0.376 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 09:44
WO-22270-65-A-1 | 0 | 32.06 | 1042.275 | 483.673 | 0.186 | 291.884 | 4.407 | 0 | 0.788 | 0.562 | 0 | 1.601 | 187.928 | 0 | 0.924 | 0 | 3.516 | 0 | 0 | 0 | 1.145 | 0 | 0.555 | 0.195 | 0.741 | 0 | 0.335 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 10:20
WO-22270-66-A-1 | 0 | 43.678 | 534.79 | 16.302 | 0.193 | 0 | 4.632 | 3.031 | 3.244 | 0.777 | 0 | 2.321 | 187.655 | 0 | 0.814 | 0 | 3.2 | 0 | 0 | 0 | 0 | 0 | 0.722 | 0.011 | 1.099 | 0 | 0.308 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-19 10:56
WO-22270-67-A-1 | 0 | 47.877 | 490.498 | -232.519 | 0.188 | 350.882 | 4.846 | 2.759 | 2.024 | 0.644 | 0 | 2.084 | 187.85 | 0 | 0.909 | 0 | 4.157 | 0 | 0 | 0 | 0 | 0 | 20.342 | 0.369 | 2.636 | 0 | 0.363 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 11:32
WO-22270-68-A-1 | 0 | 41.071 | 1114.668 | 0 | 0.182 | 0 | 4.7 | 2.365 | 3.843 | 0.874 | 0 | 2.14 | 187.807 | 0 | 1.157 | 0 | 3.686 | 0 | 0 | 0 | 0 | 0 | 0.457 | 0.253 | 1.27 | 0 | 0.383 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-19 12:08
WO-22270-69-A-1 | 0 | 90.629 | 384.049 | 1850.524 | 0.192 | 171.812 | 5.382 | 0 | 3.852 | 0.802 | 0 | 2.992 | 187.693 | 0 | 1.174 | 0 | 4.846 | 0 | 0 | 0 | 0 | 0 | 0.518 | 0.655 | 0.793 | 0 | 0.412 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 12:44
WO-22270-70-A-1 | 0 | 73.601 | 962.263 | -384.289 | 0.175 | 0 | 4.381 | 1.632 | 1.922 | 0.735 | 0 | 1.318 | 187.983 | 0 | 0.962 | 0 | 4.055 | 0 | 0 | 1.749 | 0 | 0 | 1.354 | 0.155 | 3.853 | 0 | 0.366 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-19 13:20
AB-220715-01 LRB 8_INJ001 | 0 | 9.116 | -264.025 | -300.723 | 0.17 | -0.296 | 4.321 | -0.188 | 0.733 | 0.324 | 2.078 | 0.885 | 187.552 | 0 | 0.907 | 0 | 0.33 | 0 | 0 | 0 | 0.985 | 0 | 0.692 | 0.152 | 1.451 | 0 | 0.316 | -1.484 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-19 13:56
AB-220715-01 LRB 8 | 0 | 9.295 | -210.421 | 29.482 | 0.201 | 25.296 | 4.508 | 0.346 | 0.678 | 0.359 | 2.103 | 1.442 | 187.603 | 0 | 0.877 | 0 | 0.371 | 0 | 0 | 0 | 0.981 | 0 | 0.846 | 0.207 | 1.359 | 0 | 0.335 | -0.242 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-19 14:32
WO-22270-71-A-1 | 0 | 35.578 | 140.711 | -272.815 | 0.184 | 738.168 | 5.023 | 2.742 | 1.464 | 0.85 | 0 | 0.896 | 187.673 | 0 | 0.712 | 0 | 8.605 | 0 | 0 | 0 | 0 | 0 | 5.465 | 0.188 | 2.626 | 0 | 0.322 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 15:08
WO-22270-72-A-1 | 0 | 44.232 | 190.339 | 176.56 | 0.177 | 636.91 | 4.338 | 1.019 | 1.316 | 0.652 | 0 | 0.981 | 187.778 | 0 | 0.836 | 0 | 3.996 | 0 | 0 | 0 | 0 | 0 | 0.393 | 0.147 | 1.363 | 0 | 0.305 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 15:44
WO-22270-73-A-1 | 0 | 48.91 | -64.362 | -459.024 | 0.17 | 269.895 | 0 | 3.521 | 1.954 | 0.816 | 0 | 0.921 | 187.732 | 0 | 1.086 | 0 | 3.362 | 0 | 0 | 0 | 0 | 0 | 0.46 | 0.18 | 0.576 | 0 | 0.312 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 16:20
WO-22270-74-A-1 | 0 | 52.99 | 856.849 | -72.531 | 0.173 | 419.554 | 4.292 | 2.192 | 1.892 | 0.727 | 0 | 2.049 | 187.763 | 0 | 0.88 | 0 | 6.787 | 0 | 0 | 1.274 | 0 | 0 | 2.505 | 0.23 | 2.129 | 0 | 0.332 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 16:56
WO-22270-75-A-1 | 0 | 39.104 | 201.744 | -251.403 | 0.179 | 210.064 | 4.407 | 2.943 | 1.166 | 0.885 | 0 | 1.729 | 187.918 | 0 | 0.966 | 0 | 3.346 | 0 | 0 | 2.022 | 0 | 0 | 1.793 | 0.281 | 4.871 | 0 | 0.336 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 17:32
WO-22270-76-A-1 | 0 | 103.752 | 497.725 | 0 | 0.179 | 307.607 | 4.292 | 3.26 | 2.297 | 1.211 | 0 | 3.292 | 188.311 | 0 | 1.122 | 0 | 4.155 | 0 | 0 | 0 | 0 | 0 | 0.522 | 0.298 | 1.654 | 0 | 0.319 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 18:08
WO-22270-77-A-1 | 0 | 35.759 | 1.786 | -106.899 | 0.175 | 248.929 | 4.574 | 0.62 | 1.209 | 1.108 | 0 | 2.161 | 187.938 | 0 | 0.582 | 0 | 19.796 | 0 | 0 | 1.767 | 1.755 | 0 | 0.445 | 0.251 | 0.829 | 0 | 0.318 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 18:44
WO-22270-78-A-1 | 0 | 35.054 | 507.47 | 509.707 | 0.169 | 0 | 0 | 2.357 | 0.789 | 0.783 | 0 | 1.883 | 188.028 | 0 | 0.815 | 0 | 3.982 | 0 | 0 | 1.746 | 1.926 | 0 | 0.575 | 0.158 | 0.94 | 0 | 0.304 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-19 19:20
WO-22270-79-A-1 | 0 | 51.327 | 469.012 | -236.571 | 0.172 | 123.785 | 0 | 1.421 | 1.09 | 0.796 | 0 | 1.036 | 187.933 | 0 | 0.863 | 0 | 3.306 | 0 | 0 | 0.995 | 1.998 | 0 | 0.403 | 0.072 | 1.095 | 0 | 0.318 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 19:56
WO-22270-80-A-1 | 0 | 44.766 | 229.273 | 357.704 | 0.164 | 0 | 4.443 | 2.245 | 1.521 | 0.677 | 0 | 0.867 | 187.781 | 0 | 0.85 | 0 | 3.667 | 0 | 0 | 1.982 | 2.097 | 0 | 0.983 | 0.204 | 1.857 | 0 | 0.296 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-19 20:32
AB-220715-01 LFB-H 5 | 0 | 188.365 | 202840.503 | 1234515.59 | 198.619 | 3055789.999 | 2799.149 | 211.05 | 174.405 | 168.726 | 155.005 | 164.39 | 417.344 | 0 | 204.941 | 190.537 | 199.853 | 0 | 0 | 197.117 | 205.523 | 0 | 206.05 | 176.772 | 201.716 | 0 | 203.13 | 232.48 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 21:08
AB-220715-01 LRB 9_INJ001 | 0 | 8.502 | 60.389 | 1267.874 | 0.399 | 641.222 | 7.26 | 0.111 | 0.709 | 0.406 | 2.124 | 1.658 | 187.684 | 0 | 1.081 | 0 | 0.497 | 0 | 0 | 0 | 1.066 | 0 | 0.986 | 0.186 | 1.724 | 0 | 0.698 | -1.639 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-19 21:44
AB-220715-01 LRB 9 | 0 | 7.767 | 156.527 | 24.489 | 0.216 | 309.813 | 4.813 | 0.494 | 0.742 | 0.341 | 2.23 | 0.933 | 187.609 | 0 | 1.129 | 0 | 0.364 | 0 | 0 | 0.52 | 1.009 | 0 | 1.115 | 0.228 | 1.482 | 0 | 0.364 | -1.159 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-19 22:20
AB-220715-01 LD 4 | 0 | 40.77 | 286.021 | -1.409 | 0.186 | 2852.671 | 6.915 | 3.027 | 2.589 | 0.733 | 0 | 2.064 | 187.912 | 0 | 0.839 | 0 | 3.69 | 0 | 0 | 2.485 | 2.379 | 0 | 1.152 | 0.082 | 3.311 | 0 | 0.333 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 22:56
AB-220715-01 CCC-H 5 | 0 | 356.422 | 388334.722 | 2343649.659 | 390.247 | 4984878.821 | 4802.419 | 389.461 | 338.354 | 350.896 | 315.397 | 332.343 | 604.509 | 0 | 395.002 | 356.628 | 386.242 | 0 | 0 | 376.721 | 384.448 | 0 | 395.37 | 338.568 | 376.568 | 0 | 403.856 | 386.002 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-19 23:32
AB-220715-01 LFSM-H 4 | 0 | 189.439 | 202997.037 | 1126695.618 | 201.674 | 3303844.392 | 3928.87 | 190.142 | 170.383 | 232.842 | 149.526 | 151.141 | 362.469 | 0 | 206.914 | 0 | 201.625 | 0 | 0 | 194.19 | 193.605 | 0 | 204.594 | 202.219 | 196.17 | 0 | 203.272 | 148.778 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 00:08
WO-22270-81-A-1 | 0 | 52.993 | 642.221 | 78.951 | 0.265 | 555.646 | 8.044 | 3.201 | 1.934 | 0.985 | 0 | 3.031 | 188.613 | 0 | 1.113 | 0 | 3.721 | 0 | 0 | 1.954 | 3.35 | 0 | 0.552 | 0.218 | 1.584 | 0 | 0.404 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 00:44
WO-22270-82-A-1 | 0 | 41.309 | 1672.951 | 1063.712 | 0.205 | 0 | 6.641 | 2.934 | 1.181 | 0.923 | 0 | 0.86 | 187.815 | 0 | 0.714 | 0 | 3.553 | 0 | 0 | 3.331 | 3.299 | 0 | 0.474 | 0.263 | 2.325 | 0 | 0.357 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 01:20
WO-22270-83-A-1 | 0 | 48.205 | 1269.882 | 239.84 | 0.195 | 0 | 5.685 | 2.189 | 1.222 | 0.894 | 0 | 1.866 | 187.94 | 0 | 0.769 | 0 | 3.516 | 0 | 0 | 0 | 3.92 | 0 | 0.568 | 0.344 | 0.903 | 0 | 0.337 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 01:56
WO-22270-84-A-1 | 0 | 37.294 | 112.183 | 481.908 | 0.191 | 0 | 4.693 | 0 | 2.623 | 0.787 | 0 | 0.926 | 187.662 | 0 | 0.915 | 0 | 2.515 | 0 | 0 | 0 | 10.43 | 0 | 0.444 | 0.271 | 1.798 | 0 | 0.378 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 02:32
WO-22270-85-A-1 | 0 | 37.171 | 90.487 | 244.487 | 0.185 | 33.24 | 5.408 | 1.331 | 1.233 | 1.466 | 0 | 2.173 | 187.764 | 0 | 0.852 | 0 | 3.296 | 0 | 0 | 2.354 | 8.256 | 0 | 0.342 | 0.262 | 1.847 | 0 | 0.349 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 03:08
WO-22270-86-A-1 | 0 | 41.693 | 635.704 | 53.279 | 0.184 | 0 | 4.755 | 1.607 | 1.149 | 1 | 0 | 0.937 | 187.897 | 0 | 0.729 | 0 | 3.299 | 0 | 0 | 3.729 | 8.317 | 0 | 0.451 | 0.303 | 3.486 | 0 | 0.299 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 03:44
WO-22270-87-A-1 | 0 | 39.548 | -111.161 | -463.869 | 0.177 | 239.196 | 5.035 | 0 | 0 | 0.846 | 0 | 1.704 | 187.954 | 0 | 0.647 | 0 | 3.777 | 0 | 0 | 2.319 | 29.138 | 0 | 0.522 | 0.254 | 0.651 | 0 | 0.351 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 04:20
WO-22270-88-A-1 | 0 | 55.023 | 920.01 | -135.586 | 0.174 | 83.849 | 4.619 | 2.168 | 1.205 | 0.764 | 0 | 2.057 | 188.26 | 0 | 0.884 | 0 | 4.311 | 0 | 0 | 0 | 29.571 | 0 | 0.458 | 0.22 | 0.822 | 0 | 0.362 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 04:56
WO-22270-89-A-1 | 0 | 68.22 | 1232.885 | 19.6 | 0.169 | 402.393 | 4.281 | 1.877 | 1.45 | 1.078 | 0 | 2.282 | 187.746 | 0 | 0.896 | 0 | 3.765 | 0 | 0 | 3.42 | 34.603 | 0 | 0.764 | 0.105 | 1.466 | 0 | 0.329 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 05:32
WO-22270-90-A-1 | 0 | 46.72 | 769.791 | 277.473 | 0.172 | 0 | 4.656 | 1.717 | 1.248 | 0.748 | 0 | 0.996 | 187.886 | 0 | 0.978 | 0 | 3.436 | 0 | 0 | 2.13 | 34.536 | 0 | 0 | 0.221 | 0 | 0 | 0.354 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 06:08
AB-220715-01 LRB 10_INJ001 | 0 | 10.169 | 78.436 | 467.545 | 0.175 | 0 | 4.356 | 0.364 | 0.855 | 0.33 | 0 | 0.955 | 187.712 | 0 | 0.945 | 0 | 0.39 | 0 | 0 | 0 | 0 | 0 | 1.144 | 0.173 | 1.382 | 0 | 0.316 | -0.968 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 06:44
AB-220715-01 LRB 10 | 0 | 17.776 | 203.248 | -481.869 | 0.177 | 91.877 | 4.334 | 0.581 | 0.656 | 0.382 | 0 | 1.082 | 187.804 | 0 | 1.182 | 0 | 0.442 | 0 | 0 | 0 | 1.06 | 0 | 1.02 | 0.278 | 1.482 | 0 | 0.305 | 1.344 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 07:20
WO-22270-91-A-1 | 0 | 30.101 | -151.174 | 41.871 | 0.178 | 0 | 5.171 | 0 | 0.964 | 0.79 | 0 | 1.001 | 187.829 | 0 | 0.865 | 0 | 2.934 | 0 | 0 | 1.424 | 89.122 | 0 | 0.341 | 0.183 | 2.673 | 0 | 0.305 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 07:56
WO-22270-92-A-1 | 0 | 58.989 | 413.207 | -252.14 | 0.175 | 226.358 | 4.403 | 1.92 | 1.274 | 0.775 | 0 | 1.027 | 187.98 | 0 | 0.716 | 0 | 3.82 | 0 | 0 | 3.867 | 118.451 | 0 | 0.488 | 0.184 | 1.678 | 0 | 0.314 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 08:32
WO-22270-93-A-1 | 0 | 35.932 | 363.93 | 300.949 | 0.174 | 2129.979 | 4.325 | 1.009 | 2.193 | 0.707 | 0 | 1.807 | 187.78 | 0 | 0.83 | 0 | 4.301 | 0 | 0 | 3.048 | 95.675 | 0 | 0.505 | 0.036 | 2.634 | 0 | 0.33 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 09:08
WO-22270-94-A-1 | 0 | 39.687 | 2083.552 | -172.736 | 0.171 | 0 | 4.406 | 1.973 | 1.385 | 0.72 | 0 | 0.892 | 187.733 | 0 | 0.644 | 0 | 3.73 | 0 | 0 | 3.31 | 513.964 | 0 | 0.553 | 0.014 | 4.213 | 0 | 0.324 | -1.63 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 09:44
WO-22270-95-A-1 | 0 | 48.156 | 132.114 | -385.437 | 0.171 | 56.814 | 4.569 | 1.934 | 1.772 | 1.084 | 0 | 0.933 | 187.851 | 0 | 0.954 | 0 | 3.59 | 0 | 0 | 1.419 | 296.758 | 0 | 0.443 | 0.018 | 1.156 | 0 | 0.298 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 10:20
WO-22270-96-A-1 | 0 | 30.667 | 115.319 | 355.413 | 0.172 | 69.572 | 4.278 | 0 | 0.84 | 0.912 | 0 | 1.081 | 187.768 | 0 | 0.941 | 0 | 3.898 | 0 | 0 | 1.253 | 429.411 | 0 | 0.395 | 0.337 | 2.135 | 0 | 0.34 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 10:56
WO-22270-97-A-1 | 0 | 34.976 | 1095.1 | -298.073 | 0.172 | 0 | 4.85 | 1.411 | 0.982 | 0.805 | 0 | 0.924 | 187.836 | 0 | 0.823 | 0 | 3.657 | 0 | 0 | 0 | 2.168 | 0 | 0.561 | 0.151 | 0.986 | 0 | 0.306 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 11:32
WO-22270-98-A-1 | 0 | 35.134 | 1017.672 | 656.634 | 0.172 | 0 | 4.272 | 3.012 | 0.913 | 0.602 | 0 | 0.86 | 0 | 0 | 0.931 | 0 | 4.501 | 0 | 0 | 0 | 1.918 | 0 | 0.427 | 0.235 | 0.611 | 0 | 0.313 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 12:08
WO-22270-99-A-1 | 0 | 74.339 | 2248.231 | 1226.484 | 0.171 | 0 | 6.182 | 2.588 | 1.609 | 0.856 | 0 | 2.825 | 187.718 | 0 | 0.984 | 0 | 4.353 | 0 | 0 | 0 | 0 | 0 | 1.034 | 0.217 | 0 | 0 | 0.304 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 12:44
WO-22270-100-A-1 | 0 | 64.394 | -56.538 | -403.909 | 0.17 | 0 | 0 | 0 | 2.049 | 0.622 | 0 | 0.868 | 188.207 | 0 | 0.921 | 0 | 4.313 | 0 | 0 | 1.066 | 1.921 | 0 | 0.447 | 0.224 | 1.687 | 0 | 0.3 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 13:20
AB-220715-01 LFB-H 6 | 0 | 181.741 | 196443.003 | 1149571.427 | 191.557 | 3292088.252 | 2844.398 | 182.753 | 166.991 | 181.226 | 148.72 | 161.946 | 384.475 | 0 | 187.724 | 156.313 | 188.731 | 0 | 0 | 188.302 | 191.319 | 0 | 188.34 | 167.236 | 191.836 | 0 | 201.151 | 226.894 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 13:56
AB-220715-01 LRB 11_INJ001 | 0 | 8.472 | 317.579 | 207.284 | 0.355 | 461.488 | 6.762 | 0.603 | 0.684 | 0.41 | 0 | 2.04 | 187.903 | 0 | 1.102 | 0 | 0.503 | 0 | 0 | 0 | 0 | 0 | 0.798 | 0.141 | 1.569 | 0 | 0.62 | -2.486 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 14:32
AB-220715-01 LRB 11 | 0 | 10.077 | 165.446 | 342.618 | 0.208 | 0 | 4.591 | 0.293 | 0.596 | 0.385 | 0 | 1.677 | 187.683 | 0 | 1.098 | 0 | 0.383 | 0 | 0 | 0 | 0.998 | 0 | 0.811 | 0.045 | 1.468 | 0 | 0.409 | -2.011 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 15:08
AB-220715-01 LD 5 | 0 | 97.484 | 1542.264 | -323.222 | 0.193 | 156.442 | 5.354 | 1.745 | 2.34 | 0.802 | 0 | 3.174 | 187.951 | 0 | 0.988 | 0 | 4.628 | 0 | 0 | 0 | 2.824 | 0 | 1.161 | 0.37 | 3.298 | 0 | 0.336 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 15:44
AB-220715-01 CCC-H 6 | 0 | 350.486 | 385600.225 | 2262197.668 | 367.431 | 5424472.255 | 4953.977 | 375.438 | 350.691 | 359.107 | 302.006 | 313.329 | 595.579 | 0 | 403.343 | 377.628 | 378.911 | 0 | 0 | 378.394 | 368.549 | 0 | 377.013 | 332.589 | 372.428 | 0 | 384.168 | 421.07 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 16:20
AB-220715-01 LFSM-H 5 | 0 | 201.428 | 200130.033 | 1142518.349 | 208.381 | 3197608.075 | 4363.9 | 184.913 | 171.217 | 204.164 | 144.456 | 149.42 | 365.498 | 0 | 205.917 | 0 | 198.493 | 0 | 0 | 192.215 | 184.057 | 0 | 197.638 | 196.268 | 196.1 | 0 | 202.464 | 153.964 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 16:56
WO-22270-101-A-1 | 0 | 33.731 | 1059.816 | -503.82 | 0.239 | 1008.589 | 8.518 | 3.681 | 0.982 | 0.53 | 0 | 1.145 | 187.717 | 0 | 0.972 | 0 | 4.04 | 0 | 0 | 0.779 | 2.944 | 0 | 0.434 | 0.381 | 0.731 | 0 | 0.407 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 17:32
WO-22270-102-A-1 | 0 | 24.256 | 1326.149 | -413.123 | 0.205 | 680.382 | 4.615 | 1.641 | 1.622 | 0.456 | 0 | 2.04 | 187.754 | 0 | 1.146 | 0 | 4.044 | 0 | 0 | 0 | 2.814 | 0 | 0.503 | 0.372 | 1.723 | 0 | 0.369 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 18:08
WO-22270-103-A-1 | 0 | 60.937 | -39.804 | -518.56 | 0.192 | 124.878 | 5.063 | 1.6 | 1.09 | 0.578 | 0 | 1.683 | 187.861 | 0 | 0.778 | 0 | 4.974 | 0 | 0 | 0 | 3.613 | 0 | 0.843 | 0.211 | 1.333 | 0 | 0.315 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 18:44
WO-22270-104-A-1 | 0 | 42.933 | 1198.628 | -448.666 | 0.207 | 317.913 | 5.368 | 1.636 | 1.437 | 0.719 | 0 | 1.001 | 187.697 | 0 | 0.939 | 0 | 3.97 | 0 | 0 | 0 | 8.259 | 0 | 0.365 | 0.216 | 0.625 | 0 | 0.375 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 19:20
WO-22270-105-A-1 | 0 | 93.769 | 19.379 | -319.987 | 0.184 | 110.89 | 4.965 | 1.889 | 1.197 | 0.688 | 0 | 1.069 | 187.671 | 0 | 0.593 | 0 | 4.763 | 0 | 0 | 0 | 7.025 | 0 | 0.782 | 0.222 | 0.608 | 0 | 0.361 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 19:56
WO-22270-106-A-1 | 0 | 67.49 | 3127.405 | 744.399 | 0.212 | 371.567 | 5.199 | 0 | 0.963 | 0.71 | 0 | 2.235 | 187.958 | 0 | 0.65 | 0 | 4.019 | 0 | 0 | 1.448 | 7.749 | 0 | 0.619 | 0.388 | 0.523 | 0 | 0.365 | 4.342 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-20 20:32
WO-22270-107-A-1 | 0 | 33.765 | 169.512 | -418.587 | 0.174 | 25.086 | 4.403 | 2.217 | 1.928 | 0.596 | 0 | 1.906 | 187.928 | 0 | 0.678 | 0 | 3.585 | 0 | 0 | 1.513 | 26.028 | 0 | 0.499 | 0.253 | 0.535 | 0 | 0.352 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 21:08
WO-22270-108-A-1 | 0 | 52.107 | 626.182 | 591.075 | 0.17 | 59.577 | 4.5 | 3.237 | 1.173 | 0.584 | 0 | 1.045 | 0 | 0 | 0.59 | 0 | 3.661 | 0 | 0 | 1.854 | 24.938 | 0 | 0.489 | 0.095 | 0.557 | 0 | 0.357 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 21:44
WO-22270-109-A-1 | 0 | 42.539 | 349.224 | -344.628 | 0.175 | 307.754 | 4.903 | 1.92 | 1.642 | 0.578 | 0 | 0.961 | 187.735 | 0 | 0.69 | 0 | 4.195 | 0 | 0 | 0 | 36.818 | 0 | 1.347 | 0.271 | 0.617 | 0 | 0.322 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 22:20
WO-22270-110-A-1 | 0 | 46.505 | 646.813 | -492.082 | 0.169 | 584.479 | 4.524 | 2.75 | 0.734 | 0.697 | 0 | 0.98 | 187.836 | 0 | 0.973 | 0 | 3.898 | 0 | 0 | 0 | 35.931 | 0 | 1.286 | 0.176 | 1.446 | 0 | 0.301 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 22:56
AB-220715-01 LRB 12_INJ001 | 0 | 16.116 | 1.775 | 453.877 | 0.175 | 116.664 | 4.28 | 0.578 | 0.705 | 0.443 | 0 | 1.732 | 187.671 | 0 | 1.094 | 0 | 0.332 | 0 | 0 | 0 | 0 | 0 | 0.335 | 0.258 | 0.51 | 0 | 0.346 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-20 23:32
AB-220715-01 LRB 12 | 0 | 16.842 | -140.753 | 171.064 | 0.18 | 1060.489 | 4.318 | 0.319 | 0.946 | 0.42 | 0 | 1.105 | 187.657 | 0 | 1.079 | 0 | 0.346 | 0 | 0 | 0 | 0 | 0 | 0.321 | 0.211 | 0.517 | 0 | 0.315 | -1.605 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 00:08
WO-22270-111-A-1 | 0 | 42.733 | 203.767 | -459.801 | 0.185 | 0 | 4.502 | 1.185 | 1.267 | 0.755 | 0 | 0.886 | 187.818 | 0 | 0.911 | 0 | 3.531 | 0 | 0 | 0 | 72.024 | 0 | 0.505 | 0.188 | 0.756 | 0 | 0.298 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 00:44
WO-22270-112-A-1 | 0 | 58.537 | 1266.19 | -248.505 | 0.168 | 0 | 4.417 | 2.563 | 1.418 | 0.598 | 0 | 1.181 | 187.96 | 0 | 0.682 | 0 | 3.652 | 0 | 0 | 0 | 72.186 | 0 | 0.855 | 0.243 | 0.638 | 0 | 0.315 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 01:20
WO-22270-113-A-1 | 0 | 36.305 | 730.178 | 0 | 0.201 | 0 | 4.282 | 1.551 | 1.485 | 0.621 | 0 | 2.63 | 187.671 | 0 | 0.875 | 0 | 3.505 | 0 | 0 | 0.633 | 94.178 | 0 | 0.372 | 0.313 | 0.999 | 0 | 0.305 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 01:56
WO-22270-114-A-1 | 0 | 57.432 | 808.673 | -229.463 | 0.186 | 0 | 4.303 | 0 | 1.335 | 0.655 | 0 | 2.402 | 188.007 | 0 | 0.841 | 0 | 4.205 | 0 | 0 | 1.115 | 326.568 | 0 | 0.57 | 0.257 | 1.482 | 0 | 0.313 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 02:32
WO-22270-115-A-1 | 0 | 43.75 | 1783.148 | -110.945 | 0.175 | 0 | 4.424 | 3.524 | 1.218 | 0.736 | 0 | 1.171 | 188.204 | 0 | 0.7 | 0 | 4.431 | 0 | 0 | 0 | 222.047 | 0 | 2.753 | 0.144 | 2.163 | 0 | 0.315 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 03:08
WO-22270-116-A-1 | 0 | 67.219 | 3334.089 | 68.538 | 0.185 | 335.165 | 4.367 | 0 | 2.36 | 0.911 | 0 | 0.987 | 187.883 | 0 | 0.592 | 0 | 3.651 | 0 | 0 | 0 | 239.793 | 0 | 0.482 | 0.256 | 1.595 | 0 | 0.38 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 03:44
WO-22270-117-A-1 | 0 | 31.489 | 893.006 | -483.667 | 0.178 | 0 | 0 | 2.873 | 1.803 | 0.843 | 0 | 0.964 | 187.986 | 0 | 0.989 | 0 | 27.663 | 0 | 0 | 0.705 | 2.136 | 0 | 0.507 | 0.194 | 1.126 | 0 | 0.371 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 04:20
WO-22270-118-A-1 | 0 | 47.377 | 2470.661 | -403.327 | 0.172 | 455.012 | 4.465 | 0 | 1.197 | 0.874 | 0 | 0.913 | 187.763 | 0 | 0.597 | 0 | 2.864 | 0 | 0 | 0 | 2.136 | 0 | 0.423 | 0.298 | 0.666 | 0 | 0.322 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 04:56
WO-22270-119-A-1 | 0 | 71.148 | 581.211 | 542.148 | 0.213 | 139.33 | 4.473 | 0 | 1.762 | 1.048 | 0 | 2.894 | 187.708 | 0 | 0.845 | 0 | 4.091 | 0 | 0 | 0 | 1.872 | 0 | 0.726 | 0.204 | 1.779 | 0 | 0.362 | 11.111 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 05:32
WO-22270-120-A-1 | 0 | 83.8 | 321.139 | -327.12 | 0.167 | 455.36 | 0 | 0 | 1.069 | 0.711 | 0 | 1.033 | 187.728 | 0 | 0.884 | 0 | 4.189 | 0 | 0 | 0 | 1.811 | 0 | 0.958 | 0.322 | 1.868 | 0 | 0.311 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 06:08
AB-220715-01 LFB-H 7 | 0 | 204.781 | 198140.72 | 1057135.503 | 165.166 | 2447239.965 | 2508.736 | 189.674 | 171.667 | 189.285 | 156.023 | 158.717 | 383.497 | 0 | 206.892 | 172.757 | 190.922 | 0 | 0 | 191.718 | 192.398 | 0 | 195.667 | 152.165 | 192.64 | 0 | 200.014 | 191.329 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 06:44
AB-220715-01 LRB 13_INJ001 | 0 | 12.2 | -234.346 | 273.809 | 0.345 | 1370.912 | 6.716 | 0.574 | 0.708 | 0.572 | 0 | 1.31 | 187.698 | 0 | 1.272 | 0 | 0.522 | 0 | 0 | 0 | 1.13 | 0 | 0.447 | 0.413 | 0.579 | 0 | 0.665 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 07:20
AB-220715-01 LRB 13 | 0 | 17.745 | 132.549 | 0 | 0.218 | 1747.628 | 5.293 | 0.337 | 0.858 | 0.329 | 0 | 1.057 | 187.644 | 0 | 1.183 | 0 | 0.374 | 0 | 0 | 0 | 0.956 | 0 | 0.354 | 0.267 | 0.599 | 0 | 0.465 | -2.027 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 07:56
AB-220715-01 LD 6 | 0 | 54.984 | 2628.138 | 1682.051 | 0.218 | 278.044 | 5.22 | 0 | 1.889 | 0.628 | 0 | 0.93 | 187.859 | 0 | 0.785 | 0 | 3.868 | 0 | 0 | 2.465 | 1.984 | 0 | 0.814 | 0.468 | 1.785 | 0 | 0.364 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 08:32
AB-220715-01 CCC-H 7 | 0 | 359.39 | 389323.119 | 2156754.132 | 349.245 | 4615790.133 | 4460.001 | 385.941 | 335.708 | 364.811 | 303.514 | 317.825 | 588.131 | 0 | 405.387 | 369.407 | 372.392 | 0 | 0 | 377.006 | 375.749 | 0 | 384.409 | 298.732 | 387.221 | 0 | 383.41 | 384.022 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 09:08
AB-220715-01 LFSM-H 6 | 0 | 186.26 | 204489.742 | 1159280.825 | 194.374 | 3757884.147 | 3600.105 | 202.898 | 178.901 | 221.409 | 145.053 | 146.784 | 370.14 | 0 | 202.432 | 0 | 201.023 | 0 | 0 | 197.115 | 189.642 | 0 | 195.386 | 186.132 | 200.076 | 0 | 200.886 | 191.886 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 09:44
WO-22270-121-A-1 | 0 | 46.695 | 388.276 | -219.126 | 0.228 | 307.211 | 6.823 | 1.054 | 1.007 | 0.852 | 0 | 1.38 | 187.79 | 0 | 0.887 | 0 | 3.691 | 0 | 0 | 0 | 2.834 | 0 | 0.406 | 0.44 | 0 | 0 | 0.426 | -1.099 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 10:20
WO-22270-122-A-1 | 0 | 88.084 | 2073.15 | -459.482 | 0.208 | 174.02 | 9.192 | 0 | 2.476 | 0.818 | 0 | 0.971 | 187.813 | 0 | 1.098 | 0 | 3.841 | 0 | 0 | 0 | 4.153 | 0 | 0.714 | 0.706 | 1.193 | 0 | 0.364 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 10:56
WO-22270-123-A-1 | 0 | 46.07 | 3104.927 | 31.206 | 0.192 | 0 | 5.338 | 0 | 0.914 | 1.248 | 0 | 2.967 | 187.818 | 0 | 0.632 | 0 | 4.068 | 0 | 0 | 0 | 3.574 | 0 | 0.454 | 0.383 | 0.544 | 0 | 0.411 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 11:32
WO-22270-124-A-1 | 0 | 58.248 | 2032.92 | -216.765 | 0.184 | 0 | 5.106 | 0 | 1.17 | 0.612 | 0 | 1.115 | 187.719 | 0 | 0.612 | 0 | 4.517 | 0 | 0 | 0 | 8.903 | 0 | 0.583 | 0.417 | 3.316 | 0 | 0.33 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 12:08
WO-22270-125-A-1 | 0 | 61.89 | 4734.78 | -74.305 | 0.233 | 88.921 | 4.529 | 0 | 2.165 | 0.891 | 0 | 0.877 | 188.119 | 0 | 0.564 | 0 | 4.939 | 0 | 0 | 0 | 8.16 | 0 | 0.925 | 0.512 | 1.426 | 0 | 0.396 | 0.078 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 12:44
WO-22270-126-A-1 | 0 | 44.648 | 3852.046 | -400.763 | 0.191 | 0 | 4.402 | 0 | 2.014 | 1.035 | 0 | 0.952 | 188.795 | 0 | 0.597 | 0 | 3.937 | 0 | 0 | 0 | 8.852 | 0 | 0.878 | 0.378 | 2.024 | 0 | 0.35 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 13:20
WO-22270-127-A-1 | 0 | 30.429 | 1416.199 | 1192.418 | 0.176 | 0 | 5.269 | 0 | 1.381 | 0.63 | 0 | 0.903 | 187.978 | 0 | 1.067 | 0 | 3.877 | 0 | 0 | 0 | 28.051 | 0 | 0.484 | 0.49 | 2.717 | 0 | 0.322 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 13:56
WO-22270-128-A-1 | 0 | 41.731 | 1489.561 | -254.157 | 0.182 | 0 | 4.334 | 0 | 2.627 | 0.862 | 0 | 1.558 | 188.016 | 0 | 0.83 | 0 | 3.491 | 0 | 0 | 0 | 33.692 | 0 | 0.52 | 0.306 | 1.561 | 0 | 0.336 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 14:32
WO-22270-129-A-1 | 0 | 49.417 | 2775.292 | -400.293 | 0.172 | 0 | 4.462 | 1.124 | 0.913 | 0.466 | 0 | 0.979 | 188.075 | 0 | 0.817 | 0 | 4.154 | 0 | 0 | 3.087 | 38.125 | 0 | 0.837 | 0.025 | 2.706 | 0 | 0.32 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 15:08
WO-22270-130-A-1 | 0 | 141.321 | 3730.449 | 201.143 | 0 | 0 | 0 | 0 | 1.542 | 0.742 | 0 | 0.829 | 187.988 | 0 | 0.677 | 0 | 5.842 | 0 | 0 | 2.975 | 39.499 | 0 | 1.533 | 0.3 | 21.233 | 0 | 0.333 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 15:44
AB-220715-01 LRB 14_INJ001 | 0 | 14.765 | 1040.067 | -183.837 | 0.168 | 42.386 | 4.34 | 0.484 | 0.835 | 0.445 | 0 | 1.409 | 187.638 | 0 | 1.058 | 0 | 0.356 | 0 | 0 | 0.486 | 1.109 | 0 | 0.321 | 0.184 | 0.587 | 0 | 0.324 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 16:20
AB-220715-01 LRB 14 | 0 | 19.671 | 1331.876 | -156.099 | 0.197 | 0 | 4.342 | 0.65 | 0.797 | 0.511 | 0 | 1.439 | 187.706 | 0 | 1.258 | 0 | 0.373 | 0 | 0 | 0 | 0 | 0 | 0.339 | 0.231 | 0.56 | 0 | 0.337 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 16:56
WO-22270-131-A-1 | 0 | 319.097 | 1734.18 | 204.249 | 0.172 | 0 | 4.494 | 0 | 1.28 | 0.781 | 0 | 0.743 | 188.045 | 0 | 0.864 | 0 | 3.517 | 0 | 0 | 1.2 | 142.268 | 0 | 0.796 | 0.298 | 1.455 | 0 | 0.336 | -1.639 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 17:32
WO-22270-132-A-1 | 0 | 60.479 | 2786.299 | 534.822 | 0.17 | 0 | 0 | 0 | 1.683 | 0.556 | 0 | 0 | 187.859 | 0 | 0.573 | 0 | 3.637 | 0 | 0 | 0 | 112.938 | 0 | 0.725 | 0.036 | 3.691 | 0 | 0.294 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 18:08
WO-22270-133-A-1 | 0 | 58.792 | 2230.904 | 571.219 | 0.169 | 0 | 4.358 | 0 | 1.435 | 1.299 | 0 | 3.056 | 188.927 | 0 | 0.773 | 0 | 4.59 | 0 | 0 | 1.669 | 83.818 | 0 | 2.182 | 0.032 | 2.816 | 0 | 0.311 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 18:44
WO-22270-134-A-1 | 0 | 137.963 | 3885.937 | 1275.706 | 0.166 | 433.112 | 5.014 | 0 | 2.327 | 1.098 | 0 | 1.377 | 188.853 | 0 | 0.54 | 0 | 4.942 | 0 | 0 | 5.25 | 2231.973 | 0 | 3.39 | 0.259 | 2.168 | 0 | 0.36 | 17.895 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 19:20
WO-22270-135-A-1 | 0 | 93.675 | 1434.034 | 81.724 | 0.176 | 0 | 5.045 | 0 | 1.205 | 0.905 | 0 | 2.956 | 188.61 | 0 | 0.614 | 0 | 3.738 | 0 | 0 | 1.671 | 222.327 | 0 | 0.704 | 0.397 | 0.598 | 0 | 0.314 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 19:56
WO-22270-136-A-1 | 0 | 1380.136 | 7192.438 | 232.137 | 0.189 | 0 | 4.61 | 3.41 | 1.742 | 0.611 | 0 | 0.913 | 188.042 | 0 | 0.468 | 0 | 4.151 | 0 | 0 | 3.756 | 51.868 | 0 | 0.774 | 0.054 | 1.873 | 0 | 0.301 | 27.799 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 20:32
WO-22270-137-A-1 | 0 | 244.19 | 551.055 | -215.286 | 0.186 | 0 | 0 | 11.283 | 1.387 | 0.948 | 0 | 0.983 | 187.795 | 0 | 0.449 | 0 | 2.975 | 0 | 0 | 0 | 1.438 | 0 | 0.447 | 0.015 | 1.092 | 0 | 0.328 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-21 21:08
WO-22270-138-A-1 | 0 | 722.708 | 1782.187 | 92.015 | 0.168 | 0 | 0 | 16.461 | 1.31 | 0.972 | 0 | 0.896 | 188.95 | 0 | 0.185 | 0 | 2.294 | 0 | 0 | 2.151 | 0 | 0 | 0.365 | 0.139 | 0.526 | 0 | 0.328 | 23.02 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 21:44
WO-22270-139-A-1 | 0 | 1405.259 | 1938.607 | -290.602 | 0.168 | 0 | 4.337 | 15.189 | 1.429 | 1.625 | 0 | 0.876 | 188.102 | 0 | 0.808 | 0 | 4.91 | 0 | 0 | 0 | 0 | 0 | 0.538 | 0.008 | 0.633 | 0 | 0.296 | 21.595 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 22:20
WO-22270-140-A-1 | 0 | 849.187 | 3952.986 | 20.313 | 0.168 | 505.27 | 0 | 10.723 | 1.545 | 0.744 | 0 | 0.971 | 188.247 | 0 | 0.827 | 0 | 3.954 | 0 | 0 | 0 | 0 | 0 | 0.56 | 0.063 | 0.612 | 0 | 0.276 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 22:56
AB-220715-01 LFB-H 8 | 0 | 180.431 | 192046.504 | 1076773.877 | 183.957 | 2047300.17 | 2290.255 | 182.476 | 169.13 | 199.841 | 153.04 | 153.851 | 386.009 | 0 | 200.563 | 184.848 | 183.032 | 0 | 0 | 189.842 | 190.233 | 0 | 207.329 | 159.763 | 200.802 | 0 | 200.042 | 167.174 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-21 23:32
AB-220715-01 LRB 15_INJ001 | 0 | 88.971 | 381.149 | 5342.364 | 0.959 | 8907.323 | 12.588 | 0.909 | 1.064 | 0.907 | 1.4 | 1.582 | 188.028 | 0 | 1.146 | 0 | 1.466 | 0 | 0 | 1.447 | 1.636 | 0 | 1.026 | 0.767 | 1.407 | 0 | 1.591 | 0 | 0 | 0 | 0 | 0 | C | Unknown | 2022-07-22 00:08
AB-220715-01 LRB 15 | 0 | 99.359 | 681.894 | 467.245 | 0.25 | 0 | 4.737 | 0.951 | 0.813 | 0.365 | 0 | 1.271 | 187.941 | 0 | 0.92 | 0 | 0.494 | 0 | 0 | 0 | 1.008 | 0 | 0.374 | 0.07 | 0.881 | 0 | 0.592 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 00:44
AB-220715-01 LD 7 | 0 | 1153.376 | 616.662 | 18.357 | 0.204 | 146.822 | 5.082 | 7.487 | 1.639 | 0.603 | 0 | 1.203 | 188.821 | 0 | 0.704 | 0 | 4.949 | 0 | 0 | 0 | 0 | 0 | 0.653 | 0.169 | 1.55 | 0 | 0.407 | 46.344 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 01:20
AB-220715-01 CCC-H 8 | 0 | 357.065 | 400397.555 | 2350663.417 | 389.103 | 5106299.443 | 3966.62 | 385.036 | 336.433 | 377.342 | 316.872 | 325.053 | 574.823 | 0 | 401.32 | 380.715 | 375.361 | 0 | 0 | 390.578 | 378.446 | 0 | 379.099 | 342.018 | 374.878 | 0 | 395.998 | 415.555 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 01:56
AB-220715-01 LFSM-H 7 | 0 | 872.937 | 207451.099 | 1065725.275 | 191.747 | 2848347.37 | 3079.549 | 220.478 | 176.03 | 236.177 | 153.039 | 150.528 | 364.538 | 0 | 201.623 | 0 | 201.943 | 0 | 0 | 195.526 | 188.25 | 0 | 196.726 | 176.145 | 196.063 | 0 | 204.965 | 156.285 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 02:32
WO-22270-141-A-1 | 0 | 432.865 | 1434.497 | 3331.26 | 0.55 | 7084.771 | 11.738 | 34.124 | 1.527 | 1.058 | 0 | 1.175 | 188.02 | 0 | 0.992 | 0 | 3.221 | 0 | 0 | 1.068 | 1.641 | 0 | 0.808 | 0.847 | 1.229 | 0 | 0.838 | 68.821 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 03:08
WO-22270-142-A-1 | 0 | 449.446 | 1714.475 | 239.509 | 0.321 | 0 | 5.797 | 33.507 | 1.463 | 0.945 | 0 | 1.051 | 187.917 | 0 | 0.865 | 0 | 2.606 | 0 | 0 | 0 | 0 | 0 | 0.486 | 0.41 | 1.132 | 0 | 0.492 | -1.685 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 03:44
WO-22270-143-A-1 | 0 | 1319.375 | 1620.926 | -97.186 | 0.357 | 0 | 6.583 | 26.498 | 2.015 | 1.735 | 0 | 0.984 | 188.245 | 0 | 0.786 | 0 | 3.584 | 0 | 0 | 0 | 0 | 0 | 0.608 | 0.733 | 0.601 | 0 | 0.437 | 4.384 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 04:20
WO-22270-144-A-1 | 0 | 752.207 | 3690.267 | -196.4 | 0.194 | 163.597 | 5.183 | 102.136 | 1.164 | 0.743 | 0 | 1.097 | 189.092 | 0 | 0.747 | 0 | 3.23 | 0 | 0 | 4.989 | 0 | 0 | 0.83 | 0.504 | 1.771 | 0 | 0.377 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 04:56
WO-22270-145-A-1 | 0 | 909.389 | 1765.158 | -255.025 | 0.226 | 81.577 | 4.453 | 88.027 | 0.689 | 0.617 | 0 | 0.844 | 188.287 | 0 | 0.809 | 0 | 2.649 | 0 | 0 | 2.422 | 0 | 0 | 0.525 | 0.393 | 1.335 | 0 | 0.352 | 54.026 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 05:32
WO-22270-146-A-1 | 0 | 537.5 | 339.421 | -457.809 | 0.194 | 0 | 4.289 | 92.579 | 0.882 | 0.943 | 0 | 0.965 | 189.194 | 0 | 0.883 | 0 | 2.223 | 0 | 0 | 0 | 0 | 0 | 0.591 | 0.494 | 1.95 | 0 | 0.339 | 80.268 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 06:08
WO-22270-147-A-1 | 0 | 1107.263 | 1385.91 | -28.236 | 0.187 | 0 | 0 | 324.339 | 1.916 | 0.863 | 0 | 1.445 | 189.12 | 0 | 0.57 | 0 | 3.744 | 0 | 0 | 0 | 0 | 0 | 0.489 | 0.757 | 2.685 | 0 | 0.41 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 06:44
WO-22270-148-A-1 | 0 | 977.58 | 1846.577 | 737.936 | 0.176 | 0 | 4.377 | 304.539 | 1.139 | 0.872 | 0 | 1.57 | 188.053 | 0 | 0.631 | 0 | 2.35 | 0 | 0 | 3.04 | 2.76 | 0 | 0.533 | 0.307 | 4.589 | 0 | 0.352 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 07:20
WO-22270-149-A-1 | 0 | 624.84 | 1688.909 | -125.136 | 0.179 | 0 | 4.476 | 312.784 | 0.908 | 0.747 | 0 | 0.928 | 188.119 | 0 | 0.96 | 0 | 4.147 | 0 | 0 | 0 | 2.979 | 0 | 0.671 | 0.253 | 1.944 | 0 | 0.313 | 46.838 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 07:56
WO-22270-150-A-1 | 0 | 644.971 | 359.833 | -17.626 | 0.176 | 0 | 4.301 | 267.728 | 1.28 | 0.919 | 0 | 0.854 | 188.123 | 0 | 0.974 | 0 | 3.079 | 0 | 0 | 0 | 0 | 0 | 0.73 | 0.253 | 0.552 | 0 | 0.382 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 08:32
AB-220715-01 LRB 16_INJ001 | 0 | 159.939 | 196.658 | 1089.619 | 0.192 | 1.891 | 4.362 | 2.056 | 0.682 | 0.373 | 0 | 1.009 | 187.928 | 0 | 0.884 | 0 | 0.43 | 0 | 0 | 0 | 0 | 0 | 0.56 | 0.085 | 0.858 | 0 | 0.36 | -0.437 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 09:08
AB-220715-01 LRB 16 | 0 | 133.202 | 535.015 | 1735.266 | 0.175 | 0 | 4.336 | 1.823 | 0.881 | 0.292 | 0 | 0.899 | 187.647 | 0 | 0.901 | 0 | 0.429 | 0 | 0 | 0 | 0 | 0 | 0.378 | 0.101 | 0.865 | 0 | 0.314 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 09:44
WO-22270-151-A-1 | 0 | 508.489 | 927.013 | -181.719 | 0.195 | 0 | 0 | 993.029 | 0.918 | 0.587 | 0 | 0.829 | 187.726 | 0 | 0.887 | 0 | 2.406 | 0 | 0 | 0 | 1.571 | 0 | 1.754 | 0.22 | 0.623 | 0 | 0.323 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 10:20
WO-22270-152-A-1 | 0 | 497.604 | 2413.122 | 813.222 | 0.176 | 0 | 4.416 | 1050.632 | 1.351 | 0.585 | 0 | 2.257 | 187.693 | 0 | 0.879 | 0 | 2.648 | 0 | 0 | 0 | 1.672 | 0 | 0.517 | 0.2 | 0.75 | 0 | 0.352 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 10:56
WO-22270-153-A-1 | 0 | 859.471 | 746.388 | -451.643 | 0.19 | 0 | 0 | 958.71 | 1.915 | 0.69 | 0 | 1.34 | 187.772 | 0 | 0.4 | 0 | 3.295 | 0 | 0 | 0 | 2.461 | 0 | 0.645 | 0.221 | 1.145 | 0 | 0.337 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 11:32
WO-22270-154-A-1 | 0 | 401.255 | 27.113 | 225.824 | 0.179 | 0 | 4.634 | 2069.312 | 0.893 | 0.451 | 0 | 1.161 | 187.703 | 0 | 0.885 | 0 | 4.213 | 0 | 0 | 0 | 2.409 | 0 | 5.502 | 0.058 | 0.773 | 0 | 0.318 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 12:08
WO-22270-155-A-1 | 0 | 333.591 | 604.733 | -147.316 | 0.17 | 0 | 4.405 | 3164.557 | 0.924 | 0.464 | 0 | 0.936 | 187.619 | 0 | 0.648 | 0 | 3.675 | 0 | 0 | 0 | 2.974 | 0 | 0.599 | 0.068 | 0.664 | 0 | 0.314 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 12:44
WO-22270-156-A-1 | 0 | 793.249 | 1015.968 | -511.051 | 0.225 | 0 | 0 | 2732.931 | 1.656 | 0.574 | 0 | 1.285 | 188.53 | 0 | 0.544 | 0 | 3.488 | 0 | 0 | 0 | 2.804 | 0 | 0.575 | 0.223 | 1.519 | 0 | 0.332 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 13:20
WO-22270-157-A-1 | 0 | 348.709 | 102.602 | -347.872 | 0.166 | 0 | 0 | 17.397 | 1.469 | 0.6 | 0 | 1.133 | 187.771 | 0 | 0.886 | 0 | 3.856 | 0 | 0 | 2.832 | 3.054 | 0 | 4.66 | 0.122 | 1.923 | 0 | 0.342 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 13:56
WO-22270-158-A-1 | 0 | 797.264 | 1109.268 | 252.499 | 0.178 | 0 | 0 | 7.677 | 1.107 | 0.994 | 0 | 2.294 | 188.131 | 0 | 0.61 | 0 | 2.164 | 0 | 0 | 0 | 1.143 | 0 | 1.068 | 0.065 | 1.129 | 0 | 0.301 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 14:32
WO-22270-159-A-1 | 0 | 674.126 | 829.057 | 356.948 | 0.177 | 122.794 | 4.395 | 8.152 | 1.589 | 0.622 | 0 | 1.054 | 187.663 | 0 | 0.608 | 0 | 3.551 | 0 | 0 | 0 | 0 | 0 | 0.726 | 0.067 | 1.534 | 0 | 0.307 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 15:08
WO-22270-160-A-1 | 0 | 699.334 | 3558.733 | -352.482 | 0.189 | 0 | 0 | 3.988 | 1.25 | 0.894 | 0 | 2.993 | 187.698 | 0 | 0.525 | 0 | 3.234 | 0 | 0 | 0 | 0 | 0 | 0.636 | 0.235 | 0.554 | 0 | 0.302 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 15:44
AB-220715-01 LFB-H 9 | 0 | 179.273 | 184173.045 | 1193753.672 | 181.484 | 3562575.448 | 2000.378 | 180.037 | 160.929 | 200.056 | 144.71 | 147.606 | 385.193 | 0 | 199.5 | 225.666 | 201.143 | 0 | 0 | 199.053 | 195.058 | 0 | 209.15 | 151.798 | 199.266 | 0 | 197.382 | 302.704 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 16:20
AB-220715-01 LRB 17_INJ001 | 0 | 189.576 | 860.666 | 7786.963 | 1.198 | 7931.958 | 15.366 | 0.812 | 1.121 | 1.196 | 1.497 | 1.555 | 188.161 | 0 | 1.533 | 0 | 1.538 | 0 | 0 | 1.256 | 1.817 | 0 | 1.235 | 1.184 | 1.693 | 0 | 1.839 | 18.523 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 16:56
AB-220715-01 LRB 17 | 0 | 192.221 | 324.303 | 1625.266 | 0.227 | 206.856 | 4.743 | 3.442 | 0.641 | 0.382 | 0 | 1.002 | 187.939 | 0 | 1.023 | 0 | 0.466 | 0 | 0 | 0 | 1.134 | 0 | 0.388 | 0.102 | 0.893 | 0 | 0.432 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 17:32
AB-220715-01 LD 8 | 0 | 603.311 | 1097.445 | 505.6 | 0.215 | 172.239 | 5.395 | 8.326 | 1.145 | 0.718 | 0 | 1.327 | 188.188 | 0 | 0.834 | 0 | 3.377 | 0 | 0 | 0 | 2.782 | 0 | 0.66 | 0.198 | 3.612 | 0 | 0.426 | 25.975 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 18:08
AB-220715-01 CCC-H 9 | 0 | 370.536 | 401819.365 | 2284987.744 | 376.49 | 4169724.864 | 4286.54 | 369.546 | 341.763 | 373.974 | 304.674 | 311.608 | 600.951 | 0 | 397.55 | 361.131 | 363.614 | 0 | 0 | 392.862 | 391.441 | 0 | 391.626 | 324.939 | 386.79 | 0 | 400.534 | 373.29 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 18:44
AB-220715-01 LFSM-H 8 | 0 | 545.767 | 193390.292 | 1078186.147 | 199.404 | 2273336.683 | 3525.523 | 214.18 | 178.667 | 216.133 | 144.849 | 151.908 | 365.028 | 0 | 198.021 | 0 | 187.394 | 0 | 0 | 193.345 | 184.277 | 0 | 199.472 | 168.893 | 202.262 | 0 | 205.813 | 126.079 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 19:20
WO-22270-161-A-1 | 0 | 374.117 | 1237.971 | 629.047 | 0.386 | 482.122 | 9.39 | 12.151 | 1.741 | 0.943 | 0 | 1.224 | 188.191 | 0 | 0.991 | 0 | 2.606 | 0 | 0 | 0 | 1.612 | 0 | 1.19 | 0.727 | 0.627 | 0 | 0.586 | 9.032 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 19:56
WO-22270-162-A-1 | 0 | 530.721 | 683.792 | 241.509 | 0.276 | 0 | 6.067 | 15.364 | 1.197 | 0.672 | 0 | 1.038 | 187.782 | 0 | 1.125 | 0 | 3.833 | 0 | 0 | 2.793 | 0 | 0 | 0.54 | 0.442 | 0.958 | 0 | 0.41 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 20:32
WO-22270-163-A-1 | 0 | 608.499 | -141.249 | 776.718 | 0.228 | 0 | 6.085 | 14.745 | 1.477 | 1.035 | 0 | 1.276 | 187.807 | 0 | 1.065 | 0 | 2.784 | 0 | 0 | 0 | 0 | 0 | 0.599 | 0.587 | 1.607 | 0 | 0.488 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 21:08
WO-22270-164-A-1 | 0 | 477.573 | 702.703 | -527.035 | 0.199 | 0 | 5.163 | 41.544 | 1.355 | 0.496 | 0 | 2.595 | 187.674 | 0 | 1.033 | 0 | 3.054 | 0 | 0 | 0 | 1.282 | 0 | 1.221 | 0.297 | 1.813 | 0 | 0.38 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 21:44
WO-22270-165-A-1 | 0 | 456.255 | 260.59 | -217.347 | 0.207 | 0 | 4.69 | 41.894 | 1.596 | 0.811 | 0 | 0.961 | 188.16 | 0 | 0.736 | 0 | 3.212 | 0 | 0 | 0 | 0 | 0 | 0.495 | 0.399 | 0.662 | 0 | 0.434 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 22:20
WO-22270-166-A-1 | 0 | 468.243 | 2022.892 | -171.061 | 0.177 | 0 | 4.774 | 43.64 | 1.124 | 0.617 | 0 | 1.05 | 187.76 | 0 | 0.863 | 0 | 3.827 | 0 | 0 | 0 | 0 | 0 | 0.432 | 0.292 | 0.702 | 0 | 0.361 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-22 22:56
WO-22270-167-A-1 | 0 | 926.237 | 1857.922 | 260.741 | 0.186 | 0 | 0 | 148.683 | 1.721 | 0.739 | 0 | 1.072 | 188.278 | 0 | 0.624 | 0 | 5.225 | 0 | 0 | 0 | 0 | 0 | 1.157 | 0.185 | 1.69 | 0 | 0.319 | 68.971 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-22 23:32
WO-22270-168-A-1 | 0 | 737.249 | 978.444 | -190.548 | 0.191 | 0 | 4.384 | 139.298 | 1.091 | 0.65 | 0 | 2.366 | 188.308 | 0 | 0.73 | 0 | 2.786 | 0 | 0 | 3.339 | 0 | 0 | 0.534 | 0.247 | 1.087 | 0 | 0.353 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 00:08
WO-22270-169-A-1 | 0 | 1388.371 | 4127.374 | 119.983 | 0.193 | 265.354 | 4.49 | 211.612 | 1.6 | 0.818 | 0 | 1.398 | 188.907 | 0 | 0.789 | 0 | 3.554 | 0 | 0 | 0 | 1.952 | 0 | 0.82 | 0.445 | 1.623 | 0 | 0.394 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 00:44
WO-22270-170-A-1 | 0 | 691.204 | 2073.372 | -161.839 | 0.18 | 0 | 4.516 | 223.747 | 1.209 | 0.737 | 0 | 2.394 | 188.543 | 0 | 0.582 | 0 | 2.996 | 0 | 0 | 0 | 0 | 0 | 0.459 | 0.174 | 0.549 | 0 | 0.352 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 01:20
AB-220715-01 LRB 18_INJ001 | 0 | 114.063 | 93.2 | 168.151 | 0.176 | 0 | 4.473 | 2.299 | 0.741 | 0.333 | 2.417 | 0.953 | 187.594 | 0 | 1.014 | 0 | 0.499 | 0 | 0 | 0 | 1.219 | 0 | 0.361 | 0.135 | 0.906 | 0 | 0.303 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 01:56
WO-22209-62-A-1B | 0 | 1110.491 | 2013.411 | 613.096 | 0.185 | 0 | 4.846 | 0 | 1.04 | 1.202 | 0 | 0.892 | 188.334 | 0 | 0.895 | 0 | 114.95 | 0 | 0 | 0 | 0 | 0 | 0.453 | 0.216 | 3.26 | 0 | 0.296 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 02:32
AB-220715-01 LRB 18 | 0 | 111.461 | -59.955 | 1106.857 | 0.192 | 0 | 4.314 | 1.268 | 0.778 | 0.381 | 0 | 0.933 | 188.018 | 0 | 1.142 | 0 | 0.455 | 0 | 0 | 0 | 1.199 | 0 | 0.363 | 0.089 | 0.945 | 0 | 0.332 | 12.141 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 03:08
WO-22270-171-A-1 | 0 | 363.353 | 129.772 | -388.098 | 0.17 | 657.429 | 4.75 | 448.566 | 0.883 | 0.618 | 0 | 0.835 | 187.728 | 0 | 1.015 | 0 | 2.703 | 0 | 0 | 0 | 0 | 0 | 1.36 | 0.196 | 1.172 | 0 | 0.315 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 03:44
WO-22270-172-A-1 | 0 | 388.823 | 842.025 | 598.295 | 0.172 | 0 | 4.437 | 490.756 | 0.83 | 0.609 | 0 | 2.218 | 187.709 | 0 | 0.977 | 0 | 3.409 | 0 | 0 | 0 | 1.676 | 0 | 0.561 | 0.164 | 1.077 | 0 | 0.307 | 2.457 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 04:20
WO-22270-173-A-1 | 0 | 555.787 | 1393.404 | 238.921 | 0.184 | 150.209 | 4.72 | 652.269 | 1.345 | 0.7 | 0 | 0.901 | 187.941 | 0 | 0.804 | 0 | 2.969 | 0 | 0 | 0 | 1.572 | 0 | 0.645 | 0.195 | 1.491 | 0 | 0.328 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 04:56
WO-22270-174-A-1 | 0 | 740.05 | 1611.789 | -318.849 | 0.174 | 0 | 4.334 | 2324.998 | 1.375 | 0.569 | 0 | 1.104 | 187.935 | 0 | 0.676 | 0 | 2.746 | 0 | 0 | 3.038 | 3.302 | 0 | 2.072 | 0.088 | 1.675 | 0 | 0.318 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 05:32
WO-22270-175-A-1 | 0 | 338.075 | 355.31 | -364.433 | 0.175 | 0 | 0 | 1553.208 | 0.928 | 0.703 | 0 | 1.139 | 187.774 | 0 | 0.901 | 0 | 2.174 | 0 | 0 | 0 | 1.554 | 0 | 0.439 | 0.139 | 0.839 | 0 | 0.291 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 06:08
WO-22270-176-A-1 | 0 | 873.16 | 1346.751 | 626.725 | 0.174 | 0 | 0 | 2159.757 | 3.32 | 0.703 | 0 | 1.438 | 188.086 | 0 | 0.911 | 0 | 2.838 | 0 | 0 | 0 | 3.578 | 0 | 0.586 | 0.226 | 0.633 | 0 | 0.32 | 11.94 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 06:44
WO-22270-177-A-1 | 0 | 380.104 | 588.062 | -470.427 | 0.175 | 0 | 4.969 | 14.533 | 0.98 | 0.554 | 0 | 0.977 | 187.79 | 0 | 0.786 | 0 | 1.97 | 0 | 0 | 0 | 0 | 0 | 1.46 | 0.105 | 0.592 | 0 | 0.332 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 07:20
WO-22270-178-A-1 | 0 | 712.816 | 747.72 | -397.303 | 0.175 | 0 | 0 | 11.442 | 1.21 | 0.733 | 0 | 2.858 | 187.987 | 0 | 1.206 | 0 | 2.336 | 0 | 0 | 0 | 0 | 0 | 0.663 | 0.153 | 1.784 | 0 | 0.323 | 135.129 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 07:56
WO-22270-179-A-1 | 0 | 498.837 | 246.649 | -344.837 | 0.191 | 0 | 0 | 9.067 | 1.082 | 0.897 | 0 | 0.941 | 187.734 | 0 | 0.862 | 0 | 3.117 | 0 | 0 | 0 | 0 | 0 | 0.545 | 0.064 | 0.588 | 0 | 0.326 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 08:32
WO-22270-180-A-1 | 0 | 471.512 | 503.558 | 80.265 | 0.178 | 0 | 4.363 | 7.807 | 1.39 | 0.577 | 0 | 2.126 | 187.797 | 0 | 0.858 | 0 | 3.526 | 0 | 0 | 0 | 1.518 | 0 | 0.622 | 0.113 | 1.33 | 0 | 0.301 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 09:08
WO-22209-64-A-1B | 0 | 497.715 | 249.565 | 476.014 | 0.195 | 0 | 4.428 | 0 | 1.237 | 1.769 | 0 | 0.812 | 187.997 | 0 | 0.728 | 0 | 567.441 | 0 | 0 | 0 | 0 | 0 | 0.769 | 0.043 | 7.981 | 0 | 0.326 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 09:44
AB-220715-01 LFB-H 10 | 0 | 183.313 | 192175.585 | 1136760.867 | 190.493 | 2233134.065 | 2259.241 | 181.985 | 165.749 | 195.644 | 154.851 | 164.645 | 397.271 | 0 | 197.667 | 148.473 | 192.246 | 0 | 0 | 193.519 | 193.018 | 0 | 196.216 | 148.837 | 200.351 | 0 | 200.62 | 189.03 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 10:20
AB-220715-01 LRB 19_INJ001 | 0 | 92.578 | 448.323 | 4643.06 | 0.879 | 9282.624 | 14.475 | 2.874 | 1.233 | 0.889 | 2.675 | 1.411 | 188.01 | 0 | 1.206 | 0 | 1.205 | 0 | 0 | 1.117 | 1.801 | 0 | 0.997 | 0.552 | 1.349 | 0 | 1.321 | 8.994 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 10:56
AB-220715-01 LRB 19 | 0 | 146.981 | 184.569 | -292.191 | 0.243 | 1660.244 | 5.742 | 2.143 | 0.991 | 0.589 | 2.013 | 1.059 | 187.579 | 0 | 0.956 | 0 | 0.481 | 0 | 0 | 0 | 1.098 | 0 | 0.459 | 0.267 | 0.961 | 0 | 0.439 | 25.896 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 11:32
AB-220715-01 LD 9 | 0 | 480.119 | 1791.344 | 0 | 0.194 | 0 | 5.206 | 7.444 | 1.746 | 0.798 | 0 | 2.715 | 187.866 | 0 | 0.629 | 0 | 2.987 | 0 | 0 | 0 | 0 | 0 | 0.708 | 0.3 | 0.625 | 0 | 0.427 | 39.821 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 12:08
AB-220715-01 CCC-H 10 | 0 | 385.09 | 396357.497 | 2134411.167 | 382.755 | 4321233.184 | 4306.494 | 386.392 | 339.369 | 373.044 | 312.752 | 345.56 | 614.869 | 0 | 405.738 | 368.792 | 383.207 | 0 | 0 | 394.703 | 388.865 | 0 | 383.256 | 298.051 | 398.474 | 0 | 402.651 | 384.608 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 12:44
AB-220715-01 LFSM-H 9 | 0 | 632.759 | 186689.026 | 1051632.937 | 193.501 | 3677637.883 | 4115.085 | 197.643 | 169.264 | 213.425 | 146.015 | 153.973 | 362.349 | 0 | 208.652 | 0 | 191.971 | 0 | 0 | 192.767 | 200.41 | 0 | 197.609 | 168.227 | 195.88 | 0 | 197.637 | 180.703 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 13:20
WO-22270-181-A-1 | 0 | 295.477 | 412.958 | 2691.353 | 0.71 | 8203.047 | 14.155 | 15.258 | 1.563 | 1.08 | 0 | 1.208 | 188.12 | 0 | 0.966 | 0 | 5.86 | 0 | 0 | 2.08 | 1.517 | 0 | 1.717 | 0.586 | 2.224 | 0 | 0.841 | 0 | 0 | 0 | 0 | 0 | C | Unknown | 2022-07-23 13:56
WO-22270-182-A-1 | 0 | 903.516 | 2511.945 | 457.954 | 0.269 | 499.822 | 7.563 | 13.876 | 1.902 | 0.83 | 0 | 1.079 | 187.786 | 0 | 0.879 | 0 | 4.081 | 0 | 0 | 0 | 0 | 0 | 0.692 | 0.569 | 1.034 | 0 | 0.479 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 14:32
WO-22270-183-A-1 | 0 | 305.831 | 888.062 | 55.856 | 0.234 | 67.282 | 4.935 | 16.358 | 0.822 | 0.759 | 0 | 1.153 | 187.968 | 0 | 0.864 | 0 | 3.847 | 0 | 0 | 0 | 0 | 0 | 0.573 | 0.345 | 1.842 | 0 | 0.472 | 12.809 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 15:08
WO-22270-184-A-1 | 0 | 425.218 | 4601.483 | 124.736 | 0.199 | 0 | 5.179 | 35.552 | 1.233 | 0.676 | 0 | 0.882 | 187.909 | 0 | 0.645 | 0 | 3.932 | 0 | 0 | 0 | 0 | 0 | 1.819 | 0.334 | 1.077 | 0 | 0.399 | 35.396 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 15:44
WO-22270-185-A-1 | 0 | 551.563 | 903.84 | 314.46 | 0.184 | 0 | 4.92 | 45.211 | 1.791 | 0.698 | 0 | 0.861 | 188.492 | 0 | 0.518 | 0 | 3.544 | 0 | 0 | 1.62 | 4.181 | 0 | 0.582 | 0.303 | 0.856 | 0 | 0.354 | 81.177 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 16:20
WO-22209-50-A-1B | 0 | 507.148 | 761.16 | -222.136 | 0.182 | 0 | 4.659 | 0 | 1.689 | 0.746 | 0 | 2.502 | 188.551 | 0 | 0.781 | 0 | 1038.324 | 0 | 0 | 0 | 0 | 0 | 1.284 | 0.092 | 32.495 | 0 | 0.318 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 16:56
WO-22270-186-A-1 | 0 | 942.291 | 370.595 | 705.64 | 0.19 | 0 | 4.644 | 49.661 | 1.268 | 0.689 | 0 | 3.391 | 187.894 | 0 | 0.732 | 0 | 3.091 | 0 | 0 | 0 | 2.227 | 0 | 0.7 | 0.442 | 1.14 | 0 | 0.483 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 17:32
WO-22270-187-A-1 | 0 | 264.215 | 1699.485 | -34.082 | 0.184 | 0 | 4.328 | 190.082 | 1.112 | 0.721 | 0 | 1.227 | 187.651 | 0 | 0.732 | 0 | 4.366 | 0 | 0 | 0 | 1.457 | 0 | 0.67 | 0.223 | 1.973 | 0 | 0.353 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 18:08
WO-22270-188-A-1 | 0 | 368.475 | 1408.27 | -90.525 | 0.211 | 473.302 | 4.93 | 166.615 | 1.364 | 0.714 | 0 | 0.929 | 187.935 | 0 | 0.632 | 0 | 7.36 | 0 | 0 | 1.105 | 1.564 | 0 | 0.578 | 0.255 | 4.744 | 0 | 0.383 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 18:44
WO-22270-189-A-1 | 0 | 716.649 | 4049.28 | -225.788 | 0.181 | 509.394 | 4.848 | 163.398 | 1.217 | 0.695 | 0 | 1.014 | 187.647 | 0 | 0.549 | 0 | 6.961 | 0 | 0 | 2.777 | 1.988 | 0 | 1.001 | 0.189 | 1.948 | 0 | 0.306 | 2.342 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 19:20
WO-22270-190-A-1 | 0 | 586.393 | 870.088 | 150.108 | 0.182 | 0 | 0 | 157.924 | 1.053 | 0.873 | 0 | 0.88 | 188.33 | 0 | 0.738 | 0 | 7.471 | 0 | 0 | 0 | 1.455 | 0 | 1.21 | 0.147 | 1.469 | 0 | 0.289 | 3.279 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 19:56
AB-220715-01 LRB 20_INJ001 | 0 | 105.621 | 190.516 | 439.517 | 0.171 | 311.764 | 4.68 | 2.254 | 1.062 | 0.338 | 0 | 1.871 | 187.677 | 0 | 0.943 | 0 | 0.413 | 0 | 0 | 0 | 0 | 0 | 0.395 | 0.215 | 0.643 | 0 | 0.343 | -1.08 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 20:32
AB-220715-01 LRB 20 | 0 | 122.89 | 719.619 | -451.517 | 0.178 | 435.444 | 0 | 1.802 | 0.784 | 0.361 | 2.185 | 0.947 | 187.61 | 0 | 1.019 | 0 | 0.406 | 0 | 0 | 0 | 0 | 0 | 0.65 | 0.224 | 0.678 | 0 | 0.307 | 0.728 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 21:08
WO-22270-196-A-1 | 0 | 245.233 | 1155.914 | 39.231 | 0.173 | 140.209 | 4.328 | 1495.148 | 0.888 | 0.719 | 0 | 2.261 | 187.869 | 0 | 0.841 | 0 | 7.764 | 0 | 0 | 0 | 2.292 | 0 | 0.658 | 0.132 | 1.261 | 0 | 0.299 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-23 21:44
WO-22270-191-A-1 | 0 | 342.918 | 2497.322 | -288.367 | 0.171 | 0 | 4.28 | 590.481 | 0.892 | 0.636 | 0 | 2.233 | 187.985 | 0 | 0.697 | 0 | 7.844 | 0 | 0 | 0 | 1.491 | 0 | 1.144 | 0.121 | 1.588 | 0 | 0.309 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 22:20
WO-22270-197-A-1 | 0 | 480.823 | -40.592 | -350.025 | 0.175 | 0 | 0 | 1403.188 | 1.04 | 0.603 | 0 | 1.726 | 187.977 | 0 | 0.806 | 0 | 3.32 | 0 | 0 | 0 | 2.479 | 0 | 1.063 | 0.148 | 1.949 | 0 | 0.343 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 22:56
WO-22270-192-A-1 | 0 | 260.774 | 1358.077 | 491.542 | 0.179 | 0 | 0 | 714.487 | 0.761 | 0.671 | 0 | 2.008 | 187.851 | 0 | 0.711 | 0 | 7.309 | 0 | 0 | 1.041 | 1.811 | 0 | 0.578 | 0.131 | 1.947 | 0 | 0.326 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-23 23:32
WO-22270-198-A-1 | 0 | 379.117 | 1293.838 | -395.564 | 0.173 | 0 | 4.272 | 11190.616 | 3.573 | 0.562 | 0 | 2.524 | 187.739 | 0 | 1.082 | 0 | 4.209 | 0 | 0 | 0.957 | 3333.538 | 0 | 0.556 | 0.111 | 0.887 | 0 | 0.327 | 4.961 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 00:08
WO-22270-193-A-1 | 0 | 334.998 | 354.494 | 100.364 | 0.185 | 0 | 0 | 629.943 | 0.826 | 0.573 | 0 | 1.053 | 187.674 | 0 | 0.66 | 0 | 8.219 | 0 | 0 | 0 | 1.486 | 0 | 0.646 | 0.088 | 0.916 | 0 | 0.321 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 00:44
WO-22270-199-A-1 | 0 | 592.626 | 1474.833 | 83.177 | 0.172 | 0 | 0 | 12321.404 | 2.963 | 0.733 | 0 | 2.154 | 187.77 | 0 | 0.828 | 0 | 5.232 | 0 | 0 | 0 | 3521.747 | 0 | 0.678 | 0.135 | 1.283 | 0 | 0.304 | -1.357 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 01:20
WO-22209-26-A-1B | 0 | 945.963 | 819.894 | -237.992 | 0 | 0 | 0 | 0 | 1.409 | 1.116 | 0 | 0.907 | 187.862 | 0 | 0.807 | 0 | 29.291 | 0 | 0 | 0 | 1.719 | 0 | 0.873 | 0.243 | 27.337 | 0 | 0.329 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 01:56
WO-22270-200-A-1 | 0 | 660.136 | 2288.562 | 13.805 | 0.178 | 0 | 4.291 | 11615.642 | 2.631 | 0.562 | 0 | 1.4 | 187.651 | 0 | 0.902 | 0 | 3.552 | 0 | 0 | 0 | 3344.616 | 0 | 1.058 | 0.104 | 1.155 | 0 | 0.326 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 02:32
WO-22270-194-A-1 | 0 | 195.729 | 1946.31 | 4566.076 | 0.178 | 0 | 4.458 | 1279.045 | 1.068 | 0.678 | 0 | 0.834 | 187.632 | 0 | 0.722 | 0 | 5.876 | 0 | 0 | 0.627 | 1.773 | 0 | 0.623 | 0.048 | 1.515 | 0 | 0.295 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 03:08
WO-22270-195-A-1 | 0 | 271.791 | 620.599 | -299.107 | 0.172 | 21.34 | 4.526 | 911.43 | 0.896 | 0.622 | 0 | 2.281 | 187.748 | 0 | 0.928 | 0 | 4.423 | 0 | 0 | 0 | 2.016 | 0 | 0.519 | 0.114 | 1.486 | 0 | 0.305 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 03:44
AB-220715-01 LFB-H 11 | 0 | 180.599 | 203287.697 | 1079076.758 | 180.881 | 1913705.613 | 2001.294 | 231.801 | 167.433 | 176.526 | 151.857 | 160.42 | 392.271 | 0 | 204.886 | 187.116 | 192.084 | 0 | 0 | 195.694 | 198.179 | 0 | 199.559 | 143.564 | 190.277 | 0 | 203.133 | 180.657 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 04:20
AB-220715-01 LRB 21 | 0 | 134.845 | -144.747 | -176.01 | 0.247 | 1452.803 | 5.844 | 1.799 | 0.881 | 0.4 | 0 | 2.293 | 187.912 | 0 | 0.833 | 0 | 0.455 | 0 | 0 | 0 | 1.323 | 0 | 0.413 | 0.178 | 0.894 | 0 | 0.447 | 11.457 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 04:56
AB-220715-01 LD 10 | 0 | 724.351 | 2453.882 | 275.408 | 0.205 | 70.681 | 5.027 | 5472.988 | 1.452 | 0.657 | 0 | 2.393 | 187.695 | 0 | 0.573 | 0 | 2.957 | 0 | 0 | 0 | 2201.609 | 0 | 0.723 | 0.153 | 1.37 | 0 | 0.356 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 05:32
AB-220715-01 CCC-H 11 | 0 | 360.377 | 411309.975 | 2209183.059 | 370.046 | 3728072.273 | 3494.515 | 422.132 | 351.242 | 357.395 | 319.013 | 330.388 | 615.979 | 0 | 416.753 | 343.684 | 407.738 | 0 | 0 | 383.878 | 372.456 | 0 | 397.692 | 278.725 | 389.571 | 0 | 396.039 | 358.14 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 06:08
WO-22209-16-A-1B | 0 | 1284.279 | 3668.587 | 1764.813 | 0.232 | 0 | 4.619 | 37.659 | 2.387 | 1.022 | 0 | 1.251 | 188.845 | 0 | 0.88 | 0 | 38.635 | 0 | 0 | 0 | 9.482 | 0 | 1.101 | 0.178 | 2.417 | 0 | 0.459 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 06:44
AB-220715-01 LFSM-H 10 | 0 | 1642.322 | 187709.896 | 1081933.816 | 194.187 | 2597560.901 | 3679.328 | 5413.034 | 167.184 | 216.672 | 149.14 | 147.507 | 351.761 | 0 | 197.066 | 0 | 187.19 | 0 | 0 | 187.935 | 2486.335 | 0 | 197.927 | 150.219 | 188.316 | 0 | 199.607 | 153.272 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 07:20
WO-22270-201-A-1 | 0 | 886.991 | 2818.686 | 2055.531 | 0.422 | 356.689 | 9.555 | 11785.487 | 2.236 | 1.119 | 0 | 1.709 | 188.017 | 0 | 0.838 | 0 | 3.658 | 0 | 0 | 0 | 3175.149 | 0 | 1.074 | 0.607 | 1.51 | 0 | 0.767 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 07:56
WO-22270-202-A-1 | 0 | 395.964 | 1768.526 | 189.722 | 0.257 | 75.292 | 5.169 | 12663.908 | 2.206 | 0.821 | 0 | 1.114 | 188.514 | 0 | 0.869 | 0 | 2.071 | 0 | 0 | 0 | 3534.884 | 0 | 0.541 | 0.345 | 0.69 | 0 | 0.383 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 08:32
WO-22270-203-A-1 | 0 | 342.351 | 1350.103 | 264.293 | 0.213 | 0 | 5.828 | 12265.672 | 2.256 | 0.677 | 0 | 1.007 | 187.666 | 0 | 0.818 | 0 | 3.444 | 0 | 0 | 0 | 3405.809 | 0 | 0.56 | 0.231 | 0.655 | 0 | 0.394 | 15.565 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 09:08
AB-220715-01 LRB 21_INJ001 | 0 | 106.422 | 435.685 | 3460.962 | 0.821 | 8843.147 | 12.058 | 2.938 | 1.046 | 0.919 | 0 | 2.29 | 188.236 | 0 | 1.136 | 0 | 0.736 | 0 | 0 | 1.119 | 2.423 | 0 | 0.881 | 0.282 | 1.379 | 0 | 0.957 | 3.968 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 09:44
WO-22270-204-A-1 | 0 | 856.475 | 1298.825 | -35.948 | 0.203 | 0 | 4.913 | 11125.024 | 2.686 | 0.718 | 0 | 1.336 | 188.017 | 0 | 0.797 | 0 | 1.971 | 0 | 0 | 0 | 3483.05 | 0 | 0.626 | 0.216 | 3.116 | 0 | 0.392 | 30.705 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 10:20
WO-22270-205-A-1 | 0 | 632.934 | 3274.024 | -87.21 | 0.18 | 0 | 4.354 | 11343.019 | 2.122 | 0.71 | 0 | 1.674 | 188.175 | 0 | 0.626 | 0 | 2.288 | 0 | 0 | 0 | 3455.628 | 0 | 0.737 | 0.352 | 0.526 | 0 | 0.346 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 10:56
WO-22270-206-A-1 | 0 | 719.857 | 4357.568 | 405.929 | 0.179 | 0 | 4.486 | 11381.093 | 2.199 | 0.941 | 0 | 3.293 | 188.811 | 0 | 0.876 | 0 | 4.095 | 0 | 0 | 0 | 3284.437 | 0 | 0.995 | 0.292 | 1.254 | 0 | 0.351 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 11:32
WO-22270-207-A-1 | 0 | 796.364 | 1447.769 | -161.357 | 0.185 | 0 | 0 | 10832.086 | 2.467 | 0.652 | 0 | 1.242 | 188.359 | 0 | 0.768 | 0 | 5.134 | 0 | 0 | 0 | 3592.031 | 0 | 1.019 | 0.316 | 2.265 | 0 | 0.312 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 12:08
WO-22270-208-A-1 | 0 | 405.594 | 1736.863 | 0 | 0.184 | 0 | 0 | 7278.211 | 2.764 | 0.781 | 0 | 1.031 | 188.354 | 0 | 1.023 | 0 | 2.486 | 0 | 0 | 0 | 2847.722 | 0 | 0.571 | 0.21 | 1.277 | 0 | 0.333 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 12:44
WO-22270-209-A-1 | 0 | 374.189 | 622.034 | 188.453 | 0.199 | 133.364 | 4.296 | 7487.394 | 2.036 | 0.816 | 0 | 6.395 | 188.347 | 0 | 0.794 | 0 | 6.272 | 0 | 0 | 0 | 2860.632 | 0 | 0.812 | 0.099 | 2.189 | 0 | 0.312 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 13:20
WO-22270-210-A-1 | 0 | 326.197 | 744.226 | 371.848 | 0.175 | 0 | 0 | 7969.456 | 2.025 | 0.593 | 0 | 1.962 | 187.818 | 0 | 0.609 | 0 | 6.031 | 0 | 0 | 0 | 2869.4 | 0 | 0.526 | 0.131 | 1.151 | 0 | 0.312 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 13:56
AB-220715-01 LRB 22_INJ001 | 0 | 100.5 | 122.105 | -135.885 | 0.183 | 126.266 | 4.268 | 22.399 | 1.076 | 0.345 | 0 | 1.786 | 187.615 | 0 | 0.884 | 0 | 0.462 | 0 | 0 | 0 | 8.449 | 0 | 0.737 | 0.085 | 0.813 | 0 | 0.311 | -0.51 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 14:32
AB-220715-01 LRB 22 | 0 | 113.172 | 309.636 | 568.275 | 0.205 | 0 | 4.39 | 6.275 | 0.799 | 0.441 | 0 | 2.425 | 187.634 | 0 | 0.899 | 0 | 0.496 | 0 | 0 | 1.177 | 3.437 | 0 | 0.701 | 0.152 | 1.067 | 0 | 0.338 | 0.009 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 15:08
WO-22270-211-A-1 | 0 | 208.129 | 417.085 | -215.488 | 0.175 | 57.754 | 4.326 | 7781.393 | 2.3 | 0.545 | 0 | 1.99 | 187.851 | 0 | 0.823 | 0 | 3.187 | 0 | 0 | 0 | 2748.222 | 0 | 0.51 | 0.055 | 0.993 | 0 | 0.373 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 15:44
WO-22270-212-A-1 | 0 | 262.063 | -15.73 | -295.261 | 0.177 | 0 | 4.431 | 7105.418 | 3.041 | 0.755 | 0 | 5.603 | 188.206 | 0 | 0.926 | 0 | 4.309 | 0 | 0 | 0 | 2676.981 | 0 | 0.59 | 0.108 | 1.958 | 0 | 0.312 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 16:20
WO-22270-213-A-1 | 0 | 200.2 | 1102.797 | -263 | 0.175 | 0 | 0 | 7173.014 | 2.359 | 0.615 | 0 | 2.514 | 187.737 | 0 | 0.813 | 0 | 4.681 | 0 | 0 | 0 | 2778.734 | 0 | 0.677 | 0.09 | 1.614 | 0 | 0.296 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 16:56
WO-22270-214-A-1 | 0 | 385.503 | 2133.271 | 89.812 | 0.164 | 68.858 | 0 | 12554.794 | 2.749 | 0.677 | 0 | 1.131 | 188.393 | 0 | 0.489 | 0 | 14.656 | 0 | 0 | 0 | 3346.883 | 0 | 1.638 | 0.046 | 1.669 | 0 | 0.343 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 17:32
WO-22270-215-A-1 | 0 | 969.521 | 2632.211 | 200.106 | 0.167 | 0 | 4.357 | 12151.767 | 4.025 | 0.661 | 0 | 0.993 | 189.071 | 0 | 0.527 | 0 | 2.926 | 0 | 0 | 1.526 | 3235.902 | 0 | 0.805 | 0.034 | 1.903 | 0 | 0.34 | 27.392 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 18:08
WO-22270-216-A-1 | 0 | 380.126 | 506.498 | 92.728 | 0.175 | 0 | 0 | 7184.12 | 3.068 | 0.528 | 0 | 0.951 | 188.512 | 0 | 0.902 | 0 | 8.092 | 0 | 0 | 0 | 3348.427 | 0 | 1.048 | 0.072 | 2.498 | 0 | 0.292 | 33.383 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 18:44
WO-22270-217-A-1 | 0 | 348.426 | 263.928 | 916.392 | 0.176 | 0 | 0 | 7307.135 | 2.209 | 0.754 | 0 | 2.409 | 0 | 0 | 0.834 | 0 | 4.849 | 0 | 0 | 0 | 3450.62 | 0 | 0.919 | 0.117 | 2.127 | 0 | 0.292 | 9.253 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 19:20
WO-22270-218-A-1 | 0 | 228.452 | 265.157 | 233.356 | 0.168 | 0 | 4.291 | 2491.878 | 1.585 | 0.584 | 0 | 1.886 | 187.632 | 0 | 0.949 | 0 | 2.803 | 0 | 0 | 0 | 1141.632 | 0 | 0.649 | 0.058 | 0.719 | 0 | 0.363 | 0 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 19:56
WO-22270-219-A-1 | 0 | 409.526 | 2373.125 | 373.959 | 0.17 | 286.527 | 0 | 2772.695 | 1.36 | 0.388 | 0 | 1.051 | 187.822 | 0 | 0.655 | 0 | 1.586 | 0 | 0 | 0 | 1133.005 | 0 | 0.554 | 0.045 | 1.084 | 0 | 0.315 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 20:32
WO-22270-220-A-1 | 0 | 472.426 | 988.494 | -230.532 | 0.168 | 0 | 0 | 2530.797 | 1.044 | 0.612 | 0 | 1.058 | 188.204 | 0 | 0.831 | 0 | 2.469 | 0 | 0 | 0 | 1166.087 | 0 | 0.508 | 0.052 | 0.644 | 0 | 0.306 | 17.684 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 21:08
AB-220715-01 LFB-H 12 | 0 | 169.776 | 193948.501 | 1075322.799 | 182.539 | 1827906.594 | 1835.054 | 207.993 | 165.072 | 172.498 | 145.325 | 152.715 | 381.404 | 0 | 197.609 | 144.503 | 184.98 | 0 | 0 | 190.744 | 182.251 | 0 | 192.991 | 125.822 | 186.411 | 0 | 192.954 | 180.173 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 21:44
AB-220715-01 LRB 23_INJ001 | 0 | 96.558 | 113.309 | 6826.043 | 0.917 | 7421.516 | 12.544 | 1.639 | 1.162 | 0.967 | 1.448 | 1.597 | 188.309 | 0 | 1.086 | 0 | 1.255 | 0 | 0 | 1.25 | 2.016 | 0 | 1.078 | 0.504 | 1.26 | 0 | 1.566 | -0.91 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 22:20
AB-220715-01 LRB 23 | 0 | 93.238 | -31.758 | 0 | 0.271 | 2951.116 | 5.326 | 6.134 | 0.726 | 0.347 | 0 | 2.333 | 187.589 | 0 | 1.155 | 0 | 0.472 | 0 | 0 | 1.286 | 2.845 | 0 | 0.425 | 0.224 | 0.779 | 0 | 0.543 | 1.188 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-24 22:56
AB-220715-01 LD 11 | 0 | 205.586 | 1037.342 | -173.164 | 0.186 | 160.869 | 5.301 | 2421.487 | 1.08 | 0.656 | 4.165 | 0.907 | 187.723 | 0 | 0.736 | 0 | 1.575 | 0 | 0 | 1.654 | 1091.661 | 0 | 0.555 | 0.161 | 1.364 | 0 | 0.422 | 0 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-24 23:32
AB-220715-01 CCC-H 12 | 0 | 387.532 | 368562.575 | 2167249.335 | 363.821 | 3658441.908 | 3538.283 | 418.533 | 330.585 | 347.972 | 310.509 | 319.683 | 615.692 | 0 | 410.172 | 337.382 | 381.938 | 0 | 0 | 382.884 | 384.909 | 0 | 390.787 | 242.254 | 386.724 | 0 | 392.201 | 374.953 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-25 00:08
AB-220715-01 LFSM-H 11 | 0 | 446.832 | 189558.942 | 1120806.553 | 199.866 | 3254113.065 | 2898.003 | 2977.339 | 165.797 | 216.358 | 144.818 | 151.821 | 367.706 | 0 | 201.771 | 0 | 194.326 | 0 | 0 | 194.698 | 1328.5 | 0 | 194.899 | 141.519 | 194.494 | 0 | 211.115 | 211.572 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-25 00:44
WO-22270-221-A-1 | 0 | 177.795 | 1144.229 | 605.924 | 0.247 | 753.691 | 5.817 | 2681.881 | 1.511 | 0.588 | 0 | 1.66 | 187.952 | 0 | 0.898 | 0 | 1.787 | 0 | 0 | 0 | 1085.099 | 0 | 0.525 | 0.109 | 0.693 | 0 | 0.579 | 4.277 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-25 01:20
WO-22270-222-A-1 | 0 | 261.145 | 568.837 | 1030.072 | 0.227 | 102.121 | 5.862 | 2794.788 | 1.5 | 0.739 | 0 | 1.203 | 187.81 | 0 | 0.89 | 0 | 1.519 | 0 | 0 | 0 | 1067.661 | 0 | 0.698 | 0.168 | 2.007 | 0 | 0.466 | 2.197 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-25 01:56
WO-22270-223-A-1 | 0 | 157.285 | 760.486 | 75.86 | 0.209 | 482.191 | 4.712 | 2436.532 | 0.92 | 0.6 | 0 | 2.83 | 0 | 0 | 0.883 | 0 | 1.186 | 0 | 0 | 0 | 1051.989 | 0 | 0.721 | 0.316 | 1.524 | 0 | 0.405 | 26.471 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-25 02:32
WO-22270-224-A-1 | 0 | 205.76 | -95.656 | -210.894 | 0.183 | 0 | 5.082 | 2567.789 | 1.157 | 0.579 | 0 | 0.94 | 188.189 | 0 | 0.929 | 0 | 2.39 | 0 | 0 | 0 | 1132.739 | 0 | 0.741 | 0.235 | 1.089 | 0 | 0.403 | 10.66 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-25 03:08
WO-22270-225-A-1 | 0 | 190.438 | 492.339 | -40.619 | 0.193 | 0 | 4.38 | 2686.931 | 0.964 | 0.575 | 0 | 0.968 | 188.426 | 0 | 0.949 | 0 | 2.002 | 0 | 0 | 0 | 1066.892 | 0 | 0.6 | 0.198 | 1.694 | 0 | 0.348 | 28.228 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-25 03:44
WO-22270-226-A-1 | 0 | 422.537 | 465.862 | 446.896 | 0.177 | 0 | 4.497 | 2806.729 | 1.551 | 0.433 | 0 | 2.501 | 187.79 | 0 | 0.743 | 0 | 1.745 | 0 | 0 | 0 | 1116.773 | 0 | 0.686 | 0.162 | 1.358 | 0 | 0.342 | 24.547 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-25 04:20
AB-220715-01 LRB 24_INJ001 | 0 | 73.102 | -7.809 | 0 | 0.179 | 167.865 | 4.315 | 8.854 | 0.766 | 0.493 | 2.309 | 0.875 | 187.65 | 0 | 0.991 | 0 | 0.375 | 0 | 0 | 0 | 4.418 | 0 | 0.348 | 0.135 | 0.64 | 0 | 0.329 | 3.423 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-25 04:56
AB-220715-01 LRB 24 | 0 | 105.884 | -52.849 | 0 | 0.182 | 1755.915 | 4.396 | 4.17 | 0.849 | 0.323 | 2.368 | 2.338 | 187.764 | 0 | 1.382 | 0 | 0.408 | 0 | 0 | 0 | 2.162 | 0 | 0.373 | 0.196 | 0.668 | 0 | 0.343 | 3.801 | 0 | 0 | 0 | 0 | D,Q | Unknown | 2022-07-25 05:32
AB-220715-01 CCC-H 13 | 0 | 355.433 | 396434.361 | 2226765.266 | 384.528 | 4053095.978 | 3487.488 | 376.546 | 341.506 | 359.734 | 303.614 | 323.323 | 601.413 | 0 | 400.31 | 332.365 | 387.108 | 0 | 0 | 380.562 | 374.334 | 0 | 386.732 | 246.296 | 388.839 | 0 | 396.766 | 412.81 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-25 06:08
AB-220715-01 QCS 2 | 0 | 194.411 | 209001.638 | 1163656.531 | 197.955 | 31773.153 | 33.662 | 199.036 | 171.1 | 178.018 | 158.561 | 162.451 | 413.278 | 0 | 4.614 | 161.011 | 191.063 | 0 | 0 | 201.953 | 193.156 | 0 | 209.797 | 2.703 | 196.615 | 0 | 4.913 | 229.537 | 0 | 0 | 0 | 0 |  | Unknown | 2022-07-25 06:44

## Sheet3
Batch Folder Path: | D:\TraceFinderData\Projects\PFAS Online SPE\2022-07-15\2022-07-15-AB2
Batch Name: | 2022-07-15-AB2
Last Acquisition Date: | 2022-07-26 04:21
User Name: | Quantiva
Sample Count: | 342
Area - N/F
Data File | 11-Cl-PF3OUDS | 4-2:FTS | 6-2:FTS | 8-2:FTS | 9-Cl-PF3ONS | D3-N-Et-FOSA | D3-N-Me-FOSA | FBSA | FHxA | HFPO-DA | L-PFBS | L-PFHpS | L-PFHxS | L-PFOS | M2-4:2-FTS | M2-6:2-FTS | M2-8:2-FTS | M3PFBS | M3PFHxS | M4PFHpA | M5PFHxA | M5PFPeA | M6PFDA | M8FOSA | M8PFOA | M8PFOS | M9PFNA | MPFBA | MPFDoA | MPFUDA | MPHFPO-DA | NaDONA | N-Et-FOSA | N-Me-FOSA | PFBA | PFDA | PFDoA | PFDS | PFHpA | PFHxA | PFHxDA | PFNA | PFNS | PFOA | PFODA | PFOSA | PFPeA | PFPeS | PFTeDA | PFTrDA | PFUdA | All Flags | Sample Type
AB-220715-01 Cal Std 1 1 | 95 | 6213 | 2763 | 915 | 524122 | N/F | 7075600 | 7 | 289 | N/F | 6554 | 26553 | 17730 | 26106 | 382817 | 1259222 | 1174724 | 1500040 | 3726379 | 10295093 | 4123113 | 977043 | 40182026 | 8565344 | 22704030 | 4705476 | 27045316 | 61078 | N/F | N/F | 480815 | N/F | N/F | 25014 | N/F | 190897 | N/F | N/F | 44377 | N/F | 536 | 115171 | 27576 | 103420 | 262 | 40019 | N/F | N/F | N/F | N/F | 185918 |  | Cal Std
AB-220715-01 Cal Std 2 1-inj2 | 52 | 4913 | 2295 | 180 | 928988 | N/F | 5829551 | 33 | 113 | 7460 | 8915 | 44888 | 24764 | 37728 | 311496 | 1025592 | 1028027 | 1148968 | 2883380 | 8119618 | 3078995 | 667163 | 33735557 | 7151291 | 17285957 | 4017638 | 22129693 | 35290 | N/F | N/F | 340592 | N/F | 937 | 43990 | N/F | 313576 | N/F | N/F | 73448 | 32497 | 369 | 186714 | 49467 | 155101 | 591 | 67205 | N/F | N/F | 866 | N/F | 314726 |  | Cal Std
AB-220715-01 Cal Std 3 1-inj2 | 173 | 8737 | 2289 | 19 | 2121363 | N/F | 6448212 | 14 | 20 | 12535 | 25055 | 92775 | 59911 | 78842 | 368482 | 1173468 | 1208706 | 1320162 | 3410846 | 9092210 | 3674530 | 770168 | 39237316 | 8420819 | 20100191 | 4518567 | 25080673 | 61302 | N/F | N/F | 406880 | N/F | N/F | 92908 | N/F | 706805 | N/F | N/F | 162175 | 68123 | 392 | 411787 | 103173 | 351979 | 175 | 146648 | 18605 | N/F | 1367 | N/F | 713012 |  | Cal Std
AB-220715-01 Cal Std 4 1 | 25 | 19691 | 2321 | 229 | 5664493 | N/F | 6586324 | 38 | 239 | 22962 | 63914 | 247243 | 144987 | 204672 | 359043 | 1122525 | 1171502 | 1275747 | 3262781 | 9022223 | 3599898 | 761209 | 40566614 | 8618929 | 19694899 | 4610237 | 25586736 | 73743 | N/F | N/F | 433529 | N/F | 848 | 236908 | N/F | 1787920 | N/F | 102 | 404183 | 153251 | 614 | 1080476 | 254743 | 863277 | 195 | 398101 | 44231 | N/F | 1741 | N/F | 1921138 |  | Cal Std
AB-220715-01 Cal Std 5 1 | 124 | 35656 | 1936 | 393 | 10060677 | N/F | 6104773 | 6 | 107 | 54881 | 112967 | 468870 | 266562 | 377242 | 327106 | 993803 | 1047581 | 1277711 | 3040930 | 8449225 | 3438958 | 734516 | 36873395 | 8121496 | 18617801 | 4289270 | 23616240 | 61896 | N/F | N/F | 394758 | N/F | 1252 | 461955 | N/F | 3312805 | N/F | 44 | 745073 | 281629 | 182 | 2019699 | 491260 | 1642665 | 77 | 765304 | 66523 | N/F | N/F | N/F | 3699819 |  | Cal Std
AB-220715-01 Cal Std 6 1 | 322 | 62234 | 1686 | 318 | 19645314 | N/F | 5632639 | 18 | 252 | 96138 | 212537 | 904913 | 513805 | 738357 | 298805 | 859662 | 962279 | 1138115 | 2895638 | 7656165 | 3150523 | 650813 | 35342251 | 7952548 | 17085621 | 4084789 | 22429940 | 63255 | N/F | N/F | 364862 | N/F | N/F | 860111 | N/F | 6400349 | N/F | N/F | 1393973 | 520773 | 301 | 3897545 | 964466 | 3106867 | 147 | 1539889 | 116325 | N/F | N/F | N/F | 7041520 |  | Cal Std
AB-220715-01 Cal Std 7 1 | N/F | 131304 | 1884 | 309 | 46244797 | N/F | 5809988 | 141 | N/F | 195537 | 487488 | 1930638 | 1052109 | 1669615 | 272558 | 622253 | 802721 | 1063643 | 2332530 | 6471670 | 2802723 | 603781 | 31298744 | 7466092 | 13707584 | 3718322 | 19642767 | 59957 | N/F | N/F | 334185 | 693 | N/F | 2223243 | N/F | 14103619 | N/F | N/F | 2962245 | 1208414 | 876 | 8463022 | 2235651 | 6247722 | 603 | 3686998 | 276406 | N/F | N/F | N/F | 15851584 |  | Cal Std
AB-220715-01 Cal Std 8 1 | 1139 | 287535 | 2229 | 334 | 89630124 | N/F | 5458564 | 94 | 1037 | 513219 | 1165016 | 3990860 | 2627119 | 3454143 | 288602 | 535035 | 719536 | 1192670 | 2633040 | 7809084 | 3471177 | 736491 | 29182620 | 7627492 | 15825420 | 3565013 | 20097918 | 67846 | N/F | N/F | 413789 | N/F | N/F | 4301897 | 61805 | 26678034 | N/F | N/F | 7405418 | 3137571 | 490 | 17825771 | 4079827 | 14536933 | 260 | 7677723 | 701702 | N/F | N/F | N/F | 29022193 |  | Cal Std
AB-220715-01 Cal Std 9 1 | N/F | 564951 | 2428 | 961 | 153196733 | N/F | 5079317 | 434 | N/F | 1082805 | 2336604 | 6579164 | 4733522 | 6303031 | 260099 | 415680 | 587510 | 1143888 | 2346729 | 6988392 | 3374919 | 716558 | 24297545 | 7107332 | 13421321 | 3123532 | 17013245 | 67795 | N/F | N/F | 396794 | N/F | N/F | 8129682 | 122135 | 43922722 | N/F | 206 | 13154062 | 6136372 | 1266 | 30839060 | 6794109 | 24395888 | 187 | 14342738 | 1387916 | N/F | N/F | 1686 | 47813294 |  | Cal Std
AB-220715-01 Cal Std 10 1 | N/F | 1069114 | 1682 | 929 | 272337689 | N/F | 4263423 | 508 | N/F | 2370062 | 4957446 | 10968170 | 8604258 | 11439165 | 211972 | 299937 | 464932 | 1004844 | 1657654 | 5374532 | 2835104 | 687574 | 17020941 | 6016468 | 9637873 | 2277056 | 12925865 | 72321 | N/F | N/F | 365038 | 2730573 | N/F | 16081243 | 284315 | 73814926 | N/F | 126 | 24027588 | 11738574 | 699 | 53333890 | 11451789 | 42028054 | 592 | 29148262 | 3179868 | N/F | N/F | N/F | 77629549 |  | Cal Std
AB-220715-01 Cal Std 11 1 | N/F | 1730478 | 2047 | 1553 | 399561750 | N/F | 3354460 | N/F | 11812 | 4122460 | 7985988 | 15608779 | 12929320 | 17660982 | 159048 | 243138 | 372285 | 763538 | 1116701 | 3718885 | 2281360 | 605950 | 11373297 | 4704493 | 6720492 | 1590397 | 8862935 | 65082 | N/F | N/F | 318901 | 5733955 | N/F | 26077756 | 560376 | 102250983 | N/F | 346 | 33821346 | 20092953 | 388 | 77032869 | 15636905 | 60007246 | 60 | 46882676 | 5845069 | N/F | N/F | N/F | 106223607 |  | Cal Std
AB-220715-01 Cal Std 12 1 | N/F | 2774998 | 4401 | 1395 | 555868075 | N/F | 2597638 | N/F | 40248 | 6442018 | 13129914 | 22210774 | 19072499 | 26960113 | 127641 | 205008 | 355760 | 580441 | 694484 | 2448830 | 1574116 | 497373 | 7246388 | 3762283 | 4250243 | 1079609 | 5789422 | 69192 | N/F | N/F | 237441 | 8218494 | N/F | 41856400 | 1225785 | 134355522 | N/F | 698 | 47210521 | 31244094 | 786 | 105607311 | 22074355 | 79095158 | 89 | 77083572 | 10081286 | N/F | N/F | N/F | 141216584 |  | Cal Std
AB-220715-01 Cal Std 13 1 | N/F | 4633637 | 8065 | 2020 | 776131967 | N/F | 1701861 | N/F | 10161500 | 9556766 | 19903793 | 31089706 | 26533890 | 39921460 | 84930 | 148938 | 284756 | 356739 | 387630 | 1224273 | 906084 | 314637 | 3649764 | 2432014 | 2058190 | 637062 | 3012684 | 54188 | N/F | N/F | 145507 | 19070234 | N/F | 64842405 | 2153714 | 175142790 | N/F | 941 | 60950502 | 44040422 | 271 | 136404305 | 30754577 | 99632547 | 36 | 121840134 | 16911237 | N/F | N/F | N/F | 188950269 |  | Cal Std
AB-220715-01 LFB-H 1 | N/F | 306023 | 752 | 1140 | 90979037 | N/F | 5024206 | N/F | 5821507 | 571321 | 1245558 | 3953981 | 2532598 | 3325492 | 297230 | 499166 | 638933 | 1258466 | 2609141 | 8250258 | 3578022 | 802267 | 28986281 | 7353031 | 15840003 | 3363767 | 20165202 | 72806 | N/F | N/F | 433139 | 2571376 | N/F | 3957377 | 64087 | 26690791 | N/F | N/F | 7649509 | 3387594 | 815 | 17815079 | 4012441 | 14589004 | 81 | 7584539 | 762027 | N/F | N/F | N/F | 29988087 |  | Unknown
AB-220715-01 LRB 1_INJ001 | N/F | 5883 | 1497 | 1008 | 490886 | N/F | 5235716 | N/F | 20670 | N/F | N/F | 15841 | N/F | 15616 | 350313 | 951039 | 935148 | 1346465 | 3154476 | 9585402 | 3827595 | 835320 | 37971765 | 8062435 | 19792477 | 4180363 | 24084153 | 56360 | N/F | N/F | 442707 | N/F | N/F | 21739 | N/F | 140374 | N/F | N/F | 20321 | 19069 | 631 | 73099 | 21021 | 99940 | 269 | 58030 | 7459 | N/F | N/F | N/F | 164260 |  | Unknown
AB-220715-01 LRB 1 | N/F | 5717 | 1492 | 1063 | 352036 | N/F | 4304348 | N/F | 7532 | N/F | N/F | 20944 | N/F | 18098 | 248061 | 833906 | 816597 | 889053 | 2553652 | 7187155 | 2587113 | 603704 | 31479184 | 6387473 | 15886293 | 3533089 | 19918783 | 30308 | N/F | N/F | 286866 | 8110 | N/F | 16392 | N/F | 135860 | N/F | N/F | 19860 | 29221 | 799 | 64940 | 17867 | 77820 | 733 | 36372 | 7872 | N/F | N/F | N/F | 135821 |  | Unknown
AB-220715-01 CCC-H 1 | N/F | 540432 | N/F | 1557 | 150533396 | N/F | 4418553 | N/F | 10215677 | 1072099 | 2238402 | 6389865 | 4414626 | 5777483 | 252322 | 373276 | 513966 | 1075641 | 2144318 | 6768580 | 3215096 | 702667 | 23483308 | 6577757 | 12833415 | 2836570 | 16388396 | 62643 | N/F | N/F | 403694 | 4442046 | N/F | 7074440 | 111012 | 43877030 | N/F | 820 | 13296279 | 5904108 | 649 | 29538359 | 6527615 | 23712487 | 321 | 13527594 | 1387870 | N/F | 1389 | N/F | 47427269 |  | Unknown
AB-220715-01 IBC 1 | N/F | 288619 | 1218 | 439 | 89307148 | N/F | 5246563 | N/F | 62883 | 566948 | 1249865 | 3888308 | 2638845 | 3432894 | 289657 | 473205 | 731993 | 1220172 | 2572751 | 8039672 | 3566735 | 766663 | 31908879 | 7580937 | 15612554 | 3393225 | 20318022 | 68645 | N/F | N/F | 434362 | 2610301 | N/F | 57196 | 62437 | 29237374 | N/F | N/F | 7874973 | 3269057 | 593 | 18440552 | 88210 | 14431841 | 548 | 94678 | 731729 | N/F | N/F | N/F | 34215051 |  | Unknown
AB-220715-01 QCS 1 | N/F | 5436 | 2697 | 451 | 1263554 | N/F | 5520924 | N/F | N/F | 430535 | 1260884 | 3925510 | 2714381 | 3432100 | 329456 | 526558 | 725762 | 1209903 | 2669141 | 8244516 | 3427631 | 757358 | 29652036 | 7715934 | 15921519 | 3375356 | 20459537 | 66107 | N/F | N/F | 424097 | 24623 | N/F | 23853 | 56967 | 26572877 | N/F | N/F | 7856533 | 3051759 | 507 | 18195884 | 4185892 | 14548757 | 414 | 58947 | 731269 | N/F | N/F | N/F | 31256474 |  | Unknown
WO-22270-01-A-1 | N/F | 10876 | 1062 | 639 | 80893 | N/F | 4669195 | 459 | N/F | N/F | N/F | 8640 | N/F | 7119 | 78887 | 413775 | 1018041 | 192834 | 908010 | 2494439 | 647220 | 63563 | 21056090 | 4971053 | 5930160 | 2506050 | 10683728 | 8248 | N/F | N/F | 67698 | N/F | N/F | 15106 | N/F | 426580 | N/F | 1909 | N/F | N/F | 1239 | 50034 | 5451 | 60896 | 853 | 9747 | N/F | N/F | N/F | N/F | 33126 |  | Unknown
WO-22270-02-A-1 | N/F | 7662 | 1077 | 414 | 234828 | N/F | 5247807 | 496 | N/F | N/F | N/F | 6113 | N/F | 9273 | 94003 | 449164 | 1190668 | 236621 | 1029128 | 2874284 | 786145 | 84736 | 23902714 | 6119673 | 6874699 | 2904969 | 12208610 | 9251 | N/F | N/F | 83138 | N/F | N/F | 12905 | N/F | 439266 | N/F | 1337 | N/F | N/F | 303 | 30694 | 3632 | 76246 | 1114 | 8672 | N/F | N/F | N/F | N/F | 33909 |  | Unknown
WO-22270-03-A-1 | N/F | 16616 | 229 | 333 | 201594 | N/F | 5593529 | 363 | N/F | N/F | N/F | 3310 | N/F | 6154 | 89055 | 440861 | 1113762 | 238397 | 1021694 | 2845686 | 797719 | 87570 | 24275846 | 5818673 | 6865251 | 2919310 | 12613090 | 10607 | N/F | N/F | 81746 | N/F | N/F | 14538 | N/F | 419329 | N/F | 1742 | N/F | N/F | 1158 | 27931 | 3843 | 35972 | 756 | 7450 | N/F | N/F | N/F | N/F | 34882 |  | Unknown
WO-22270-04-A-1 | N/F | 11347 | 1274 | 417 | 59421 | N/F | 5108541 | 338 | 1200 | N/F | N/F | N/F | N/F | 7378 | 81390 | 411752 | 1036812 | 213958 | 971558 | 2661510 | 700096 | 71859 | 22475000 | 5665387 | 6380733 | 2739134 | 11465262 | 8158 | N/F | N/F | 78456 | N/F | N/F | 14253 | N/F | 396647 | N/F | 1298 | N/F | N/F | 1355 | 31531 | 4409 | 38856 | 53 | 5810 | N/F | N/F | 748 | N/F | 26437 |  | Unknown
WO-22270-05-A-1 | N/F | 15709 | 444 | 268 | 69855 | N/F | 5293106 | 407 | N/F | N/F | N/F | N/F | N/F | 9272 | 96706 | 419242 | 1044860 | 256964 | 1055212 | 2888134 | 832899 | 88943 | 24352740 | 6533380 | 6786453 | 2931235 | 12838486 | 9053 | N/F | N/F | 89911 | N/F | N/F | 13199 | N/F | 514693 | N/F | 1364 | 15733 | N/F | 473 | 25115 | 17247 | 90859 | 141 | 5437 | N/F | N/F | N/F | N/F | 32769 |  | Unknown
WO-22270-06-A-1 | N/F | 16683 | 634 | 222 | 130221 | N/F | 5005474 | 326 | N/F | 2086 | N/F | 3390 | N/F | 9636 | 82398 | 391009 | 970489 | 217242 | 979994 | 2656633 | 709091 | 80012 | 23800787 | 6200326 | 6504078 | 2718954 | 12100289 | 10949 | N/F | N/F | 73672 | N/F | N/F | 14258 | N/F | 411032 | N/F | 1883 | N/F | N/F | 796 | 45206 | 2197 | 88619 | 428 | 5388 | N/F | N/F | 1042 | 1145 | 23522 |  | Unknown
WO-22270-07-A-1 | N/F | 10525 | 666 | 544 | 35374 | N/F | 5229662 | 232 | 1578 | N/F | N/F | N/F | N/F | 3377 | 85099 | 392957 | 972787 | 226617 | 997478 | 2815174 | 774955 | 84223 | 23786530 | 5672081 | 6580585 | 2727936 | 12143878 | 8642 | N/F | N/F | 85596 | N/F | N/F | 14424 | N/F | 408982 | N/F | 1358 | N/F | N/F | 1267 | 14597 | 2391 | 63181 | 431 | 3820 | N/F | N/F | 579 | N/F | 24166 |  | Unknown
WO-22270-08-A-1 | N/F | 9358 | 669 | 209 | 109032 | N/F | 4550828 | 404 | 899 | N/F | N/F | 4195 | N/F | 4473 | 66226 | 350324 | 924377 | 174942 | 860282 | 2259842 | 574949 | 60247 | 20956659 | 4674506 | 5709385 | 2416625 | 10567001 | 6088 | N/F | N/F | 64185 | N/F | N/F | 9200 | N/F | 413652 | N/F | 918 | N/F | N/F | 649 | 28784 | 13801 | 69581 | 208 | 3813 | N/F | N/F | 1167 | N/F | 17907 |  | Unknown
WO-22270-09-A-1 | N/F | 13998 | 604 | 210 | 49896 | 1762 | 5580216 | 128 | 422 | N/F | N/F | 3354 | N/F | 28643 | 93119 | 412521 | 1084480 | 252013 | 1014121 | 2867767 | 841657 | 88334 | 25413825 | 5993878 | 6892188 | 2912696 | 12991287 | 8956 | N/F | N/F | 91451 | N/F | N/F | 13129 | N/F | 594363 | N/F | 1650 | N/F | N/F | 812 | 44806 | 11091 | 49396 | 161 | 5664 | N/F | N/F | 1001 | N/F | 20985 |  | Unknown
WO-22270-10-A-1 | N/F | 13130 | 552 | 574 | 25908 | N/F | 5359057 | 153 | 823 | N/F | N/F | 3297 | N/F | 5212 | 84521 | 421836 | 1099498 | 241047 | 1015479 | 2731655 | 763305 | 76864 | 24903677 | 5851126 | 6737953 | 2862004 | 12581059 | 7559 | N/F | N/F | 84932 | N/F | 933 | 12582 | N/F | 714121 | N/F | 1134 | N/F | N/F | 655 | 31778 | 10246 | 81511 | 894 | 3089 | N/F | N/F | N/F | N/F | 19629 |  | Unknown
AB-220715-01 LRB 2_INJ001 | N/F | 7792 | 872 | 381 | 116350 | N/F | 4845215 | 40 | 1730 | N/F | N/F | 2158 | 2252 | 7052 | 289916 | 924230 | 919673 | 1198731 | 2987314 | 8874445 | 3489215 | 680960 | 36852981 | 7601214 | 18792228 | 4077662 | 23446218 | 55088 | 860 | N/F | 399197 | N/F | 509 | 14823 | N/F | 42664 | N/F | 196 | N/F | N/F | 929 | 49176 | 5913 | 116728 | 322 | 3564 | 7090 | N/F | 1010 | N/F | 63295 |  | Unknown
AB-220715-01 LRB 2 | N/F | 10681 | 1446 | 49 | 116186 | N/F | 4899378 | 57 | 1401 | N/F | N/F | 2066 | N/F | 9981 | 321096 | 834721 | 858736 | 1203801 | 2966182 | 9035214 | 3479659 | 716768 | 36367310 | 7678793 | 18880621 | 3931466 | 23536400 | 63071 | N/F | N/F | 408133 | N/F | N/F | 14942 | N/F | 29212 | N/F | 87 | N/F | N/F | 811 | 59295 | 1830 | 114871 | 539 | 5763 | 8880 | N/F | 2366 | N/F | 61780 |  | Unknown
WO-22270-11-A-1 | N/F | 8293 | 224 | 397 | 89285 | N/F | 5345726 | 33 | N/F | N/F | N/F | 7019 | N/F | 23095 | 89283 | 389337 | 1037725 | 243227 | 973170 | 2826319 | 752337 | 78960 | 24750178 | 5748661 | 6499004 | 2756593 | 12415592 | 7880 | N/F | 1106 | 86169 | 1425 | 1294 | 14976 | N/F | 502631 | N/F | 1522 | N/F | N/F | 215 | 13437 | 9939 | 25023 | 163 | 3392 | N/F | N/F | N/F | 717 | 23088 |  | Unknown
WO-22270-12-A-1 | N/F | 13090 | 498 | 337 | 18192 | N/F | 5119672 | 153 | 943 | N/F | N/F | 2187 | N/F | 4687 | 80508 | 358982 | 1038679 | 224022 | 943192 | 2643229 | 751211 | 78407 | 23411615 | 5545342 | 6244197 | 2676764 | 11980648 | 14360 | 528 | N/F | 78773 | N/F | 1087 | 12745 | N/F | 312938 | N/F | 489 | 8140 | N/F | 1094 | 11503 | 1351 | 65549 | 267 | 3340 | N/F | N/F | N/F | N/F | 18823 |  | Unknown
WO-22270-13-A-1 | N/F | 8954 | 610 | 415 | 61257 | N/F | 4780860 | 123 | 570 | 881 | N/F | 2440 | 245 | 3344 | 69047 | 348247 | 913652 | 163457 | 806418 | 2210712 | 552437 | 58929 | 20831902 | 4775930 | 5559627 | 2342145 | 10373051 | 11567 | N/F | N/F | 60443 | N/F | 779 | 15684 | N/F | 390873 | N/F | 1044 | N/F | N/F | 868 | 12824 | 7422 | N/F | 131 | 4218 | N/F | 1370 | 633 | N/F | 17660 |  | Unknown
WO-22270-14-A-1 | N/F | 7551 | 445 | 65 | 11869 | 711 | 5157693 | 191 | 654 | N/F | N/F | 9344 | N/F | 3523 | 86344 | 353793 | 921153 | 224289 | 957616 | 2691158 | 756615 | 83915 | 24688316 | 5599538 | 6661242 | 2682247 | 12409492 | 9159 | N/F | N/F | 86858 | N/F | 642 | 18848 | N/F | 453618 | 1177 | 1671 | N/F | N/F | 569 | 7285 | 8209 | N/F | 63 | 3337 | N/F | N/F | 1432 | N/F | 15441 |  | Unknown
WO-22270-15-A-1 | N/F | 10570 | 611 | 157 | 26733 | N/F | 4918455 | 123 | 739 | N/F | N/F | 4385 | N/F | 21942 | 75145 | 391737 | 1082151 | 215373 | 945971 | 2624491 | 722193 | 75596 | 23345937 | 5491737 | 6359552 | 2761557 | 12095536 | 7200 | 2104 | 735 | 77998 | 680 | 975 | 18252 | N/F | 247490 | N/F | 1289 | N/F | N/F | 412 | 28213 | 8623 | 78161 | 76 | 3445 | N/F | N/F | N/F | N/F | 19585 |  | Unknown
WO-22270-16-A-1 | N/F | 15600 | 804 | 217 | 18424 | 1813 | 5130370 | 213 | 1325 | N/F | N/F | 4153 | N/F | 5732 | 88173 | 369409 | 1013619 | 253091 | 1007387 | 2898935 | 866499 | 95847 | 25899008 | 5731327 | 6896612 | 2856034 | 12993054 | 9043 | N/F | N/F | 93335 | N/F | 977 | 16026 | N/F | 538801 | 2027 | 2341 | N/F | N/F | 1087 | 25661 | 10556 | 67306 | 425 | 3098 | N/F | N/F | N/F | N/F | 13617 |  | Unknown
WO-22270-17-A-1 | N/F | 9187 | 369 | 154 | 33247 | N/F | 4454366 | 25 | 594 | N/F | N/F | 3686 | N/F | 3024 | 57500 | 291558 | 746674 | 212397 | 874969 | 2515303 | 702222 | 73140 | 22950166 | 5062689 | 6171068 | 2464575 | 11191368 | 8750 | 1027 | N/F | 70790 | 1625 | 1065 | 14121 | N/F | 466056 | N/F | 1867 | 7328 | N/F | 354 | 33885 | 1042 | 102402 | 362 | 2266 | N/F | N/F | N/F | N/F | 15425 |  | Unknown
WO-22270-18-A-1 | N/F | 8469 | 626 | 257 | 16882 | N/F | 4531250 | 43 | 614 | N/F | N/F | 7784 | 2915 | 2998 | 47793 | 262827 | 722196 | 209350 | 862988 | 2453817 | 684538 | 76786 | 23325324 | 5075605 | 5952927 | 2473151 | 11429261 | 7583 | N/F | N/F | 71362 | 2196 | N/F | 19187 | N/F | 181072 | N/F | 1487 | N/F | N/F | 1066 | 7863 | 9075 | 31050 | 612 | 2670 | N/F | N/F | 614 | N/F | 14008 |  | Unknown
WO-22270-19-A-1 | N/F | 7755 | 454 | 58 | 21264 | N/F | 4783308 | 174 | 631 | 1063 | N/F | 6718 | N/F | 3632 | 54380 | 279775 | 818919 | 215765 | 877577 | 2635070 | 726015 | 83001 | 24026812 | 5309774 | 6127998 | 2551859 | 11774606 | 7571 | 1177 | N/F | 72915 | N/F | 662 | 14857 | N/F | 364715 | N/F | 1216 | N/F | N/F | 1737 | 12722 | 9538 | 44323 | 674 | 2721 | N/F | N/F | N/F | 847 | 16054 |  | Unknown
WO-22270-20-A-1 | N/F | 14231 | 102 | 55 | 18679 | N/F | 4461897 | 21 | 253 | 1630 | N/F | 6807 | N/F | 4907 | 48444 | 262537 | 778530 | 187792 | 807537 | 2560391 | 709476 | 78552 | 22800078 | 4996141 | 6092727 | 2279685 | 11580470 | 9344 | N/F | N/F | 75300 | 1543 | 348 | 12995 | N/F | 524991 | N/F | 1597 | N/F | N/F | 1492 | 10690 | 8208 | 61642 | 486 | 3346 | N/F | 984 | N/F | N/F | 13574 |  | Unknown
AB-220715-01 LFB-H 2 | 896 | 252043 | 245983 | 758 | 79704811 | 966 | 4443096 | 1504 | 5713070 | 509179 | 1042375 | 3606648 | 2316370 | 2958261 | 250243 | 440030 | 573433 | 1098509 | 2424591 | 7886486 | 3271639 | 712883 | 28517778 | 6848818 | 15597174 | 3134823 | 20010453 | 64930 | N/F | N/F | 394898 | 2406724 | 579 | 3538610 | 50530 | 25872300 | N/F | 41 | 7260851 | 2912201 | 356 | 17230755 | 3689535 | 14002239 | 883 | 6864001 | 643839 | N/F | 985 | 778 | 29871586 |  | Unknown
AB-220715-01 LRB 3_INJ001 | N/F | 13673 | 1821 | 578 | 130012 | 1555 | 4899041 | 199 | 7686 | N/F | N/F | 4746 | 3717 | 9276 | 313332 | 833076 | 836168 | 1180945 | 2944018 | 9321519 | 3688857 | 769517 | 38568092 | 7941740 | 20033002 | 4024004 | 24842549 | 73454 | N/F | N/F | 413998 | N/F | 697 | 14953 | N/F | 58750 | N/F | 303 | N/F | 5896 | 1561 | 27444 | 7128 | 81119 | 206 | 19873 | 6066 | N/F | 1069 | N/F | 60331 |  | Unknown
AB-220715-01 LRB 3 | N/F | 8076 | 1232 | 384 | 36846 | 1573 | 4314636 | 314 | 2716 | N/F | N/F | 7580 | 1320 | 4748 | 295515 | 777538 | 771788 | 1175836 | 2801679 | 8854880 | 3571226 | 739760 | 34668358 | 7252767 | 18187751 | 3696160 | 22675639 | 36802 | 1242 | N/F | 416397 | N/F | 1166 | 14368 | N/F | 33155 | N/F | 300 | N/F | 4629 | 796 | 21535 | 10534 | 71098 | 277 | 5269 | 4284 | 611 | 1173 | N/F | 39352 |  | Unknown
AB-220715-01 LD 1 | 1822 | 12574 | 641 | 294 | 18370 | N/F | 4562558 | 168 | 864 | 1525 | N/F | 3294 | 5362 | 3262 | 60389 | 284936 | 795556 | 229663 | 936028 | 2787213 | 798127 | 88657 | 24291523 | 5436427 | 6343479 | 2637392 | 12213570 | 9334 | N/F | N/F | 89095 | N/F | 652 | 15361 | N/F | 368248 | N/F | 1922 | 7588 | N/F | 1552 | 20805 | 4112 | 77477 | 471 | 3844 | N/F | 446 | 925 | N/F | 16797 |  | Unknown
AB-220715-01 CCC-H 2 | N/F | 467853 | 604647 | 14065 | 127515851 | 1591 | 3998196 | 1997433 | 9777030 | 972544 | 1925808 | 5482636 | 3910010 | 5275461 | 222585 | 309806 | 449591 | 936774 | 1927568 | 6825645 | 3073192 | 632809 | 22522730 | 6363199 | 12833118 | 2652217 | 16922867 | 50312 | 436 | N/F | 368167 | 4254599 | 588 | 6219407 | 81297 | 40824483 | N/F | 103 | 12664272 | 5568877 | 1336 | 29431341 | 5608014 | 22916200 | 504 | 12394320 | 1221444 | 875 | 1869 | N/F | 47313622 |  | Unknown
AB-220715-01 LFSM-H 1 | N/F | 55498 | 186178 | 19779 | 55437628 | 943 | 3964591 | 238038 | 2037615 | 102854 | 209664 | 1530943 | 783331 | 1975351 | 50831 | 186872 | 499014 | 216376 | 821320 | 2530217 | 709783 | 84115 | 20039375 | 4783404 | 5718369 | 2247881 | 10492766 | 11665 | N/F | N/F | 82309 | 774935 | 434 | 3080721 | N/F | 18488225 | 4044 | 1309 | 2384371 | 656296 | 976 | 9162237 | 3057498 | 5229081 | 161 | 4879962 | 76786 | 1065 | N/F | N/F | 23144799 |  | Unknown
WO-22270-21 | N/F | 7756 | 621 | 425 | 94821 | N/F | 4426569 | 388 | 3352 | 1582 | N/F | 9139 | N/F | 5813 | 60704 | 284753 | 800756 | 195198 | 830131 | 2504947 | 654006 | 63196 | 21910933 | 4706025 | 6014678 | 2358870 | 10956068 | 6573 | N/F | N/F | 74382 | 2480 | 1059 | 12609 | N/F | 368962 | N/F | 415 | 13649 | N/F | 556 | 20296 | 6359 | 116215 | 412 | 12785 | N/F | N/F | N/F | 1459 | 51124 |  | Unknown
WO-22270-22 | N/F | 11559 | 565 | 121 | 43801 | N/F | 4818465 | 96 | 1163 | N/F | N/F | 7992 | N/F | 2817 | 63535 | 313244 | 830640 | 232969 | 951987 | 2782723 | 782289 | 76751 | 24552358 | 5243368 | 6548169 | 2605651 | 12533477 | 10036 | N/F | N/F | 85811 | N/F | 543 | 13336 | N/F | 346931 | N/F | 1911 | N/F | N/F | 1486 | 20623 | 4005 | 48181 | 525 | 7928 | N/F | N/F | N/F | N/F | 22030 |  | Unknown
WO-22270-23 | N/F | 14379 | 345 | 878 | 28925 | N/F | 4403653 | 132 | 1688 | N/F | N/F | 7117 | N/F | 4054 | 66999 | 309178 | 787061 | 228669 | 951131 | 2828235 | 774613 | 88796 | 23873966 | 5287515 | 6553624 | 2536246 | 12070045 | 10071 | N/F | N/F | 83242 | N/F | 832 | 14023 | N/F | 432762 | N/F | 1704 | 6310 | 4808 | 2978 | 21653 | 8951 | 94240 | 542 | 3833 | N/F | N/F | N/F | N/F | 20387 |  | Unknown
WO-22270-24 | N/F | 11612 | 782 | 43 | 23402 | N/F | 4397284 | 73 | 1119 | N/F | N/F | 7938 | N/F | 5494 | 67265 | 290184 | 793912 | 205118 | 864524 | 2532880 | 697914 | 75091 | 22565786 | 4732208 | 6063512 | 2380645 | 11585880 | 8155 | N/F | N/F | 75539 | N/F | 524 | 13513 | N/F | 333513 | 3414 | 2587 | N/F | N/F | 1872 | 17721 | 5086 | 56769 | 836 | 2538 | N/F | N/F | 795 | N/F | 20015 |  | Unknown
WO-22270-25 | N/F | 9891 | 264 | 258 | 21921 | 840 | 5175826 | 55 | 971 | N/F | N/F | 3384 | N/F | 17479 | 75905 | 344207 | 967060 | 198886 | 926178 | 2667921 | 762129 | 70567 | 24376114 | 5252399 | 6436247 | 2578343 | 12191895 | 8796 | N/F | N/F | 80841 | 1289 | 932 | 12786 | N/F | 479098 | N/F | 1643 | N/F | N/F | 163 | 17036 | 4193 | N/F | 426 | 4088 | N/F | 692 | 655 | 2311 | 20260 |  | Unknown
WO-22270-26 | N/F | 12563 | 804 | 155 | N/F | 889 | 4240706 | 101 | 663 | N/F | N/F | 4028 | N/F | 24770 | 60302 | 287378 | 774001 | 162614 | 803516 | 2348661 | 581597 | 60936 | 21456251 | 4630088 | 5598039 | 2253364 | 10809945 | 7011 | N/F | N/F | 63660 | N/F | 699 | 16966 | N/F | 360105 | 5507 | 1449 | 7112 | N/F | 1516 | 18358 | 8807 | 63768 | 350 | 2050 | N/F | 578 | N/F | N/F | 13669 |  | Unknown
WO-22270-27 | N/F | 10830 | 701 | 104 | 13108 | N/F | 4605715 | 100 | 314 | N/F | N/F | 6301 | N/F | 4677 | 82304 | 324466 | 851466 | 212558 | 871868 | 2738199 | 780466 | 78593 | 24393292 | 5288119 | 6437946 | 2467084 | 12334450 | 10699 | N/F | N/F | 88315 | 1968 | 1116 | 16220 | N/F | 356538 | 5273 | 1763 | 7639 | N/F | 955 | 12699 | 5702 | 36875 | 42 | 3624 | N/F | N/F | 883 | N/F | 13913 |  | Unknown
WO-22270-28 | N/F | 10457 | 524 | 99 | 9957 | N/F | 5043883 | 20 | 414 | N/F | N/F | 3150 | N/F | 3969 | 74707 | 337992 | 984642 | 178048 | 868396 | 2574578 | 711838 | 64290 | 23251013 | 5357991 | 6198911 | 2520075 | 11956537 | 7360 | 612 | N/F | 77701 | N/F | 302 | 15171 | N/F | 474742 | 3959 | 2054 | N/F | N/F | 904 | 12326 | 969 | N/F | 220 | 2026 | N/F | N/F | 1066 | 2137 | 13004 |  | Unknown
WO-22270-29 | N/F | 12872 | 556 | 834 | 12842 | N/F | 6084251 | 187 | 506 | N/F | N/F | 3253 | N/F | 3277 | 83419 | 389494 | 1502784 | 220808 | 958138 | 2927337 | 844415 | 79396 | 25150085 | 5750699 | 6971406 | 2761177 | 13663824 | 9079 | N/F | N/F | 89823 | N/F | 754 | 12068 | N/F | 524541 | 3807 | 618 | N/F | N/F | 569 | 18288 | 5094 | 17113 | 351 | 2913 | N/F | 933 | N/F | N/F | 15681 |  | Unknown
WO-22270-30 | N/F | 12519 | 485 | 578 | 9564 | N/F | 5138125 | 17 | 323 | 1867 | N/F | 5059 | N/F | 19147 | 87621 | 392561 | 1192287 | 204675 | 918954 | 2781792 | 768432 | 73479 | 23382076 | 5713404 | 6664518 | 2658465 | 12720413 | 8079 | 711 | 324 | 83667 | N/F | 428 | 12627 | N/F | 426600 | N/F | 1882 | 11901 | 4216 | 1328 | 33975 | 4366 | 74997 | 175 | 1832 | N/F | 914 | 701 | N/F | 16340 |  | Unknown
AB-220715-01 LRB 4_INJ001 | N/F | 10895 | 1752 | 298 | 10929 | N/F | 4511804 | 13 | 259 | N/F | N/F | 2175 | 1307 | 20244 | 292870 | 775819 | 805956 | 1024715 | 2806334 | 8987422 | 3438556 | 618581 | 37059342 | 7521692 | 19118635 | 3894796 | 24146421 | 69085 | 608 | N/F | 389709 | N/F | 432 | 17977 | N/F | 23003 | N/F | 64 | N/F | 5329 | 922 | 38587 | 6892 | 78641 | 284 | 1408 | 4441 | 490 | 604 | N/F | 33367 |  | Unknown
AB-220715-01 LRB 4 | N/F | 12928 | 1156 | 280 | 12806 | N/F | 4103688 | 21 | 845 | N/F | N/F | 3381 | 980 | 4413 | 288647 | 766217 | 748056 | 1052684 | 2720683 | 8801455 | 3423926 | 625109 | 35065625 | 6775832 | 18050520 | 3495357 | 22767222 | 40255 | 478 | N/F | 414308 | N/F | 353 | 14157 | N/F | 22050 | N/F | N/F | N/F | 5627 | 811 | 16074 | 4214 | 65411 | 561 | 1194 | 5380 | N/F | N/F | N/F | 35713 |  | Unknown
WO-22270-31 | N/F | 17740 | 916 | 412 | 9709 | 1706 | 4530340 | N/F | 330 | 602 | N/F | 6884 | N/F | 3651 | 83273 | 313712 | 849440 | 214161 | 928348 | 2812408 | 809761 | 78127 | 24969566 | 5163807 | 6415707 | 2588435 | 12305829 | 8390 | 1154 | 1004 | 76706 | 1273 | 662 | 12344 | N/F | 327601 | 4336 | 1659 | N/F | N/F | 2393 | 44651 | 5015 | 40426 | 101 | 3216 | N/F | 553 | 1475 | N/F | 13481 |  | Unknown
WO-22270-32 | N/F | 8511 | 548 | N/F | 13906 | 790 | 3272033 | 6 | 118 | N/F | N/F | 1287 | N/F | 15285 | 39896 | 233516 | 616483 | 84477 | 572347 | 1490609 | 328226 | 27873 | 15464622 | 2917629 | 3991210 | 1568949 | 7756948 | 4741 | 875 | N/F | 36053 | 1472 | 621 | 9931 | N/F | 273842 | N/F | 1140 | N/F | N/F | 2665 | 14962 | 3843 | 28871 | 409 | 1416 | N/F | 331 | 665 | 1588 | 10723 |  | Unknown
WO-22270-33 | 1074 | 10015 | 561 | 550 | 8965 | N/F | 4280618 | 40 | 544 | 460 | N/F | 2555 | N/F | 3430 | 61573 | 273477 | 801845 | 140821 | 760076 | 2215063 | 565262 | 47704 | 20700734 | 4532416 | 5467538 | 2116313 | 10249490 | 7519 | N/F | N/F | 52877 | 1248 | 637 | 12670 | N/F | 297063 | 5248 | 1245 | N/F | N/F | 2296 | 8208 | 3137 | 52635 | 445 | 1879 | N/F | 234 | N/F | N/F | 10936 |  | Unknown
WO-22270-34 | N/F | 14094 | 1280 | 461 | 88736 | N/F | 4357754 | N/F | 157 | N/F | N/F | 10865 | N/F | 6554 | 51881 | 283283 | 937403 | 126717 | 692041 | 2007401 | 507631 | 40635 | 19171427 | 4009787 | 5012065 | 1973366 | 9767518 | 4927 | N/F | N/F | 54088 | 3472 | 737 | 12501 | N/F | 383528 | 5189 | 2090 | 13647 | N/F | 1569 | 33204 | 12105 | 66563 | 597 | 2207 | N/F | 229 | 1055 | N/F | 29760 |  | Unknown
WO-22270-35 | N/F | 10898 | 454 | 840 | 9144 | 918 | 4152130 | 9 | 18 | N/F | N/F | 4961 | N/F | 15657 | 76175 | 311573 | 795677 | 176728 | 846287 | 2622592 | 725329 | 70253 | 23300048 | 5055111 | 6213151 | 2480709 | 11749688 | 7517 | N/F | N/F | 79629 | 3910 | 830 | 12379 | N/F | 302652 | 5544 | 2568 | 16655 | N/F | 1248 | 71218 | 5537 | 20044 | 536 | 2096 | 750 | 472 | 959 | N/F | 54433 |  | Unknown
WO-22270-36 | N/F | 9492 | 355 | 280 | 10104 | 654 | 5184774 | 51 | 373 | 448 | N/F | 3041 | N/F | 18991 | 97537 | 368302 | 948395 | 268439 | 1045386 | 3246348 | 1022167 | 89486 | 27583948 | 5987705 | 7413062 | 2946160 | 14258175 | 9213 | 641 | N/F | 105468 | N/F | 426 | 13406 | N/F | 441391 | N/F | 2795 | N/F | N/F | 1030 | 17637 | 5518 | 42763 | 193 | 2588 | N/F | 227 | 899 | N/F | 15427 |  | Unknown
WO-22270-37 | N/F | 10332 | 563 | 396 | 8872 | N/F | 4063981 | N/F | 38 | N/F | 401 | 2797 | N/F | 5351 | 47222 | 232043 | 680580 | 161519 | 787237 | 2434280 | 651708 | 62361 | 21995949 | 4596267 | 5639306 | 2273046 | 11143067 | 7128 | N/F | N/F | 73754 | 1053 | 508 | 11361 | N/F | 312906 | 1394 | 458 | N/F | N/F | 1523 | 10275 | 3223 | 29492 | 656 | 1358 | N/F | 204 | 1120 | N/F | 11783 |  | Unknown
WO-22270-38 | N/F | 9640 | 461 | 176 | 7387 | 1443 | 3944586 | 7 | 127 | N/F | 795 | 3203 | N/F | 4662 | 44240 | 220292 | 665978 | 167345 | 789416 | 2453308 | 670769 | 59037 | 23508771 | 4810667 | 5871523 | 2368565 | 11099051 | 7073 | 2048 | N/F | 79210 | 1112 | 218 | 11781 | N/F | 235073 | 2134 | 1872 | N/F | N/F | 20007 | 11910 | 6624 | 40824 | 585 | 2417 | N/F | 160 | N/F | N/F | 10112 |  | Unknown
WO-22270-39 | 1655 | 10560 | 274 | 412 | 9472 | 750 | 4072945 | 12 | 350 | N/F | N/F | 3299 | N/F | 18152 | 42236 | 200878 | 662794 | 162958 | 774833 | 2416852 | 658063 | 64087 | 21946613 | 4608560 | 5599197 | 2214400 | 10961914 | 7799 | 3147 | N/F | 74250 | 1582 | 312 | 12309 | N/F | 245235 | N/F | 1437 | N/F | N/F | 2054 | 9257 | 3638 | 44465 | 2770 | 2254 | N/F | 650 | N/F | N/F | 10059 |  | Unknown
WO-22270-40 | N/F | 10673 | 99 | 171 | 5949 | N/F | 3709133 | 33 | 111 | 756 | N/F | 3111 | N/F | 4335 | 44639 | 202213 | 639394 | 158309 | 777921 | 2411900 | 641009 | 65390 | 22015251 | 4605304 | 5748344 | 2242114 | 10864387 | 7929 | 1044 | N/F | 71697 | 1287 | 147 | 10549 | N/F | 337269 | N/F | 1844 | N/F | N/F | 1581 | 57909 | 1994 | 13094 | 867 | 1880 | N/F | 175 | 1236 | N/F | 26301 |  | Unknown
AB-220715-01 LFB-H 3 | 422 | 240472 | 372235 | 471820 | 73421517 | N/F | 4016559 | 1336060 | 5302509 | 460723 | 866543 | 3210719 | 2134741 | 2895909 | 230664 | 369808 | 511786 | 875068 | 2211219 | 7663213 | 3161735 | 599800 | 28117518 | 6474161 | 14776874 | 2985029 | 19755156 | 58803 | 901 | N/F | 354933 | 2359615 | 214 | 3172879 | 45308 | 26206636 | N/F | 128 | 7206593 | 2870834 | 1520 | 17690084 | 3444990 | 13902298 | 701 | 6625200 | 551694 | 683 | 1863 | N/F | 30375905 |  | Unknown
AB-220715-01 LRB 5_INJ001 | N/F | 13357 | 558 | 973 | 118786 | 577 | 4455679 | 459 | 4726 | 1056 | N/F | 4603 | 12713 | 8094 | 281552 | 727887 | 751049 | 929593 | 2745820 | 9074474 | 3521842 | 633950 | 37502477 | 7360423 | 19081863 | 3774492 | 24113664 | 69278 | 433 | N/F | 409478 | 973 | 496 | 15467 | N/F | 52804 | N/F | 40 | N/F | 4546 | 936 | 22563 | 6570 | 43889 | 301 | 15998 | 7499 | 233 | 1021 | N/F | 53833 |  | Unknown
AB-220715-01 LRB 5 | N/F | 14108 | 1924 | 368 | 40026 | N/F | 3951425 | 233 | 1280 | 996 | 1036 | 1998 | 12933 | 19809 | 271440 | 690852 | 678063 | 931713 | 2582623 | 8614914 | 3314814 | 594412 | 33609010 | 6444273 | 17477041 | 3390735 | 21661304 | 41800 | N/F | 1039 | 370035 | N/F | 679 | 14415 | N/F | 37057 | N/F | 24 | N/F | 4856 | 546 | 16399 | 4003 | 48039 | 141 | 4633 | 7217 | 121 | N/F | N/F | 27916 |  | Unknown
AB-220715-01 LD 2 | N/F | 11852 | 492 | 535 | 11993 | 1027 | 3955903 | 40 | 911 | N/F | 537 | 3921 | 7081 | 4946 | 47321 | 239917 | 700116 | 164230 | 782533 | 2368011 | 637486 | 61350 | 22153123 | 4388743 | 5642949 | 2166130 | 10589898 | 7489 | N/F | N/F | 67770 | 2383 | 224 | 11420 | N/F | 248641 | N/F | 1786 | 11702 | N/F | 1870 | 38898 | 5194 | 104476 | 301 | 1885 | N/F | 310 | 1857 | N/F | 13803 |  | Unknown
AB-220715-01 CCC-H 3 | 1718 | 401257 | 546110 | 776055 | 118154697 | 1398 | 3691444 | 2721153 | 9427031 | 917700 | 1587692 | 5156356 | 3651431 | 4813582 | 190093 | 278290 | 394786 | 789700 | 1820963 | 6730117 | 2920353 | 551744 | 22653173 | 5760940 | 12414542 | 2451267 | 16675537 | 55846 | N/F | 1006 | 349251 | 4060807 | 300 | 5700770 | 89542 | 41381237 | 1110 | 84 | 12759033 | 5329685 | 230351 | 29292748 | 5272417 | 22885479 | 369 | 11720831 | 1030184 | 38296 | 734 | N/F | 48056493 |  | Unknown
AB-220715-01 LFSM-H 2 | N/F | 47675 | 157010 | 381884 | 47775354 | N/F | 3357558 | 282064 | 1842856 | 99018 | 166741 | 1285860 | 710040 | 1655044 | 43158 | 159953 | 388410 | 168355 | 752646 | 2447482 | 704358 | 65708 | 18025841 | 4180270 | 5276322 | 1847622 | 9470779 | 8264 | 1897 | N/F | 81085 | 744907 | 668 | 2673604 | N/F | 16596031 | N/F | 1870 | 2350666 | 614991 | 142709 | 8331626 | 2623067 | 4922539 | 724 | 4157203 | 61288 | 171 | 622 | N/F | 21911280 |  | Unknown
WO-22270-41 | N/F | 12565 | 459 | 1118 | 25476 | 520 | 4519159 | 88 | 1341 | N/F | N/F | 3072 | N/F | 4431 | 62234 | 261319 | 928886 | 150874 | 743759 | 2339634 | 601965 | 56194 | 21698381 | 4412153 | 5644774 | 2172604 | 11178511 | 8915 | N/F | 782 | 66208 | N/F | 482 | 13246 | N/F | 242436 | N/F | 1673 | N/F | N/F | 1945 | 21006 | 5577 | 18115 | 337 | 3623 | N/F | 205 | N/F | 2033 | 19271 |  | Unknown
WO-22270-42 | N/F | 9275 | 352 | 374 | 19191 | 507 | 4171824 | 119 | 697 | 1304 | 630 | 3997 | N/F | 3003 | 55602 | 250951 | 719660 | 183430 | 829097 | 2572875 | 748662 | 72721 | 23151283 | 4592944 | 6032911 | 2347910 | 11485926 | 7287 | N/F | N/F | 72490 | 2897 | 362 | 12105 | N/F | 156052 | 3236 | 1974 | N/F | N/F | 4346 | 9615 | 7467 | 27799 | 371 | 2981 | N/F | 527 | N/F | N/F | 16443 |  | Unknown
WO-22270-43 | 2532 | 13148 | 421 | 134 | 15301 | 627 | 4017411 | 122 | 388 | N/F | N/F | 2838 | N/F | 3819 | 48685 | 211757 | 630766 | 169870 | 810068 | 2487020 | 684181 | 66134 | 21964913 | 4527415 | 5623487 | 2210459 | 10691478 | 7526 | N/F | N/F | 73276 | 1030 | 799 | 10338 | N/F | 251107 | N/F | 515 | N/F | N/F | 144935 | 20229 | 6955 | 115769 | 2700 | 2471 | N/F | 696 | N/F | N/F | 12806 |  | Unknown
WO-22270-44 | N/F | 11537 | 1248 | 631 | 14772 | N/F | 4298749 | 68 | 229 | 1029 | N/F | 3821 | N/F | 4225 | 65987 | 275173 | 735004 | 187714 | 867535 | 2683705 | 757318 | 71330 | 24105423 | 4851054 | 6108574 | 2357439 | 12037762 | 10678 | 624 | N/F | 80655 | 3863 | 320 | 13043 | N/F | 193193 | 4666 | 2012 | N/F | N/F | 1254 | 10846 | 5933 | 29625 | 793 | 1318 | N/F | 191 | N/F | N/F | 13334 |  | Unknown
WO-22270-45 | 1032 | 12749 | 361 | 1095 | 8594 | 594 | 5249589 | 139 | 238 | 2225 | 967 | 2591 | N/F | 3510 | 68467 | 322921 | 1364450 | 206195 | 926799 | 2782733 | 801893 | 77952 | 24106814 | 5265961 | 6624417 | 2543666 | 12634452 | 8379 | N/F | N/F | 86272 | 1055 | 235 | 12904 | N/F | 179624 | N/F | 1447 | 38462 | N/F | 106613 | 68047 | 4015 | N/F | 228 | 2910 | N/F | 284 | 1153 | N/F | 68099 |  | Unknown
WO-22270-52 | 612 | 12612 | 531 | 550 | 8385 | 1111 | 3359440 | 11 | 107 | 1094 | 387 | 1671 | N/F | 14240 | 38718 | 203258 | 615855 | 86810 | 593628 | 1737365 | 351720 | 32302 | 17035994 | 3287481 | 4543967 | 1681002 | 8766585 | 4926 | 728 | N/F | 40552 | 2265 | 248 | 12366 | N/F | 207236 | 9812 | 1409 | N/F | N/F | 149277 | 15925 | 3707 | 21915 | 129 | 1410 | N/F | 197 | N/F | N/F | 9239 |  | Unknown
WO-22270-46 | N/F | 11804 | 937 | 381 | 10168 | N/F | 4333466 | N/F | 115 | 1516 | 608 | 3013 | N/F | 12402 | 67660 | 276508 | 743856 | 189053 | 865596 | 2734436 | 752980 | 71468 | 24586598 | 5026202 | 6296925 | 2375957 | 12022917 | 8136 | N/F | N/F | 86130 | 1029 | 168 | 12004 | N/F | 340406 | N/F | 1565 | N/F | N/F | 1694 | 17288 | 4834 | 28606 | 491 | 2563 | N/F | 527 | 957 | N/F | 15249 |  | Unknown
WO-22270-47 | N/F | 11397 | 657 | 633 | 10705 | 1422 | 4380525 | 70 | 121 | 674 | N/F | 3627 | N/F | 4012 | 62698 | 272901 | 744937 | 156945 | 810284 | 2543616 | 630850 | 57044 | 22479422 | 4696323 | 5846099 | 2286942 | 11703113 | 12054 | 1254 | N/F | 71626 | 2102 | 131 | 11830 | N/F | 179050 | 4216 | 1681 | N/F | N/F | 2072 | 10525 | 4204 | 24328 | 310 | 1749 | N/F | 152 | 1442 | N/F | 9166 |  | Unknown
WO-22270-53 | 1033 | 16786 | 886 | 306 | 9476 | 1010 | 4611598 | N/F | 60 | 1222 | 1199 | 3861 | N/F | 3284 | 80325 | 312897 | 819325 | 200515 | 915032 | 2871834 | 829180 | 78538 | 27991159 | 5511633 | 7017773 | 2757852 | 14229850 | 9140 | 1300 | 1171 | 90937 | 1002 | 82 | 14512 | N/F | 440712 | N/F | 1058 | N/F | N/F | 145375 | 28328 | 4219 | 44875 | 367 | 3349 | N/F | 246 | 1030 | N/F | 11412 |  | Unknown
WO-22270-48 | N/F | 12145 | 773 | 441 | 10044 | N/F | 4387633 | 31 | 899 | 994 | N/F | 3065 | N/F | 19183 | 74234 | 288110 | 743081 | 182187 | 873671 | 2638257 | 767247 | 67919 | 24901790 | 4945002 | 6441919 | 2495139 | 12535265 | 6969 | N/F | 679 | 81609 | 1407 | 165 | 13379 | N/F | 441632 | 4945 | 2195 | N/F | N/F | 110414 | 7722 | 4253 | 6358 | 262 | 1878 | N/F | 175 | 882 | 2240 | 10264 |  | Unknown
WO-22270-54 | 679 | 12893 | 689 | 841 | 11258 | 2239 | 5722559 | N/F | 135 | 3100 | N/F | 4070 | N/F | 15640 | 110268 | 467600 | 1110748 | 298970 | 1426202 | 4239213 | 1206695 | 118308 | 36975477 | 6821550 | 10975723 | 3872401 | 21571853 | 11737 | 711 | N/F | 129241 | 725 | 294 | 16967 | N/F | 523828 | N/F | 2667 | N/F | 3451 | 123039 | 12964 | 4672 | 28466 | 147 | 2288 | N/F | 336 | 1104 | N/F | 17731 |  | Unknown
WO-22270-49 | 1264 | 11146 | 617 | 323 | 6523 | 504 | 4288525 | 134 | 206 | N/F | N/F | 2792 | N/F | 4685 | 61272 | 269611 | 787090 | 156686 | 805874 | 2424454 | 598947 | 57519 | 22711009 | 4555655 | 5872741 | 2184481 | 11118381 | 7700 | N/F | N/F | 65892 | 1800 | 344 | 11221 | N/F | 452371 | 4118 | 1869 | N/F | N/F | 3086 | 12772 | 4594 | 13784 | 451 | 2677 | 7135 | 452 | 554 | N/F | 11863 |  | Unknown
WO-22270-55 | 715 | 9757 | 414 | 190 | 8115 | N/F | 2067748 | 18 | N/F | 1220 | 542 | 2122 | N/F | 14848 | 47145 | 196169 | 375140 | 115484 | 548155 | 1739005 | 423319 | 44005 | 11295160 | 2250704 | 4001568 | 1054358 | 6579510 | 7978 | 738 | N/F | 53133 | 1697 | 339 | 5678 | N/F | 146500 | 4173 | 1070 | N/F | N/F | 144241 | 85275 | 3794 | N/F | 380 | 1760 | N/F | 119 | 631 | N/F | 44005 |  | Unknown
WO-22270-50 | N/F | 13897 | 987 | 291 | 8718 | 798 | 4116789 | 17 | 72 | N/F | 285 | 3136 | N/F | 4410 | 77530 | 272500 | 718502 | 194929 | 837207 | 2654686 | 796369 | 78761 | 23587410 | 4722789 | 6125561 | 2303412 | 11990486 | 9505 | N/F | N/F | 83917 | 2230 | 379 | 13126 | N/F | 451250 | N/F | 1142 | N/F | N/F | 136992 | 52288 | 6475 | 16042 | 1227 | 1873 | N/F | 307 | 1252 | 824 | 12905 |  | Unknown
WO-22270-56 | 1557 | 16672 | 489 | 230 | 8261 | 882 | 4227742 | 113 | 223 | N/F | 918 | 1880 | N/F | 3700 | 71072 | 275863 | 768504 | 191057 | 818452 | 2705364 | 739990 | 78589 | 24097262 | 4729537 | 6054762 | 2354598 | 11905052 | 8377 | 557 | N/F | 79209 | 1228 | 464 | 11971 | N/F | 476727 | N/F | 1119 | N/F | N/F | 148882 | 29472 | 2411 | 42409 | 264 | 2793 | N/F | 16 | 636 | N/F | 16061 |  | Unknown
WO-22270-57 | N/F | 10762 | 393 | 389 | 7715 | 1370 | 5400712 | 25 | 128 | 2035 | 1152 | 4374 | N/F | 8763 | 91133 | 349112 | 971209 | 250975 | 1125337 | 3549627 | 986250 | 102889 | 33342342 | 6327514 | 8768564 | 3228336 | 18264750 | 12298 | 567 | N/F | 102772 | 3292 | 62 | 12430 | N/F | 612723 | 7707 | 1802 | N/F | N/F | 112010 | 36393 | 1753 | 26218 | 518 | 2905 | N/F | 233 | 943 | 621 | 12681 |  | Unknown
AB-220715-01 LRB 6_INJ001 | N/F | 13566 | 1307 | 394 | 14897 | N/F | 4184959 | 26 | 459 | 2520 | N/F | 715 | 12970 | 12681 | 249987 | 695868 | 713617 | 926924 | 2633926 | 8440018 | 3234487 | 626722 | 35663352 | 6910705 | 18386672 | 3520878 | 23132124 | 68749 | 1275 | N/F | 380763 | 1974 | 366 | 15253 | N/F | 27488 | N/F | N/F | N/F | 4021 | 113868 | 10939 | 3655 | 17963 | 442 | 2655 | 2743 | 341 | 1219 | N/F | 17146 |  | Unknown
AB-220715-01 LRB 6 | 1392 | 12406 | 1290 | 206 | 14463 | 536 | 3869761 | N/F | 176 | 4467 | 349 | 980 | 8861 | 6323 | 276686 | 672169 | 692694 | 938199 | 2505670 | 8616998 | 3297290 | 591494 | 33771748 | 6396599 | 17770400 | 3461344 | 22861691 | 37349 | N/F | N/F | 384775 | 988 | 254 | 14337 | N/F | 26795 | 1417 | N/F | N/F | 4741 | 1020 | 10319 | 3799 | 21478 | 106 | 2593 | N/F | 160 | 2326 | 815 | 15253 |  | Unknown
WO-22270-58 | 898 | 12819 | 937 | 455 | 5681 | 860 | 3919999 | N/F | N/F | N/F | 924 | 3091 | N/F | 5986 | 38222 | 217849 | 733040 | 81253 | 530893 | 1687259 | 366671 | 31073 | 17839994 | 3375777 | 4361498 | 1713680 | 9103504 | 3990 | 713 | N/F | 38250 | 1348 | 153 | 11054 | N/F | 317313 | 7307 | 679 | N/F | N/F | 126396 | 5811 | 4677 | 23486 | 239 | 1543 | N/F | 193 | 1518 | N/F | 7999 |  | Unknown
WO-22270-51 | N/F | 11462 | 1622 | 777 | 8293 | 866 | 4338525 | N/F | 189 | N/F | N/F | 3832 | N/F | 17029 | 78355 | 301735 | 865355 | 224738 | 922328 | 2877685 | 795112 | 78775 | 26825806 | 5294649 | 6789733 | 2656259 | 13508564 | 7845 | N/F | N/F | 91409 | 569 | 92 | 10647 | N/F | 336000 | N/F | 2124 | N/F | N/F | 9698 | 42603 | 3979 | 122297 | 153 | 1786 | N/F | 251 | 977 | N/F | 16648 |  | Unknown
WO-22270-59 | 1303 | 14327 | 553 | 261 | 6719 | N/F | 4070067 | 5 | 196 | N/F | 613 | 1792 | N/F | 12926 | 67376 | 239761 | 675464 | 180885 | 772260 | 2613869 | 738467 | 74023 | 23459572 | 4540862 | 6053205 | 2337281 | 12371284 | 8397 | 1133 | 2276 | 82960 | 1148 | 205 | 13305 | N/F | 327341 | N/F | 1586 | N/F | N/F | 139908 | 12858 | 2508 | 13899 | 291 | 2291 | N/F | 42 | N/F | N/F | 12870 |  | Unknown
WO-22270-60 | N/F | 12674 | 545 | 842 | 7227 | 960 | 4143868 | 19 | 126 | N/F | 1832 | 2765 | N/F | 4973 | 56129 | 248651 | 778452 | 145756 | 774035 | 2401626 | 626941 | 56797 | 23786204 | 4497406 | 5882129 | 2250888 | 11885493 | 7086 | N/F | 1965 | 62466 | 1338 | 41 | 12588 | N/F | 370656 | 9277 | 1576 | N/F | N/F | 157440 | 71277 | 1943 | 90623 | 404 | 2062 | N/F | 169 | 392 | N/F | 43884 |  | Unknown
AB-220715-01 LFB-H 4 | 1503 | 216664 | 340216 | 468646 | 70555392 | 1037 | 3471558 | 1386398 | 5018141 | 476936 | 854112 | 3102696 | 2057543 | 2705853 | 219256 | 336367 | 445453 | 823884 | 2084483 | 7765345 | 3179654 | 569850 | 27461333 | 5928938 | 14682995 | 2784852 | 19333922 | 60308 | 1444 | 615 | 364744 | 2394438 | 223 | 2809455 | 47640 | 25579333 | 987 | 93 | 7400010 | 3012016 | 185394 | 17701061 | 3189357 | 13826704 | 592 | 6067246 | 551317 | 1274113 | 1064 | N/F | 30232399 |  | Unknown
AB-220715-01 LRB 7_INJ001 | N/F | 9934 | 1221 | 743 | 106248 | 1591 | 3797787 | 543 | 7331 | 1479 | 949 | 4920 | 16982 | 14978 | 183324 | 465772 | 542135 | 922069 | 2444509 | 8294839 | 3371818 | 644923 | 33799083 | 7495353 | 17384755 | 3432322 | 21251175 | 67495 | 1239 | 599 | 400382 | 3092 | 162 | 17510 | N/F | 49198 | N/F | 12 | N/F | 5548 | 150344 | 31085 | 6214 | 99307 | 1105 | 17551 | 2780 | 387 | N/F | N/F | 49880 |  | Unknown
AB-220715-01 LRB 7 | N/F | 9068 | 737 | 434 | 28369 | 743 | 3080807 | 144 | 1126 | 2441 | N/F | 2148 | 15741 | 5139 | 157534 | 382542 | 440949 | 836401 | 2057854 | 7324006 | 3069010 | 602445 | 26605293 | 5935861 | 14208774 | 2763779 | 17045869 | 30987 | 1341 | 1487 | 347450 | 1490 | N/F | 15152 | N/F | 25132 | N/F | 39 | N/F | 4199 | 100117 | 40911 | 3762 | 78467 | 1174 | 4406 | N/F | 113 | N/F | N/F | 42877 |  | Unknown
AB-220715-01 LD 3 | N/F | 9727 | 715 | 485 | 15610 | 847 | 3720395 | 47 | 723 | N/F | 802 | 4682 | N/F | 4370 | 42024 | 147793 | 476243 | 178949 | 729556 | 2370771 | 695541 | 69273 | 22038283 | 5065795 | 5371566 | 2242228 | 10928132 | 9064 | 903 | N/F | 76012 | 946 | 252 | 10715 | N/F | 317139 | N/F | 1214 | 18163 | N/F | 140134 | 12068 | 6173 | 43604 | 4729 | 2416 | 3766 | 288 | N/F | N/F | 26057 |  | Unknown
AB-220715-01 CCC-H 4 | 806 | 222550 | 365786 | 604598 | 106354285 | 1077 | 3214531 | 2666056 | 9206694 | 794479 | 1350899 | 4408793 | 3107148 | 4300134 | 117532 | 196078 | 296213 | 688206 | 1527722 | 5407967 | 2491474 | 473808 | 18661633 | 5490930 | 9983058 | 2197810 | 12885821 | 42768 | 412 | N/F | 309516 | 3252020 | 56 | 5012152 | 64162 | 33456766 | N/F | 792 | 10070390 | 4486862 | 215472 | 22318412 | 4804286 | 18232010 | 1245 | 11143862 | 906840 | 158 | N/F | N/F | 39915006 |  | Unknown
AB-220715-01 LFSM-H 3 | N/F | 36407 | 125954 | 414069 | 57346629 | 625 | 3847670 | 349916 | 2624158 | 101590 | 179279 | 1617761 | 799544 | 2144933 | 36865 | 123413 | 411469 | 171966 | 840934 | 2431705 | 760210 | 73609 | 20032416 | 5307830 | 5587522 | 2356133 | 10882993 | 8005 | N/F | N/F | 80393 | 749930 | 494 | 3032026 | N/F | 18801662 | N/F | 1647 | 2319848 | 658540 | 157806 | 9664692 | 2992562 | 5209723 | 30697 | 5498947 | 68732 | 564937 | N/F | N/F | 23249965 |  | Unknown
WO-22270-61 | N/F | 8528 | 436 | 187 | 33456 | 1084 | 3907553 | 131 | 1085 | N/F | 1245 | 4960 | N/F | 19785 | 37760 | 172665 | 570884 | 159381 | 784412 | 2398705 | 629715 | 65566 | 23853721 | 5469572 | 5632908 | 2417223 | 11059184 | 8100 | 392 | N/F | 76889 | 679 | 243 | 10777 | N/F | 352623 | N/F | 3135 | N/F | N/F | 137398 | 11958 | 6632 | 14656 | 1779 | 4778 | N/F | 297 | N/F | N/F | 24071 |  | Unknown
WO-22270-62 | N/F | 7653 | 261 | 353 | 18271 | 561 | 4556279 | 36 | 640 | 848 | 962 | 4196 | N/F | 19162 | 34080 | 154958 | 998949 | 142858 | 704598 | 2125865 | 558706 | 59803 | 21352912 | 4488000 | 5277855 | 1890861 | 8851652 | 7449 | 1507 | N/F | 70819 | 1303 | 422 | 10320 | N/F | 302395 | N/F | 1777 | N/F | N/F | 122233 | 13698 | 1860 | 21000 | 2819 | 1925 | N/F | 247 | N/F | N/F | 14575 |  | Unknown
WO-22270-63 | N/F | 7161 | 146 | 328 | 12254 | N/F | 3353568 | 15 | 579 | N/F | 1087 | 5279 | N/F | 4162 | 25119 | 122773 | 479574 | 97773 | 562587 | 1651627 | 424919 | 38546 | 16824197 | 3891033 | 4194261 | 1618097 | 7870391 | 5932 | 780 | N/F | 45185 | 1701 | 247 | 8106 | N/F | 249797 | N/F | 1442 | N/F | N/F | 117439 | 10332 | 4586 | 29950 | 1468 | 2559 | N/F | 111 | 1454 | 996 | 12859 |  | Unknown
WO-22270-64 | N/F | 8247 | 553 | N/F | 9112 | 520 | 2733288 | 41 | 214 | 795 | 1124 | 3252 | N/F | 2237 | 22957 | 113298 | 318198 | 85349 | 525246 | 1505081 | 355301 | 33003 | 15389981 | 3429698 | 3735298 | 1461250 | 7084669 | 5328 | 574 | N/F | 39411 | 989 | 58 | 8600 | N/F | 240933 | N/F | 673 | N/F | N/F | 136085 | 100894 | 6670 | N/F | 4997 | 2194 | N/F | 112 | N/F | N/F | 10773 |  | Unknown
WO-22270-65 | N/F | 8776 | 1401 | 507 | 13177 | 415 | 4217569 | 30 | 92 | N/F | 364 | 2346 | N/F | 16131 | 48780 | 193054 | 568489 | 210409 | 928365 | 2859405 | 889358 | 88180 | 27040132 | 6026398 | 6960461 | 2805077 | 13445178 | 10197 | 805 | N/F | 94969 | 2297 | 116 | 13004 | N/F | 421107 | N/F | 1492 | N/F | 1868 | 129789 | 18975 | 4182 | 15194 | 979 | 2610 | N/F | 366 | N/F | N/F | 19084 |  | Unknown
WO-22270-66 | 1718 | 9186 | 696 | 244 | 11499 | 672 | 3632128 | N/F | 180 | 1778 | 2545 | 2901 | N/F | 20719 | 37834 | 145998 | 499692 | 156805 | 677177 | 2262051 | 693512 | 68368 | 21093445 | 4742478 | 5210331 | 2073449 | 9956516 | 9090 | 1181 | N/F | 74964 | 631 | 43 | 9661 | N/F | 297182 | N/F | 1124 | N/F | N/F | 128198 | 21512 | 529 | 20167 | 3006 | 1412 | N/F | 225 | 2010 | N/F | 12781 |  | Unknown
WO-22270-67 | N/F | 8396 | 603 | 114 | 8886 | 599 | 3237595 | 23 | 220 | 1309 | 1196 | 2053 | N/F | 15382 | 31621 | 132487 | 419572 | 132533 | 642445 | 1923433 | 540893 | 55026 | 19195880 | 4373277 | 4549314 | 1789267 | 8858423 | 6707 | 717 | N/F | 59699 | 1228 | 67 | 9794 | N/F | 356678 | N/F | 899 | N/F | N/F | 125356 | 799413 | 5326 | 50601 | 1842 | 2508 | N/F | 240 | N/F | N/F | 18803 |  | Unknown
WO-22270-68 | 530 | 6806 | 953 | N/F | 7334 | 514 | 3081227 | N/F | 151 | 1014 | 2175 | 3052 | N/F | 15670 | 29763 | 125264 | 426377 | 110015 | 600447 | 1795732 | 493220 | 48021 | 18173454 | 4165740 | 4280878 | 1755705 | 8346155 | 6750 | 1132 | 735 | 52466 | 1015 | 137 | 12284 | N/F | 297539 | N/F | 323 | N/F | N/F | 121219 | 8084 | 3557 | 20037 | 5915 | 2814 | N/F | 47 | N/F | N/F | 12048 |  | Unknown
WO-22270-69 | N/F | 8516 | 363 | 583 | 6462 | 539 | 2197280 | 8 | 211 | N/F | 1117 | 1852 | N/F | 16773 | 17138 | 90159 | 281018 | 56380 | 412816 | 1172252 | 270748 | 23222 | 12658951 | 2683042 | 3107400 | 1202969 | 5936933 | 4206 | 1531 | 480 | 28093 | 467 | 112 | 8906 | N/F | 276103 | N/F | 631 | N/F | N/F | 122422 | 7372 | 6055 | 7548 | 6479 | 2195 | N/F | 33 | 2672 | 797 | 10235 |  | Unknown
WO-22270-70 | 589 | 9125 | 846 | 51 | 5341 | 493 | 2947739 | N/F | 36 | 693 | 854 | 2218 | N/F | 6699 | 22543 | 123191 | 363737 | 101322 | 563025 | 1671924 | 420469 | 41393 | 17148751 | 3840491 | 4183401 | 1639413 | 7799028 | 5216 | N/F | 1501 | 47787 | 1555 | 178 | 9525 | N/F | 310369 | N/F | 739 | 11162 | N/F | 148268 | 38971 | 2170 | 70540 | 2658 | 2252 | N/F | 209 | N/F | N/F | 36753 |  | Unknown
AB-220715-01 LRB 8_INJ001 | 1940 | 8888 | 333 | 113 | 8882 | 567 | 3616106 | 27 | 154 | 1081 | 1272 | 1403 | 17968 | 5161 | 159490 | 389331 | 529744 | 902082 | 2411403 | 8382177 | 3527356 | 665890 | 31665544 | 7118950 | 17191201 | 3354216 | 20522310 | 79005 | 735 | N/F | 434366 | 735 | 133 | 10913 | N/F | 19957 | N/F | 85 | N/F | 4703 | 135545 | 41497 | 3937 | 95064 | 2325 | 2389 | 7023 | 388 | 932 | N/F | 57680 |  | Unknown
AB-220715-01 LRB 8 | N/F | 8320 | 398 | 206 | 18875 | 659 | 2974855 | 20 | 522 | 2086 | 884 | 1812 | 15835 | 13815 | 146726 | 356459 | 411697 | 816716 | 2081068 | 7496657 | 3015102 | 591282 | 25841839 | 6033865 | 14127326 | 2866857 | 17494277 | 35119 | 650 | N/F | 347082 | 1152 | N/F | 8632 | N/F | 21196 | N/F | 24 | N/F | 3964 | 71862 | 47534 | 4222 | 72045 | 1222 | 2609 | 5094 | 350 | N/F | N/F | 31297 |  | Unknown
WO-22270-71 | 779 | 7557 | 429 | 175 | 9018 | 2076 | 4337036 | 62 | 321 | 1547 | 804 | 3363 | N/F | 3231 | 37984 | 151120 | 737289 | 140433 | 688630 | 2146971 | 606222 | 62675 | 20858411 | 4884731 | 5108915 | 2019373 | 9865279 | 8916 | 225 | N/F | 70935 | 690 | 523 | 9814 | N/F | 822733 | N/F | 1473 | N/F | N/F | 127394 | 231361 | 3120 | 56578 | 1916 | 1808 | N/F | 152 | N/F | N/F | 15039 |  | Unknown
WO-22270-72 | 619 | 8491 | 461 | 291 | 7095 | 553 | 3551904 | 46 | 34 | 734 | 748 | 2206 | N/F | 4182 | 34545 | 149554 | 463809 | 154375 | 676054 | 2175315 | 618722 | 63479 | 20577747 | 4741897 | 5013576 | 1987691 | 9800270 | 7587 | 231 | N/F | 70231 | 1081 | 92 | 9748 | N/F | 366751 | N/F | 749 | N/F | N/F | 107666 | 6675 | 2488 | 25655 | 1013 | 1328 | N/F | 81 | N/F | N/F | 13043 |  | Unknown
WO-22270-73 | 661 | 7541 | 231 | 31 | 4864 | 524 | 3082300 | 17 | N/F | 1516 | 1073 | 2832 | N/F | 3110 | 27815 | 126217 | 404727 | 124508 | 615152 | 1896728 | 507276 | 48718 | 18861487 | 4229281 | 4622658 | 1778582 | 8907123 | 6110 | 779 | 951 | 56268 | 831 | 200 | 11436 | N/F | 280113 | N/F | 750 | N/F | N/F | 129252 | 8770 | 2714 | 6497 | 1488 | 1335 | N/F | 226 | 686 | 1170 | 10272 |  | Unknown
WO-22270-74 | 590 | 7780 | 792 | 164 | 5337 | 547 | 3318136 | 29 | 12 | 1022 | 1013 | 2321 | N/F | 14702 | 26535 | 124794 | 399227 | 122767 | 598123 | 1856010 | 511395 | 45379 | 18839817 | 4076075 | 4366331 | 1752307 | 8820378 | 7115 | N/F | 809 | 56226 | 891 | 18 | 9670 | N/F | 582455 | N/F | 713 | 8115 | N/F | 124582 | 89636 | 3377 | 38117 | 1108 | 1705 | N/F | 350 | N/F | N/F | 22051 |  | Unknown
WO-22270-75 | N/F | 7901 | 473 | 110 | 8118 | 1096 | 3335552 | 19 | 78 | 1874 | 752 | 3844 | N/F | 13655 | 36234 | 150869 | 428711 | 189479 | 743176 | 2339942 | 754436 | 83220 | 21679778 | 5001727 | 5477571 | 2099428 | 10411323 | 8457 | 942 | 653 | 80966 | 1765 | 71 | 10830 | N/F | 320347 | N/F | 868 | 18721 | N/F | 131126 | 72539 | 4666 | 118686 | 1504 | 2180 | N/F | 215 | N/F | 908 | 48403 |  | Unknown
WO-22270-76 | N/F | 8354 | 403 | N/F | 4052 | 364 | 2096551 | 14 | 6 | 713 | 654 | 3248 | N/F | 16895 | 14709 | 87946 | 264520 | 61407 | 414464 | 1167594 | 264233 | 26837 | 11850992 | 2540321 | 2991368 | 1075264 | 5806209 | 4541 | 386 | N/F | 28247 | 1862 | 67 | 8073 | N/F | 220081 | N/F | 427 | N/F | N/F | 128838 | 7309 | 2690 | 19413 | 4611 | 892 | N/F | 103 | N/F | N/F | 26763 |  | Unknown
WO-22270-77 | 801 | 8146 | 314 | 248 | 7231 | N/F | 4154030 | 24 | 167 | 700 | 814 | 5068 | N/F | 19500 | 40741 | 145771 | 651043 | 193017 | 724842 | 2397560 | 763299 | 86615 | 22956505 | 5293461 | 5720353 | 2155191 | 10925884 | 9382 | 651 | 3841203 | 89528 | 1931 | 375 | 7308 | N/F | 2110610 | N/F | 973 | 16217 | 3827 | 140541 | 9976 | 4459 | 14858 | 6607 | 1837 | N/F | 148 | 4583 | N/F | 15653 |  | Unknown
WO-22270-78 | N/F | 8391 | 682 | 462 | 5509 | 1313 | 3699256 | N/F | N/F | 1887 | 371 | 3362 | N/F | 16060 | 42786 | 147140 | 504829 | 213636 | 775811 | 2544211 | 831775 | 104731 | 23688616 | 5532173 | 5776537 | 2165997 | 11367131 | 11120 | N/F | N/F | 97909 | 2346 | 60 | 9851 | N/F | 420719 | N/F | 1171 | 16951 | 4851 | 125913 | 17021 | 3043 | 18039 | 6593 | 1537 | N/F | 704 | N/F | N/F | 12977 |  | Unknown
WO-22270-79 | 1091 | 11316 | 665 | 132 | 6789 | 625 | 3746846 | 17 | N/F | 1218 | 718 | 3357 | N/F | 5471 | 39816 | 149632 | 491960 | 204186 | 756442 | 2503476 | 816834 | 99216 | 23328471 | 5429378 | 5728536 | 2258374 | 11034958 | 11439 | N/F | N/F | 92966 | 1911 | 108 | 10667 | N/F | 340294 | N/F | 906 | 7548 | 5042 | 136957 | 8018 | 1593 | 22072 | 4521 | 1875 | N/F | 216 | N/F | N/F | 17641 |  | Unknown
WO-22270-80 | N/F | 11070 | 508 | 424 | 4389 | 431 | 3896630 | N/F | 112 | 1913 | 1429 | 2717 | N/F | 3215 | 44513 | 155299 | 540683 | 235939 | 784201 | 2601534 | 876904 | 112711 | 24003414 | 5637572 | 5798240 | 2245614 | 11573539 | 11307 | 1454 | N/F | 103221 | 1263 | 63 | 10909 | N/F | 390842 | N/F | 939 | 20313 | 5827 | 129938 | 38564 | 3864 | 43172 | 7328 | 1345 | N/F | 117 | N/F | N/F | 24660 |  | Unknown
AB-220715-01 LFB-H 5 | 420 | 111982 | 196278 | 305790 | 55658364 | 414 | 2745957 | 1383392 | 4858328 | 429462 | 727330 | 2406752 | 1700250 | 2191323 | 109154 | 197025 | 286216 | 707518 | 1746852 | 6345696 | 2509330 | 543401 | 20292643 | 5351752 | 11001705 | 2265980 | 13828390 | 50507 | 754 | 743 | 307686 | 1930004 | N/F | 2177492 | 45638 | 19004852 | N/F | 612 | 6059930 | 2455116 | 157348 | 12777293 | 2535358 | 10454486 | 31475 | 5454048 | 534838 | 1550301 | N/F | 2106 | 23217998 |  | Unknown
AB-220715-01 LRB 9_INJ001 | N/F | 6898 | 878 | 676 | 90524 | 707 | 3232848 | 426 | 6204 | 1565 | 1063 | 2813 | 17436 | 17785 | 131630 | 359261 | 430506 | 838718 | 2252011 | 7418345 | 2983127 | 607329 | 29943930 | 6482352 | 15068280 | 2921618 | 18439011 | 69981 | 901 | N/F | 351091 | 2161 | 7 | 11932 | N/F | 42210 | N/F | 123 | N/F | 5141 | 145415 | 61688 | 4445 | 102768 | 4143 | 14637 | 5730 | 742 | N/F | N/F | 39976 |  | Unknown
AB-220715-01 LRB 9 | N/F | 6187 | 967 | 198 | 22611 | 995 | 2854269 | 152 | 1117 | 2415 | 1182 | 1426 | 16617 | 4877 | 127754 | 331799 | 400328 | 809751 | 1976682 | 6810433 | 2903771 | 556892 | 26409809 | 5926337 | 13921341 | 2677379 | 17044426 | 48600 | 1424 | 2477 | 345555 | 1205 | 103 | 11073 | N/F | 20834 | N/F | 22 | 4847 | 4203 | 99478 | 66876 | 4691 | 79033 | 4612 | 3444 | 5033 | 656 | N/F | N/F | 38388 |  | Unknown
AB-220715-01 LD 4 | 787 | 9929 | 492 | 203 | 9337 | N/F | 3371119 | 228 | 1319 | 1955 | 2210 | 2638 | N/F | 17149 | 43730 | 138563 | 430349 | 178794 | 672550 | 2294548 | 716808 | 83682 | 21008614 | 4560525 | 5339548 | 2022776 | 10202374 | 8799 | N/F | N/F | 82518 | 1696 | N/F | 9291 | N/F | 344332 | N/F | 961 | 23514 | 5730 | 124757 | 41744 | 1581 | 76384 | 6456 | 1927 | N/F | 218 | N/F | 1608 | 25873 |  | Unknown
AB-220715-01 CCC-H 5 | 1172 | 214490 | 344224 | 548232 | 99461190 | 527 | 2936305 | 2782037 | 8952549 | 862627 | 1454306 | 4263034 | 2963542 | 4047002 | 110818 | 180671 | 270352 | 727228 | 1486649 | 5634422 | 2693442 | 567500 | 18565659 | 5306374 | 10341770 | 2060162 | 12768315 | 62264 | 514 | N/F | 335321 | 3291555 | 47 | 4489163 | 98549 | 33619431 | N/F | 621 | 10292609 | 4937353 | 223687 | 22650407 | 4460795 | 18357758 | 65069 | 10758132 | 1083022 | 2988149 | N/F | 740 | 39887512 |  | Unknown
AB-220715-01 LFSM-H 4 | N/F | 39296 | 112519 | 334592 | 48480791 | 583 | 3426306 | 360623 | 2165387 | 116974 | 198809 | 1348886 | 665615 | 1727331 | 38088 | 112861 | 343129 | 197978 | 709155 | 2391526 | 796480 | 88530 | 18544192 | 4643841 | 5148850 | 1943838 | 9525681 | 12178 | N/F | N/F | 92993 | 687604 | 197 | 2743178 | N/F | 17521458 | N/F | 2089 | 2249845 | 733926 | 154873 | 8739401 | 2652123 | 4758056 | 15459 | 4735905 | 83428 | 522033 | N/F | 1167 | 22429132 |  | Unknown
WO-22270-81 | N/F | 8305 | 641 | 199 | 24432 | 459 | 3111289 | 34 | 1315 | 1405 | 1047 | 3504 | N/F | 23967 | 28323 | 121074 | 366521 | 123221 | 585225 | 1890612 | 501260 | 50803 | 18293922 | 4205561 | 4423039 | 1690896 | 8639728 | 6348 | 1245 | 1020 | 56564 | 3770 | 358 | 11874 | N/F | 302441 | N/F | 670 | 14503 | 6331 | 129422 | 12073 | 3130 | 27237 | 7777 | 3288 | N/F | 367 | N/F | N/F | 16563 |  | Unknown
WO-22270-82 | N/F | 9428 | 1841 | 1510 | 16569 | 876 | 5036626 | N/F | 1321 | 2071 | 800 | 4478 | N/F | 3291 | 40997 | 177948 | 1082853 | 197317 | 816846 | 2710178 | 800165 | 89713 | 24731305 | 6124641 | 6277768 | 2372708 | 12553357 | 10423 | 551 | 890 | 89715 | 1529 | 30 | 11445 | N/F | 389467 | N/F | 1855 | 38908 | 9913 | 140968 | 13121 | 5010 | 60602 | 6780 | 3339 | N/F | 416 | N/F | N/F | 16749 |  | Unknown
WO-22270-83 | N/F | 12084 | 1356 | 350 | 13535 | 1860 | 3904656 | N/F | 835 | 1717 | 883 | 4139 | N/F | 16992 | 45208 | 161946 | 513506 | 205795 | 789113 | 2669831 | 844460 | 95023 | 24460084 | 5757290 | 6020994 | 2323868 | 11784212 | 10377 | 1001 | N/F | 94527 | 2042 | 154 | 9702 | N/F | 380924 | N/F | 1758 | N/F | 12967 | 136701 | 17280 | 6342 | 17751 | 3247 | 2554 | N/F | 518 | N/F | N/F | 10618 |  | Unknown
WO-22270-84 | N/F | 8232 | 420 | 447 | 11714 | 1058 | 3662467 | N/F | 233 | N/F | 2285 | 3281 | N/F | 3946 | 39531 | 155791 | 502287 | 181820 | 750904 | 2546463 | 772173 | 90614 | 23008855 | 5423355 | 5720144 | 2220971 | 11441580 | 10508 | N/F | N/F | 91522 | 723 | N/F | 11169 | N/F | 250306 | N/F | 1229 | N/F | 35871 | 143166 | 10404 | 4790 | 41011 | 1320 | 3532 | N/F | 181 | N/F | N/F | 12801 |  | Unknown
WO-22270-85 | N/F | 8213 | 389 | 389 | 10466 | 609 | 3742549 | 7 | 621 | 1144 | 888 | 7525 | N/F | 20926 | 39565 | 150225 | 566827 | 203653 | 758643 | 2643643 | 780033 | 101864 | 24050284 | 5436740 | 5866912 | 2293834 | 11662114 | 10741 | 949 | 2955 | 91493 | 1202 | 278 | 10502 | N/F | 349758 | N/F | 1872 | 25406 | 28135 | 133755 | 5274 | 4864 | 43425 | 6310 | 2734 | N/F | 441 | 1339 | N/F | 14726 |  | Unknown
WO-22270-86 | 775 | 9313 | 895 | 534 | 10473 | 809 | 5056926 | N/F | 261 | 1240 | 731 | 4685 | N/F | 4299 | 40135 | 170124 | 1026824 | 189036 | 766801 | 2512472 | 757330 | 82981 | 23675098 | 5729521 | 6104626 | 2335104 | 11998524 | 10626 | 1253 | 563 | 86556 | 1870 | 275 | 11785 | N/F | 344562 | N/F | 1857 | 40915 | 27535 | 142258 | 11327 | 5469 | 92389 | 1085 | 1459 | N/F | 134 | N/F | N/F | 13852 |  | Unknown
WO-22270-87 | 975 | 9242 | 266 | 72 | 8556 | 780 | 5251589 | 26 | 428 | N/F | N/F | 3652 | N/F | 15370 | 41923 | 166079 | 996984 | 202143 | 753415 | 2590736 | 795368 | 89935 | 22623919 | 5981619 | 6274669 | 2416998 | 12290455 | 10305 | N/F | N/F | 86474 | 2193 | 76 | 10567 | N/F | 380055 | N/F | 1528 | 24468 | 108027 | 143887 | 15496 | 4433 | 11044 | 4720 | 3067 | N/F | 568 | N/F | 2251 | 14997 |  | Unknown
WO-22270-88 | N/F | 10272 | 945 | 181 | 6947 | 500 | 4014703 | 10 | 177 | 1501 | 760 | 3098 | N/F | 18592 | 33766 | 141890 | 508439 | 181172 | 741640 | 2504749 | 709023 | 80438 | 23474682 | 5391922 | 5546647 | 2202924 | 11225127 | 9324 | 804 | 25220582 | 83236 | 3233 | 20 | 11768 | N/F | 453097 | N/F | 1699 | N/F | 97764 | 144953 | 10936 | 4039 | 14225 | 6776 | 3066 | N/F | 360 | 5138 | N/F | 10300 |  | Unknown
WO-22270-89 | 573 | 8969 | 952 | 189 | 4256 | 580 | 3134655 | 29 | 7 | 759 | 643 | 3878 | N/F | 16214 | 23877 | 116216 | 385850 | 113973 | 574467 | 1852675 | 420888 | 46694 | 18307018 | 4044882 | 4370898 | 1660729 | 8539778 | 7368 | 276 | N/F | 47149 | 834 | 202 | 9326 | N/F | 306446 | N/F | 1543 | 27394 | 68153 | 133721 | 20029 | 1669 | 24498 | 6108 | 1633 | N/F | 158 | N/F | N/F | 31980 |  | Unknown
WO-22270-90 | 908 | 7161 | 687 | 261 | 4525 | 453 | 3091452 | N/F | 145 | 903 | 580 | 2379 | N/F | 3320 | 27620 | 115994 | 365302 | 130434 | 587176 | 1982942 | 524788 | 55299 | 18276354 | 3819111 | 4568482 | 1514821 | 7750832 | 7240 | 572 | 1226 | 59977 | 1359 | 197 | 10187 | N/F | 277744 | N/F | 1340 | 16901 | 84809 | 132586 | N/F | 3162 | N/F | 5156 | 2013 | N/F | 301 | N/F | N/F | 10142 |  | Unknown
AB-220715-01 LRB 10_INJ001 | N/F | 7674 | 825 | 320 | 7950 | 644 | 2869361 | N/F | 187 | 1804 | 1589 | 1256 | N/F | 4648 | 124925 | 325823 | 364032 | 746001 | 1996278 | 7035897 | 2761186 | 519819 | 26542230 | 5577756 | 14133799 | 2386365 | 16210796 | 51251 | 902 | 658 | 294329 | 2325 | 19 | 9080 | N/F | 24094 | N/F | N/F | N/F | N/F | 124659 | 65741 | 3691 | 73607 | 4204 | 1872 | 5750 | 886 | N/F | N/F | 45985 |  | Unknown
AB-220715-01 LRB 10 | 1363 | 8723 | 778 | 19 | 7780 | 1508 | 2141288 | 22 | 98 | 1595 | 483 | 1500 | N/F | 5824 | 85019 | 247579 | 326823 | 504579 | 1417805 | 5090765 | 1855802 | 375550 | 20692741 | 4535551 | 10464064 | 2159872 | 13478997 | 18497 | 1095 | 734 | 210863 | 2457 | N/F | 8742 | N/F | 23909 | N/F | N/F | N/F | 3141 | 78968 | 47147 | 4414 | 59437 | 1486 | 1286 | 4009 | 190 | N/F | N/F | 26474 |  | Unknown
WO-22270-91 | N/F | 7002 | 219 | 239 | 8069 | N/F | 3812237 | N/F | 473 | N/F | 553 | 3280 | N/F | 4881 | 41354 | 155800 | 468885 | 199719 | 746753 | 2566212 | 748555 | 85774 | 23113498 | 5284956 | 5809917 | 2198104 | 11308729 | 11147 | N/F | 1326 | 84669 | 1477 | 106 | 10886 | N/F | 296888 | N/F | 1646 | 13086 | 316158 | 121358 | 5055 | 3385 | 65642 | 2083 | 1491 | N/F | 332 | N/F | N/F | 16417 |  | Unknown
WO-22270-92 | 781 | 12486 | 626 | 215 | 7929 | 1041 | 4747567 | 28 | 84 | 1461 | 974 | 3192 | N/F | 5649 | 38339 | 150126 | 840361 | 211782 | 746938 | 2698378 | 832541 | 96188 | 24576698 | 5678325 | 5952482 | 2382566 | 12500813 | 11620 | N/F | 729 | 89135 | 2199 | 144 | 10830 | N/F | 417804 | N/F | 1988 | 45749 | 468269 | 138780 | 13864 | 3607 | 39305 | 2097 | 1848 | N/F | 347 | N/F | N/F | 19522 |  | Unknown
WO-22270-93 | N/F | 9196 | 648 | 392 | 7428 | 1534 | 4163346 | 251 | 40 | 1035 | 2250 | 2990 | N/F | 16394 | 45781 | 164928 | 533461 | 224370 | 805398 | 2774581 | 863242 | 103146 | 24694734 | 5666683 | 6257337 | 2352554 | 12289407 | 12938 | 1408 | 670471 | 99668 | 1354 | 183 | 11335 | N/F | 475445 | N/F | 1213 | 36010 | 391620 | 146214 | 14555 | 1058 | 69532 | 5926 | 2302 | N/F | 270 | N/F | 1692 | 35299 |  | Unknown
WO-22270-94 | N/F | 7849 | 2148 | 413 | 6136 | 975 | 5311255 | N/F | 68 | 1340 | 891 | 2734 | N/F | 3458 | 35484 | 173837 | 1271845 | 169459 | 715334 | 2450242 | 666687 | 78552 | 22499349 | 5566900 | 5846242 | 2195247 | 12149788 | 10572 | N/F | 22544708 | 80065 | 1060 | 80 | 10624 | N/F | 372922 | N/F | 1100 | 34928 | 1634577 | 131701 | 16992 | 619 | 108505 | 6438 | 2115 | 870 | 463 | 4841 | N/F | 35962 |  | Unknown
WO-22270-95 | 1093 | 9120 | 362 | 60 | 5304 | 965 | 3276199 | 7 | 122 | 1057 | 1064 | 4188 | N/F | 3386 | 34155 | 129496 | 424027 | 141093 | 616178 | 2109594 | 565736 | 63022 | 20794691 | 4462463 | 4906830 | 1862901 | 9458542 | 7724 | N/F | N/F | 64144 | 1331 | N/F | 10489 | N/F | 331029 | N/F | 1891 | 10709 | 800072 | 112670 | 8579 | 625 | 20322 | 1245 | 1098 | N/F | 134 | N/F | N/F | 9938 |  | Unknown
WO-22270-96 | N/F | 7304 | 394 | 361 | 6595 | N/F | 3634137 | 12 | 11 | N/F | 432 | 4036 | N/F | 6001 | 42375 | 145047 | 462509 | 211412 | 748224 | 2733224 | 813660 | 100463 | 23826961 | 5443547 | 5844232 | 2230558 | 11541020 | 12916 | 771 | 32251827 | 97388 | 1216 | 37 | 11443 | N/F | 413690 | N/F | 1438 | 11665 | 1666284 | 129368 | 7982 | 6074 | 51197 | 7520 | 2501 | N/F | 25 | 7000 | N/F | 12556 |  | Unknown
WO-22270-97 | N/F | 7703 | 1100 | 93 | 6423 | 905 | 3550060 | N/F | 346 | 1210 | 605 | 3355 | N/F | 3907 | 39361 | 146435 | 429484 | 210196 | 743371 | 2621815 | 843725 | 107670 | 23791887 | 5256730 | 5748496 | 2214819 | 11492106 | 12260 | N/F | N/F | 92828 | 1490 | 151 | 9562 | N/F | 386211 | N/F | 1103 | N/F | 5893 | 139773 | 16496 | 2946 | 19195 | 7907 | 1503 | N/F | 445 | 3407 | N/F | 14296 |  | Unknown
WO-22270-98 | N/F | 7374 | 937 | 429 | 6318 | 1070 | 3337603 | N/F | 8 | 2142 | 532 | 2140 | N/F | 2917 | 37518 | 131347 | 411539 | 215239 | 750296 | 2577667 | 805018 | 106408 | 23127370 | 5209000 | 5652528 | 2104134 | 11383458 | 10973 | 1200 | 32172240 | 90782 | N/F | 310 | 10383 | N/F | 466992 | N/F | 1613 | N/F | 4660 | 122663 | 9495 | 4229 | 8883 | 6621 | 1673 | N/F | 398 | 8015 | N/F | 13964 |  | Unknown
WO-22270-99 | N/F | 7009 | 1051 | 393 | 2769 | 532 | 2020279 | N/F | 383 | 639 | 439 | 2034 | N/F | 12759 | 17146 | 79828 | 256072 | 66716 | 412706 | 1239030 | 287145 | 32303 | 12534084 | 2466790 | 3190425 | 984422 | 5970633 | 6270 | 1106 | 417 | 30747 | 540 | 288 | 6705 | N/F | 244377 | N/F | 761 | N/F | N/F | 153933 | 21247 | 2136 | N/F | 6889 | 684 | N/F | N/F | 4505 | N/F | 7031 |  | Unknown
WO-22270-100 | N/F | 9741 | 225 | 42 | 4286 | 879 | 2945899 | N/F | N/F | N/F | 1128 | 1706 | N/F | 2278 | 27446 | 120132 | 333347 | 122929 | 565199 | 1890414 | 505037 | 57395 | 17880141 | 3864311 | 4410993 | 1587015 | 8549173 | 6952 | 1729 | N/F | 54327 | 2393 | N/F | 9058 | N/F | 345252 | N/F | 647 | 6357 | 2931 | 140567 | 7895 | 3131 | 29320 | 8353 | 987 | N/F | 165 | N/F | N/F | 24362 |  | Unknown
AB-220715-01 LFB-H 6 | 758 | 93352 | 177684 | 267333 | 52254895 | 713 | 2920160 | 1177742 | 4405881 | 320835 | 581243 | 2258951 | 1424820 | 2101300 | 94289 | 184156 | 268701 | 590622 | 1526333 | 5483197 | 2239385 | 454298 | 21637634 | 5221880 | 10675315 | 2205900 | 14485110 | 39912 | 568 | 3330155 | 265343 | 1604932 | 128 | 2120976 | 30479 | 19135615 | N/F | 115 | 5001643 | 2039056 | 178607 | 12232426 | 2556969 | 9646803 | 32291 | 5269778 | 412715 | 1244965 | 5234 | N/F | 23725077 |  | Unknown
AB-220715-01 LRB 11_INJ001 | N/F | 6558 | 1200 | 262 | 71967 | 505 | 2992995 | 344 | 5532 | 2907 | 987 | 2742 | N/F | 23425 | 125522 | 323862 | 401357 | 880847 | 2134600 | 7838529 | 3189079 | 56292 | 29425554 | 6370089 | 15551575 | 2808866 | 19371096 | 76916 | 575 | 37322337 | 377121 | 4827 | 282 | 11294 | N/F | 42416 | N/F | 4215 | N/F | N/F | 144120 | 48396 | 3445 | 94678 | 7401 | 11879 | 3351 | 683 | 7618 | 41498 | 36454 |  | Unknown
AB-220715-01 LRB 11 | 1697 | 8369 | 940 | 292 | 20516 | 342 | 2972327 | N/F | 727 | 1964 | 511 | 2164 | N/F | 17159 | 137342 | 317843 | 379481 | 853202 | 2004282 | 7611678 | 3153962 | 726873 | 27909548 | 5997042 | 15029356 | 2767757 | 18225009 | 53079 | N/F | 32040963 | 347007 | 2148 | 116 | 11166 | N/F | 24449 | N/F | 1486 | N/F | 4407 | 109124 | 46611 | 1374 | 84364 | 5010 | 4832 | 3453 | 362 | 6976 | N/F | 41219 |  | Unknown
AB-220715-01 LD 5 | N/F | 8281 | 874 | 42 | 6063 | 479 | 1986878 | 9 | 205 | 459 | 722 | 1870 | N/F | 16381 | 15506 | 90056 | 215599 | 66288 | 416502 | 1251945 | 270600 | 32221 | 12616466 | 2376266 | 3149240 | 1090966 | 6088699 | 4897 | N/F | N/F | 30142 | 1093 | 141 | 6619 | N/F | 262286 | N/F | 419 | N/F | 2738 | 120360 | 25155 | 3509 | 44862 | 6470 | 1045 | N/F | N/F | 3523 | 1552 | 21553 |  | Unknown
AB-220715-01 CCC-H 6 | N/F | 196917 | 308770 | 513602 | 96291946 | N/F | 2898925 | 2555889 | 8104449 | 709682 | 1282249 | 4087929 | 2657766 | 3921642 | 103456 | 163210 | 262392 | 618538 | 1392965 | 5330652 | 2363626 | 514750 | 18987306 | 5039727 | 10085369 | 2118416 | 13396215 | 52567 | 811 | 4566785 | 286156 | 3141215 | 347 | 4525629 | 87716 | 33730016 | 4903 | 916 | 9781006 | 4153258 | 251355 | 22660272 | 4480858 | 17705680 | 81763 | 9719129 | 995560 | 2449170 | 7012 | N/F | 41422884 |  | Unknown
AB-220715-01 LFSM-H 5 | N/F | 22399 | 78727 | 232164 | 34753341 | 1060 | 2425621 | 185907 | 1329522 | 60744 | 108395 | 850590 | 462375 | 1184609 | 20425 | 80095 | 234791 | 107413 | 510075 | 1644704 | 440231 | 51458 | 14363915 | 3269054 | 3657926 | 1348555 | 7062576 | 6486 | N/F | 18145210 | 49652 | 496955 | 107 | 1932644 | N/F | 13360699 | N/F | 1930001 | 1531504 | 385577 | 153142 | 6259048 | 1993524 | 3379075 | 35937 | 3320598 | 45943 | 331309 | 5236 | 1755 | 19208271 |  | Unknown
WO-22270-101 | N/F | 7780 | 1093 | 17 | 27790 | 536 | 3835357 | 121 | 2583 | 2655 | 618 | 1900 | N/F | 7610 | 41175 | 148820 | 439884 | 214921 | 839518 | 2897469 | 875364 | 108862 | 26517096 | 5629952 | 6443524 | 2482525 | 13267462 | 12883 | N/F | 35893482 | 94777 | 1086 | 612 | 12543 | N/F | 478153 | N/F | 1915 | 5694 | 9359 | 129595 | 11487 | 7569 | 13752 | 5261 | 4485 | N/F | 90 | 7319 | N/F | 20653 |  | Unknown
WO-22270-102 | 1302 | 5468 | 1173 | 48 | 15037 | 900 | 3311374 | 77 | 203 | 1290 | 1377 | 1201 | N/F | 17781 | 39713 | 135655 | 411209 | 206879 | 723586 | 2616679 | 821723 | 102235 | 23878042 | 5166919 | 5790764 | 2132395 | 11273981 | 11977 | N/F | 25279006 | 88647 | 1141 | 37 | 13065 | N/F | 430960 | N/F | 1461 | N/F | 8274 | 138971 | 13257 | 6669 | 39453 | 6840 | 3128 | N/F | 646 | 8080 | N/F | 15372 |  | Unknown
WO-22270-103 | N/F | 17675 | 411 | 16 | 17273 | 426 | 4474384 | 17 | 572 | 1557 | 910 | 3028 | N/F | 19696 | 52570 | 210505 | 635785 | 258852 | 1140323 | 3590592 | 1025468 | 122453 | 32941178 | 6517029 | 9321004 | 3157958 | 18248804 | 11421 | N/F | 32517906 | 109029 | 2597 | 126 | 11278 | N/F | 738232 | N/F | 2598 | N/F | 14244 | 105429 | 49312 | 5477 | 46365 | 3644 | 2183 | N/F | 296 | 7297 | 1166 | 15304 |  | Unknown
WO-22270-104 | N/F | 8195 | 963 | 33 | 13699 | 264 | 3153089 | 23 | 457 | 994 | 815 | 2394 | N/F | 4205 | 34321 | 119995 | 387220 | 146485 | 627835 | 2098333 | 594156 | 68013 | 20829860 | 4519697 | 4923165 | 1894019 | 9773223 | 7143 | 1092 | 26242891 | 68480 | 756 | N/F | 9908 | N/F | 368749 | N/F | 1487 | N/F | 21439 | 135476 | 5438 | 3532 | 8042 | 7071 | 2866 | N/F | 390 | 5697 | 983 | 10976 |  | Unknown
WO-22270-105 | N/F | 17011 | 346 | 208 | 8404 | 1922 | 4766902 | 10 | 291 | 1066 | 622 | 2374 | N/F | 4921 | 33101 | 154351 | 1055037 | 149845 | 667558 | 2284706 | 594750 | 66095 | 20703722 | 4675678 | 5264972 | 1880352 | 10364697 | 7358 | N/F | 27092787 | 65824 | 704 | 11 | 8593 | N/F | 443481 | N/F | 576 | N/F | 17952 | 136349 | 25151 | 3600 | 8182 | 4802 | 2643 | N/F | 166 | 4673 | N/F | 9214 |  | Unknown
WO-22270-106 | N/F | 13257 | 2867 | 1176 | 16756 | 2007 | 5150078 | 34 | 494 | N/F | 470 | 2591 | N/F | 20023 | 35669 | 164080 | 1051736 | 169669 | 693100 | 2492907 | 757298 | 1671 | 23012653 | 5140291 | 5855602 | 2110216 | 11728390 | 9382 | 186976 | 28027075 | 83047 | 2063 | 252 | 10414 | N/F | 412716 | N/F | 1404 | 12995 | 25480 | 147826 | 19912 | 6687 | 6750 | 7557 | 2991 | 3305 | 111 | 3315 | 12213 | 10647 |  | Unknown
WO-22270-107 | 503 | 8028 | 500 | 114 | 7274 | 892 | 5234659 | 7 | 87 | 1766 | 1876 | 2178 | N/F | 16985 | 42447 | 167909 | 1022189 | 221668 | 776790 | 2807323 | 868187 | 117322 | 24780276 | 5387118 | 6231741 | 2249895 | 12348570 | 12559 | N/F | 31670741 | 96248 | 2055 | N/F | 11158 | N/F | 393938 | N/F | 1758 | 15532 | 105019 | 125386 | 14272 | 4852 | 7556 | 4448 | 2782 | N/F | 62 | 4739 | N/F | 9548 |  | Unknown
WO-22270-108 | 1749 | 11423 | 901 | 1261 | 5788 | 746 | 4990634 | 10 | 132 | 2182 | 739 | 2000 | N/F | 5400 | 39607 | 172729 | 1279879 | 184261 | 738575 | 2679015 | 789601 | 102087 | 22216407 | 5568319 | 6080203 | 2179844 | 12446477 | 10918 | N/F | 29316635 | 86995 | N/F | 69 | 8935 | N/F | 361029 | N/F | 1566 | 19250 | 91402 | 129729 | 13852 | 1881 | 8002 | 5237 | 3020 | N/F | 166 | 5626 | N/F | 12194 |  | Unknown
WO-22270-109 | 596 | 9235 | 565 | 136 | 6666 | 1139 | 4391587 | 28 | 319 | 1376 | 1197 | 1840 | N/F | 3990 | 39024 | 146405 | 772287 | 176680 | 691625 | 2452194 | 714474 | 84760 | 23125269 | 4826802 | 5583680 | 2009241 | 11210476 | 9062 | N/F | 26090607 | 83904 | 1019 | 229 | 9574 | N/F | 433820 | N/F | 1279 | N/F | 123250 | 124799 | 55680 | 4812 | 8920 | 5824 | 1775 | N/F | 185 | 6523 | N/F | 12767 |  | Unknown
WO-22270-110 | N/F | 9725 | 697 | 29 | 5268 | 812 | 3766804 | 51 | 132 | 1680 | 237 | 2323 | N/F | 4289 | 37678 | 131146 | 604494 | 167364 | 639255 | 2422685 | 717074 | 89531 | 22851904 | 4833379 | 5485504 | 2044642 | 10807877 | 9122 | 1591 | 28846263 | 76826 | 1425 | 400 | 12332 | N/F | 396788 | N/F | 1368 | N/F | 120661 | 130326 | 50700 | 3235 | 30208 | 4335 | 1275 | N/F | 458 | 5800 | 1145 | 12901 |  | Unknown
AB-220715-01 LRB 12_INJ001 | 486 | 9746 | 610 | 319 | 8604 | 787 | 2783402 | 77 | 40 | 2183 | 804 | 2673 | N/F | 16940 | 104101 | 282971 | 367593 | 648203 | 1718292 | 6322451 | 2598668 | 40953 | 27455193 | 5781197 | 13146927 | 2596459 | 17583313 | 55209 | 1124 | 34292524 | 289548 | 1754 | 40 | 10418 | N/F | 17547 | N/F | 369 | N/F | N/F | 136448 | 7426 | 5456 | 14392 | 8259 | 2823 | N/F | 134 | 6755 | 11049 | 10183 |  | Unknown
AB-220715-01 LRB 12 | N/F | 10332 | 395 | 212 | 9696 | 420 | 2494929 | 242 | 103 | 1590 | 1744 | 2208 | N/F | 6887 | 105914 | 270566 | 341357 | 653605 | 1613514 | 6247041 | 2474200 | 514414 | 24969885 | 5143456 | 12028132 | 2430888 | 15855749 | 24568 | 582 | 31796499 | 272721 | 1482 | 502 | 9195 | N/F | 17519 | N/F | 677 | N/F | N/F | 96961 | 5700 | 4142 | 13572 | 4227 | 1702 | 2049 | 45 | 7824 | 1122 | 10192 |  | Unknown
WO-22270-111 | 826 | 7144 | 382 | 31 | 8150 | 1530 | 3227456 | N/F | 107 | 834 | 728 | 2494 | N/F | 2783 | 30058 | 121368 | 406464 | 159741 | 607635 | 2185624 | 629552 | 71022 | 20700911 | 4402708 | 4971085 | 1801407 | 9759566 | 8953 | 653 | 30426687 | 72226 | 1221 | 51 | 9791 | N/F | 323831 | N/F | 708 | N/F | 214478 | 102581 | 11546 | 3104 | 11205 | 7150 | 1090 | N/F | 648 | 6126 | N/F | 10521 |  | Unknown
WO-22270-112 | 855 | 11443 | 1212 | 265 | 5052 | 1092 | 4799181 | N/F | 80 | 1706 | 917 | 1860 | N/F | 6653 | 35402 | 145063 | 1023248 | 168350 | 659295 | 2444290 | 724965 | 86105 | 22186783 | 5029725 | 5706817 | 2029749 | 11568769 | 8481 | N/F | 26945968 | 82708 | 2022 | 356 | 10316 | N/F | 359615 | N/F | 1381 | N/F | 247544 | 133925 | 31902 | 4175 | 9679 | 6479 | 1682 | N/F | 57 | 4926 | N/F | 11627 |  | Unknown
WO-22270-113 | 888 | 8638 | 812 | N/F | 12296 | 639 | 3115066 | N/F | 14 | 1234 | 1095 | 2019 | N/F | 22242 | 42574 | 141736 | 373101 | 187370 | 671160 | 2541811 | 828045 | 88615 | 21945106 | 4602464 | 5569755 | 1882586 | 10515390 | 10972 | N/F | 31623082 | 88396 | 746 | N/F | 9017 | N/F | 340596 | N/F | 1140 | 3195 | 369729 | 120342 | 6166 | 5220 | 18946 | 5345 | 1302 | N/F | 168 | 7099 | 1879 | 10675 |  | Unknown
WO-22270-114 | 1733 | 8496 | 725 | 182 | 8207 | 1193 | 3725906 | N/F | 16 | N/F | 664 | 1934 | N/F | 18438 | 26781 | 118616 | 660687 | 133810 | 588942 | 1973235 | 509911 | 4079 | 20575282 | 4140382 | 4792766 | 1760651 | 9550675 | 7648 | N/F | 28763999 | 58578 | 1870 | 28 | 10289 | N/F | 386865 | N/F | 1091 | 7104 | 793736 | 122090 | 14087 | 4076 | 27204 | 3982 | 1342 | N/F | 214 | 5251 | 6787 | 13248 |  | Unknown
WO-22270-115 | N/F | 9971 | 1667 | 352 | 7114 | 614 | 4214793 | N/F | 88 | 2335 | 795 | 2707 | N/F | 6824 | 41002 | 153119 | 931989 | 186150 | 684544 | 2544702 | 760625 | 70308 | 21189412 | 5084282 | 6071450 | 2120464 | 11883155 | 10171 | N/F | 28615700 | 86584 | 3279 | 85 | 9351 | N/F | 420953 | N/F | 775 | N/F | 804229 | 130067 | 134021 | 2518 | 53993 | 7007 | 1680 | N/F | 393 | 4618 | N/F | 14157 |  | Unknown
WO-22270-116 | 370 | 13573 | 2629 | 574 | 10181 | 1447 | 5118525 | 37 | 58 | N/F | 2244 | 3800 | N/F | 4800 | 36663 | 142183 | 1077206 | 203753 | 705787 | 2678440 | 764118 | 5131 | 23720312 | 5224879 | 5794834 | 2244457 | 12331419 | 10969 | 158199 | 30307098 | 83753 | 1713 | N/F | 9218 | N/F | 384390 | N/F | 1318 | N/F | 872699 | 142850 | 13338 | 4697 | 35997 | 8773 | 3451 | N/F | 184 | 4594 | 14547 | 20536 |  | Unknown
WO-22270-117 | N/F | 6647 | 815 | 17 | 6447 | 853 | 2969305 | N/F | N/F | 1754 | 1385 | 3053 | N/F | 3559 | 37591 | 124895 | 312576 | 179288 | 633138 | 2334350 | 724122 | 65549 | 18305780 | 4357596 | 5166087 | 1775753 | 9898041 | 9589 | 671 | 29085631 | 77324 | 1934 | 56 | 9905 | N/F | 2358530 | N/F | 1825 | 3755 | 4947 | 147306 | 11813 | 2821 | 20667 | 6984 | 2667 | N/F | 482 | 6538 | 1828 | 18040 |  | Unknown
WO-22270-118 | 1004 | 9648 | 1868 | 107 | 5639 | N/F | 4617915 | 42 | 100 | N/F | 717 | 3359 | N/F | 3273 | 36708 | 131000 | 852029 | 172905 | 660948 | 2484436 | 695421 | N/F | 22323639 | 4580744 | 5271343 | 1922212 | 10991257 | 9560 | N/F | 29292081 | 75347 | 1074 | N/F | 8403 | N/F | 279353 | N/F | 608 | N/F | 4750 | 134924 | 8997 | 5079 | 9643 | 7538 | 1675 | N/F | 280 | 7569 | 16117 | 14987 |  | Unknown
WO-22270-119 | 794 | 10156 | 511 | 441 | 11946 | 8467 | 3300312 | 11 | 78 | N/F | 978 | 3769 | N/F | 20063 | 25942 | 102191 | 467353 | 130696 | 579807 | 2031972 | 526616 | N/F | 18653659 | 3685317 | 4337529 | 1501090 | 8736161 | 6734 | 345 | 26572196 | 57455 | 703 | N/F | 9167 | N/F | 340759 | N/F | 994 | N/F | 2934 | 130554 | 19007 | 3004 | 30699 | 6715 | 2097 | 4432 | 382 | 5669 | 17897 | 11313 |  | Unknown
WO-22270-120 | N/F | 13425 | 464 | 91 | 4221 | 1119 | 3566552 | 42 | N/F | N/F | 559 | 2342 | N/F | 4412 | 29187 | 124746 | 475819 | 164834 | 625207 | 2233921 | 700236 | N/F | 19981202 | 4411901 | 5097827 | 1830436 | 10252495 | 9573 | N/F | 28645941 | 71437 | 902 | 97 | 10452 | N/F | 374246 | N/F | 904 | N/F | 3697 | 132565 | 32982 | 4884 | 38221 | 6052 | 1370 | N/F | 216 | 7732 | 13548 | 16714 |  | Unknown
AB-220715-01 LFB-H 7 | 978 | 93569 | 157739 | 217365 | 43323859 | 28268 | 2467414 | 1147400 | 4025961 | 340943 | 595780 | 2084872 | 1321371 | 1980230 | 83936 | 162087 | 237570 | 588835 | 1348651 | 5292971 | 2320535 | 9920 | 19325253 | 4954955 | 10225391 | 2121381 | 14186538 | 52308 | N/F | 25294292 | 271714 | 1529659 | 244963 | 1975252 | 43462 | 17289254 | N/F | 2076662 | 4915900 | 2124906 | 176266 | 12446934 | 2077126 | 9278982 | 37898 | 4972102 | 457899 | 1184998 | 6762 | 36946697 | 22834740 |  | Unknown
AB-220715-01 LRB 13_INJ001 | N/F | 7883 | 262 | 236 | 60667 | 1145 | 2710589 | 654 | 4389 | 2242 | 856 | 4312 | N/F | 10045 | 108854 | 261738 | 331664 | 677789 | 1653649 | 6162095 | 2577659 | 9296 | 25459165 | 5377769 | 12666893 | 2487994 | 16585527 | 51800 | N/F | 34450912 | 298463 | 1951 | 45 | 12011 | N/F | 38896 | N/F | 3080 | N/F | 5229 | 141037 | 15298 | 7838 | 17956 | 7136 | 11246 | N/F | 393 | 7257 | 52089 | 41995 |  | Unknown
AB-220715-01 LRB 13 | N/F | 8739 | 626 | N/F | 16337 | 662 | 1995875 | 291 | 1662 | 1468 | 1292 | 882 | N/F | 4832 | 85307 | 223761 | 277068 | 602438 | 1417471 | 5592209 | 2317303 | 87616 | 19935037 | 4161700 | 10266887 | 1896730 | 13284930 | 18202 | 82355 | 27098095 | 246703 | 1157 | 18 | 8156 | N/F | 16657 | N/F | 426 | N/F | 2774 | 80629 | 6713 | 4101 | 15535 | 3461 | 4517 | 1171 | 474 | 5304 | 18300 | 15883 |  | Unknown
AB-220715-01 LD 6 | N/F | 7780 | 1540 | 887 | 12794 | 3420775 | 3272663 | 20 | 352 | N/F | 1000 | 1598 | N/F | 2676 | 25592 | 102465 | 459732 | 121477 | 520854 | 1889293 | 527612 | 1208 | 18592197 | 3523377 | 4320638 | 1486179 | 8631675 | 6918 | 33089542 | 26544965 | 56706 | 1197 | 571 | 8339 | N/F | 320260 | N/F | 596 | 19181 | 3223 | 152423 | 22188 | 6443 | 30719 | 9102 | 2031 | N/F | 545 | 4664 | 19691 | 12960 |  | Unknown
AB-220715-01 CCC-H 7 | N/F | 192409 | 273859 | 392149 | 84780714 | 208258 | 2422180 | 2368855 | 7566830 | 728256 | 1190381 | 3794292 | 2440505 | 3685206 | 98591 | 143374 | 210135 | 599962 | 1272681 | 5150143 | 2451488 | 13192 | 16990268 | 4860963 | 9574526 | 1962335 | 13051808 | 57256 | 24852574 | 22383436 | 285666 | 2927674 | 699532 | 3800535 | 93615 | 29662858 | N/F | 3540329 | 9415086 | 4391961 | 240664 | 22511073 | 3598328 | 17477003 | 78024 | 9355863 | 990913 | 2402540 | 5193 | 69204424 | 38991038 |  | Unknown
AB-220715-01 LFSM-H 6 | N/F | 29203 | 88546 | 363425 | 40294448 | 2174 | 3446732 | 265530 | 1552368 | 91104 | 151097 | 990717 | 498624 | 1446391 | 28785 | 88168 | 362227 | 143272 | 547779 | 2057003 | 623203 | 1340 | 17823352 | 3875227 | 4488239 | 1676334 | 9111106 | 7883 | N/F | 23802449 | 67886 | 625660 | 283616 | 2699715 | N/F | 16790034 | N/F | 2213197 | 1964343 | 562460 | 177466 | 7982366 | 2345309 | 4230287 | 42973 | 3905607 | 69205 | 420087 | 4955 | 34499758 | 21692791 |  | Unknown
WO-22270-121 | N/F | 10243 | 527 | 125 | 18186 | N/F | 3346866 | 38 | 1550 | 1010 | 680 | 3253 | N/F | 8165 | 39530 | 130133 | 439426 | 224847 | 663452 | 2607042 | 872626 | 6491 | 18940059 | 4276007 | 5454144 | 1834652 | 10044035 | 12286 | 1281 | 28037079 | 94550 | 1226 | 374 | 9850 | N/F | 310527 | N/F | 1568 | N/F | 8873 | 142331 | 7412 | 6196 | N/F | 5746 | 3804 | 1306 | 55 | 4516 | 4736 | 18785 |  | Unknown
WO-22270-122 | N/F | 8578 | 988 | 17 | 8019 | 395 | 1962344 | 10 | 1166 | N/F | 894 | 1998 | N/F | 2232 | 17755 | 80274 | 221875 | 76465 | 432530 | 1402745 | 341178 | N/F | 13693187 | 2599912 | 3497743 | 1091594 | 6324605 | 5276 | 10858558 | 22223242 | 37659 | 847 | 192 | 7371 | N/F | 234144 | N/F | 1279 | N/F | 5618 | 123086 | 13440 | 7040 | 15085 | 5220 | 1511 | N/F | 128 | 5410 | 20522 | 11932 |  | Unknown
WO-22270-123 | N/F | 8318 | 2228 | 517 | 10452 | 614616 | 4742762 | N/F | 536 | N/F | 465 | 4764 | N/F | 26394 | 32527 | 128336 | 1031737 | 187488 | 585518 | 2343953 | 716815 | N/F | 21031285 | 4403054 | 5333152 | 1912900 | 11064215 | 9628 | 25733920 | 27745914 | 73169 | 1309 | 456 | 9267 | N/F | 381921 | N/F | 633 | N/F | 9823 | 130187 | 10571 | 6036 | 6673 | 8039 | 3599 | N/F | 94 | 3786 | 22525 | 13481 | D,Q | Unknown
WO-22270-124 | N/F | 12360 | 1735 | 243 | 8863 | 3418 | 4502979 | N/F | 468 | N/F | 774 | 2052 | N/F | 5754 | 38426 | 143232 | 847099 | 194237 | 698798 | 2698175 | 796284 | 3584 | 20479034 | 4615263 | 5699289 | 1990011 | 11270126 | 11548 | N/F | 29015502 | 89594 | 968 | N/F | 8455 | N/F | 415043 | N/F | 1152 | N/F | 31180 | 152101 | 17305 | 6374 | 81684 | 10025 | 1872 | N/F | 357 | 4677 | 20952 | 12617 |  | Unknown
WO-22270-125 | 1225 | 13423 | 7145 | 532 | 20382 | 4291 | 5032037 | 14 | 146 | N/F | 1854 | 3524 | N/F | 2895 | 39321 | 281848 | 1297419 | 188027 | 675536 | 2529167 | 780730 | 3257 | 19296311 | 5132345 | 6037692 | 1941943 | 10993672 | 12097 | 5478569 | 26748912 | 76656 | 2872 | 420 | 8501 | N/F | 429275 | N/F | 933 | N/F | 27801 | 120977 | 33743 | 7287 | 32679 | 6943 | 3805 | 1930 | 136 | 4899 | 26663 | 29774 |  | Unknown
WO-22270-126 | N/F | 6507 | 2567 | 81 | 8040 | 661 | 3806806 | N/F | 53 | N/F | 1051 | 3399 | N/F | 2943 | 26233 | 122073 | 636822 | 117206 | 531254 | 1943876 | 528099 | 751 | 18499001 | 3614134 | 4521694 | 1525797 | 8799705 | 7859 | 982148 | 27128926 | 54211 | 4482 | 642 | 6918 | N/F | 324594 | N/F | 790 | N/F | 20550 | 127694 | 25164 | 5251 | 37246 | 7510 | 1829 | N/F | 348 | 6551 | 21066 | 12203 |  | Unknown
WO-22270-127 | N/F | 5302 | 1034 | 458 | 4998 | 1442 | 2750708 | N/F | 438 | N/F | 763 | 1819 | N/F | 2359 | 30992 | 113753 | 304044 | 145717 | 589442 | 2298372 | 625802 | 3620 | 17297313 | 3645695 | 4825909 | 1437646 | 8553778 | 10110 | 584 | 27287211 | 70416 | 1775 | 59 | 10012 | N/F | 298617 | N/F | 981 | N/F | 81746 | 116200 | 9339 | 6266 | 55508 | 7689 | 1341 | N/F | 81 | 5379 | 19944 | 9109 |  | Unknown
WO-22270-128 | N/F | 6631 | 1055 | 100 | 6814 | 378 | 3121770 | N/F | 33 | N/F | 1831 | 3002 | N/F | 8992 | 28550 | 111707 | 393039 | 145476 | 602592 | 2268452 | 626511 | 2659 | 19982121 | 3898364 | 4834988 | 1634387 | 9499258 | 8183 | N/F | 28858281 | 63632 | 1919 | 58 | 8495 | N/F | 308845 | N/F | 846 | N/F | 98720 | 128830 | 11881 | 4655 | 29249 | 7461 | 1709 | N/F | 430 | 7588 | 18071 | 11308 |  | Unknown
WO-22270-129 | N/F | 9793 | 2037 | 106 | 5890 | 891 | 4363666 | N/F | 105 | 867 | 463 | 1091 | N/F | 4164 | 35759 | 129362 | 827472 | 187501 | 625168 | 2399759 | 745632 | 23632 | 23228761 | 4588869 | 5481662 | 1991819 | 11240230 | 9689 | N/F | 30378953 | 77786 | 2421 | 63 | 11657 | N/F | 431246 | N/F | 846 | 31601 | 133282 | 141067 | 30092 | 816 | 62780 | 7855 | 1653 | N/F | 238 | 4870 | N/F | 11223 |  | Unknown
WO-22270-130 | N/F | 14532 | 2338 | 495 | N/F | 4459 | 3950125 | N/F | N/F | N/F | 565 | 1880 | N/F | 1705 | 18842 | 114440 | 764223 | 91354 | 470343 | 1717563 | 416689 | 1154 | 17402469 | 3322756 | 4352324 | 1412516 | 8791448 | 7022 | N/F | 24236317 | 50714 | 1636 | 40 | 8402 | N/F | 460842 | N/F | 892 | 21686 | 77218 | 138307 | 50982 | 3982 | 430240 | 9130 | 1401 | N/F | 248 | 6006 | 16405 | 12223 |  | Unknown
AB-220715-01 LRB 14_INJ001 | N/F | 7677 | 1702 | 101 | 5562 | 468 | 2482971 | 35 | 136 | 1840 | 1215 | 2326 | N/F | 10254 | 88950 | 234898 | 322193 | 605013 | 1479328 | 5808207 | 2390547 | 5671 | 24648982 | 4996868 | 11851395 | 2218538 | 15909323 | 49023 | 37095111 | 32739726 | 265948 | 1289 | 379 | 8951 | N/F | 18502 | N/F | 338 | 3158 | 4604 | 143390 | 5670 | 3621 | 17286 | 9071 | 1895 | N/F | 120 | 5147 | 19384 | 15288 |  | Unknown
AB-220715-01 LRB 14 | 421 | 7296 | 1730 | 84 | 10567 | 1837976 | 1845814 | N/F | 97 | 1476 | 784 | 2455 | N/F | 8467 | 64644 | 199445 | 249140 | 439049 | 1164324 | 4588548 | 1665274 | N/F | 19358553 | 3749485 | 9277931 | 1765251 | 12816098 | 17287 | 30284513 | 25021746 | 184050 | 1485 | 171 | 8083 | N/F | 16077 | N/F | 464 | N/F | N/F | 83227 | 5603 | 3490 | 12316 | 3280 | 1671 | N/F | 315 | 5124 | 17433 | 8296 | D,Q | Unknown
WO-22270-131 | 485 | 40760 | 1080 | 346 | 4312 | 1487 | 3345422 | N/F | 85 | N/F | 574 | 2375 | N/F | 1047 | 23513 | 101404 | 531293 | 123781 | 550218 | 2033265 | 517798 | N/F | 17528152 | 3573179 | 4487682 | 1494266 | 9084723 | 7442 | 31167389 | 24998725 | 56755 | 1881 | 33 | 9544 | N/F | 273071 | N/F | 1137 | 8153 | 350150 | 110637 | 22620 | 3988 | 24909 | 6627 | 1570 | 609 | 129 | 5109 | 17801 | 11775 | D,Q | Unknown
WO-22270-132 | N/F | 11594 | 2742 | 1215 | 5351 | 4213661 | 5291859 | N/F | N/F | N/F | 1264 | 1618 | N/F | N/F | 34739 | 173513 | 1297582 | 180112 | 653695 | 2637207 | 766261 | N/F | 21597146 | 4614516 | 5903689 | 1980270 | 11965642 | 10844 | 31592473 | 27560248 | 82526 | 1637 | 207 | 9130 | N/F | 348548 | N/F | 732 | N/F | 410810 | 127795 | 25998 | 925 | 95039 | 6827 | 1057 | N/F | 67 | 3264 | 16765 | 14274 | D,Q | Unknown
WO-22270-133 | N/F | 8044 | 1438 | 747 | 4184 | 2328089 | 4036385 | N/F | 37 | N/F | 668 | 4609 | N/F | 23418 | 24781 | 109917 | 772092 | 120387 | 538697 | 2093334 | 537745 | 589 | 18901668 | 3804378 | 4592238 | 1634970 | 9484061 | 7134 | 30951840 | 26388231 | 53059 | 5017 | 28 | 10086 | N/F | 389545 | N/F | 775 | 13163 | 213496 | 136200 | 82659 | 757 | 54975 | 9798 | 1184 | N/F | 447 | 5119 | 16608 | 14105 |  | Unknown
WO-22270-134 | 634 | 22076 | 2585 | 1543 | 3774 | 2350021 | 4630474 | 34 | 264 | N/F | 1372 | 3914 | N/F | 7484 | 29312 | 121983 | 978151 | 126845 | 566135 | 2118182 | 503867 | 815 | 19266677 | 3860438 | 5122490 | 1689109 | 9798107 | 8036 | 28298022 | 25782986 | 62531 | 5307 | N/F | 7402 | N/F | 428910 | N/F | 347 | 50138 | 5370496 | 108394 | 138525 | 3852 | 45655 | 6028 | 2157 | 7752 | 136 | 5626 | 12767 | 27645 |  | Unknown
WO-22270-135 | N/F | 10932 | 962 | 352 | 4920 | 3978846 | 3779774 | N/F | 229 | N/F | 417 | 2636 | N/F | 19849 | 21293 | 104846 | 646508 | 99331 | 494603 | 1792899 | 420742 | 728 | 18105377 | 3172844 | 4251316 | 1445158 | 8965474 | 6475 | 29960619 | 25520988 | 49385 | 3615 | 786 | 7124 | N/F | 300841 | N/F | 364 | 11290 | 445424 | 127293 | 18616 | 5370 | 6418 | 6970 | 1038 | N/F | 118 | 4349 | 14622 | 9140 |  | Unknown
WO-22270-136 | 25040 | 159225 | 4582 | 871 | 10941 | 6469830 | 6284887 | N/F | 144 | 1455 | 894 | 1514 | N/F | 3679 | 21297 | 122537 | 1290263 | 121399 | 516680 | 2079913 | 591592 | N/F | 22119472 | 4962161 | 5855521 | 2165432 | 13194586 | 3981 | 34135354 | 26642322 | 55484 | 2440 | 679 | 8280 | N/F | 410301 | 126530 | 1115 | 34152 | 144584 | 135676 | 31567 | 1222 | 44053 | 4545 | 1311 | 5620 | N/F | 360776 | 488809 | 15690 |  | Unknown
WO-22270-137 | 10686 | 37878 | 484 | 64 | 6043 | 2293590 | 1738739 | N/F | N/F | 5182 | 718 | 2710 | N/F | 2740 | 28521 | 99835 | 222982 | 136130 | 477072 | 2235862 | 659323 | 600 | 14943000 | 2940610 | 4941408 | 1297714 | 9263209 | 6665 | 8797599 | 12713582 | 66307 | 1126 | 647 | 2164 | N/F | 194761 | 46031 | 440 | N/F | 2307 | 27028 | 8562 | 422 | 18969 | 1976 | 1168 | N/F | 63 | 831424 | 402768 | 14273 |  | Unknown
WO-22270-138 | 2198 | 80680 | 952 | 263 | 3486 | 2776532 | 2726384 | N/F | N/F | 6052 | 487 | 2449 | N/F | 2247 | 20592 | 87499 | 475667 | 101071 | 416637 | 1801389 | 520661 | 596 | 13390151 | 3202624 | 4391642 | 1403595 | 9369027 | 3857 | 4500446 | 8289063 | 53890 | 4874 | 240 | 608 | N/F | 131743 | N/F | 279 | 15538 | N/F | 3308 | 5231 | 1541 | 5138 | 933 | 1269 | 4614 | 333 | 14032 | 10631 | 5081 |  | Unknown
WO-22270-139 | 4065 | 93691 | 780 | 57 | 2539 | 2607462 | 2055113 | N/F | 18 | 3132 | 295 | 3929 | N/F | 1518 | 12308 | 66938 | 254224 | 53525 | 350065 | 1293108 | 323209 | 1096 | 14070998 | 2331221 | 3271597 | 1022183 | 6537586 | 3358 | 25647005 | 22168034 | 30142 | 1513 | 75 | 5418 | N/F | 311132 | 17330 | 284 | N/F | N/F | 13302 | 8700 | 325 | 5471 | 1687 | 547 | 3800 | 28 | 27797 | 22856 | 9140 |  | Unknown
WO-22270-140 | 5175 | 115120 | 2149 | 228 | 3229 | 2810226 | 2695366 | 31 | N/F | 3967 | 649 | 1862 | N/F | 2735 | 25012 | 99852 | 462866 | 104719 | 463804 | 1782722 | 495760 | 670 | 17201517 | 3206086 | 4299628 | 1339615 | 8340094 | 6284 | 29305932 | 23883446 | 53287 | 2465 | 121 | 7305 | N/F | 303248 | 13875 | 257 | N/F | N/F | 53824 | 11949 | 1068 | 6763 | 1750 | 436 | N/F | 183 | 30156 | 23463 | 24759 |  | Unknown
AB-220715-01 LFB-H 8 | 41960077 | 82761 | 133600 | 238536 | 45947449 | 1867084 | 2428242 | 967091 | 3234682 | 290850 | 522117 | 1897952 | 1117295 | 1827269 | 84195 | 141629 | 255958 | 523805 | 1162800 | 4783389 | 2042640 | N/F | 20685732 | 4351996 | 8931373 | 2019840 | 12567876 | 52700 | 34422581 | 25511094 | 240908 | 1353206 | 1520412 | 1884393 | 46394 | 17740796 | 31817093 | 2082675 | 4399065 | 1849319 | 117100 | 11684793 | 2334818 | 8448635 | 10761 | 4367669 | 404436 | 1002477 | 27019269 | 32079761 | 21801379 |  | Unknown
AB-220715-01 LRB 15_INJ001 | 505691 | 47643 | 979 | 1837 | 237603 | 2389937 | 2663934 | 2683 | 13403 | 2249 | 1764 | 8172 | 4872 | 13351 | 97640 | 243868 | 360597 | 524588 | 1528681 | 5858095 | 2322692 | N/F | 25839013 | 4906561 | 12555336 | 2366627 | 15948317 | 33463 | 37704073 | 33275338 | 231321 | 5099 | 10872 | 10502 | N/F | 153971 | 444900 | 19790 | 30517 | 10325 | 106091 | 56189 | 14402 | 66864 | 7404 | 33092 | N/F | 3634 | 191832 | 344056 | 273704 | C | Unknown
AB-220715-01 LRB 15 | 240320 | 52125 | 1290 | 302 | 28497 | 2340833 | 2477558 | N/F | 723 | 2518 | 987 | 1391 | N/F | 8617 | 95789 | 234891 | 344095 | 525317 | 1521114 | 5984993 | 2179518 | N/F | 26677815 | 4618474 | 11929136 | 2262813 | 15116895 | 23381 | 36337608 | 32216398 | 251665 | 4047 | 2380 | 7605 | N/F | 37245 | 302707 | 7321 | N/F | 3144 | 101504 | 9024 | 1776 | 33914 | 4912 | 7959 | N/F | 231 | 168123 | 300571 | 99026 | D,Q | Unknown
AB-220715-01 LD 7 | 61927 | 128092 | 468 | 216 | 8271 | 2418492 | 2304726 | 11 | 211 | 2414 | 523 | 1267 | N/F | 4090 | 20498 | 90500 | 441457 | 77348 | 442798 | 1586895 | 370515 | 2199 | 14886295 | 2544183 | 3675123 | 1199582 | 7328787 | 6424 | 26427335 | 21641575 | 45449 | 3716 | 768 | 5148 | N/F | 331890 | 128981 | 1667 | N/F | N/F | 52477 | 13544 | 2031 | 22048 | 1045 | 2019 | 14443 | 141 | 107814 | 159126 | 24505 |  | Unknown
AB-220715-01 CCC-H 8 | 71551754 | 163649 | 238936 | 420933 | 80292227 | 1473347 | 2186247 | 1889168 | 5703583 | 603501 | 974031 | 3258977 | 2116612 | 3204232 | 84399 | 121635 | 206957 | 489858 | 1056800 | 4246032 | 2077923 | N/F | 15691752 | 3501116 | 7967798 | 1667999 | 11021968 | 41276 | 28511691 | 19334403 | 237285 | 2355443 | 2505872 | 3395912 | 69395 | 27614293 | 49644592 | 3085030 | 8041984 | 3749471 | 81433 | 18747354 | 3809028 | 14080172 | 7415 | 6959968 | 771694 | 1862004 | 44127122 | 55259551 | 34201767 |  | Unknown
AB-220715-01 LFSM-H 7 | 37105095 | 139840 | 87836 | 344174 | 33535600 | 2004697 | 2618416 | 221558 | 1324591 | 87769 | 127874 | 954060 | 475136 | 1251619 | 29557 | 86215 | 373138 | 123238 | 494490 | 1961354 | 621772 | N/F | 15804340 | 3113149 | 4360705 | 1414281 | 8084466 | 8678 | 27019135 | 21492600 | 60200 | 589238 | 1789709 | 2042724 | N/F | 14956285 | 27153692 | 1786605 | 1857875 | 557035 | 34595 | 7131542 | 1967561 | 4027523 | 1711 | 3201341 | 62369 | 370014 | 17490655 | 25555270 | 19703675 |  | Unknown
WO-22270-141 | 575929 | 84001 | 1121 | 1537 | 85465 | 2579865 | 2489714 | 552 | 3539 | 16313 | 889 | 3773 | N/F | 5574 | 35756 | 122088 | 457927 | 145861 | 573003 | 2303238 | 683065 | 699 | 19782291 | 3631663 | 4969822 | 1717122 | 9793481 | 8651 | 32647379 | 26500673 | 71295 | 1989 | 11584 | 8340 | N/F | 280697 | 719559 | 24313 | 7759 | 3053 | 40801 | 24917 | 12133 | 22293 | 944 | 10752 | 28208 | 1687 | 390343 | 716273 | 240809 |  | Unknown
WO-22270-142 | 219943 | 74990 | 1195 | 295 | 32220 | 2568325 | 2482730 | N/F | 583 | 14348 | 719 | 2843 | N/F | 3797 | 30746 | 113243 | 433527 | 125794 | 502274 | 2037922 | 546595 | N/F | 19080449 | 3402434 | 4796678 | 1508916 | 9330821 | 6102 | 32489542 | 27207567 | 63843 | 1541 | 2400 | 7089 | N/F | 215670 | 530483 | 5330 | N/F | N/F | 58831 | 10243 | 5836 | 19308 | 895 | 4160 | 487 | 139 | 433071 | 934919 | 61093 | D,Q | Unknown
WO-22270-143 | 36101 | 110233 | 878 | 117 | 28103 | 1961016 | 1827300 | N/F | 479 | 5707 | 556 | 4391 | N/F | 2301 | 15423 | 87021 | 300838 | 61997 | 362414 | 1301505 | 297470 | 853 | 14059759 | 2311155 | 3388263 | 1086537 | 6503170 | 5425 | 25268723 | 20352708 | 31973 | 1938 | 602 | 4663 | N/F | 223461 | 118065 | 650 | N/F | N/F | 19964 | 10710 | 7495 | 5156 | 745 | 2185 | 1922 | 89 | 145691 | 229957 | 11740 |  | Unknown
WO-22270-144 | 50150 | 130273 | 2711 | 239 | 9371 | 3326905 | 3257402 | 19 | 440 | 47864 | 574 | 2300 | N/F | 4609 | 31948 | 133975 | 786429 | 145314 | 573430 | 2245959 | 686914 | N/F | 20363449 | 3862700 | 5437062 | 1655493 | 10268993 | 10473 | 29628833 | 26788780 | 70659 | 6624 | 795 | 7820 | N/F | 289769 | 200162 | 1394 | 50315 | N/F | 22252 | 27127 | 7572 | 38297 | 845 | 2489 | N/F | 362 | 482097 | 718507 | 18652 |  | Unknown
WO-22270-145 | 14646 | 147215 | 1273 | 126 | 15453 | 3202739 | 2718872 | 12 | 87 | 45090 | 165 | 1629 | N/F | 2070 | 29870 | 117916 | 496409 | 143294 | 547070 | 2173396 | 643791 | 562 | 21179565 | 3903414 | 4886714 | 1600079 | 9773329 | 10870 | 32218106 | 26419173 | 77164 | 2952 | 442 | 7175 | N/F | 243683 | 62290 | 688 | 21613 | N/F | 24955 | 12445 | 6230 | 24354 | 872 | 2015 | 28201 | N/F | 190291 | 255661 | 15720 |  | Unknown
WO-22270-146 | 12313 | 71512 | 352 | 29 | 6999 | 2443134 | 2130717 | N/F | 10 | 29507 | 208 | 2446 | N/F | 2461 | 24527 | 92427 | 372452 | 90916 | 433524 | 1722900 | 460155 | 822 | 16305276 | 2881326 | 4034522 | 1225828 | 7903729 | 8906 | 24455743 | 24486686 | 48029 | 5230 | 477 | 6231 | N/F | 155044 | 42607 | 608 | N/F | N/F | 14092 | 12418 | 5948 | 31822 | 504 | 1313 | 33629 | 56 | 190866 | 206992 | 13336 |  | Unknown
WO-22270-147 | 10509 | 80758 | 694 | 121 | 5021 | 1770695 | 1576669 | N/F | N/F | 58147 | 401 | 1692 | N/F | 5035 | 13461 | 77591 | 270031 | 47744 | 339378 | 1133748 | 246120 | 1267 | 11583296 | 2265564 | 3225824 | 1042155 | 6775621 | 5007 | 7112602 | 9955180 | 27133 | 3999 | 315 | 2703 | N/F | 192780 | 22789 | 454 | N/F | N/F | 1728 | 7539 | 6376 | 36620 | 812 | 1835 | N/F | 83 | 87073 | 117758 | 9943 |  | Unknown
WO-22270-148 | 13344 | 191064 | 1426 | 870 | 5624 | 3343703 | 3170322 | N/F | 64 | 157426 | 629 | 2882 | N/F | 9240 | 36067 | 127360 | 782452 | 165268 | 569266 | 2418578 | 774954 | N/F | 21295071 | 4031105 | 5704854 | 1659164 | 10865694 | 9707 | 30180524 | 26972536 | 78227 | 2424 | 1534 | 6181 | N/F | 215138 | 41573 | 434 | 31305 | 7606 | 33030 | 14250 | 4970 | 116007 | 835 | 2099 | N/F | 11 | 231521 | 224666 | 15863 | D,Q | Unknown
WO-22270-149 | 10026 | 110124 | 1233 | 155 | 6379 | 2541509 | 2806006 | N/F | 111 | 150437 | 415 | 2401 | N/F | 3012 | 32501 | 118281 | 422855 | 169894 | 594572 | 2415749 | 737088 | 1191 | 20994098 | 3793380 | 5450588 | 1681935 | 10149306 | 9394 | 34033054 | 28283258 | 72788 | 2591 | N/F | 9045 | N/F | 389076 | 33901 | 822 | N/F | 8006 | 85493 | 19585 | 4112 | 42824 | 1570 | 1230 | 21330 | 175 | 265915 | 180891 | 12164 |  | Unknown
WO-22270-150 | 9078 | 100230 | 400 | 183 | 5143 | 2488460 | 2323430 | N/F | 18 | 114556 | 574 | 2723 | N/F | 2008 | 28659 | 102217 | 399978 | 123721 | 499943 | 2141597 | 588086 | 602 | 18687443 | 3432329 | 4809651 | 1482777 | 8941046 | 8914 | 32135818 | 27064942 | 64735 | 2303 | 42 | 7619 | N/F | 252746 | 29455 | 623 | N/F | N/F | 73821 | 19612 | 3650 | 6208 | 3045 | 2301 | N/F | 30 | 178534 | 164917 | 15186 |  | Unknown
AB-220715-01 LRB 16_INJ001 | 12725 | 70519 | 628 | 450 | 11012 | 2071304 | 2230968 | 11 | 132 | 3328 | 456 | 1139 | N/F | 4618 | 80865 | 201897 | 317824 | 412519 | 1165125 | 4796231 | 1839962 | 1907 | 24724590 | 4145597 | 10435240 | 2040059 | 14346854 | 31362 | 36647315 | 29843384 | 192514 | 3440 | 95 | 6540 | N/F | 27172 | 48869 | 1133 | N/F | N/F | 107690 | 20526 | 1915 | 28522 | 10151 | 2319 | 4272 | 75 | 258269 | 229153 | 19935 |  | Unknown
AB-220715-01 LRB 16 | 14788 | 58229 | 955 | 589 | 7048 | 1999097 | 2140762 | N/F | 99 | 2890 | 908 | 388 | N/F | 3389 | 80058 | 200163 | 298191 | 398216 | 1207377 | 4825536 | 1841664 | N/F | 23758425 | 4080586 | 10238552 | 2094631 | 13706642 | 28700 | 33160647 | 29602926 | 183473 | 1183 | 658 | 6412 | N/F | 25924 | 42249 | 611 | N/F | N/F | 79166 | 8413 | 2097 | 28349 | 6368 | 1333 | N/F | 149 | 425162 | 291585 | 42717 | D,Q | Unknown
WO-22270-151 | 7370 | 75554 | 697 | 109 | 8172 | 2042724 | 2119261 | N/F | N/F | 375117 | 290 | 1369 | N/F | 1687 | 27389 | 104109 | 343248 | 115917 | 501542 | 1904176 | 565315 | N/F | 17809073 | 3099270 | 4398704 | 1396366 | 8451782 | 6017 | 31128854 | 24697936 | 57239 | 773 | 158 | 6233 | N/F | 184634 | 25573 | 411 | N/F | 2336 | 82564 | 57405 | 3071 | 7151 | 1970 | 1150 | N/F | N/F | 222136 | 126347 | 14480 | D,Q | Unknown
WO-22270-152 | 8657 | 81871 | 1563 | 448 | 5309 | 2336261 | 2127264 | N/F | 72 | 441864 | 743 | 1456 | N/F | 14896 | 30326 | 111869 | 380712 | 147061 | 535569 | 2171600 | 658354 | N/F | 19950926 | 3349826 | 4858332 | 1549125 | 9492805 | 9329 | 30673222 | 26557310 | 63729 | 730 | 218 | 6194 | N/F | 229472 | 30181 | 390 | N/F | 3039 | 67866 | 11739 | 3163 | 10813 | 2016 | 1744 | N/F | N/F | 131567 | 64111 | 19700 | D,Q | Unknown
WO-22270-153 | 7298 | 60621 | 393 | 24 | 5492 | 1701362 | 1564749 | N/F | N/F | 139674 | 367 | 1108 | N/F | 4432 | 13014 | 67719 | 289275 | 43772 | 309984 | 1052335 | 226704 | 907 | 12714269 | 2069753 | 2952713 | 1051780 | 6427292 | 4068 | 8207649 | 14212626 | 22075 | 622 | 163 | 1654 | N/F | 184796 | 18893 | 95 | N/F | 1900 | 4817 | 11665 | 2199 | 12071 | 591 | 915 | N/F | N/F | 89022 | 51438 | 33975 |  | Unknown
WO-22270-154 | 12664 | 71973 | 260 | 254 | 5691 | 2261118 | 2299737 | N/F | 165 | 898816 | 306 | 853 | N/F | 4798 | 33043 | 114107 | 379144 | 130265 | 527595 | 2077924 | 633498 | 476 | 18892054 | 3274452 | 4786316 | 1516944 | 9308754 | 7703 | 30639980 | 25007841 | 65835 | 757 | 196 | 6746 | N/F | 355975 | 28161 | 517 | N/F | 5154 | 61549 | 219849 | 1100 | 11180 | 1746 | 1138 | N/F | 95 | 159251 | 66842 | 24528 |  | Unknown
WO-22270-155 | 8327 | 86666 | 1184 | 216 | 6589 | 2757769 | 3129704 | N/F | 109 | 1940827 | 540 | 1937 | N/F | 4566 | 47830 | 231690 | 624293 | 212998 | 1124693 | 3992764 | 1072043 | N/F | 29629816 | 4792469 | 10742136 | 2483686 | 17632076 | 9634 | 31468489 | 31276463 | 92967 | 1012 | 400 | 6309 | N/F | 483522 | 21344 | 223 | N/F | 11617 | 9572 | 28329 | 1943 | 19536 | 1380 | 1559 | N/F | 262 | 58735 | 38304 | 14125 | D,Q | Unknown
WO-22270-156 | 5650 | 126718 | 780 | 25 | 13526 | 3039166 | 3437986 | N/F | N/F | 865942 | 776 | 1254 | N/F | 5525 | 29471 | 109388 | 788804 | 113207 | 477127 | 1908518 | 545367 | 1945 | 19334524 | 3479976 | 4447850 | 1419354 | 9283976 | 6618 | 30123706 | 25117097 | 48029 | 3510 | 545 | 5543 | N/F | 298517 | 20506 | 36 | N/F | 5468 | 37581 | 13930 | 3376 | 26027 | 1114 | 1448 | N/F | 44 | 151890 | 54341 | 39135 |  | Unknown
WO-22270-157 | 6313 | 60042 | 309 | 68 | 3363 | 2440098 | 2535751 | N/F | N/F | 7802 | 733 | 1546 | N/F | 4511 | 31705 | 116444 | 390497 | 127295 | 544228 | 2158416 | 593474 | 777 | 19537435 | 3281876 | 5196545 | 1504434 | 8946909 | 7115 | 33235078 | 27039566 | 65854 | 1090 | 255 | 7453 | N/F | 335444 | 21675 | 273 | 25756 | 6657 | 51130 | 177514 | 2018 | 40313 | 565 | 1544 | N/F | 104 | 94032 | 40485 | 40286 |  | Unknown
WO-22270-158 | 7691 | 122801 | 850 | 498 | 5292 | 2940924 | 2731400 | N/F | N/F | 3226 | 449 | 3011 | N/F | 14211 | 28416 | 112130 | 719429 | 124087 | 496865 | 2017716 | 611183 | N/F | 18063766 | 3312252 | 4676702 | 1445449 | 9566553 | 7034 | 20541693 | 21948470 | 59343 | 2268 | 280 | 5101 | N/F | 166705 | 21515 | 326 | N/F | 1278 | 8449 | 35506 | 1138 | 18772 | 4616 | 874 | N/F | 46 | 131663 | 44244 | 13867 | D,Q | Unknown
WO-22270-159 | 7627 | 109970 | 862 | 387 | 7261 | 2739398 | 2961268 | 9 | 56 | 3675 | 743 | 2118 | N/F | 5084 | 30087 | 138831 | 495058 | 115004 | 701554 | 2369286 | 591800 | N/F | 25609527 | 4010470 | 7154057 | 2006936 | 13331964 | 5910 | 29263325 | 30028896 | 63918 | 910 | 539 | 5514 | N/F | 402999 | 21513 | 388 | N/F | N/F | 5823 | 29008 | 1661 | 42393 | 569 | 1176 | N/F | 20 | 48453 | 28922 | 15111 |  | Unknown
WO-22270-160 | 7049 | 75945 | 1596 | 45 | 5199 | 1775674 | 1506043 | N/F | N/F | 1129 | 321 | 1921 | N/F | 14375 | 20030 | 81480 | 266946 | 72027 | 366427 | 1480547 | 378715 | N/F | 12729966 | 2417975 | 3403835 | 1030482 | 6645223 | 5944 | 10546442 | 14907093 | 37585 | 526 | 505 | 2321 | N/F | 181425 | 14552 | 731 | N/F | N/F | 4494 | 11769 | 2329 | 4423 | 691 | 642 | N/F | N/F | 99724 | 26202 | 22332 | D,Q | Unknown
AB-220715-01 LFB-H 9 | 30319282 | 78129 | 122040 | 219772 | 37325000 | 1408186 | 1858966 | 767079 | 2546116 | 251436 | 414913 | 1634635 | 908452 | 1443127 | 79993 | 134892 | 212725 | 437559 | 1000400 | 4073788 | 1841312 | 3694 | 17287027 | 3262654 | 8235689 | 1663172 | 11264146 | 24022 | 25085889 | 20351319 | 211075 | 1242675 | 1259489 | 1434965 | 25157 | 16294504 | 21535008 | 1468070 | 3928616 | 1709479 | 6960 | 10564742 | 1853562 | 7730878 | 790 | 3230802 | 329348 | 791170 | 6392993 | 15621679 | 18182280 |  | Unknown
AB-220715-01 LRB 17_INJ001 | 514617 | 88811 | 1307 | 2207 | 262345 | 2037892 | 2143829 | 2177 | 15255 | 1873 | 1629 | 9907 | 4870 | 11071 | 86018 | 205239 | 306042 | 440566 | 1283922 | 5236737 | 1982393 | N/F | 24594162 | 4027139 | 10618169 | 2018735 | 14436278 | 30471 | 34299972 | 31065468 | 206198 | 5388 | 12909 | 11665 | N/F | 154766 | 413021 | 26092 | 22428 | 10521 | 107546 | 64440 | 20926 | 70871 | 5258 | 32178 | 30257 | 3376 | 167236 | 219854 | 340948 |  | Unknown
AB-220715-01 LRB 17 | 134814 | 91508 | 783 | 559 | 20980 | 2161197 | 2162359 | 58 | 637 | 5154 | 419 | 1435 | N/F | 4821 | 87418 | 209558 | 297338 | 482924 | 1362703 | 5235691 | 1894534 | 2331 | 23348606 | 4310166 | 10012992 | 2164904 | 14326661 | 26613 | 34862689 | 30359741 | 195002 | 3387 | 1198 | 7496 | N/F | 29550 | 186663 | 3877 | N/F | 3875 | 100185 | 9421 | 2079 | 29019 | 5411 | 3957 | N/F | 304 | 165314 | 138999 | 50245 |  | Unknown
AB-220715-01 LD 8 | 41436 | 68227 | 624 | 271 | 8508 | 2072463 | 2144285 | 12 | 275 | 2244 | 300 | 1491 | N/F | 4278 | 20853 | 82944 | 297844 | 78067 | 392269 | 1462805 | 349396 | N/F | 13275971 | 2557496 | 3429914 | 1034197 | 6126156 | 6367 | 26544928 | 19344985 | 38257 | 1813 | 696 | 5871 | N/F | 198115 | 96624 | 1115 | N/F | 3465 | 84806 | 11517 | 2084 | 53943 | 2745 | 2282 | 8465 | 65 | 207477 | 149077 | 21195 |  | Unknown
AB-220715-01 CCC-H 9 | 70864251 | 166242 | 247337 | 398246 | 78987400 | 1458367 | 2208073 | 1931914 | 5848146 | 599342 | 1030682 | 3278063 | 2064707 | 3122077 | 82629 | 125467 | 201430 | 510230 | 1072569 | 4332367 | 1971418 | N/F | 16032009 | 3554542 | 8055316 | 1695888 | 10457366 | 51690 | 27716630 | 18446645 | 245512 | 2541952 | 2551914 | 3397584 | 82766 | 27329670 | 51739431 | 3240393 | 8253523 | 3679669 | 160843 | 18375133 | 3695689 | 14687515 | 24595 | 7147165 | 870112 | 1817444 | 47642794 | 54798454 | 32583449 |  | Unknown
AB-220715-01 LFSM-H 8 | 35421212 | 71756 | 71941 | 245142 | 31933295 | 1757891 | 2180831 | 159622 | 1170507 | 68901 | 102562 | 816895 | 420590 | 1156594 | 24239 | 75736 | 262702 | 97378 | 462711 | 1659169 | 479856 | N/F | 14810303 | 3050140 | 3756214 | 1294954 | 7389918 | 7833 | 25417726 | 19974113 | 48644 | 508961 | 1485139 | 1670928 | N/F | 13004847 | 25959675 | 1764465 | 1554075 | 420789 | 70091 | 6609977 | 1767581 | 3579067 | 5895 | 3149538 | 45697 | 309971 | 22019946 | 26459234 | 17938028 |  | Unknown
WO-22270-161 | 454888 | 51350 | 742 | 337 | 38099 | 1986392 | 2019447 | 22 | 1467 | 4113 | 712 | 2380 | N/F | 4573 | 25280 | 90301 | 330552 | 96858 | 421825 | 1505230 | 412731 | 981 | 16116943 | 2994895 | 3674705 | 1294928 | 7362581 | 4753 | 28572131 | 21517596 | 49045 | 1948 | 6605 | 6752 | N/F | 182171 | 696407 | 15416 | N/F | 1787 | 65756 | 31372 | 8530 | 6047 | 1850 | 5073 | 2682 | 637 | 619865 | 881519 | 153728 |  | Unknown
WO-22270-162 | 233216 | 79567 | 507 | 248 | 23189 | 2331250 | 2109942 | N/F | 611 | 5724 | 474 | 1714 | N/F | 3584 | 27638 | 92255 | 362497 | 114259 | 500477 | 1830214 | 487440 | N/F | 16928868 | 3135963 | 4109872 | 1468628 | 8525979 | 5357 | 28589800 | 25150482 | 54481 | 898 | 1993 | 8149 | N/F | 288812 | 478690 | 5407 | 21493 | N/F | 52940 | 11432 | 5562 | 13179 | 1753 | 2541 | N/F | 169 | 529921 | 729200 | 54215 | D,Q | Unknown
WO-22270-163 | 61314 | 52091 | 116 | 263 | 10208 | 1672142 | 1536856 | N/F | 346 | 3097 | 316 | 2213 | N/F | 3966 | 15786 | 79985 | 229617 | 54525 | 345906 | 1149852 | 273405 | 1388 | 12857812 | 2325081 | 2898261 | 1032507 | 5966876 | 3741 | 24504190 | 19001604 | 30667 | 689 | 727 | 5579 | N/F | 156096 | 167409 | 1306 | N/F | N/F | 60066 | 9584 | 5538 | 18172 | 1560 | 2799 | N/F | 115 | 373145 | 436123 | 19697 |  | Unknown
WO-22270-164 | 52296 | 87690 | 671 | 8 | 10025 | 2626230 | 2813938 | N/F | 418 | 20883 | 736 | 1119 | N/F | 18398 | 33841 | 119909 | 425434 | 145041 | 564007 | 2236290 | 666467 | N/F | 19781444 | 3816615 | 5072580 | 1585522 | 9497668 | 8580 | 32264373 | 26196559 | 75186 | 689 | 106 | 9871 | N/F | 265157 | 174526 | 468 | N/F | 1834 | 77452 | 41781 | 4479 | 36720 | 1481 | 2522 | N/F | 39 | 352723 | 339437 | 16091 | D,Q | Unknown
WO-22270-165 | 28070 | 74718 | 390 | 103 | 10876 | 2160422 | 2037933 | N/F | 166 | 18987 | 857 | 2307 | N/F | 2947 | 30179 | 114002 | 361304 | 131824 | 505746 | 2046580 | 556366 | N/F | 17130650 | 3127881 | 4585177 | 1487528 | 8954291 | 4607 | 28547043 | 23654570 | 67796 | 2325 | 265 | 4805 | N/F | 242310 | 89464 | 1126 | N/F | N/F | 71158 | 10220 | 5113 | 8295 | 2082 | 2917 | N/F | 369 | 283246 | 319686 | 19570 | D,Q | Unknown
WO-22270-166 | 22860 | 89312 | 1406 | 186 | 5704 | 2725002 | 2604755 | N/F | 212 | 18950 | 491 | 1704 | N/F | 3956 | 35152 | 116542 | 571070 | 132070 | 571861 | 2081678 | 591614 | 1444 | 20176247 | 3617881 | 4721342 | 1578553 | 9307376 | 6403 | 32476254 | 25849404 | 64993 | 953 | 711 | 7417 | N/F | 343588 | 74555 | 697 | N/F | N/F | 77237 | 7977 | 4505 | 9444 | 2314 | 2032 | N/F | 72 | 262154 | 256537 | 13327 |  | Unknown
WO-22270-167 | 14732 | 98836 | 1176 | 430 | 5648 | 2798596 | 2800951 | N/F | N/F | 37554 | 536 | 1657 | N/F | 3200 | 19690 | 104460 | 614421 | 74108 | 416676 | 1539979 | 344472 | N/F | 16788719 | 3056745 | 3611791 | 1214396 | 8053607 | 3531 | 27618267 | 22126715 | 38149 | 2155 | 1559 | 5383 | N/F | 396046 | 50441 | 346 | N/F | N/F | 20702 | 33127 | 2486 | 24058 | 1177 | 1071 | 11539 | 82 | 217016 | 185585 | 12013 |  | Unknown
WO-22270-168 | 19277 | 111019 | 962 | 279 | 8070 | 3425960 | 3555450 | N/F | 50 | 57210 | 470 | 1655 | N/F | 15800 | 27777 | 138586 | 903088 | 133297 | 509001 | 1951951 | 576787 | N/F | 18385424 | 3786516 | 4653025 | 1540673 | 9125726 | 7327 | 28966846 | 24199721 | 62015 | 2885 | 634 | 8290 | N/F | 223342 | 50765 | 782 | 28099 | N/F | 50359 | 11998 | 3522 | 17749 | 1004 | 1973 | N/F | 357 | 317569 | 285695 | 15495 | D,Q | Unknown
WO-22270-169 | 15747 | 129191 | 2221 | 541 | 6454 | 3251637 | 3030011 | 14 | 51 | 44090 | 471 | 1792 | N/F | 5296 | 17178 | 99263 | 935816 | 72228 | 387787 | 1305458 | 318093 | 912 | 14113607 | 2790302 | 3405037 | 1161526 | 7248648 | 5124 | 25512240 | 19715255 | 31504 | 3667 | 587 | 7767 | N/F | 222308 | 37634 | 1231 | N/F | 1894 | 76201 | 18854 | 4666 | 21597 | 1326 | 2038 | N/F | 88 | 288974 | 223936 | 12841 |  | Unknown
WO-22270-170 | 12562 | 99301 | 1438 | 274 | 5780 | 3311183 | 3116038 | N/F | 88 | 79811 | 454 | 1986 | N/F | 15434 | 26498 | 116824 | 819148 | 107664 | 501441 | 1835729 | 491147 | N/F | 17762304 | 3443669 | 4397285 | 1480671 | 8815748 | 8010 | 28478833 | 24417699 | 53943 | 3513 | 540 | 5492 | N/F | 233304 | 34502 | 456 | N/F | N/F | 61507 | 8628 | 2484 | 5625 | 1362 | 1778 | N/F | 107 | 230905 | 173685 | 11498 | D,Q | Unknown
AB-220715-01 LRB 18_INJ001 | 10350 | 58167 | 609 | 221 | 7751 | 2283502 | 2471648 | N/F | 285 | 4044 | 704 | 930 | 13574 | 4283 | 93252 | 233752 | 356241 | 482731 | 1416177 | 5197232 | 1918687 | 3151 | 24380115 | 4685267 | 11009099 | 2214346 | 15008419 | 29718 | 36131084 | 31226055 | 214031 | 824 | 260 | 8483 | N/F | 34632 | 40527 | 308 | N/F | 4709 | 117657 | 8082 | 2742 | 32618 | 7523 | 1268 | N/F | 279 | 233574 | 177962 | 21549 |  | Unknown
210451-d2-reprep-shaken | 8903 | 119216 | 1097 | 388 | 5147 | 2904826 | 2578303 | N/F | 166 | N/F | 262 | 3010 | N/F | 1783 | 19813 | 91270 | 386639 | 81465 | 388031 | 1685855 | 407490 | 1677 | 15793284 | 2772252 | 4145766 | 1129293 | 8313996 | 5468 | 28165428 | 22997577 | 43172 | 2652 | 470 | 7659 | N/F | 8501198 | 19793 | 749 | N/F | N/F | 114727 | 7923 | 2677 | 58318 | 5592 | 656 | N/F | N/F | 110652 | 63242 | 11886 |  | Unknown
AB-220715-01 LRB 18 | 13718 | 51526 | 390 | 419 | 11047 | 2076112 | 2104985 | N/F | 68 | 2420 | 734 | 1316 | N/F | 3732 | 84515 | 210157 | 292534 | 438976 | 1260583 | 4690238 | 1770614 | N/F | 21369570 | 4256085 | 9609305 | 2051760 | 13137166 | 15606 | 32285653 | 28180034 | 200118 | 3829 | 494 | 8272 | N/F | 25981 | 34180 | 248 | N/F | 4176 | 98528 | 7173 | 1710 | 30230 | 6856 | 1778 | 10997 | 28 | 234892 | 142608 | 14113 |  | Unknown
WO-22270-171 | 10230 | 75049 | 312 | 50 | 4262 | 2335439 | 2198294 | 55 | 219 | 200208 | 337 | 1642 | N/F | 1926 | 38037 | 112197 | 364912 | 146873 | 549928 | 2147506 | 642829 | 984 | 18517951 | 3454792 | 4696955 | 1549442 | 9109177 | 8750 | 30973664 | 25645065 | 67583 | 833 | 447 | 7555 | N/F | 217723 | 25911 | 121 | N/F | N/F | 67908 | 45737 | 2878 | 19811 | 1737 | 1147 | N/F | 147 | 218310 | 132844 | 8593 |  | Unknown
WO-22270-172 | 9152 | 71905 | 727 | 367 | 4623 | 2228636 | 2240503 | N/F | 77 | 226030 | 278 | 1556 | N/F | 14459 | 34064 | 115792 | 370298 | 140379 | 534822 | 2116280 | 624566 | N/F | 18818518 | 3503523 | 4712013 | 1540644 | 9120655 | 7624 | 30940338 | 23825739 | 69748 | 768 | 431 | 7374 | N/F | 283647 | 26172 | 495 | N/F | 2896 | 78559 | 13092 | 2500 | 17754 | 1533 | 1019 | 2036 | 270 | 206796 | 107433 | 11751 | D,Q | Unknown
WO-22270-173 | 9452 | 105234 | 1035 | 417 | 6829 | 2978223 | 2855709 | 12 | 213 | 300289 | 765 | 1957 | N/F | 2509 | 34909 | 115272 | 612041 | 152186 | 535398 | 2166115 | 663709 | 1938 | 20252525 | 3712945 | 4820123 | 1539161 | 9704164 | 7022 | 31655033 | 25841038 | 69738 | 1637 | 261 | 7477 | N/F | 263382 | 26324 | 331 | N/F | 2746 | 91204 | 17588 | 3137 | 27570 | 2379 | 1471 | N/F | 744 | 215880 | 99241 | 12594 |  | Unknown
WO-22270-174 | 7675 | 131266 | 2504 | 230 | 4929 | 3612097 | 3660921 | N/F | 27 | 848586 | 650 | 1347 | N/F | 4336 | 32719 | 249185 | 1160423 | 124965 | 521543 | 2029520 | 519485 | N/F | 17439954 | 3952504 | 4975969 | 1533210 | 9471485 | 5824 | 28412921 | 23723366 | 55322 | 1666 | 202 | 7778 | N/F | 208597 | 22834 | 315 | 26242 | 6442 | 86761 | 77864 | 1382 | 32792 | 2041 | 1367 | N/F | 37 | 158201 | 72391 | 20492 | D,Q | Unknown
WO-22270-175 | 10729 | 65677 | 450 | 61 | 5202 | 2398714 | 2275071 | N/F | N/F | 662854 | 331 | 2002 | N/F | 4690 | 35768 | 115601 | 384726 | 129497 | 543406 | 2122488 | 594416 | N/F | 18830491 | 3471206 | 4637843 | 1545338 | 9280381 | 6686 | 30332074 | 26278341 | 64679 | 986 | 341 | 6819 | N/F | 174743 | 21806 | 155 | N/F | 2410 | 66510 | 8258 | 2168 | 12266 | 2435 | 733 | N/F | 89 | 102596 | 54781 | 11404 | D,Q | Unknown
WO-22270-176 | 6273 | 84757 | 713 | 285 | 3574 | 2016722 | 1756199 | N/F | N/F | 477838 | 1079 | 1386 | N/F | 5335 | 17910 | 81447 | 280251 | 64699 | 376408 | 1376487 | 308150 | N/F | 14235878 | 2612632 | 3435999 | 1113208 | 6876162 | 4826 | 27025433 | 21697185 | 33535 | 1546 | 8 | 5326 | N/F | 176403 | 20359 | 460 | N/F | 4228 | 76627 | 10643 | 2512 | 5745 | 2423 | 932 | 3357 | 153 | 98580 | 47231 | 11021 |  | Unknown
WO-22270-177 | 7539 | 78026 | 671 | 27 | 6115 | 2616987 | 2578011 | N/F | 339 | 7277 | 408 | 1586 | N/F | 3809 | 37809 | 133344 | 401985 | 142206 | 645123 | 2473796 | 688421 | 1487 | 21958673 | 3899051 | 5828726 | 1831898 | 11134728 | 6742 | 34124614 | 29440950 | 73075 | 1309 | 146 | 6572 | N/F | 182687 | 20864 | 266 | N/F | N/F | 63191 | 60919 | 2008 | 8634 | 681 | 1630 | N/F | 34 | 101784 | 41845 | 11364 |  | Unknown
WO-22270-178 | 8086 | 72943 | 416 | 35 | 3559 | 1901999 | 1671047 | N/F | N/F | 2525 | 262 | 1288 | N/F | 13948 | 18875 | 71500 | 267865 | 62119 | 328193 | 1229164 | 311949 | N/F | 13680992 | 2378705 | 2867754 | 1060428 | 6148355 | 3037 | 25643746 | 19854001 | 31875 | 1074 | N/F | 6980 | N/F | 137333 | 18984 | 83 | N/F | N/F | 60858 | 11636 | 1712 | 20372 | 2081 | 892 | 18950 | 21 | 123552 | 47848 | 16835 |  | Unknown
WO-22270-179 | 5946 | 58955 | 315 | 60 | 6725 | 2217106 | 2124782 | N/F | N/F | 2837 | 319 | 2367 | N/F | 2372 | 21784 | 93854 | 341678 | 91873 | 449059 | 1618004 | 411207 | N/F | 16153915 | 2963710 | 3815560 | 1270501 | 7580320 | 7275 | 28800320 | 23226701 | 44654 | 692 | 218 | 6047 | N/F | 221329 | 17078 | 234 | N/F | N/F | 65301 | 10328 | 1008 | 5573 | 2126 | 1153 | N/F | 392 | 93064 | 45012 | 31965 | D,Q | Unknown
WO-22270-180 | 7146 | 75050 | 524 | 189 | 5399 | 2338431 | 2254397 | N/F | 40 | 3259 | 609 | 1351 | N/F | 13118 | 29334 | 113564 | 347165 | 115193 | 509927 | 1975791 | 543370 | 622 | 18072983 | 3246147 | 4484080 | 1482883 | 9321170 | 6772 | 30790723 | 25551444 | 59021 | 1030 | 152 | 6380 | N/F | 282275 | 24502 | 173 | N/F | 2108 | 80175 | 15954 | 1751 | 22248 | 2089 | 853 | N/F | 120 | 104948 | 42488 | 29425 |  | Unknown
210453-d2-reprep | 8640 | 62513 | 313 | 264 | 6614 | 2562909 | 2371990 | N/F | 46 | N/F | 329 | 4717 | N/F | 1259 | 23151 | 92870 | 298614 | 74984 | 380524 | 1739384 | 393360 | N/F | 13756687 | 2729733 | 4102904 | 1135938 | 8234356 | 6654 | 28508518 | 21987613 | 41538 | 1570 | 978 | 5521 | N/F | 36603694 | 24008 | 259 | N/F | N/F | 120200 | 19521 | 653 | 149080 | 7812 | 1063 | N/F | 70 | 123418 | 63514 | 14660 | D,Q | Unknown
AB-220715-01 LFB-H 10 | 39725586 | 94571 | 138584 | 233492 | 42132782 | 1688084 | 2176229 | 942232 | 3161751 | 278478 | 490428 | 1824989 | 1110535 | 1732321 | 94706 | 146815 | 237331 | 502091 | 1142118 | 4594567 | 2024045 | 3002 | 19219979 | 4022973 | 8528689 | 1788550 | 12295535 | 47073 | 30849628 | 23540197 | 231281 | 1365503 | 1451158 | 1664423 | 34438 | 17314369 | 28979289 | 1846690 | 4307415 | 1859407 | 133057 | 10818129 | 2020472 | 8049585 | 12920 | 4049144 | 407238 | 948622 | 26605827 | 32870799 | 21500403 |  | Unknown
AB-220715-01 LRB 19_INJ001 | 504690 | 56071 | 1066 | 1682 | 210154 | 2497754 | 2649689 | 2748 | 14750 | 5189 | 2280 | 7454 | 16066 | 10772 | 110492 | 245240 | 374577 | 523247 | 1433477 | 5673945 | 2083916 | 4661 | 27205917 | 5174139 | 12104147 | 2324384 | 16362010 | 32898 | 39810080 | 34283040 | 228759 | 4750 | 9317 | 11065 | N/F | 128694 | 659581 | 21462 | 20461 | 10905 | 142594 | 55520 | 11051 | 61130 | 7424 | 27880 | 18506 | 4448 | 609600 | 874784 | 242861 |  | Unknown
AB-220715-01 LRB 19 | 108562 | 72779 | 643 | 67 | 23495 | 2054776 | 2181846 | 285 | 2046 | 4041 | 1337 | 3373 | 8642 | 5104 | 90755 | 210782 | 304150 | 455894 | 1226736 | 5085319 | 1989883 | N/F | 23794591 | 4300225 | 10277703 | 1992401 | 13901883 | 18754 | 33400239 | 28685978 | 226224 | 651 | 2083 | 7001 | N/F | 31781 | 208111 | 2668 | N/F | 3730 | 86210 | 13625 | 4884 | 33090 | 4614 | 4115 | 24865 | 307 | 346446 | 346628 | 39985 |  | Unknown
AB-220715-01 LD 9 | 40772 | 66339 | 1069 | N/F | 7654 | 2322582 | 2205038 | N/F | 281 | 2369 | 648 | 1950 | N/F | 16559 | 25466 | 97876 | 338383 | 87728 | 437075 | 1677253 | 427754 | 1164 | 16173618 | 2971139 | 4044534 | 1344566 | 7983247 | 6853 | 19664981 | 19992938 | 44848 | 1144 | 1285 | 4280 | N/F | 211734 | 96771 | 837 | N/F | N/F | 10670 | 16718 | 3692 | 6618 | 1068 | 2658 | 13391 | 64 | 262460 | 233626 | 18612 |  | Unknown
AB-220715-01 CCC-H 10 | 68644616 | 174355 | 224417 | 393470 | 72610027 | 1434971 | 2012785 | 1894444 | 5774867 | 568949 | 995393 | 3052473 | 1979024 | 3133013 | 83396 | 115407 | 213050 | 496252 | 1001249 | 4174848 | 1937687 | N/F | 16156602 | 3656733 | 7871819 | 1533437 | 11186151 | 48911 | 27926845 | 19222363 | 222916 | 2567670 | 2522446 | 3160907 | 79847 | 29026980 | 49987879 | 3229065 | 7990749 | 3592858 | 40365 | 19235370 | 3413907 | 14786797 | 2261 | 7391520 | 847744 | 1797970 | 35996371 | 51167032 | 33466968 |  | Unknown
AB-220715-01 LFSM-H 9 | 27164147 | 76329 | 64968 | 197232 | 26738027 | 1545947 | 1731979 | 145288 | 1106989 | 58714 | 87396 | 696884 | 366311 | 1011650 | 22246 | 70844 | 216694 | 87609 | 399748 | 1477259 | 388732 | N/F | 13002454 | 2605984 | 3360966 | 1117379 | 6758566 | 4407 | 17821948 | 16600286 | 44911 | 448532 | 1366225 | 1398315 | N/F | 11696560 | 17222211 | 1351237 | 1379538 | 370839 | 11142 | 5988743 | 1545673 | 3101264 | 1216 | 2583888 | 36490 | 254195 | 7202326 | 12275479 | 15394190 |  | Unknown
WO-22270-181 | 558093 | 55529 | 474 | 1054 | 99527 | 2541559 | 2903610 | 524 | 4373 | 6965 | 809 | 3586 | N/F | 4919 | 34583 | 113576 | 376135 | 128274 | 530007 | 2134145 | 637872 | 1240 | 18956485 | 3302136 | 4816138 | 1431554 | 8673978 | 7099 | 31064721 | 26373722 | 66730 | 2295 | 14386 | 9424 | N/F | 503638 | 662380 | 24261 | 17675 | 2471 | 31120 | 57475 | 8142 | 44214 | 675 | 9817 | N/F | 1087 | 561464 | 645039 | 258790 |  | Unknown
WO-22270-182 | 206978 | 101392 | 1472 | 564 | 17706 | 2969994 | 2675088 | 29 | 912 | 4101 | 648 | 1910 | N/F | 3184 | 20706 | 101785 | 648459 | 77988 | 404719 | 1552925 | 398398 | 996 | 16459906 | 2850392 | 3981976 | 1188611 | 8167641 | 6042 | 28196325 | 22726913 | 43052 | 881 | 3270 | 7786 | N/F | 299920 | 332537 | 5882 | N/F | N/F | 76707 | 16544 | 6883 | 14191 | 4005 | 3296 | N/F | 100 | 481480 | 517108 | 60018 |  | Unknown
WO-22270-183 | 94721 | 55543 | 725 | 205 | 15674 | 2552126 | 2414103 | 8 | 285 | 6959 | 269 | 2187 | N/F | 4589 | 33425 | 111515 | 393390 | 138842 | 528202 | 2116849 | 608931 | 443 | 19253640 | 3482020 | 4724255 | 1473230 | 9552472 | 8715 | 31010478 | 26567073 | 62339 | 1703 | 798 | 6886 | N/F | 329735 | 201755 | 2311 | N/F | N/F | 76787 | 14249 | 5012 | 34840 | 2205 | 3899 | 6405 | 78 | 349117 | 413846 | 28012 |  | Unknown
WO-22270-184 | 39634 | 77187 | 2671 | 390 | 9243 | 3045376 | 3107179 | N/F | 408 | 16754 | 588 | 1793 | N/F | 2222 | 33445 | 108132 | 670962 | 134816 | 517852 | 2138816 | 639985 | N/F | 20139215 | 3527137 | 4647348 | 1462737 | 9434538 | 8644 | 31957941 | 25600649 | 70329 | 1463 | 728 | 6225 | N/F | 352874 | 105457 | 1496 | N/F | N/F | 86619 | 66840 | 5080 | 17512 | 1623 | 2663 | 15167 | 18 | 288500 | 312140 | 18848 |  | Unknown
WO-22270-185 | 26032 | 132927 | 1871 | 1140 | 9694 | 4114102 | 4439424 | N/F | 374 | 27259 | 1325 | 2701 | N/F | 3072 | 44432 | 284415 | 1528768 | 173171 | 741466 | 3012761 | 814638 | 525 | 23150911 | 4995261 | 8056361 | 2206331 | 14661485 | 8066 | 31879476 | 27876175 | 90282 | 6127 | 2843 | 6718 | N/F | 363564 | 77254 | 979 | 18229 | 13524 | 42552 | 22413 | 5335 | 21957 | 865 | 2651 | 30786 | 47 | 208887 | 233683 | 16360 |  | Unknown
210439-d2-reprep | 8393 | 77531 | 565 | 64 | 5433 | 2748932 | 2476931 | N/F | 143 | N/F | 681 | 1770 | N/F | 14281 | 28180 | 96116 | 226583 | 96543 | 438961 | 1940787 | 514912 | 1392 | 13086565 | 3045019 | 4449372 | 1291088 | 9071610 | 6448 | 30817615 | 25281866 | 50146 | 3584 | 996 | 6272 | N/F | 63726070 | 22597 | 160 | N/F | N/F | 112266 | 42454 | 1077 | 676202 | 6269 | 1052 | N/F | 132 | 129924 | 84239 | 18722 |  | Unknown
WO-22270-186 | 17864 | 94535 | 330 | 464 | 5582 | 2517533 | 2315006 | N/F | 87 | 11289 | 318 | 1277 | N/F | 17415 | 18513 | 83338 | 427778 | 69596 | 358339 | 1340149 | 326720 | N/F | 13512307 | 2417070 | 3342417 | 1068896 | 6822586 | 5478 | 26634056 | 20612123 | 34076 | 1015 | 764 | 5417 | N/F | 183481 | 56030 | 444 | N/F | 2374 | 76422 | 14043 | 4439 | 13588 | 2503 | 2845 | N/F | 177 | 198907 | 183684 | 13584 | D,Q | Unknown
WO-22270-187 | 17312 | 78141 | 1598 | 273 | 8616 | 3105591 | 3137435 | N/F | 56 | 154304 | 902 | 2595 | N/F | 6979 | 54399 | 152561 | 613865 | 247399 | 678660 | 3081640 | 1155541 | N/F | 23880077 | 4428297 | 6336298 | 1965195 | 12985380 | 20413 | 36863822 | 30522292 | 122710 | 750 | 762 | 7349 | N/F | 467064 | 53312 | 346 | N/F | 4148 | 118115 | 25029 | 4170 | 50651 | 4851 | 2321 | N/F | 117 | 295254 | 245126 | 13431 | D,Q | Unknown
WO-22270-188 | 25306 | 130219 | 2530 | 592 | 16763 | 4410106 | 3966564 | 96 | 575 | 138516 | 1383 | 2872 | N/F | 3859 | 65085 | 279624 | 1495305 | 269362 | 761956 | 3426436 | 1236410 | N/F | 24832194 | 5412626 | 7920627 | 2152651 | 14474228 | 20853 | 35133627 | 30346587 | 125616 | 2655 | 1076 | 7757 | N/F | 834435 | 50264 | 2114 | 12160 | 5070 | 103190 | 21897 | 4890 | 166862 | 6044 | 3648 | N/F | 465 | 185414 | 162019 | 19642 |  | Unknown
WO-22270-189 | 13289 | 202250 | 4075 | 338 | 8319 | 4195390 | 4444594 | 80 | 426 | 130511 | 982 | 2598 | N/F | 4657 | 52056 | 185304 | 1211984 | 230129 | 719147 | 3039471 | 1042060 | N/F | 26099177 | 4691945 | 6857499 | 2026103 | 14163833 | 16292 | 35332446 | 30898688 | 120679 | 792 | 745 | 7261 | N/F | 828208 | 39865 | 1245 | 35460 | 6383 | 90873 | 48333 | 3936 | 54022 | 3025 | 1352 | 4266 | 205 | 178957 | 144277 | 20596 | D,Q | Unknown
WO-22270-190 | 13196 | 129742 | 848 | 480 | 7324 | 3385725 | 3380513 | N/F | N/F | 94096 | 573 | 2816 | N/F | 2611 | 40796 | 132285 | 794217 | 173696 | 555196 | 2360636 | 765946 | N/F | 22583744 | 4021777 | 5309754 | 1732972 | 10823491 | 10518 | 33708414 | 27803523 | 90012 | 3380 | 760 | 7999 | N/F | 770711 | 29169 | 575 | N/F | 2743 | 84349 | 47088 | 2736 | 29831 | 4638 | 808 | 3200 | 41 | 154598 | 119433 | 13868 | D,Q | Unknown
AB-220715-01 LRB 20_INJ001 | 15132 | 56595 | 727 | 285 | 5963 | 2301883 | 2441407 | 145 | 649 | 4696 | 1660 | 949 | N/F | 15497 | 97905 | 235916 | 333220 | 495819 | 1365573 | 5549776 | 2217661 | 4936 | 25699741 | 4284905 | 11680942 | 2111009 | 15617144 | 46383 | 37831393 | 32251869 | 252520 | 1617 | 176 | 7712 | N/F | 26104 | 43746 | 58 | N/F | N/F | 127902 | 10756 | 4332 | 20108 | 6724 | 2033 | 4970 | 13 | 260293 | 195700 | 16641 |  | Unknown
AB-220715-01 LRB 20 | 11362 | 56621 | 1139 | 21 | 6284 | 1664744 | 1727037 | 85 | N/F | 3194 | 726 | 925 | 8515 | 3211 | 84316 | 200638 | 255559 | 425025 | 1047795 | 4725506 | 1842133 | 3885 | 19889324 | 3421158 | 9592561 | 1689166 | 12643900 | 19932 | 29509342 | 24939790 | 204594 | 833 | 674 | 5962 | N/F | 19596 | 30109 | 789 | N/F | N/F | 81536 | 23237 | 3484 | 18105 | 3571 | 1006 | 3765 | 188 | 243674 | 125983 | 16151 |  | Unknown
WO-22270-196 | 9559 | 74913 | 1193 | 323 | 5913 | 3475175 | 3609015 | 33 | 56 | 1205951 | 558 | 2455 | N/F | 18645 | 56168 | 152660 | 635473 | 240146 | 643972 | 3077886 | 1139282 | N/F | 25598630 | 4480390 | 6522938 | 1933863 | 13141993 | 20721 | 36274402 | 31830044 | 122240 | 1855 | 250 | 9976 | N/F | 908718 | 28550 | 352 | N/F | 8632 | 107972 | 24622 | 2822 | 30238 | 4913 | 1128 | N/F | N/F | 121927 | 56077 | 15345 |  | Unknown
WO-22270-191 | 11371 | 114386 | 2845 | 240 | 6041 | 4168593 | 4169182 | N/F | 19 | 493174 | 625 | 2253 | N/F | 20085 | 61418 | 197773 | 1066820 | 266195 | 718456 | 3340064 | 1247643 | N/F | 25991201 | 4756345 | 7343352 | 2119930 | 14534261 | 21024 | 36443751 | 32417895 | 126505 | 2741 | 1023 | 9191 | N/F | 932421 | 30458 | 898 | N/F | 4682 | 113135 | 58931 | 2669 | 45380 | 3891 | 1427 | N/F | 39 | 150400 | 96171 | 17359 | D,Q | Unknown
WO-22270-197 | 8544 | 108285 | 223 | 91 | 5127 | 3060822 | 2907200 | N/F | N/F | 661577 | 451 | 1524 | N/F | 10039 | 41507 | 114366 | 533290 | 140258 | 532702 | 2301723 | 685938 | 655 | 20959273 | 3613718 | 5169965 | 1546636 | 10400087 | 10025 | 32629540 | 27735261 | 71453 | 1900 | 882 | 7637 | N/F | 307083 | 26348 | 653 | N/F | 5812 | 110655 | 38383 | 2551 | 40740 | 5413 | 1710 | N/F | 289 | 111271 | 47637 | 18723 |  | Unknown
WO-22270-192 | 13755 | 89051 | 1404 | 534 | 8024 | 3255030 | 3194752 | N/F | N/F | 642411 | 479 | 2437 | N/F | 17061 | 62809 | 159453 | 593436 | 303874 | 712567 | 3138719 | 1275267 | N/F | 27126592 | 4732495 | 6666580 | 2093905 | 13334156 | 24153 | 37008033 | 32957844 | 136209 | 1805 | 152 | 7224 | N/F | 905056 | 31473 | 383 | 10172 | 6731 | 121922 | 20139 | 2967 | 52485 | 5302 | 1839 | N/F | 98 | 163229 | 103489 | 17224 | D,Q | Unknown
WO-22270-198 | 10937 | 70483 | 1127 | 55 | 4930 | 2716667 | 2716111 | N/F | 6 | 4957461 | 3088 | 1487 | N/F | 18341 | 34243 | 132773 | 415753 | 169954 | 588901 | 2632540 | 634144 | 1138 | 22458245 | 3893802 | 5584033 | 1639334 | 11065748 | 11645 | 35158316 | 28517614 | 67161 | 1036 | 306 | 10045 | N/F | 422744 | 24109 | 1170 | 7455 | 10095978 | 88421 | 15621 | 2156 | 16044 | 2019 | 1533 | 4428 | 340 | 114429 | 49156 | 12408 |  | Unknown
WO-22270-193 | 10275 | 87391 | 516 | 302 | 8211 | 2751426 | 3256161 | N/F | N/F | 443882 | 410 | 1573 | N/F | 4575 | 48029 | 132788 | 537847 | 209770 | 600249 | 2709957 | 956387 | N/F | 22468485 | 4265959 | 5772170 | 1812329 | 11873065 | 16422 | 36378824 | 26626199 | 106735 | 785 | 573 | 6719 | N/F | 845512 | 29150 | 772 | N/F | 3564 | 118873 | 21558 | 1784 | 17370 | 3532 | 1548 | N/F | N/F | 135098 | 89654 | 14078 | D,Q | Unknown
WO-22270-199 | 7735 | 108890 | 1194 | 234 | 4712 | 2702542 | 2648514 | N/F | N/F | 5251817 | 2362 | 2229 | N/F | 14472 | 33880 | 127379 | 428635 | 162087 | 567704 | 2477762 | 604404 | N/F | 21670250 | 3858263 | 5411755 | 1606146 | 10627487 | 13829 | 32825112 | 26677468 | 64619 | 1133 | 54 | 7189 | N/F | 511934 | 21091 | 130 | N/F | 10165882 | 80677 | 20830 | 2446 | 25640 | 3219 | 1061 | 1309 | 36 | 119372 | 45944 | 12191 | D,Q | Unknown
210415-d2-reprep | 7853 | 108223 | 513 | 80 | N/F | 2693180 | 2227796 | N/F | N/F | N/F | 459 | 2718 | N/F | 1844 | 21111 | 83175 | 298085 | 85097 | 385010 | 1687173 | 424919 | N/F | 16095196 | 2766395 | 3955236 | 1107166 | 8137461 | 8466 | 28553540 | 23616166 | 46123 | 1104 | 696 | 5862 | N/F | 2196621 | 25042 | 609 | N/F | 2057 | 122848 | 23075 | 3036 | 504866 | 6185 | 1120 | N/F | 21 | 147347 | 89679 | 9849 | D,Q | Unknown
WO-22270-200 | 12249 | 129741 | 1612 | 196 | 5551 | 2822996 | 2538817 | N/F | 14 | 4731481 | 1915 | 1368 | N/F | 6931 | 36247 | 120612 | 404347 | 151874 | 541995 | 2337775 | 599146 | 1032 | 19825367 | 3522961 | 5338460 | 1516460 | 9989664 | 9850 | 33040892 | 29166318 | 61754 | 633 | 565 | 7614 | N/F | 312091 | 19905 | 218 | N/F | 9570492 | 92934 | 36644 | 1801 | 22090 | 3646 | 1366 | N/F | N/F | 106847 | 42409 | 15896 |  | Unknown
WO-22270-194 | 10482 | 71986 | 2249 | 2424 | 8461 | 3164550 | 3459614 | N/F | 196 | 1176117 | 1013 | 2838 | N/F | 2792 | 67543 | 192429 | 547904 | 299005 | 815770 | 3675290 | 1414189 | N/F | 29615948 | 4979431 | 8333244 | 2263229 | 15929215 | 23998 | 41782086 | 35639961 | 139349 | 863 | 118 | 7965 | N/F | 789105 | 28655 | 476 | 4519 | 7212 | 101307 | 27304 | 1516 | 48620 | 3805 | 1155 | N/F | 119 | 113023 | 61368 | 15283 | D,Q | Unknown
WO-22270-195 | 8073 | 64268 | 629 | 81 | 4864 | 2492256 | 2449810 | 7 | 151 | 538074 | 409 | 1719 | N/F | 15896 | 43499 | 121157 | 374565 | 172405 | 569531 | 2495069 | 813259 | 1537 | 21702220 | 3889121 | 5337636 | 1629132 | 10417731 | 12327 | 34253435 | 28763560 | 89450 | 1027 | 446 | 7593 | N/F | 430315 | 26220 | 281 | N/F | 5092 | 110642 | 13009 | 2114 | 30408 | 3408 | 1094 | N/F | N/F | 133442 | 63509 | 10859 |  | Unknown
AB-220715-01 LFB-H 11 | 8300103 | 102578 | 151891 | 165795 | 37091052 | 1339382 | 1510707 | 1012549 | 3279981 | 411277 | 544623 | 1747135 | 1155467 | 1564650 | 104259 | 152134 | 177524 | 551945 | 1211979 | 5429895 | 2370944 | 3251 | 12123754 | 2815373 | 10268475 | 1658269 | 12889642 | 59029 | 2439918 | 6507855 | 268344 | 1604874 | 1104598 | 1197644 | 52513 | 10912514 | 1308668 | 441616 | 5147848 | 2236546 | 2805 | 11534321 | 1229182 | 9203621 | 280 | 2869232 | 488593 | 1062719 | 84104 | 142683 | 5596511 |  | Unknown
AB-220715-01 LRB 21 | 19127 | 72520 | 305 | 97 | 22973 | 2018130 | 1993054 | 296 | 2300 | 3767 | 1047 | 1435 | N/F | 18413 | 98501 | 212014 | 300489 | 459063 | 1190488 | 5381324 | 2093452 | 4562 | 22951723 | 4078338 | 10613237 | 1873834 | 14466924 | 22186 | 32214217 | 28860293 | 241550 | 3371 | 528 | 5444 | N/F | 27883 | 27888 | 507 | N/F | 6170 | 95559 | 11141 | 3277 | 30829 | 6430 | 4054 | 14949 | 449 | 90782 | 31055 | 14516 |  | Unknown
AB-220715-01 LD 10 | 11695 | 121897 | 1982 | 518 | 11995 | 3239830 | 3591831 | 8 | 334 | 2325449 | 795 | 1782 | N/F | 17694 | 31041 | 139851 | 727302 | 140679 | 538589 | 2342617 | 627072 | N/F | 21373739 | 3718049 | 5347087 | 1698249 | 11200934 | 8432 | 31359191 | 27209242 | 64412 | 812 | 728 | 6200 | N/F | 276811 | 22220 | 755 | N/F | 6592735 | 103695 | 24217 | 2671 | 27542 | 4237 | 1997 | N/F | 57 | 90434 | 30085 | 13490 |  | Unknown
AB-220715-01 CCC-H 11 | 62552188 | 207935 | 243513 | 389966 | 73359033 | 1506161 | 1945719 | 2119115 | 6083579 | 799752 | 1106896 | 3087642 | 2131782 | 3129270 | 106256 | 120679 | 204008 | 533110 | 1057160 | 5033259 | 2516157 | 4721 | 16211907 | 3906454 | 9109447 | 1602482 | 12641679 | 63416 | 24572230 | 18986827 | 286852 | 2979081 | 2596559 | 3138564 | 97016 | 30991789 | 47238687 | 3043485 | 9369287 | 4468238 | 2595 | 22557576 | 3201908 | 16729028 | 587 | 7766560 | 1025063 | 2082237 | 9435820 | 31931568 | 33784552 |  | Unknown
210405-d2-reprep | 17900 | 121431 | 1659 | 728 | 13230 | 3294712 | 2922934 | N/F | 163 | 19678 | 1213 | 2944 | N/F | 4692 | 17453 | 82440 | 363929 | 108593 | 467768 | 2254105 | 652479 | N/F | 18631992 | 3757811 | 5062016 | 1271826 | 9339887 | 11790 | 30737034 | 26027307 | 78051 | 5214 | 1924 | 8522 | N/F | 3359472 | 42546 | 1142 | N/F | 27354 | 123174 | 36059 | 2662 | 51075 | 6611 | 3959 | N/F | 64 | 102206 | 94166 | 22796 | D,Q | Unknown
AB-220715-01 LFSM-H 10 | 33053985 | 205488 | 78188 | 442677 | 29238898 | 2339082 | 2863278 | 157275 | 1077678 | 1855928 | 96435 | 753422 | 398517 | 1055766 | 23100 | 84798 | 472744 | 97878 | 425695 | 1809408 | 423311 | N/F | 15765296 | 2919915 | 4128931 | 1217569 | 7908652 | 6755 | 26227227 | 20577475 | 51976 | 517654 | 2004048 | 2183224 | N/F | 13828310 | 25060463 | 1565779 | 1647284 | 5026234 | 150556 | 7019115 | 1672745 | 3662555 | 34056 | 2924044 | 47637 | 261838 | 25631714 | 29181725 | 18640318 |  | Unknown
WO-22270-201 | 312025 | 125584 | 1668 | 830 | 47016 | 2599212 | 2414262 | 27 | 1847 | 4156516 | 1183 | 3393 | N/F | 8864 | 26124 | 104482 | 368656 | 115076 | 479193 | 2042148 | 503451 | N/F | 18531576 | 3044015 | 4412096 | 1387673 | 8828642 | 7649 | 30722272 | 24511707 | 53468 | 1756 | 5019 | 6648 | N/F | 300939 | 414824 | 11902 | N/F | 7634346 | 100381 | 33011 | 8242 | 25639 | 5839 | 7917 | N/F | 186 | 291308 | 401806 | 113185 |  | Unknown
WO-22270-202 | 101140 | 56850 | 1152 | 238 | 21530 | 2531157 | 2535540 | 10 | 286 | 4121898 | 1114 | 2454 | N/F | 4629 | 26447 | 106538 | 373671 | 110319 | 528127 | 2111930 | 453729 | 418 | 20517245 | 3936619 | 4885475 | 1603564 | 11059800 | 10213 | 34087430 | 29226073 | 49345 | 3798 | 866 | 7285 | N/F | 180415 | 183489 | 2445 | N/F | 7660050 | 98722 | 14872 | 5347 | 9480 | 5078 | 2657 | N/F | 464 | 210371 | 226280 | 29170 |  | Unknown
WO-22270-203 | 37735 | 55936 | 1008 | 254 | 12645 | 2518716 | 2578022 | N/F | 694 | 5006172 | 1661 | 2040 | N/F | 3532 | 30084 | 115064 | 361149 | 159799 | 587666 | 2376762 | 637910 | 1511 | 20235970 | 4040754 | 5152350 | 1567207 | 11351161 | 11191 | 34267592 | 27898593 | 61876 | 669 | 1012 | 6892 | N/F | 308249 | 108198 | 1231 | N/F | 10376157 | 104773 | 16260 | 3645 | 9158 | 7815 | 2953 | 9618 | 31 | 166601 | 192487 | 13852 |  | Unknown
AB-220715-01 LRB 21_INJ001 | 56952 | 57243 | 1098 | 1070 | 176228 | 2297486 | 2189545 | 3001 | 11701 | 5716 | 1551 | 7340 | N/F | 20760 | 98287 | 256407 | 308502 | 476769 | 1347453 | 5928264 | 2165647 | 5387 | 25827315 | 4537931 | 11944023 | 2116050 | 16532369 | 37697 | 40632069 | 33874755 | 247394 | 6750 | 6376 | 8550 | N/F | 65437 | 39367 | 2617 | 21458 | 17770 | 133636 | 47524 | 5579 | 62002 | 8094 | 16145 | 12642 | 3180 | 92429 | 35941 | 37570 |  | Unknown
WO-22270-204 | 17832 | 70008 | 763 | 123 | 8223 | 2241633 | 2150724 | N/F | 141 | 2942196 | 934 | 1408 | N/F | 5061 | 15081 | 89635 | 277149 | 72212 | 369779 | 1566810 | 310477 | N/F | 16437401 | 2840280 | 3656935 | 1206656 | 7737605 | 8216 | 30122983 | 23599311 | 40094 | 1453 | 386 | 5573 | N/F | 136873 | 86023 | 598 | N/F | 5164732 | 108136 | 13374 | 2782 | 48956 | 5594 | 2048 | 12677 | 253 | 179928 | 168606 | 12649 |  | Unknown
WO-22270-205 | 15074 | 108560 | 2050 | 132 | 5614 | 2456603 | 2269103 | N/F | 37 | 4759559 | 1300 | 1962 | N/F | 8896 | 31630 | 112654 | 331253 | 135280 | 524284 | 2318626 | 553318 | 2302 | 19113363 | 3302608 | 5052642 | 1438636 | 9676819 | 9178 | 33850264 | 27441481 | 63613 | 2620 | 834 | 4381 | N/F | 187568 | 44727 | 246 | N/F | 9131872 | 104820 | 21548 | 5072 | 5896 | 5492 | 1611 | N/F | 170 | 153193 | 123758 | 10804 |  | Unknown
WO-22270-206 | 9450 | 85810 | 2372 | 236 | 4433 | 2487020 | 2282657 | N/F | 68 | 3845542 | 902 | 2281 | N/F | 18253 | 21988 | 100941 | 285885 | 89706 | 405642 | 1939750 | 430233 | 849 | 17110554 | 2645828 | 4363613 | 1161523 | 8297687 | 6202 | 31380877 | 25625571 | 51225 | 4381 | 63 | 6613 | N/F | 312911 | 23821 | 140 | N/F | 6748674 | 127915 | 28084 | 3819 | 20076 | 7971 | 1352 | N/F | 109 | 115861 | 51787 | 11526 |  | Unknown
WO-22270-207 | 5656 | 70531 | 891 | 112 | 4494 | 2359358 | 2073401 | N/F | N/F | 2851255 | 724 | 1083 | N/F | 3571 | 16339 | 96343 | 336241 | 62211 | 331950 | 1463601 | 288327 | N/F | 14515480 | 2327833 | 3735302 | 981687 | 7687885 | 7888 | 26318461 | 21979217 | 39905 | 2460 | 174 | 5142 | N/F | 336239 | 17927 | 456 | N/F | 4946374 | 123029 | 26850 | 3482 | 35013 | 9012 | 743 | N/F | 39 | 74391 | 36076 | 7504 | D,Q | Unknown
WO-22270-208 | 8909 | 66061 | 1135 | N/F | 5979 | 2534823 | 2372244 | N/F | N/F | 3047778 | 1633 | 2014 | N/F | 3187 | 30005 | 106441 | 349290 | 121904 | 466090 | 2172200 | 530208 | N/F | 19465500 | 3179151 | 4888982 | 1331510 | 9756775 | 7325 | 33714520 | 27843778 | 63483 | 3201 | 674 | 8227 | N/F | 209115 | 25463 | 285 | N/F | 7210798 | 135077 | 14443 | 3224 | 23029 | 8176 | 1341 | N/F | 226 | 116617 | 43068 | 12722 | D,Q | Unknown
WO-22270-209 | 8357 | 80740 | 690 | 239 | 9629 | 2888848 | 2587611 | 17 | 18 | 3831700 | 1566 | 2479 | N/F | 52412 | 39740 | 132766 | 374695 | 172114 | 537979 | 2707808 | 673820 | N/F | 22370756 | 3639080 | 5839527 | 1542083 | 11268998 | 11136 | 34597397 | 30503898 | 77582 | 3791 | 594 | 6684 | N/F | 637576 | 27913 | 963 | N/F | 9205464 | 109800 | 28858 | 1957 | 52630 | 6194 | 1164 | N/F | 54 | 96436 | 41347 | 13893 |  | Unknown
WO-22270-210 | 8490 | 64695 | 720 | 320 | 5265 | 2874013 | 2644704 | N/F | N/F | 4286981 | 1489 | 1589 | N/F | 12614 | 36511 | 124293 | 402229 | 164860 | 571580 | 2783024 | 719747 | N/F | 23175947 | 3795063 | 5975251 | 1600983 | 12077200 | 13982 | 36592000 | 30984253 | 81550 | 1467 | 210 | 4933 | N/F | 634337 | 24013 | 884 | N/F | 9863046 | 130234 | 15470 | 2547 | 24611 | 6962 | 1211 | N/F | N/F | 89340 | 35803 | 13116 | D,Q | Unknown
AB-220715-01 LRB 22_INJ001 | 11291 | 56332 | 675 | 118 | 8959 | 2202590 | 2170194 | 66 | 17 | 42480 | 1639 | 981 | N/F | 14276 | 102358 | 245910 | 332309 | 477532 | 1304177 | 6359957 | 2460376 | 6040 | 27996557 | 4183363 | 12532133 | 2087106 | 17969396 | 44556 | 39671586 | 35426456 | 280454 | 1134 | 286 | 6361 | N/F | 34913 | 32126 | 407 | N/F | 91011 | 135162 | 40008 | 2161 | 31607 | 7918 | 1298 | 5923 | 148 | 116496 | 49318 | 19013 |  | Unknown
AB-220715-01 LRB 22 | 16166 | 45826 | 689 | 231 | 10304 | 1571959 | 1489872 | N/F | 156 | 8101 | 617 | 1418 | N/F | 15527 | 74040 | 188098 | 238767 | 343182 | 922690 | 4543262 | 1697926 | 4263 | 19412130 | 3053423 | 9138883 | 1463968 | 12911343 | 20648 | 29270139 | 25550953 | 179572 | 961 | 166 | 4451 | N/F | 27297 | 27868 | 535 | 17719 | 22152 | 86822 | 26680 | 2418 | 34000 | 6039 | 1372 | 3229 | 315 | 113610 | 40269 | 14669 |  | Unknown
WO-22270-211 | 11391 | 42656 | 554 | 103 | 5220 | 2677206 | 2601572 | 8 | 32 | 3669083 | 1659 | 1305 | N/F | 12844 | 37653 | 132177 | 357338 | 155620 | 546469 | 2662864 | 688156 | 2203 | 22430735 | 3889188 | 5715933 | 1597403 | 11448530 | 9728 | 35545385 | 31362643 | 71482 | 1547 | 487 | 7008 | N/F | 314661 | 27444 | 251 | N/F | 9031795 | 124718 | 13818 | 1259 | 19279 | 6975 | 2420 | N/F | 115 | 95874 | 30307 | 15591 |  | Unknown
WO-22270-212 | 7669 | 46684 | 240 | 78 | 5300 | 2486375 | 2419707 | N/F | 73 | 3137236 | 1865 | 2031 | N/F | 43893 | 32765 | 116012 | 358441 | 124086 | 494734 | 2339850 | 604138 | N/F | 21237452 | 3878303 | 5130742 | 1497004 | 10811392 | 11841 | 33654171 | 28488882 | 66935 | 2779 | 237 | 7485 | N/F | 409709 | 26068 | 475 | N/F | 7723499 | 116204 | 16919 | 1989 | 40653 | 7874 | 1238 | N/F | 231 | 88624 | 32020 | 12209 | D,Q | Unknown
WO-22270-213 | 6285 | 43971 | 1046 | 105 | 5609 | 2770240 | 2843749 | N/F | N/F | 4053827 | 1725 | 1679 | N/F | 19171 | 40341 | 138558 | 427921 | 156739 | 566561 | 2715015 | 718723 | 4317 | 22977323 | 4053729 | 5825835 | 1722576 | 12507360 | 14404 | 37145700 | 30265495 | 85676 | 1073 | 153 | 7546 | N/F | 483397 | 25420 | 199 | N/F | 9537727 | 119441 | 24451 | 1860 | 36698 | 7955 | 967 | N/F | 123 | 84445 | 28264 | 10146 |  | Unknown
WO-22270-214 | 6522 | 108209 | 2894 | 621 | 4072 | 4539473 | 4771740 | 21 | N/F | 7378664 | 3003 | 2367 | N/F | 6054 | 51702 | 229641 | 1126258 | 225661 | 682200 | 3498357 | 926666 | 2438 | 26389497 | 4734977 | 7941049 | 2028914 | 15341442 | 21435 | 38816731 | 32974346 | 89100 | 5438 | 239 | 6689 | N/F | 1789912 | 25413 | 347 | N/F | 14812187 | 129451 | 96237 | 1324 | 52113 | 6391 | 2232 | N/F | 9 | 84787 | 29131 | 11917 |  | Unknown
WO-22270-215 | 5402 | 135273 | 2433 | 702 | 3391 | 4176153 | 4060069 | N/F | 36 | 4546002 | 2076 | 1602 | N/F | 3091 | 25747 | 161656 | 1085132 | 99597 | 480468 | 2269269 | 530274 | N/F | 19718899 | 3500936 | 5368983 | 1424552 | 10532768 | 7520 | 34149628 | 27595557 | 56715 | 6454 | 638 | 6280 | N/F | 252462 | 22651 | 284 | 12693 | 8194978 | 133687 | 26662 | 817 | 41157 | 8066 | 1608 | 10478 | 8 | 86045 | 32263 | 11696 |  | Unknown
WO-22270-216 | 6195 | 73877 | 529 | 206 | 5304 | 2726022 | 2530439 | N/F | N/F | 4057707 | 2204 | 1228 | N/F | 3059 | 35796 | 114146 | 372582 | 145052 | 545555 | 2624182 | 695961 | 1882 | 21788958 | 3462874 | 5124682 | 1591366 | 11166318 | 11818 | 36370572 | 30596928 | 85626 | 3975 | 52 | 7587 | N/F | 807030 | 21239 | 305 | N/F | 11129646 | 119383 | 40456 | 1485 | 53662 | 5025 | 746 | 19663 | N/F | 75444 | 26490 | 19377 |  | Unknown
WO-22270-217 | 4961 | 61103 | 436 | 469 | 4827 | 2764257 | 2458533 | N/F | N/F | 3736563 | 1297 | 2111 | N/F | 14497 | 32291 | 126558 | 370237 | 128228 | 515257 | 2459430 | 628020 | N/F | 20589884 | 3213941 | 5442785 | 1379006 | 10309220 | 9822 | 34627258 | 30356599 | 77522 | N/F | 106 | 6725 | N/F | 449416 | 19824 | 70 | N/F | 10349730 | 107665 | 31383 | 2050 | 47461 | 5187 | 703 | 5640 | 182 | 79191 | 23647 | 15123 |  | Unknown
WO-22270-218 | 5976 | 54756 | 486 | 224 | 3844 | 2262345 | 2236572 | N/F | 19 | 1694263 | 1077 | 1615 | N/F | 11841 | 44056 | 141010 | 331739 | 167267 | 596399 | 2791465 | 832272 | N/F | 22102196 | 3572074 | 6062329 | 1592824 | 10781459 | 12976 | 35667119 | 28942155 | 103059 | 628 | 482 | 7118 | N/F | 270234 | 25330 | 330 | N/F | 4535970 | 128451 | 19757 | 1294 | 12592 | 7771 | 2053 | N/F | 75 | 73645 | 23161 | 11889 | D,Q | Unknown
WO-22270-219 | 5598 | 101909 | 2236 | 359 | 4349 | 2728844 | 3087515 | 43 | N/F | 1997080 | 1026 | 700 | N/F | 4046 | 45844 | 162319 | 450121 | 200937 | 635989 | 3227442 | 985580 | N/F | 25175785 | 3672166 | 6853538 | 1608310 | 12721918 | 14625 | 36859882 | 31183282 | 109178 | 1707 | 36 | 6307 | N/F | 164155 | 25429 | 195 | N/F | 5330895 | 116486 | 17854 | 1239 | 26057 | 8005 | 1216 | N/F | N/F | 78119 | 23486 | 13433 |  | Unknown
WO-22270-220 | 5652 | 110541 | 930 | 126 | 3611 | 3049007 | 2627370 | N/F | N/F | 1504121 | 516 | 1610 | N/F | 3812 | 43123 | 132933 | 458514 | 159008 | 548267 | 2684486 | 794512 | 1937 | 21250524 | 3493487 | 5814175 | 1490656 | 11275577 | 11184 | 33822810 | 30032673 | 90086 | 3141 | 141 | 7163 | N/F | 226547 | 21441 | 292 | N/F | 4422984 | 120749 | 13532 | 1156 | 10021 | 6094 | 1006 | 10682 | 116 | 69977 | 20350 | 12322 |  | Unknown
AB-220715-01 LFB-H 12 | 36381095 | 98671 | 148393 | 216699 | 39316537 | 1772012 | 2105501 | 829757 | 3168897 | 361754 | 451190 | 1659065 | 1074141 | 1563996 | 106636 | 155772 | 232840 | 463823 | 1177803 | 5712844 | 2498642 | 12454 | 21976360 | 4028401 | 11098232 | 1741780 | 15111322 | 50643 | 35157116 | 26995358 | 262976 | 1642505 | 1404723 | 1609855 | 36228 | 19048452 | 31989001 | 1682250 | 5278846 | 2166880 | 174172 | 13076749 | 1951907 | 9744951 | 34687 | 3899485 | 418087 | 899203 | 30475713 | 36043861 | 23477223 |  | Unknown
AB-220715-01 LRB 23_INJ001 | 402647 | 54320 | 693 | 2106 | 193873 | 2098694 | 2145973 | 3324 | 14301 | 3888 | 1877 | 7688 | 4589 | 11674 | 102683 | 256403 | 330224 | 476218 | 1315732 | 6234700 | 2491489 | 7660 | 27014804 | 4192262 | 12901297 | 2037344 | 17624904 | 49716 | 39662450 | 33305233 | 267466 | 8009 | 9942 | 7963 | N/F | 134202 | 477765 | 18222 | 26506 | 15596 | 130912 | 66260 | 10061 | 59737 | 5168 | 27729 | 5708 | 4107 | 398280 | 468384 | 215321 |  | Unknown
AB-220715-01 LRB 23 | 93460 | 53457 | 426 | N/F | 24551 | 1659891 | 1571594 | 446 | 1618 | 10433 | 618 | 813 | N/F | 16281 | 104603 | 213783 | 254963 | 453195 | 1061182 | 5575634 | 2186411 | 11598 | 20842471 | 3299747 | 10513831 | 1617515 | 14217666 | 16664 | 30332378 | 27464005 | 236105 | 750 | 2016 | 6253 | N/F | 26992 | 162845 | 3314 | 24680 | 22345 | 80100 | 11760 | 3647 | 24824 | 5061 | 4881 | 3494 | 748 | 257419 | 227917 | 38219 |  | Unknown
AB-220715-01 LD 11 | 46827 | 54489 | 1121 | 147 | 7820 | 2691807 | 2966040 | 24 | 700 | 1736444 | 665 | 2014 | 12575 | 2778 | 48689 | 154958 | 452702 | 192560 | 610756 | 3095658 | 968559 | 12544 | 23851746 | 3831322 | 6647475 | 1667184 | 12323242 | 13637 | 35827354 | 31028782 | 108695 | 1150 | 1405 | 6992 | N/F | 154288 | 105028 | 1200 | 19240 | 5047543 | 119340 | 17365 | 3123 | 34040 | 4356 | 3328 | N/F | N/F | 158137 | 155769 | 23642 |  | Unknown
AB-220715-01 CCC-H 12 | 67132820 | 194827 | 240422 | 377617 | 68418576 | 1564101 | 1987061 | 1717726 | 5829225 | 738343 | 906823 | 2905206 | 2004732 | 2871604 | 92603 | 132950 | 201369 | 464159 | 1021654 | 5283209 | 2381097 | 16789 | 18249143 | 3864453 | 9563423 | 1520143 | 13445640 | 52383 | 31317683 | 22353940 | 267099 | 3125450 | 2581997 | 3154617 | 78787 | 32677781 | 55545293 | 2935265 | 9809087 | 4370039 | 209188 | 23575348 | 3129772 | 17434277 | 63595 | 7608544 | 885608 | 1776807 | 55762356 | 64199347 | 39515268 |  | Unknown
AB-220715-01 LFSM-H 11 | 33877302 | 92892 | 94439 | 275479 | 32206137 | 2047813 | 2503716 | 303454 | 1592427 | 1772733 | 152159 | 943257 | 485040 | 1163110 | 38308 | 101426 | 283991 | 155733 | 533729 | 2570384 | 794392 | 2503 | 18420698 | 2928486 | 5431109 | 1303000 | 10005229 | 10404 | 31492016 | 24496569 | 90253 | 747005 | 1786691 | 1954672 | N/F | 16774095 | 30204012 | 1667092 | 2424442 | 5038635 | 149980 | 8743862 | 1840911 | 4975936 | 27292 | 3101917 | 100473 | 386227 | 28960602 | 33131987 | 22199581 |  | Unknown
WO-22270-221 | 327231 | 48422 | 1134 | 309 | 20959 | 2204623 | 2213568 | 56 | 1022 | 1910886 | 1185 | 1762 | N/F | 10425 | 49986 | 146188 | 310052 | 197478 | 643498 | 2999326 | 945716 | 3690 | 21968010 | 3757084 | 6540357 | 1709687 | 11936740 | 7905 | 37118808 | 29994751 | 108003 | 2276 | 2974 | 6606 | N/F | 163902 | 539891 | 9007 | N/F | 4898854 | 121602 | 15234 | 2078 | 12803 | 6726 | 6235 | 2762 | 202 | 549754 | 739675 | 100230 |  | Unknown
WO-22270-222 | 117573 | 54135 | 666 | 345 | 13671 | 2048660 | 1886591 | 13 | 892 | 1500100 | 914 | 2177 | N/F | 4755 | 38128 | 134996 | 252600 | 153941 | 547460 | 2586570 | 802223 | N/F | 19031010 | 3300496 | 5659326 | 1394304 | 10703960 | 10046 | 33625122 | 27343427 | 81361 | 1356 | 1495 | 5568 | N/F | 118063 | 256473 | 2769 | N/F | 4088731 | 115643 | 21939 | 2578 | 46159 | 6538 | 3596 | 2565 | 246 | 374812 | 451662 | 32889 | D,Q | Unknown
WO-22270-223 | 56357 | 44046 | 854 | 159 | 12160 | 2306835 | 2242020 | 66 | 301 | 1690883 | 491 | 1706 | N/F | 21117 | 51354 | 145398 | 295444 | 195432 | 601342 | 2940838 | 956808 | 7540 | 21877461 | 3756828 | 6492293 | 1625551 | 11896518 | 14152 | 36519222 | 30167351 | 105189 | N/F | 1174 | 6563 | N/F | 101586 | 162486 | 1431 | N/F | 4804981 | 123881 | 25650 | 5251 | 38151 | 7315 | 2949 | 19131 | 211 | 340521 | 363697 | 18826 |  | Unknown
WO-22270-224 | 28983 | 51410 | 250 | 84 | 6567 | 2233609 | 2103128 | N/F | 520 | 1802416 | 773 | 1647 | N/F | 2872 | 45899 | 148616 | 289846 | 197599 | 617437 | 3042827 | 912321 | 2671 | 20876279 | 3569848 | 6240995 | 1546054 | 11117897 | 14602 | 36850187 | 29761044 | 106397 | 3302 | 909 | 6525 | N/F | 214795 | 110638 | 995 | N/F | 4933483 | 117826 | 24928 | 3819 | 23863 | 7666 | 2768 | 9313 | 120 | 305987 | 318724 | 14545 |  | Unknown
WO-22270-225 | 22627 | 48756 | 721 | 171 | 8893 | 2584551 | 2494007 | N/F | 80 | 1727911 | 495 | 1529 | N/F | 3217 | 47010 | 158089 | 390749 | 178774 | 581314 | 2928703 | 939533 | 2848 | 22968960 | 3562636 | 6278527 | 1586849 | 11814640 | 10087 | 36259506 | 30970398 | 97477 | 4454 | 255 | 7930 | N/F | 194548 | 82943 | 420 | N/F | 4785112 | 118832 | 19043 | 3608 | 41925 | 6742 | 1783 | 14436 | 694 | 274154 | 272325 | 16941 |  | Unknown
WO-22270-226 | 15222 | 110917 | 630 | 272 | 5599 | 2417391 | 2608656 | N/F | 163 | 1949239 | 1286 | 895 | N/F | 17495 | 48364 | 142178 | 316378 | 206238 | 607152 | 3028326 | 982839 | 1728 | 21528776 | 3614861 | 6307692 | 1582656 | 11399572 | 15159 | 34828417 | 30742043 | 105271 | 1415 | 358 | 6222 | N/F | 156407 | 56994 | 623 | N/F | 5239862 | 116938 | 22761 | 2827 | 32118 | 5852 | 1692 | 19175 | 243 | 243343 | 233486 | 11284 |  | Unknown
AB-220715-01 LRB 24_INJ001 | 15379 | 44414 | 565 | N/F | 7990 | 2225709 | 2290815 | 96 | 102 | 16337 | 787 | 2672 | 12140 | 3083 | 110469 | 267599 | 338576 | 491306 | 1362996 | 6449184 | 2594506 | N/F | 27937892 | 4208453 | 13271086 | 2089775 | 18250216 | 52145 | 40894237 | 35837291 | 262966 | 1558 | 631 | 7656 | N/F | 23395 | 53067 | 348 | N/F | 46009 | 131271 | 8731 | 3142 | 22634 | 5314 | 1689 | 16204 | 257 | 288790 | 235985 | 14950 | D,Q | Unknown
AB-220715-01 LRB 24 | 10088 | 38928 | 369 | N/F | 6529 | 1683999 | 1610896 | 204 | 146 | 4986 | 589 | 573 | 9220 | 15961 | 67176 | 195250 | 275126 | 281158 | 993605 | 4386362 | 1527220 | 9287 | 20882693 | 3206341 | 9800208 | 1581659 | 13529513 | 12666 | 30643917 | 26524847 | 159622 | 2005 | 435 | 7826 | N/F | 20739 | 38331 | 450 | N/F | 10622 | 96302 | 8013 | 3249 | 18008 | 6804 | 1511 | 4152 | 40 | 314319 | 199915 | 12003 |  | Unknown
AB-220715-01 CCC-H 13 | 67004759 | 191770 | 252607 | 388877 | 71610730 | 1548163 | 2075909 | 1807160 | 5949231 | 717339 | 948636 | 3012853 | 1965897 | 2876286 | 99354 | 129878 | 201832 | 469969 | 1024841 | 5578768 | 2465554 | 14631 | 18796639 | 3741444 | 10213630 | 1505355 | 13836837 | 49744 | 31652205 | 23151359 | 288393 | 3226640 | 2646721 | 3216405 | 73797 | 34114036 | 58666982 | 2870342 | 10294949 | 4400492 | 205025 | 24009307 | 3277793 | 18721523 | 50506 | 7452157 | 924004 | 1833859 | 55550190 | 66674518 | 41008050 |  | Unknown
AB-220715-01 QCS 2 | 39795487 | 128633 | 179751 | 268643 | 44786197 | 1883348 | 2188869 | 17882 | 63492 | 474535 | 609328 | 1938473 | 1327902 | 1748193 | 121508 | 175128 | 266751 | 604229 | 1333423 | 7234142 | 3117107 | 20796 | 25271347 | 4232125 | 12800782 | 1829471 | 16710888 | 62716 | 34739142 | 29282534 | 360443 | 2205882 | 39271 | 38025 | 49099 | 22625606 | 35249778 | 55636 | 7078152 | 2865614 | 146614 | 15721824 | 48501 | 11856062 | 22248 | 99151 | 655909 | 1223248 | 1308224 | 1735494 | 27986677 |  | Unknown

## Sheet4
Batch Folder Path: | D:\TraceFinderData\Projects\PFAS Online SPE\2022-07-15\2022-07-15-AB2
Batch Name: | 2022-07-15-AB2
Last Acquisition Date: | 2022-07-26 04:21
User Name: | Quantiva
Sample Count: | 342
Area - Zero
Data File | 11-Cl-PF3OUDS | 4-2:FTS | 6-2:FTS | 8-2:FTS | 9-Cl-PF3ONS | D3-N-Et-FOSA | D3-N-Me-FOSA | FBSA | FHxA | HFPO-DA | L-PFBS | L-PFHpS | L-PFHxS | L-PFOS | M2-4:2-FTS | M2-6:2-FTS | M2-8:2-FTS | M3PFBS | M3PFHxS | M4PFHpA | M5PFHxA | M5PFPeA | M6PFDA | M8FOSA | M8PFOA | M8PFOS | M9PFNA | MPFBA | MPFDoA | MPFUDA | MPHFPO-DA | NaDONA | N-Et-FOSA | N-Me-FOSA | PFBA | PFDA | PFDoA | PFDS | PFHpA | PFHxA | PFHxDA | PFNA | PFNS | PFOA | PFODA | PFOSA | PFPeA | PFPeS | PFTeDA | PFTrDA | PFUdA | All Flags | Sample Type
AB-220715-01 Cal Std 1 1 | 95 | 6213 | 2763 | 915 | 524122 | 0 | 7075600 | 7 | 289 | 0 | 6554 | 26553 | 17730 | 26106 | 382817 | 1259222 | 1174724 | 1500040 | 3726379 | 10295093 | 4123113 | 977043 | 40182026 | 8565344 | 22704030 | 4705476 | 27045316 | 61078 | 0 | 0 | 480815 | 0 | 0 | 25014 | 0 | 190897 | 0 | 0 | 44377 | 0 | 536 | 115171 | 27576 | 103420 | 262 | 40019 | 0 | 0 | 0 | 0 | 185918 |  | Cal Std
AB-220715-01 Cal Std 2 1 | 52 | 4913 | 2295 | 180 | 928988 | 0 | 5829551 | 33 | 113 | 7460 | 8915 | 44888 | 24764 | 37728 | 311496 | 1025592 | 1028027 | 1148968 | 2883380 | 8119618 | 3078995 | 667163 | 33735557 | 7151291 | 17285957 | 4017638 | 22129693 | 35290 | 0 | 0 | 340592 | 0 | 937 | 43990 | 0 | 313576 | 0 | 0 | 73448 | 32497 | 369 | 186714 | 49467 | 155101 | 591 | 67205 | 0 | 0 | 866 | 0 | 314726 |  | Cal Std
AB-220715-01 Cal Std 3 1 | 173 | 8737 | 2289 | 19 | 2121363 | 0 | 6448212 | 14 | 20 | 12535 | 25055 | 92775 | 59911 | 78842 | 368482 | 1173468 | 1208706 | 1320162 | 3410846 | 9092210 | 3674530 | 770168 | 39237316 | 8420819 | 20100191 | 4518567 | 25080673 | 61302 | 0 | 0 | 406880 | 0 | 0 | 92908 | 0 | 706805 | 0 | 0 | 162175 | 68123 | 392 | 411787 | 103173 | 351979 | 175 | 146648 | 18605 | 0 | 1367 | 0 | 713012 |  | Cal Std
AB-220715-01 Cal Std 4 1 | 25 | 19691 | 2321 | 229 | 5664493 | 0 | 6586324 | 38 | 239 | 22962 | 63914 | 247243 | 144987 | 204672 | 359043 | 1122525 | 1171502 | 1275747 | 3262781 | 9022223 | 3599898 | 761209 | 40566614 | 8618929 | 19694899 | 4610237 | 25586736 | 73743 | 0 | 0 | 433529 | 0 | 848 | 236908 | 0 | 1787920 | 0 | 102 | 404183 | 153251 | 614 | 1080476 | 254743 | 863277 | 195 | 398101 | 44231 | 0 | 1741 | 0 | 1921138 |  | Cal Std
AB-220715-01 Cal Std 5 1 | 124 | 35656 | 1936 | 393 | 10060677 | 0 | 6104773 | 6 | 107 | 54881 | 112967 | 468870 | 266562 | 377242 | 327106 | 993803 | 1047581 | 1277711 | 3040930 | 8449225 | 3438958 | 734516 | 36873395 | 8121496 | 18617801 | 4289270 | 23616240 | 61896 | 0 | 0 | 394758 | 0 | 1252 | 461955 | 0 | 3312805 | 0 | 44 | 745073 | 281629 | 182 | 2019699 | 491260 | 1642665 | 77 | 765304 | 66523 | 0 | 0 | 0 | 3699819 |  | Cal Std
AB-220715-01 Cal Std 6 1 | 322 | 62234 | 1686 | 318 | 19645314 | 0 | 5632639 | 18 | 252 | 96138 | 212537 | 904913 | 513805 | 738357 | 298805 | 859662 | 962279 | 1138115 | 2895638 | 7656165 | 3150523 | 650813 | 35342251 | 7952548 | 17085621 | 4084789 | 22429940 | 63255 | 0 | 0 | 364862 | 0 | 0 | 860111 | 0 | 6400349 | 0 | 0 | 1393973 | 520773 | 301 | 3897545 | 964466 | 3106867 | 147 | 1539889 | 116325 | 0 | 0 | 0 | 7041520 |  | Cal Std
AB-220715-01 Cal Std 7 1 | 0 | 131304 | 1884 | 309 | 46244797 | 0 | 5809988 | 141 | 0 | 195537 | 487488 | 1930638 | 1052109 | 1669615 | 272558 | 622253 | 802721 | 1063643 | 2332530 | 6471670 | 2802723 | 603781 | 31298744 | 7466092 | 13707584 | 3718322 | 19642767 | 59957 | 0 | 0 | 334185 | 693 | 0 | 2223243 | 0 | 14103619 | 0 | 0 | 2962245 | 1208414 | 876 | 8463022 | 2235651 | 6247722 | 603 | 3686998 | 276406 | 0 | 0 | 0 | 15851584 |  | Cal Std
AB-220715-01 Cal Std 8 1 | 1139 | 287535 | 2229 | 334 | 89630124 | 0 | 5458564 | 94 | 1037 | 513219 | 1165016 | 3990860 | 2627119 | 3454143 | 288602 | 535035 | 719536 | 1192670 | 2633040 | 7809084 | 3471177 | 736491 | 29182620 | 7627492 | 15825420 | 3565013 | 20097918 | 67846 | 0 | 0 | 413789 | 0 | 0 | 4301897 | 61805 | 26678034 | 0 | 0 | 7405418 | 3137571 | 490 | 17825771 | 4079827 | 14536933 | 260 | 7677723 | 701702 | 0 | 0 | 0 | 29022193 |  | Cal Std
AB-220715-01 Cal Std 9 1 | 0 | 564951 | 2428 | 961 | 153196733 | 0 | 5079317 | 434 | 0 | 1082805 | 2336604 | 6579164 | 4733522 | 6303031 | 260099 | 415680 | 587510 | 1143888 | 2346729 | 6988392 | 3374919 | 716558 | 24297545 | 7107332 | 13421321 | 3123532 | 17013245 | 67795 | 0 | 0 | 396794 | 0 | 0 | 8129682 | 122135 | 43922722 | 0 | 206 | 13154062 | 6136372 | 1266 | 30839060 | 6794109 | 24395888 | 187 | 14342738 | 1387916 | 0 | 0 | 1686 | 47813294 |  | Cal Std
AB-220715-01 Cal Std 10 1 | 0 | 1069114 | 1682 | 929 | 272337689 | 0 | 4263423 | 508 | 0 | 2370062 | 4957446 | 10968170 | 8604258 | 11439165 | 211972 | 299937 | 464932 | 1004844 | 1657654 | 5374532 | 2835104 | 687574 | 17020941 | 6016468 | 9637873 | 2277056 | 12925865 | 72321 | 0 | 0 | 365038 | 2730573 | 0 | 16081243 | 284315 | 73814926 | 0 | 126 | 24027588 | 11738574 | 699 | 53333890 | 11451789 | 42028054 | 592 | 29148262 | 3179868 | 0 | 0 | 0 | 77629549 |  | Cal Std
AB-220715-01 Cal Std 11 1 | 0 | 1730478 | 2047 | 1553 | 399561750 | 0 | 3354460 | 0 | 11812 | 4122460 | 7985988 | 15608779 | 12929320 | 17660982 | 159048 | 243138 | 372285 | 763538 | 1116701 | 3718885 | 2281360 | 605950 | 11373297 | 4704493 | 6720492 | 1590397 | 8862935 | 65082 | 0 | 0 | 318901 | 5733955 | 0 | 26077756 | 560376 | 102250983 | 0 | 346 | 33821346 | 20092953 | 388 | 77032869 | 15636905 | 60007246 | 60 | 46882676 | 5845069 | 0 | 0 | 0 | 106223607 |  | Cal Std
AB-220715-01 Cal Std 12 1 | 0 | 2774998 | 4401 | 1395 | 555868075 | 0 | 2597638 | 0 | 40248 | 6442018 | 13129914 | 22210774 | 19072499 | 26960113 | 127641 | 205008 | 355760 | 580441 | 694484 | 2448830 | 1574116 | 497373 | 7246388 | 3762283 | 4250243 | 1079609 | 5789422 | 69192 | 0 | 0 | 237441 | 8218494 | 0 | 41856400 | 1225785 | 134355522 | 0 | 698 | 47210521 | 31244094 | 786 | 105607311 | 22074355 | 79095158 | 89 | 77083572 | 10081286 | 0 | 0 | 0 | 141216584 |  | Cal Std
AB-220715-01 Cal Std 13 1 | 0 | 4633637 | 8065 | 2020 | 776131967 | 0 | 1701861 | 0 | 10161500 | 9556766 | 19903793 | 31089706 | 26533890 | 39921460 | 84930 | 148938 | 284756 | 356739 | 387630 | 1224273 | 906084 | 314637 | 3649764 | 2432014 | 2058190 | 637062 | 3012684 | 54188 | 0 | 0 | 145507 | 19070234 | 0 | 64842405 | 2153714 | 175142790 | 0 | 941 | 60950502 | 44040422 | 271 | 136404305 | 30754577 | 99632547 | 36 | 121840134 | 16911237 | 0 | 0 | 0 | 188950269 |  | Cal Std
AB-220715-01 LFB-H 1 | 0 | 306023 | 752 | 1140 | 90979037 | 0 | 5024206 | 0 | 5821507 | 571321 | 1245558 | 3953981 | 2532598 | 3325492 | 297230 | 499166 | 638933 | 1258466 | 2609141 | 8250258 | 3578022 | 802267 | 28986281 | 7353031 | 15840003 | 3363767 | 20165202 | 72806 | 0 | 0 | 433139 | 2571376 | 0 | 3957377 | 64087 | 26690791 | 0 | 0 | 7649509 | 3387594 | 815 | 17815079 | 4012441 | 14589004 | 81 | 7584539 | 762027 | 0 | 0 | 0 | 29988087 |  | Unknown
AB-220715-01 LRB 1_INJ001 | 0 | 5883 | 1497 | 1008 | 490886 | 0 | 5235716 | 0 | 20670 | 0 | 0 | 15841 | 0 | 15616 | 350313 | 951039 | 935148 | 1346465 | 3154476 | 9585402 | 3827595 | 835320 | 37971765 | 8062435 | 19792477 | 4180363 | 24084153 | 56360 | 0 | 0 | 442707 | 0 | 0 | 21739 | 0 | 140374 | 0 | 0 | 20321 | 19069 | 631 | 73099 | 21021 | 99940 | 269 | 58030 | 7459 | 0 | 0 | 0 | 164260 |  | Unknown
AB-220715-01 LRB 1 | 0 | 5717 | 1492 | 1063 | 352036 | 0 | 4304348 | 0 | 7532 | 0 | 0 | 20944 | 0 | 18098 | 248061 | 833906 | 816597 | 889053 | 2553652 | 7187155 | 2587113 | 603704 | 31479184 | 6387473 | 15886293 | 3533089 | 19918783 | 30308 | 0 | 0 | 286866 | 8110 | 0 | 16392 | 0 | 135860 | 0 | 0 | 19860 | 29221 | 799 | 64940 | 17867 | 77820 | 733 | 36372 | 7872 | 0 | 0 | 0 | 135821 |  | Unknown
AB-220715-01 CCC-H 1 | 0 | 540432 | 0 | 1557 | 150533396 | 0 | 4418553 | 0 | 10215677 | 1072099 | 2238402 | 6389865 | 4414626 | 5777483 | 252322 | 373276 | 513966 | 1075641 | 2144318 | 6768580 | 3215096 | 702667 | 23483308 | 6577757 | 12833415 | 2836570 | 16388396 | 62643 | 0 | 0 | 403694 | 4442046 | 0 | 7074440 | 111012 | 43877030 | 0 | 820 | 13296279 | 5904108 | 649 | 29538359 | 6527615 | 23712487 | 321 | 13527594 | 1387870 | 0 | 1389 | 0 | 47427269 |  | Unknown
AB-220715-01 IBC 1 | 0 | 288619 | 1218 | 439 | 89307148 | 0 | 5246563 | 0 | 62883 | 566948 | 1249865 | 3888308 | 2638845 | 3432894 | 289657 | 473205 | 731993 | 1220172 | 2572751 | 8039672 | 3566735 | 766663 | 31908879 | 7580937 | 15612554 | 3393225 | 20318022 | 68645 | 0 | 0 | 434362 | 2610301 | 0 | 57196 | 62437 | 29237374 | 0 | 0 | 7874973 | 3269057 | 593 | 18440552 | 88210 | 14431841 | 548 | 94678 | 731729 | 0 | 0 | 0 | 34215051 |  | Unknown
AB-220715-01 QCS 1 | 0 | 5436 | 2697 | 451 | 1263554 | 0 | 5520924 | 0 | 0 | 430535 | 1260884 | 3925510 | 2714381 | 3432100 | 329456 | 526558 | 725762 | 1209903 | 2669141 | 8244516 | 3427631 | 757358 | 29652036 | 7715934 | 15921519 | 3375356 | 20459537 | 66107 | 0 | 0 | 424097 | 24623 | 0 | 23853 | 56967 | 26572877 | 0 | 0 | 7856533 | 3051759 | 507 | 18195884 | 4185892 | 14548757 | 414 | 58947 | 731269 | 0 | 0 | 0 | 31256474 |  | Unknown
WO-22270-01-A-1 | 0 | 10876 | 1062 | 639 | 80893 | 0 | 4669195 | 459 | 0 | 0 | 0 | 8640 | 0 | 7119 | 78887 | 413775 | 1018041 | 192834 | 908010 | 2494439 | 647220 | 63563 | 21056090 | 4971053 | 5930160 | 2506050 | 10683728 | 8248 | 0 | 0 | 67698 | 0 | 0 | 15106 | 0 | 426580 | 0 | 1909 | 0 | 0 | 1239 | 50034 | 5451 | 60896 | 853 | 9747 | 0 | 0 | 0 | 0 | 33126 |  | Unknown
WO-22270-02-A-1 | 0 | 7662 | 1077 | 414 | 234828 | 0 | 5247807 | 496 | 0 | 0 | 0 | 6113 | 0 | 9273 | 94003 | 449164 | 1190668 | 236621 | 1029128 | 2874284 | 786145 | 84736 | 23902714 | 6119673 | 6874699 | 2904969 | 12208610 | 9251 | 0 | 0 | 83138 | 0 | 0 | 12905 | 0 | 439266 | 0 | 1337 | 0 | 0 | 303 | 30694 | 3632 | 76246 | 1114 | 8672 | 0 | 0 | 0 | 0 | 33909 |  | Unknown
WO-22270-03-A-1 | 0 | 16616 | 229 | 333 | 201594 | 0 | 5593529 | 363 | 0 | 0 | 0 | 3310 | 0 | 6154 | 89055 | 440861 | 1113762 | 238397 | 1021694 | 2845686 | 797719 | 87570 | 24275846 | 5818673 | 6865251 | 2919310 | 12613090 | 10607 | 0 | 0 | 81746 | 0 | 0 | 14538 | 0 | 419329 | 0 | 1742 | 0 | 0 | 1158 | 27931 | 3843 | 35972 | 756 | 7450 | 0 | 0 | 0 | 0 | 34882 |  | Unknown
WO-22270-04-A-1 | 0 | 11347 | 1274 | 417 | 59421 | 0 | 5108541 | 338 | 1200 | 0 | 0 | 0 | 0 | 7378 | 81390 | 411752 | 1036812 | 213958 | 971558 | 2661510 | 700096 | 71859 | 22475000 | 5665387 | 6380733 | 2739134 | 11465262 | 8158 | 0 | 0 | 78456 | 0 | 0 | 14253 | 0 | 396647 | 0 | 1298 | 0 | 0 | 1355 | 31531 | 4409 | 38856 | 53 | 5810 | 0 | 0 | 748 | 0 | 26437 |  | Unknown
WO-22270-05-A-1 | 0 | 15709 | 444 | 268 | 69855 | 0 | 5293106 | 407 | 0 | 0 | 0 | 0 | 0 | 9272 | 96706 | 419242 | 1044860 | 256964 | 1055212 | 2888134 | 832899 | 88943 | 24352740 | 6533380 | 6786453 | 2931235 | 12838486 | 9053 | 0 | 0 | 89911 | 0 | 0 | 13199 | 0 | 514693 | 0 | 1364 | 15733 | 0 | 473 | 25115 | 17247 | 90859 | 141 | 5437 | 0 | 0 | 0 | 0 | 32769 |  | Unknown
WO-22270-06-A-1 | 0 | 16683 | 634 | 222 | 130221 | 0 | 5005474 | 326 | 0 | 2086 | 0 | 3390 | 0 | 9636 | 82398 | 391009 | 970489 | 217242 | 979994 | 2656633 | 709091 | 80012 | 23800787 | 6200326 | 6504078 | 2718954 | 12100289 | 10949 | 0 | 0 | 73672 | 0 | 0 | 14258 | 0 | 411032 | 0 | 1883 | 0 | 0 | 796 | 45206 | 2197 | 88619 | 428 | 5388 | 0 | 0 | 1042 | 1145 | 23522 |  | Unknown
WO-22270-07-A-1 | 0 | 10525 | 666 | 544 | 35374 | 0 | 5229662 | 232 | 1578 | 0 | 0 | 0 | 0 | 3377 | 85099 | 392957 | 972787 | 226617 | 997478 | 2815174 | 774955 | 84223 | 23786530 | 5672081 | 6580585 | 2727936 | 12143878 | 8642 | 0 | 0 | 85596 | 0 | 0 | 14424 | 0 | 408982 | 0 | 1358 | 0 | 0 | 1267 | 14597 | 2391 | 63181 | 431 | 3820 | 0 | 0 | 579 | 0 | 24166 |  | Unknown
WO-22270-08-A-1 | 0 | 9358 | 669 | 209 | 109032 | 0 | 4550828 | 404 | 899 | 0 | 0 | 4195 | 0 | 4473 | 66226 | 350324 | 924377 | 174942 | 860282 | 2259842 | 574949 | 60247 | 20956659 | 4674506 | 5709385 | 2416625 | 10567001 | 6088 | 0 | 0 | 64185 | 0 | 0 | 9200 | 0 | 413652 | 0 | 918 | 0 | 0 | 649 | 28784 | 13801 | 69581 | 208 | 3813 | 0 | 0 | 1167 | 0 | 17907 |  | Unknown
WO-22270-09-A-1 | 0 | 13998 | 604 | 210 | 49896 | 1762 | 5580216 | 128 | 422 | 0 | 0 | 3354 | 0 | 28643 | 93119 | 412521 | 1084480 | 252013 | 1014121 | 2867767 | 841657 | 88334 | 25413825 | 5993878 | 6892188 | 2912696 | 12991287 | 8956 | 0 | 0 | 91451 | 0 | 0 | 13129 | 0 | 594363 | 0 | 1650 | 0 | 0 | 812 | 44806 | 11091 | 49396 | 161 | 5664 | 0 | 0 | 1001 | 0 | 20985 |  | Unknown
WO-22270-10-A-1 | 0 | 13130 | 552 | 574 | 25908 | 0 | 5359057 | 153 | 823 | 0 | 0 | 3297 | 0 | 5212 | 84521 | 421836 | 1099498 | 241047 | 1015479 | 2731655 | 763305 | 76864 | 24903677 | 5851126 | 6737953 | 2862004 | 12581059 | 7559 | 0 | 0 | 84932 | 0 | 933 | 12582 | 0 | 714121 | 0 | 1134 | 0 | 0 | 655 | 31778 | 10246 | 81511 | 894 | 3089 | 0 | 0 | 0 | 0 | 19629 |  | Unknown
AB-220715-01 LRB 2_INJ001 | 0 | 7792 | 872 | 381 | 116350 | 0 | 4845215 | 40 | 1730 | 0 | 0 | 2158 | 2252 | 7052 | 289916 | 924230 | 919673 | 1198731 | 2987314 | 8874445 | 3489215 | 680960 | 36852981 | 7601214 | 18792228 | 4077662 | 23446218 | 55088 | 860 | 0 | 399197 | 0 | 509 | 14823 | 0 | 42664 | 0 | 196 | 0 | 0 | 929 | 49176 | 5913 | 116728 | 322 | 3564 | 7090 | 0 | 1010 | 0 | 63295 |  | Unknown
AB-220715-01 LRB 2 | 0 | 10681 | 1446 | 49 | 116186 | 0 | 4899378 | 57 | 1401 | 0 | 0 | 2066 | 0 | 9981 | 321096 | 834721 | 858736 | 1203801 | 2966182 | 9035214 | 3479659 | 716768 | 36367310 | 7678793 | 18880621 | 3931466 | 23536400 | 63071 | 0 | 0 | 408133 | 0 | 0 | 14942 | 0 | 29212 | 0 | 87 | 0 | 0 | 811 | 59295 | 1830 | 114871 | 539 | 5763 | 8880 | 0 | 2366 | 0 | 61780 |  | Unknown
WO-22270-11-A-1 | 0 | 8293 | 224 | 397 | 89285 | 0 | 5345726 | 33 | 0 | 0 | 0 | 7019 | 0 | 23095 | 89283 | 389337 | 1037725 | 243227 | 973170 | 2826319 | 752337 | 78960 | 24750178 | 5748661 | 6499004 | 2756593 | 12415592 | 7880 | 0 | 1106 | 86169 | 1425 | 1294 | 14976 | 0 | 502631 | 0 | 1522 | 0 | 0 | 215 | 13437 | 9939 | 25023 | 163 | 3392 | 0 | 0 | 0 | 717 | 23088 |  | Unknown
WO-22270-12-A-1 | 0 | 13090 | 498 | 337 | 18192 | 0 | 5119672 | 153 | 943 | 0 | 0 | 2187 | 0 | 4687 | 80508 | 358982 | 1038679 | 224022 | 943192 | 2643229 | 751211 | 78407 | 23411615 | 5545342 | 6244197 | 2676764 | 11980648 | 14360 | 528 | 0 | 78773 | 0 | 1087 | 12745 | 0 | 312938 | 0 | 489 | 8140 | 0 | 1094 | 11503 | 1351 | 65549 | 267 | 3340 | 0 | 0 | 0 | 0 | 18823 |  | Unknown
WO-22270-13-A-1 | 0 | 8954 | 610 | 415 | 61257 | 0 | 4780860 | 123 | 570 | 881 | 0 | 2440 | 245 | 3344 | 69047 | 348247 | 913652 | 163457 | 806418 | 2210712 | 552437 | 58929 | 20831902 | 4775930 | 5559627 | 2342145 | 10373051 | 11567 | 0 | 0 | 60443 | 0 | 779 | 15684 | 0 | 390873 | 0 | 1044 | 0 | 0 | 868 | 12824 | 7422 | 0 | 131 | 4218 | 0 | 1370 | 633 | 0 | 17660 |  | Unknown
WO-22270-14-A-1 | 0 | 7551 | 445 | 65 | 11869 | 711 | 5157693 | 191 | 654 | 0 | 0 | 9344 | 0 | 3523 | 86344 | 353793 | 921153 | 224289 | 957616 | 2691158 | 756615 | 83915 | 24688316 | 5599538 | 6661242 | 2682247 | 12409492 | 9159 | 0 | 0 | 86858 | 0 | 642 | 18848 | 0 | 453618 | 1177 | 1671 | 0 | 0 | 569 | 7285 | 8209 | 0 | 63 | 3337 | 0 | 0 | 1432 | 0 | 15441 |  | Unknown
WO-22270-15-A-1 | 0 | 10570 | 611 | 157 | 26733 | 0 | 4918455 | 123 | 739 | 0 | 0 | 4385 | 0 | 21942 | 75145 | 391737 | 1082151 | 215373 | 945971 | 2624491 | 722193 | 75596 | 23345937 | 5491737 | 6359552 | 2761557 | 12095536 | 7200 | 2104 | 735 | 77998 | 680 | 975 | 18252 | 0 | 247490 | 0 | 1289 | 0 | 0 | 412 | 28213 | 8623 | 78161 | 76 | 3445 | 0 | 0 | 0 | 0 | 19585 |  | Unknown
WO-22270-16-A-1 | 0 | 15600 | 804 | 217 | 18424 | 1813 | 5130370 | 213 | 1325 | 0 | 0 | 4153 | 0 | 5732 | 88173 | 369409 | 1013619 | 253091 | 1007387 | 2898935 | 866499 | 95847 | 25899008 | 5731327 | 6896612 | 2856034 | 12993054 | 9043 | 0 | 0 | 93335 | 0 | 977 | 16026 | 0 | 538801 | 2027 | 2341 | 0 | 0 | 1087 | 25661 | 10556 | 67306 | 425 | 3098 | 0 | 0 | 0 | 0 | 13617 |  | Unknown
WO-22270-17-A-1 | 0 | 9187 | 369 | 154 | 33247 | 0 | 4454366 | 25 | 594 | 0 | 0 | 3686 | 0 | 3024 | 57500 | 291558 | 746674 | 212397 | 874969 | 2515303 | 702222 | 73140 | 22950166 | 5062689 | 6171068 | 2464575 | 11191368 | 8750 | 1027 | 0 | 70790 | 1625 | 1065 | 14121 | 0 | 466056 | 0 | 1867 | 7328 | 0 | 354 | 33885 | 1042 | 102402 | 362 | 2266 | 0 | 0 | 0 | 0 | 15425 |  | Unknown
WO-22270-18-A-1 | 0 | 8469 | 626 | 257 | 16882 | 0 | 4531250 | 43 | 614 | 0 | 0 | 7784 | 2915 | 2998 | 47793 | 262827 | 722196 | 209350 | 862988 | 2453817 | 684538 | 76786 | 23325324 | 5075605 | 5952927 | 2473151 | 11429261 | 7583 | 0 | 0 | 71362 | 2196 | 0 | 19187 | 0 | 181072 | 0 | 1487 | 0 | 0 | 1066 | 7863 | 9075 | 31050 | 612 | 2670 | 0 | 0 | 614 | 0 | 14008 |  | Unknown
WO-22270-19-A-1 | 0 | 7755 | 454 | 58 | 21264 | 0 | 4783308 | 174 | 631 | 1063 | 0 | 6718 | 0 | 3632 | 54380 | 279775 | 818919 | 215765 | 877577 | 2635070 | 726015 | 83001 | 24026812 | 5309774 | 6127998 | 2551859 | 11774606 | 7571 | 1177 | 0 | 72915 | 0 | 662 | 14857 | 0 | 364715 | 0 | 1216 | 0 | 0 | 1737 | 12722 | 9538 | 44323 | 674 | 2721 | 0 | 0 | 0 | 847 | 16054 |  | Unknown
WO-22270-20-A-1 | 0 | 14231 | 102 | 55 | 18679 | 0 | 4461897 | 21 | 253 | 1630 | 0 | 6807 | 0 | 4907 | 48444 | 262537 | 778530 | 187792 | 807537 | 2560391 | 709476 | 78552 | 22800078 | 4996141 | 6092727 | 2279685 | 11580470 | 9344 | 0 | 0 | 75300 | 1543 | 348 | 12995 | 0 | 524991 | 0 | 1597 | 0 | 0 | 1492 | 10690 | 8208 | 61642 | 486 | 3346 | 0 | 984 | 0 | 0 | 13574 |  | Unknown
AB-220715-01 LFB-H 2 | 896 | 252043 | 245983 | 758 | 79704811 | 966 | 4443096 | 1504 | 5713070 | 509179 | 1042375 | 3606648 | 2316370 | 2958261 | 250243 | 440030 | 573433 | 1098509 | 2424591 | 7886486 | 3271639 | 712883 | 28517778 | 6848818 | 15597174 | 3134823 | 20010453 | 64930 | 0 | 0 | 394898 | 2406724 | 579 | 3538610 | 50530 | 25872300 | 0 | 41 | 7260851 | 2912201 | 356 | 17230755 | 3689535 | 14002239 | 883 | 6864001 | 643839 | 0 | 985 | 778 | 29871586 |  | Unknown
AB-220715-01 LRB 3_INJ001 | 0 | 13673 | 1821 | 578 | 130012 | 1555 | 4899041 | 199 | 7686 | 0 | 0 | 4746 | 3717 | 9276 | 313332 | 833076 | 836168 | 1180945 | 2944018 | 9321519 | 3688857 | 769517 | 38568092 | 7941740 | 20033002 | 4024004 | 24842549 | 73454 | 0 | 0 | 413998 | 0 | 697 | 14953 | 0 | 58750 | 0 | 303 | 0 | 5896 | 1561 | 27444 | 7128 | 81119 | 206 | 19873 | 6066 | 0 | 1069 | 0 | 60331 |  | Unknown
AB-220715-01 LRB 3 | 0 | 8076 | 1232 | 384 | 36846 | 1573 | 4314636 | 314 | 2716 | 0 | 0 | 7580 | 1320 | 4748 | 295515 | 777538 | 771788 | 1175836 | 2801679 | 8854880 | 3571226 | 739760 | 34668358 | 7252767 | 18187751 | 3696160 | 22675639 | 36802 | 1242 | 0 | 416397 | 0 | 1166 | 14368 | 0 | 33155 | 0 | 300 | 0 | 4629 | 796 | 21535 | 10534 | 71098 | 277 | 5269 | 4284 | 611 | 1173 | 0 | 39352 |  | Unknown
AB-220715-01 LD 1 | 1822 | 12574 | 641 | 294 | 18370 | 0 | 4562558 | 168 | 864 | 1525 | 0 | 3294 | 5362 | 3262 | 60389 | 284936 | 795556 | 229663 | 936028 | 2787213 | 798127 | 88657 | 24291523 | 5436427 | 6343479 | 2637392 | 12213570 | 9334 | 0 | 0 | 89095 | 0 | 652 | 15361 | 0 | 368248 | 0 | 1922 | 7588 | 0 | 1552 | 20805 | 4112 | 77477 | 471 | 3844 | 0 | 446 | 925 | 0 | 16797 |  | Unknown
AB-220715-01 CCC-H 2 | 0 | 467853 | 604647 | 14065 | 127515851 | 1591 | 3998196 | 1997433 | 9777030 | 972544 | 1925808 | 5482636 | 3910010 | 5275461 | 222585 | 309806 | 449591 | 936774 | 1927568 | 6825645 | 3073192 | 632809 | 22522730 | 6363199 | 12833118 | 2652217 | 16922867 | 50312 | 436 | 0 | 368167 | 4254599 | 588 | 6219407 | 81297 | 40824483 | 0 | 103 | 12664272 | 5568877 | 1336 | 29431341 | 5608014 | 22916200 | 504 | 12394320 | 1221444 | 875 | 1869 | 0 | 47313622 |  | Unknown
AB-220715-01 LFSM-H 1 | 0 | 55498 | 186178 | 19779 | 55437628 | 943 | 3964591 | 238038 | 2037615 | 102854 | 209664 | 1530943 | 783331 | 1975351 | 50831 | 186872 | 499014 | 216376 | 821320 | 2530217 | 709783 | 84115 | 20039375 | 4783404 | 5718369 | 2247881 | 10492766 | 11665 | 0 | 0 | 82309 | 774935 | 434 | 3080721 | 0 | 18488225 | 4044 | 1309 | 2384371 | 656296 | 976 | 9162237 | 3057498 | 5229081 | 161 | 4879962 | 76786 | 1065 | 0 | 0 | 23144799 |  | Unknown
WO-22270-21-A-1 | 0 | 7756 | 621 | 425 | 94821 | 0 | 4426569 | 388 | 3352 | 1582 | 0 | 9139 | 0 | 5813 | 60704 | 284753 | 800756 | 195198 | 830131 | 2504947 | 654006 | 63196 | 21910933 | 4706025 | 6014678 | 2358870 | 10956068 | 6573 | 0 | 0 | 74382 | 2480 | 1059 | 12609 | 0 | 368962 | 0 | 415 | 13649 | 0 | 556 | 20296 | 6359 | 116215 | 412 | 12785 | 0 | 0 | 0 | 1459 | 51124 |  | Unknown
WO-22270-22-A-1 | 0 | 11559 | 565 | 121 | 43801 | 0 | 4818465 | 96 | 1163 | 0 | 0 | 7992 | 0 | 2817 | 63535 | 313244 | 830640 | 232969 | 951987 | 2782723 | 782289 | 76751 | 24552358 | 5243368 | 6548169 | 2605651 | 12533477 | 10036 | 0 | 0 | 85811 | 0 | 543 | 13336 | 0 | 346931 | 0 | 1911 | 0 | 0 | 1486 | 20623 | 4005 | 48181 | 525 | 7928 | 0 | 0 | 0 | 0 | 22030 |  | Unknown
WO-22270-23-A-1 | 0 | 14379 | 345 | 878 | 28925 | 0 | 4403653 | 132 | 1688 | 0 | 0 | 7117 | 0 | 4054 | 66999 | 309178 | 787061 | 228669 | 951131 | 2828235 | 774613 | 88796 | 23873966 | 5287515 | 6553624 | 2536246 | 12070045 | 10071 | 0 | 0 | 83242 | 0 | 832 | 14023 | 0 | 432762 | 0 | 1704 | 6310 | 4808 | 2978 | 21653 | 8951 | 94240 | 542 | 3833 | 0 | 0 | 0 | 0 | 20387 |  | Unknown
WO-22270-24-A-1 | 0 | 11612 | 782 | 43 | 23402 | 0 | 4397284 | 73 | 1119 | 0 | 0 | 7938 | 0 | 5494 | 67265 | 290184 | 793912 | 205118 | 864524 | 2532880 | 697914 | 75091 | 22565786 | 4732208 | 6063512 | 2380645 | 11585880 | 8155 | 0 | 0 | 75539 | 0 | 524 | 13513 | 0 | 333513 | 3414 | 2587 | 0 | 0 | 1872 | 17721 | 5086 | 56769 | 836 | 2538 | 0 | 0 | 795 | 0 | 20015 |  | Unknown
WO-22270-25-A-1 | 0 | 9891 | 264 | 258 | 21921 | 840 | 5175826 | 55 | 971 | 0 | 0 | 3384 | 0 | 17479 | 75905 | 344207 | 967060 | 198886 | 926178 | 2667921 | 762129 | 70567 | 24376114 | 5252399 | 6436247 | 2578343 | 12191895 | 8796 | 0 | 0 | 80841 | 1289 | 932 | 12786 | 0 | 479098 | 0 | 1643 | 0 | 0 | 163 | 17036 | 4193 | 0 | 426 | 4088 | 0 | 692 | 655 | 2311 | 20260 |  | Unknown
WO-22270-26-A-1 | 0 | 12563 | 804 | 155 | 0 | 889 | 4240706 | 101 | 663 | 0 | 0 | 4028 | 0 | 24770 | 60302 | 287378 | 774001 | 162614 | 803516 | 2348661 | 581597 | 60936 | 21456251 | 4630088 | 5598039 | 2253364 | 10809945 | 7011 | 0 | 0 | 63660 | 0 | 699 | 16966 | 0 | 360105 | 5507 | 1449 | 7112 | 0 | 1516 | 18358 | 8807 | 63768 | 350 | 2050 | 0 | 578 | 0 | 0 | 13669 |  | Unknown
WO-22270-27-A-1 | 0 | 10830 | 701 | 104 | 13108 | 0 | 4605715 | 100 | 314 | 0 | 0 | 6301 | 0 | 4677 | 82304 | 324466 | 851466 | 212558 | 871868 | 2738199 | 780466 | 78593 | 24393292 | 5288119 | 6437946 | 2467084 | 12334450 | 10699 | 0 | 0 | 88315 | 1968 | 1116 | 16220 | 0 | 356538 | 5273 | 1763 | 7639 | 0 | 955 | 12699 | 5702 | 36875 | 42 | 3624 | 0 | 0 | 883 | 0 | 13913 |  | Unknown
WO-22270-28-A-1 | 0 | 10457 | 524 | 99 | 9957 | 0 | 5043883 | 20 | 414 | 0 | 0 | 3150 | 0 | 3969 | 74707 | 337992 | 984642 | 178048 | 868396 | 2574578 | 711838 | 64290 | 23251013 | 5357991 | 6198911 | 2520075 | 11956537 | 7360 | 612 | 0 | 77701 | 0 | 302 | 15171 | 0 | 474742 | 3959 | 2054 | 0 | 0 | 904 | 12326 | 969 | 0 | 220 | 2026 | 0 | 0 | 1066 | 2137 | 13004 |  | Unknown
WO-22270-29-A-1 | 0 | 12872 | 556 | 834 | 12842 | 0 | 6084251 | 187 | 506 | 0 | 0 | 3253 | 0 | 3277 | 83419 | 389494 | 1502784 | 220808 | 958138 | 2927337 | 844415 | 79396 | 25150085 | 5750699 | 6971406 | 2761177 | 13663824 | 9079 | 0 | 0 | 89823 | 0 | 754 | 12068 | 0 | 524541 | 3807 | 618 | 0 | 0 | 569 | 18288 | 5094 | 17113 | 351 | 2913 | 0 | 933 | 0 | 0 | 15681 |  | Unknown
WO-22270-30-A-1 | 0 | 12519 | 485 | 578 | 9564 | 0 | 5138125 | 17 | 323 | 1867 | 0 | 5059 | 0 | 19147 | 87621 | 392561 | 1192287 | 204675 | 918954 | 2781792 | 768432 | 73479 | 23382076 | 5713404 | 6664518 | 2658465 | 12720413 | 8079 | 711 | 324 | 83667 | 0 | 428 | 12627 | 0 | 426600 | 0 | 1882 | 11901 | 4216 | 1328 | 33975 | 4366 | 74997 | 175 | 1832 | 0 | 914 | 701 | 0 | 16340 |  | Unknown
AB-220715-01 LRB 4_INJ001 | 0 | 10895 | 1752 | 298 | 10929 | 0 | 4511804 | 13 | 259 | 0 | 0 | 2175 | 1307 | 20244 | 292870 | 775819 | 805956 | 1024715 | 2806334 | 8987422 | 3438556 | 618581 | 37059342 | 7521692 | 19118635 | 3894796 | 24146421 | 69085 | 608 | 0 | 389709 | 0 | 432 | 17977 | 0 | 23003 | 0 | 64 | 0 | 5329 | 922 | 38587 | 6892 | 78641 | 284 | 1408 | 4441 | 490 | 604 | 0 | 33367 |  | Unknown
AB-220715-01 LRB 4 | 0 | 12928 | 1156 | 280 | 12806 | 0 | 4103688 | 21 | 845 | 0 | 0 | 3381 | 980 | 4413 | 288647 | 766217 | 748056 | 1052684 | 2720683 | 8801455 | 3423926 | 625109 | 35065625 | 6775832 | 18050520 | 3495357 | 22767222 | 40255 | 478 | 0 | 414308 | 0 | 353 | 14157 | 0 | 22050 | 0 | 0 | 0 | 5627 | 811 | 16074 | 4214 | 65411 | 561 | 1194 | 5380 | 0 | 0 | 0 | 35713 |  | Unknown
WO-22270-31-A-1 | 0 | 17740 | 916 | 412 | 9709 | 1706 | 4530340 | 0 | 330 | 602 | 0 | 6884 | 0 | 3651 | 83273 | 313712 | 849440 | 214161 | 928348 | 2812408 | 809761 | 78127 | 24969566 | 5163807 | 6415707 | 2588435 | 12305829 | 8390 | 1154 | 1004 | 76706 | 1273 | 662 | 12344 | 0 | 327601 | 4336 | 1659 | 0 | 0 | 2393 | 44651 | 5015 | 40426 | 101 | 3216 | 0 | 553 | 1475 | 0 | 13481 |  | Unknown
WO-22270-32-A-1 | 0 | 8511 | 548 | 0 | 13906 | 790 | 3272033 | 6 | 118 | 0 | 0 | 1287 | 0 | 15285 | 39896 | 233516 | 616483 | 84477 | 572347 | 1490609 | 328226 | 27873 | 15464622 | 2917629 | 3991210 | 1568949 | 7756948 | 4741 | 875 | 0 | 36053 | 1472 | 621 | 9931 | 0 | 273842 | 0 | 1140 | 0 | 0 | 2665 | 14962 | 3843 | 28871 | 409 | 1416 | 0 | 331 | 665 | 1588 | 10723 |  | Unknown
WO-22270-33-A-1 | 1074 | 10015 | 561 | 550 | 8965 | 0 | 4280618 | 40 | 544 | 460 | 0 | 2555 | 0 | 3430 | 61573 | 273477 | 801845 | 140821 | 760076 | 2215063 | 565262 | 47704 | 20700734 | 4532416 | 5467538 | 2116313 | 10249490 | 7519 | 0 | 0 | 52877 | 1248 | 637 | 12670 | 0 | 297063 | 5248 | 1245 | 0 | 0 | 2296 | 8208 | 3137 | 52635 | 445 | 1879 | 0 | 234 | 0 | 0 | 10936 |  | Unknown
WO-22270-34-A-1 | 0 | 14094 | 1280 | 461 | 88736 | 0 | 4357754 | 0 | 157 | 0 | 0 | 10865 | 0 | 6554 | 51881 | 283283 | 937403 | 126717 | 692041 | 2007401 | 507631 | 40635 | 19171427 | 4009787 | 5012065 | 1973366 | 9767518 | 4927 | 0 | 0 | 54088 | 3472 | 737 | 12501 | 0 | 383528 | 5189 | 2090 | 13647 | 0 | 1569 | 33204 | 12105 | 66563 | 597 | 2207 | 0 | 229 | 1055 | 0 | 29760 |  | Unknown
WO-22270-35-A-1 | 0 | 10898 | 454 | 840 | 9144 | 918 | 4152130 | 9 | 18 | 0 | 0 | 4961 | 0 | 15657 | 76175 | 311573 | 795677 | 176728 | 846287 | 2622592 | 725329 | 70253 | 23300048 | 5055111 | 6213151 | 2480709 | 11749688 | 7517 | 0 | 0 | 79629 | 3910 | 830 | 12379 | 0 | 302652 | 5544 | 2568 | 16655 | 0 | 1248 | 71218 | 5537 | 20044 | 536 | 2096 | 750 | 472 | 959 | 0 | 54433 |  | Unknown
WO-22270-36-A-1 | 0 | 9492 | 355 | 280 | 10104 | 654 | 5184774 | 51 | 373 | 448 | 0 | 3041 | 0 | 18991 | 97537 | 368302 | 948395 | 268439 | 1045386 | 3246348 | 1022167 | 89486 | 27583948 | 5987705 | 7413062 | 2946160 | 14258175 | 9213 | 641 | 0 | 105468 | 0 | 426 | 13406 | 0 | 441391 | 0 | 2795 | 0 | 0 | 1030 | 17637 | 5518 | 42763 | 193 | 2588 | 0 | 227 | 899 | 0 | 15427 |  | Unknown
WO-22270-37-A-1 | 0 | 10332 | 563 | 396 | 8872 | 0 | 4063981 | 0 | 38 | 0 | 401 | 2797 | 0 | 5351 | 47222 | 232043 | 680580 | 161519 | 787237 | 2434280 | 651708 | 62361 | 21995949 | 4596267 | 5639306 | 2273046 | 11143067 | 7128 | 0 | 0 | 73754 | 1053 | 508 | 11361 | 0 | 312906 | 1394 | 458 | 0 | 0 | 1523 | 10275 | 3223 | 29492 | 656 | 1358 | 0 | 204 | 1120 | 0 | 11783 |  | Unknown
WO-22270-38-A-1 | 0 | 9640 | 461 | 176 | 7387 | 1443 | 3944586 | 7 | 127 | 0 | 795 | 3203 | 0 | 4662 | 44240 | 220292 | 665978 | 167345 | 789416 | 2453308 | 670769 | 59037 | 23508771 | 4810667 | 5871523 | 2368565 | 11099051 | 7073 | 2048 | 0 | 79210 | 1112 | 218 | 11781 | 0 | 235073 | 2134 | 1872 | 0 | 0 | 20007 | 11910 | 6624 | 40824 | 585 | 2417 | 0 | 160 | 0 | 0 | 10112 |  | Unknown
WO-22270-39-A-1 | 1655 | 10560 | 274 | 412 | 9472 | 750 | 4072945 | 12 | 350 | 0 | 0 | 3299 | 0 | 18152 | 42236 | 200878 | 662794 | 162958 | 774833 | 2416852 | 658063 | 64087 | 21946613 | 4608560 | 5599197 | 2214400 | 10961914 | 7799 | 3147 | 0 | 74250 | 1582 | 312 | 12309 | 0 | 245235 | 0 | 1437 | 0 | 0 | 2054 | 9257 | 3638 | 44465 | 2770 | 2254 | 0 | 650 | 0 | 0 | 10059 |  | Unknown
WO-22270-40-A-1 | 0 | 10673 | 99 | 171 | 5949 | 0 | 3709133 | 33 | 111 | 756 | 0 | 3111 | 0 | 4335 | 44639 | 202213 | 639394 | 158309 | 777921 | 2411900 | 641009 | 65390 | 22015251 | 4605304 | 5748344 | 2242114 | 10864387 | 7929 | 1044 | 0 | 71697 | 1287 | 147 | 10549 | 0 | 337269 | 0 | 1844 | 0 | 0 | 1581 | 57909 | 1994 | 13094 | 867 | 1880 | 0 | 175 | 1236 | 0 | 26301 |  | Unknown
AB-220715-01 LFB-H 3 | 422 | 240472 | 372235 | 471820 | 73421517 | 0 | 4016559 | 1336060 | 5302509 | 460723 | 866543 | 3210719 | 2134741 | 2895909 | 230664 | 369808 | 511786 | 875068 | 2211219 | 7663213 | 3161735 | 599800 | 28117518 | 6474161 | 14776874 | 2985029 | 19755156 | 58803 | 901 | 0 | 354933 | 2359615 | 214 | 3172879 | 45308 | 26206636 | 0 | 128 | 7206593 | 2870834 | 1520 | 17690084 | 3444990 | 13902298 | 701 | 6625200 | 551694 | 683 | 1863 | 0 | 30375905 |  | Unknown
AB-220715-01 LRB 5_INJ001 | 0 | 13357 | 558 | 973 | 118786 | 577 | 4455679 | 459 | 4726 | 1056 | 0 | 4603 | 12713 | 8094 | 281552 | 727887 | 751049 | 929593 | 2745820 | 9074474 | 3521842 | 633950 | 37502477 | 7360423 | 19081863 | 3774492 | 24113664 | 69278 | 433 | 0 | 409478 | 973 | 496 | 15467 | 0 | 52804 | 0 | 40 | 0 | 4546 | 936 | 22563 | 6570 | 43889 | 301 | 15998 | 7499 | 233 | 1021 | 0 | 53833 |  | Unknown
AB-220715-01 LRB 5 | 0 | 14108 | 1924 | 368 | 40026 | 0 | 3951425 | 233 | 1280 | 996 | 1036 | 1998 | 12933 | 19809 | 271440 | 690852 | 678063 | 931713 | 2582623 | 8614914 | 3314814 | 594412 | 33609010 | 6444273 | 17477041 | 3390735 | 21661304 | 41800 | 0 | 1039 | 370035 | 0 | 679 | 14415 | 0 | 37057 | 0 | 24 | 0 | 4856 | 546 | 16399 | 4003 | 48039 | 141 | 4633 | 7217 | 121 | 0 | 0 | 27916 |  | Unknown
AB-220715-01 LD 2 | 0 | 11852 | 492 | 535 | 11993 | 1027 | 3955903 | 40 | 911 | 0 | 537 | 3921 | 7081 | 4946 | 47321 | 239917 | 700116 | 164230 | 782533 | 2368011 | 637486 | 61350 | 22153123 | 4388743 | 5642949 | 2166130 | 10589898 | 7489 | 0 | 0 | 67770 | 2383 | 224 | 11420 | 0 | 248641 | 0 | 1786 | 11702 | 0 | 1870 | 38898 | 5194 | 104476 | 301 | 1885 | 0 | 310 | 1857 | 0 | 13803 |  | Unknown
AB-220715-01 CCC-H 3 | 1718 | 401257 | 546110 | 776055 | 118154697 | 1398 | 3691444 | 2721153 | 9427031 | 917700 | 1587692 | 5156356 | 3651431 | 4813582 | 190093 | 278290 | 394786 | 789700 | 1820963 | 6730117 | 2920353 | 551744 | 22653173 | 5760940 | 12414542 | 2451267 | 16675537 | 55846 | 0 | 1006 | 349251 | 4060807 | 300 | 5700770 | 89542 | 41381237 | 1110 | 84 | 12759033 | 5329685 | 230351 | 29292748 | 5272417 | 22885479 | 369 | 11720831 | 1030184 | 38296 | 734 | 0 | 48056493 |  | Unknown
AB-220715-01 LFSM-H 2 | 0 | 47675 | 157010 | 381884 | 47775354 | 0 | 3357558 | 282064 | 1842856 | 99018 | 166741 | 1285860 | 710040 | 1655044 | 43158 | 159953 | 388410 | 168355 | 752646 | 2447482 | 704358 | 65708 | 18025841 | 4180270 | 5276322 | 1847622 | 9470779 | 8264 | 1897 | 0 | 81085 | 744907 | 668 | 2673604 | 0 | 16596031 | 0 | 1870 | 2350666 | 614991 | 142709 | 8331626 | 2623067 | 4922539 | 724 | 4157203 | 61288 | 171 | 622 | 0 | 21911280 |  | Unknown
WO-22270-41-A-1 | 0 | 12565 | 459 | 1118 | 25476 | 520 | 4519159 | 88 | 1341 | 0 | 0 | 3072 | 0 | 4431 | 62234 | 261319 | 928886 | 150874 | 743759 | 2339634 | 601965 | 56194 | 21698381 | 4412153 | 5644774 | 2172604 | 11178511 | 8915 | 0 | 782 | 66208 | 0 | 482 | 13246 | 0 | 242436 | 0 | 1673 | 0 | 0 | 1945 | 21006 | 5577 | 18115 | 337 | 3623 | 0 | 205 | 0 | 2033 | 19271 |  | Unknown
WO-22270-42-A-1 | 0 | 9275 | 352 | 374 | 19191 | 507 | 4171824 | 119 | 697 | 1304 | 630 | 3997 | 0 | 3003 | 55602 | 250951 | 719660 | 183430 | 829097 | 2572875 | 748662 | 72721 | 23151283 | 4592944 | 6032911 | 2347910 | 11485926 | 7287 | 0 | 0 | 72490 | 2897 | 362 | 12105 | 0 | 156052 | 3236 | 1974 | 0 | 0 | 4346 | 9615 | 7467 | 27799 | 371 | 2981 | 0 | 527 | 0 | 0 | 16443 |  | Unknown
WO-22270-43-A-1 | 2532 | 13148 | 421 | 134 | 15301 | 627 | 4017411 | 122 | 388 | 0 | 0 | 2838 | 0 | 3819 | 48685 | 211757 | 630766 | 169870 | 810068 | 2487020 | 684181 | 66134 | 21964913 | 4527415 | 5623487 | 2210459 | 10691478 | 7526 | 0 | 0 | 73276 | 1030 | 799 | 10338 | 0 | 251107 | 0 | 515 | 0 | 0 | 144935 | 20229 | 6955 | 115769 | 2700 | 2471 | 0 | 696 | 0 | 0 | 12806 |  | Unknown
WO-22270-44-A-1 | 0 | 11537 | 1248 | 631 | 14772 | 0 | 4298749 | 68 | 229 | 1029 | 0 | 3821 | 0 | 4225 | 65987 | 275173 | 735004 | 187714 | 867535 | 2683705 | 757318 | 71330 | 24105423 | 4851054 | 6108574 | 2357439 | 12037762 | 10678 | 624 | 0 | 80655 | 3863 | 320 | 13043 | 0 | 193193 | 4666 | 2012 | 0 | 0 | 1254 | 10846 | 5933 | 29625 | 793 | 1318 | 0 | 191 | 0 | 0 | 13334 |  | Unknown
WO-22270-45-A-1 | 1032 | 12749 | 361 | 1095 | 8594 | 594 | 5249589 | 139 | 238 | 2225 | 967 | 2591 | 0 | 3510 | 68467 | 322921 | 1364450 | 206195 | 926799 | 2782733 | 801893 | 77952 | 24106814 | 5265961 | 6624417 | 2543666 | 12634452 | 8379 | 0 | 0 | 86272 | 1055 | 235 | 12904 | 0 | 179624 | 0 | 1447 | 38462 | 0 | 106613 | 68047 | 4015 | 0 | 228 | 2910 | 0 | 284 | 1153 | 0 | 68099 |  | Unknown
WO-22270-52-A-1 | 612 | 12612 | 531 | 550 | 8385 | 1111 | 3359440 | 11 | 107 | 1094 | 387 | 1671 | 0 | 14240 | 38718 | 203258 | 615855 | 86810 | 593628 | 1737365 | 351720 | 32302 | 17035994 | 3287481 | 4543967 | 1681002 | 8766585 | 4926 | 728 | 0 | 40552 | 2265 | 248 | 12366 | 0 | 207236 | 9812 | 1409 | 0 | 0 | 149277 | 15925 | 3707 | 21915 | 129 | 1410 | 0 | 197 | 0 | 0 | 9239 |  | Unknown
WO-22270-46-A-1 | 0 | 11804 | 937 | 381 | 10168 | 0 | 4333466 | 0 | 115 | 1516 | 608 | 3013 | 0 | 12402 | 67660 | 276508 | 743856 | 189053 | 865596 | 2734436 | 752980 | 71468 | 24586598 | 5026202 | 6296925 | 2375957 | 12022917 | 8136 | 0 | 0 | 86130 | 1029 | 168 | 12004 | 0 | 340406 | 0 | 1565 | 0 | 0 | 1694 | 17288 | 4834 | 28606 | 491 | 2563 | 0 | 527 | 957 | 0 | 15249 |  | Unknown
WO-22270-47-A-1 | 0 | 11397 | 657 | 633 | 10705 | 1422 | 4380525 | 70 | 121 | 674 | 0 | 3627 | 0 | 4012 | 62698 | 272901 | 744937 | 156945 | 810284 | 2543616 | 630850 | 57044 | 22479422 | 4696323 | 5846099 | 2286942 | 11703113 | 12054 | 1254 | 0 | 71626 | 2102 | 131 | 11830 | 0 | 179050 | 4216 | 1681 | 0 | 0 | 2072 | 10525 | 4204 | 24328 | 310 | 1749 | 0 | 152 | 1442 | 0 | 9166 |  | Unknown
WO-22270-53-A-1 | 1033 | 16786 | 886 | 306 | 9476 | 1010 | 4611598 | 0 | 60 | 1222 | 1199 | 3861 | 0 | 3284 | 80325 | 312897 | 819325 | 200515 | 915032 | 2871834 | 829180 | 78538 | 27991159 | 5511633 | 7017773 | 2757852 | 14229850 | 9140 | 1300 | 1171 | 90937 | 1002 | 82 | 14512 | 0 | 440712 | 0 | 1058 | 0 | 0 | 145375 | 28328 | 4219 | 44875 | 367 | 3349 | 0 | 246 | 1030 | 0 | 11412 |  | Unknown
WO-22270-48-A-1 | 0 | 12145 | 773 | 441 | 10044 | 0 | 4387633 | 31 | 899 | 994 | 0 | 3065 | 0 | 19183 | 74234 | 288110 | 743081 | 182187 | 873671 | 2638257 | 767247 | 67919 | 24901790 | 4945002 | 6441919 | 2495139 | 12535265 | 6969 | 0 | 679 | 81609 | 1407 | 165 | 13379 | 0 | 441632 | 4945 | 2195 | 0 | 0 | 110414 | 7722 | 4253 | 6358 | 262 | 1878 | 0 | 175 | 882 | 2240 | 10264 |  | Unknown
WO-22270-54-A-1 | 679 | 12893 | 689 | 841 | 11258 | 2239 | 5722559 | 0 | 135 | 3100 | 0 | 4070 | 0 | 15640 | 110268 | 467600 | 1110748 | 298970 | 1426202 | 4239213 | 1206695 | 118308 | 36975477 | 6821550 | 10975723 | 3872401 | 21571853 | 11737 | 711 | 0 | 129241 | 725 | 294 | 16967 | 0 | 523828 | 0 | 2667 | 0 | 3451 | 123039 | 12964 | 4672 | 28466 | 147 | 2288 | 0 | 336 | 1104 | 0 | 17731 |  | Unknown
WO-22270-49-A-1 | 1264 | 11146 | 617 | 323 | 6523 | 504 | 4288525 | 134 | 206 | 0 | 0 | 2792 | 0 | 4685 | 61272 | 269611 | 787090 | 156686 | 805874 | 2424454 | 598947 | 57519 | 22711009 | 4555655 | 5872741 | 2184481 | 11118381 | 7700 | 0 | 0 | 65892 | 1800 | 344 | 11221 | 0 | 452371 | 4118 | 1869 | 0 | 0 | 3086 | 12772 | 4594 | 13784 | 451 | 2677 | 7135 | 452 | 554 | 0 | 11863 |  | Unknown
WO-22270-55-A-1 | 715 | 9757 | 414 | 190 | 8115 | 0 | 2067748 | 18 | 0 | 1220 | 542 | 2122 | 0 | 14848 | 47145 | 196169 | 375140 | 115484 | 548155 | 1739005 | 423319 | 44005 | 11295160 | 2250704 | 4001568 | 1054358 | 6579510 | 7978 | 738 | 0 | 53133 | 1697 | 339 | 5678 | 0 | 146500 | 4173 | 1070 | 0 | 0 | 144241 | 85275 | 3794 | 0 | 380 | 1760 | 0 | 119 | 631 | 0 | 44005 |  | Unknown
WO-22270-50-A-1 | 0 | 13897 | 987 | 291 | 8718 | 798 | 4116789 | 17 | 72 | 0 | 285 | 3136 | 0 | 4410 | 77530 | 272500 | 718502 | 194929 | 837207 | 2654686 | 796369 | 78761 | 23587410 | 4722789 | 6125561 | 2303412 | 11990486 | 9505 | 0 | 0 | 83917 | 2230 | 379 | 13126 | 0 | 451250 | 0 | 1142 | 0 | 0 | 136992 | 52288 | 6475 | 16042 | 1227 | 1873 | 0 | 307 | 1252 | 824 | 12905 |  | Unknown
WO-22270-56-A-1 | 1557 | 16672 | 489 | 230 | 8261 | 882 | 4227742 | 113 | 223 | 0 | 918 | 1880 | 0 | 3700 | 71072 | 275863 | 768504 | 191057 | 818452 | 2705364 | 739990 | 78589 | 24097262 | 4729537 | 6054762 | 2354598 | 11905052 | 8377 | 557 | 0 | 79209 | 1228 | 464 | 11971 | 0 | 476727 | 0 | 1119 | 0 | 0 | 148882 | 29472 | 2411 | 42409 | 264 | 2793 | 0 | 16 | 636 | 0 | 16061 |  | Unknown
WO-22270-57-A-1 | 0 | 10762 | 393 | 389 | 7715 | 1370 | 5400712 | 25 | 128 | 2035 | 1152 | 4374 | 0 | 8763 | 91133 | 349112 | 971209 | 250975 | 1125337 | 3549627 | 986250 | 102889 | 33342342 | 6327514 | 8768564 | 3228336 | 18264750 | 12298 | 567 | 0 | 102772 | 3292 | 62 | 12430 | 0 | 612723 | 7707 | 1802 | 0 | 0 | 112010 | 36393 | 1753 | 26218 | 518 | 2905 | 0 | 233 | 943 | 621 | 12681 |  | Unknown
AB-220715-01 LRB 6_INJ001 | 0 | 13566 | 1307 | 394 | 14897 | 0 | 4184959 | 26 | 459 | 2520 | 0 | 715 | 12970 | 12681 | 249987 | 695868 | 713617 | 926924 | 2633926 | 8440018 | 3234487 | 626722 | 35663352 | 6910705 | 18386672 | 3520878 | 23132124 | 68749 | 1275 | 0 | 380763 | 1974 | 366 | 15253 | 0 | 27488 | 0 | 0 | 0 | 4021 | 113868 | 10939 | 3655 | 17963 | 442 | 2655 | 2743 | 341 | 1219 | 0 | 17146 |  | Unknown
AB-220715-01 LRB 6 | 1392 | 12406 | 1290 | 206 | 14463 | 536 | 3869761 | 0 | 176 | 4467 | 349 | 980 | 8861 | 6323 | 276686 | 672169 | 692694 | 938199 | 2505670 | 8616998 | 3297290 | 591494 | 33771748 | 6396599 | 17770400 | 3461344 | 22861691 | 37349 | 0 | 0 | 384775 | 988 | 254 | 14337 | 0 | 26795 | 1417 | 0 | 0 | 4741 | 1020 | 10319 | 3799 | 21478 | 106 | 2593 | 0 | 160 | 2326 | 815 | 15253 |  | Unknown
WO-22270-58-A-1 | 898 | 12819 | 937 | 455 | 5681 | 860 | 3919999 | 0 | 0 | 0 | 924 | 3091 | 0 | 5986 | 38222 | 217849 | 733040 | 81253 | 530893 | 1687259 | 366671 | 31073 | 17839994 | 3375777 | 4361498 | 1713680 | 9103504 | 3990 | 713 | 0 | 38250 | 1348 | 153 | 11054 | 0 | 317313 | 7307 | 679 | 0 | 0 | 126396 | 5811 | 4677 | 23486 | 239 | 1543 | 0 | 193 | 1518 | 0 | 7999 |  | Unknown
WO-22270-51-A-1 | 0 | 11462 | 1622 | 777 | 8293 | 866 | 4338525 | 0 | 189 | 0 | 0 | 3832 | 0 | 17029 | 78355 | 301735 | 865355 | 224738 | 922328 | 2877685 | 795112 | 78775 | 26825806 | 5294649 | 6789733 | 2656259 | 13508564 | 7845 | 0 | 0 | 91409 | 569 | 92 | 10647 | 0 | 336000 | 0 | 2124 | 0 | 0 | 9698 | 42603 | 3979 | 122297 | 153 | 1786 | 0 | 251 | 977 | 0 | 16648 |  | Unknown
WO-22270-59-A-1 | 1303 | 14327 | 553 | 261 | 6719 | 0 | 4070067 | 5 | 196 | 0 | 613 | 1792 | 0 | 12926 | 67376 | 239761 | 675464 | 180885 | 772260 | 2613869 | 738467 | 74023 | 23459572 | 4540862 | 6053205 | 2337281 | 12371284 | 8397 | 1133 | 2276 | 82960 | 1148 | 205 | 13305 | 0 | 327341 | 0 | 1586 | 0 | 0 | 139908 | 12858 | 2508 | 13899 | 291 | 2291 | 0 | 42 | 0 | 0 | 12870 |  | Unknown
WO-22270-60-A-1 | 0 | 12674 | 545 | 842 | 7227 | 960 | 4143868 | 19 | 126 | 0 | 1832 | 2765 | 0 | 4973 | 56129 | 248651 | 778452 | 145756 | 774035 | 2401626 | 626941 | 56797 | 23786204 | 4497406 | 5882129 | 2250888 | 11885493 | 7086 | 0 | 1965 | 62466 | 1338 | 41 | 12588 | 0 | 370656 | 9277 | 1576 | 0 | 0 | 157440 | 71277 | 1943 | 90623 | 404 | 2062 | 0 | 169 | 392 | 0 | 43884 |  | Unknown
AB-220715-01 LFB-H 4 | 1503 | 216664 | 340216 | 468646 | 70555392 | 1037 | 3471558 | 1386398 | 5018141 | 476936 | 854112 | 3102696 | 2057543 | 2705853 | 219256 | 336367 | 445453 | 823884 | 2084483 | 7765345 | 3179654 | 569850 | 27461333 | 5928938 | 14682995 | 2784852 | 19333922 | 60308 | 1444 | 615 | 364744 | 2394438 | 223 | 2809455 | 47640 | 25579333 | 987 | 93 | 7400010 | 3012016 | 185394 | 17701061 | 3189357 | 13826704 | 592 | 6067246 | 551317 | 1274113 | 1064 | 0 | 30232399 |  | Unknown
AB-220715-01 LRB 7_INJ001 | 0 | 9934 | 1221 | 743 | 106248 | 1591 | 3797787 | 543 | 7331 | 1479 | 949 | 4920 | 16982 | 14978 | 183324 | 465772 | 542135 | 922069 | 2444509 | 8294839 | 3371818 | 644923 | 33799083 | 7495353 | 17384755 | 3432322 | 21251175 | 67495 | 1239 | 599 | 400382 | 3092 | 162 | 17510 | 0 | 49198 | 0 | 12 | 0 | 5548 | 150344 | 31085 | 6214 | 99307 | 1105 | 17551 | 2780 | 387 | 0 | 0 | 49880 |  | Unknown
AB-220715-01 LRB 7 | 0 | 9068 | 737 | 434 | 28369 | 743 | 3080807 | 144 | 1126 | 2441 | 0 | 2148 | 15741 | 5139 | 157534 | 382542 | 440949 | 836401 | 2057854 | 7324006 | 3069010 | 602445 | 26605293 | 5935861 | 14208774 | 2763779 | 17045869 | 30987 | 1341 | 1487 | 347450 | 1490 | 0 | 15152 | 0 | 25132 | 0 | 39 | 0 | 4199 | 100117 | 40911 | 3762 | 78467 | 1174 | 4406 | 0 | 113 | 0 | 0 | 42877 |  | Unknown
AB-220715-01 LD 3 | 0 | 9727 | 715 | 485 | 15610 | 847 | 3720395 | 47 | 723 | 0 | 802 | 4682 | 0 | 4370 | 42024 | 147793 | 476243 | 178949 | 729556 | 2370771 | 695541 | 69273 | 22038283 | 5065795 | 5371566 | 2242228 | 10928132 | 9064 | 903 | 0 | 76012 | 946 | 252 | 10715 | 0 | 317139 | 0 | 1214 | 18163 | 0 | 140134 | 12068 | 6173 | 43604 | 4729 | 2416 | 3766 | 288 | 0 | 0 | 26057 |  | Unknown
AB-220715-01 CCC-H 4 | 806 | 222550 | 365786 | 604598 | 106354285 | 1077 | 3214531 | 2666056 | 9206694 | 794479 | 1350899 | 4408793 | 3107148 | 4300134 | 117532 | 196078 | 296213 | 688206 | 1527722 | 5407967 | 2491474 | 473808 | 18661633 | 5490930 | 9983058 | 2197810 | 12885821 | 42768 | 412 | 0 | 309516 | 3252020 | 56 | 5012152 | 64162 | 33456766 | 0 | 792 | 10070390 | 4486862 | 215472 | 22318412 | 4804286 | 18232010 | 1245 | 11143862 | 906840 | 158 | 0 | 0 | 39915006 |  | Unknown
AB-220715-01 LFSM-H 3 | 0 | 36407 | 125954 | 414069 | 57346629 | 625 | 3847670 | 349916 | 2624158 | 101590 | 179279 | 1617761 | 799544 | 2144933 | 36865 | 123413 | 411469 | 171966 | 840934 | 2431705 | 760210 | 73609 | 20032416 | 5307830 | 5587522 | 2356133 | 10882993 | 8005 | 0 | 0 | 80393 | 749930 | 494 | 3032026 | 0 | 18801662 | 0 | 1647 | 2319848 | 658540 | 157806 | 9664692 | 2992562 | 5209723 | 30697 | 5498947 | 68732 | 564937 | 0 | 0 | 23249965 |  | Unknown
WO-22270-61-A-1 | 0 | 8528 | 436 | 187 | 33456 | 1084 | 3907553 | 131 | 1085 | 0 | 1245 | 4960 | 0 | 19785 | 37760 | 172665 | 570884 | 159381 | 784412 | 2398705 | 629715 | 65566 | 23853721 | 5469572 | 5632908 | 2417223 | 11059184 | 8100 | 392 | 0 | 76889 | 679 | 243 | 10777 | 0 | 352623 | 0 | 3135 | 0 | 0 | 137398 | 11958 | 6632 | 14656 | 1779 | 4778 | 0 | 297 | 0 | 0 | 24071 |  | Unknown
WO-22270-62-A-1 | 0 | 7653 | 261 | 353 | 18271 | 561 | 4556279 | 36 | 640 | 848 | 962 | 4196 | 0 | 19162 | 34080 | 154958 | 998949 | 142858 | 704598 | 2125865 | 558706 | 59803 | 21352912 | 4488000 | 5277855 | 1890861 | 8851652 | 7449 | 1507 | 0 | 70819 | 1303 | 422 | 10320 | 0 | 302395 | 0 | 1777 | 0 | 0 | 122233 | 13698 | 1860 | 21000 | 2819 | 1925 | 0 | 247 | 0 | 0 | 14575 |  | Unknown
WO-22270-63-A-1 | 0 | 7161 | 146 | 328 | 12254 | 0 | 3353568 | 15 | 579 | 0 | 1087 | 5279 | 0 | 4162 | 25119 | 122773 | 479574 | 97773 | 562587 | 1651627 | 424919 | 38546 | 16824197 | 3891033 | 4194261 | 1618097 | 7870391 | 5932 | 780 | 0 | 45185 | 1701 | 247 | 8106 | 0 | 249797 | 0 | 1442 | 0 | 0 | 117439 | 10332 | 4586 | 29950 | 1468 | 2559 | 0 | 111 | 1454 | 996 | 12859 |  | Unknown
WO-22270-64-A-1 | 0 | 8247 | 553 | 0 | 9112 | 520 | 2733288 | 41 | 214 | 795 | 1124 | 3252 | 0 | 2237 | 22957 | 113298 | 318198 | 85349 | 525246 | 1505081 | 355301 | 33003 | 15389981 | 3429698 | 3735298 | 1461250 | 7084669 | 5328 | 574 | 0 | 39411 | 989 | 58 | 8600 | 0 | 240933 | 0 | 673 | 0 | 0 | 136085 | 100894 | 6670 | 0 | 4997 | 2194 | 0 | 112 | 0 | 0 | 10773 |  | Unknown
WO-22270-65-A-1 | 0 | 8776 | 1401 | 507 | 13177 | 415 | 4217569 | 30 | 92 | 0 | 364 | 2346 | 0 | 16131 | 48780 | 193054 | 568489 | 210409 | 928365 | 2859405 | 889358 | 88180 | 27040132 | 6026398 | 6960461 | 2805077 | 13445178 | 10197 | 805 | 0 | 94969 | 2297 | 116 | 13004 | 0 | 421107 | 0 | 1492 | 0 | 1868 | 129789 | 18975 | 4182 | 15194 | 979 | 2610 | 0 | 366 | 0 | 0 | 19084 |  | Unknown
WO-22270-66-A-1 | 1718 | 9186 | 696 | 244 | 11499 | 672 | 3632128 | 0 | 180 | 1778 | 2545 | 2901 | 0 | 20719 | 37834 | 145998 | 499692 | 156805 | 677177 | 2262051 | 693512 | 68368 | 21093445 | 4742478 | 5210331 | 2073449 | 9956516 | 9090 | 1181 | 0 | 74964 | 631 | 43 | 9661 | 0 | 297182 | 0 | 1124 | 0 | 0 | 128198 | 21512 | 529 | 20167 | 3006 | 1412 | 0 | 225 | 2010 | 0 | 12781 |  | Unknown
WO-22270-67-A-1 | 0 | 8396 | 603 | 114 | 8886 | 599 | 3237595 | 23 | 220 | 1309 | 1196 | 2053 | 0 | 15382 | 31621 | 132487 | 419572 | 132533 | 642445 | 1923433 | 540893 | 55026 | 19195880 | 4373277 | 4549314 | 1789267 | 8858423 | 6707 | 717 | 0 | 59699 | 1228 | 67 | 9794 | 0 | 356678 | 0 | 899 | 0 | 0 | 125356 | 799413 | 5326 | 50601 | 1842 | 2508 | 0 | 240 | 0 | 0 | 18803 |  | Unknown
WO-22270-68-A-1 | 530 | 6806 | 953 | 0 | 7334 | 514 | 3081227 | 0 | 151 | 1014 | 2175 | 3052 | 0 | 15670 | 29763 | 125264 | 426377 | 110015 | 600447 | 1795732 | 493220 | 48021 | 18173454 | 4165740 | 4280878 | 1755705 | 8346155 | 6750 | 1132 | 735 | 52466 | 1015 | 137 | 12284 | 0 | 297539 | 0 | 323 | 0 | 0 | 121219 | 8084 | 3557 | 20037 | 5915 | 2814 | 0 | 47 | 0 | 0 | 12048 |  | Unknown
WO-22270-69-A-1 | 0 | 8516 | 363 | 583 | 6462 | 539 | 2197280 | 8 | 211 | 0 | 1117 | 1852 | 0 | 16773 | 17138 | 90159 | 281018 | 56380 | 412816 | 1172252 | 270748 | 23222 | 12658951 | 2683042 | 3107400 | 1202969 | 5936933 | 4206 | 1531 | 480 | 28093 | 467 | 112 | 8906 | 0 | 276103 | 0 | 631 | 0 | 0 | 122422 | 7372 | 6055 | 7548 | 6479 | 2195 | 0 | 33 | 2672 | 797 | 10235 |  | Unknown
WO-22270-70-A-1 | 589 | 9125 | 846 | 51 | 5341 | 493 | 2947739 | 0 | 36 | 693 | 854 | 2218 | 0 | 6699 | 22543 | 123191 | 363737 | 101322 | 563025 | 1671924 | 420469 | 41393 | 17148751 | 3840491 | 4183401 | 1639413 | 7799028 | 5216 | 0 | 1501 | 47787 | 1555 | 178 | 9525 | 0 | 310369 | 0 | 739 | 11162 | 0 | 148268 | 38971 | 2170 | 70540 | 2658 | 2252 | 0 | 209 | 0 | 0 | 36753 |  | Unknown
AB-220715-01 LRB 8_INJ001 | 1940 | 8888 | 333 | 113 | 8882 | 567 | 3616106 | 27 | 154 | 1081 | 1272 | 1403 | 17968 | 5161 | 159490 | 389331 | 529744 | 902082 | 2411403 | 8382177 | 3527356 | 665890 | 31665544 | 7118950 | 17191201 | 3354216 | 20522310 | 79005 | 735 | 0 | 434366 | 735 | 133 | 10913 | 0 | 19957 | 0 | 85 | 0 | 4703 | 135545 | 41497 | 3937 | 95064 | 2325 | 2389 | 7023 | 388 | 932 | 0 | 57680 |  | Unknown
AB-220715-01 LRB 8 | 0 | 8320 | 398 | 206 | 18875 | 659 | 2974855 | 20 | 522 | 2086 | 884 | 1812 | 15835 | 13815 | 146726 | 356459 | 411697 | 816716 | 2081068 | 7496657 | 3015102 | 591282 | 25841839 | 6033865 | 14127326 | 2866857 | 17494277 | 35119 | 650 | 0 | 347082 | 1152 | 0 | 8632 | 0 | 21196 | 0 | 24 | 0 | 3964 | 71862 | 47534 | 4222 | 72045 | 1222 | 2609 | 5094 | 350 | 0 | 0 | 31297 |  | Unknown
WO-22270-71-A-1 | 779 | 7557 | 429 | 175 | 9018 | 2076 | 4337036 | 62 | 321 | 1547 | 804 | 3363 | 0 | 3231 | 37984 | 151120 | 737289 | 140433 | 688630 | 2146971 | 606222 | 62675 | 20858411 | 4884731 | 5108915 | 2019373 | 9865279 | 8916 | 225 | 0 | 70935 | 690 | 523 | 9814 | 0 | 822733 | 0 | 1473 | 0 | 0 | 127394 | 231361 | 3120 | 56578 | 1916 | 1808 | 0 | 152 | 0 | 0 | 15039 |  | Unknown
WO-22270-72-A-1 | 619 | 8491 | 461 | 291 | 7095 | 553 | 3551904 | 46 | 34 | 734 | 748 | 2206 | 0 | 4182 | 34545 | 149554 | 463809 | 154375 | 676054 | 2175315 | 618722 | 63479 | 20577747 | 4741897 | 5013576 | 1987691 | 9800270 | 7587 | 231 | 0 | 70231 | 1081 | 92 | 9748 | 0 | 366751 | 0 | 749 | 0 | 0 | 107666 | 6675 | 2488 | 25655 | 1013 | 1328 | 0 | 81 | 0 | 0 | 13043 |  | Unknown
WO-22270-73-A-1 | 661 | 7541 | 231 | 31 | 4864 | 524 | 3082300 | 17 | 0 | 1516 | 1073 | 2832 | 0 | 3110 | 27815 | 126217 | 404727 | 124508 | 615152 | 1896728 | 507276 | 48718 | 18861487 | 4229281 | 4622658 | 1778582 | 8907123 | 6110 | 779 | 951 | 56268 | 831 | 200 | 11436 | 0 | 280113 | 0 | 750 | 0 | 0 | 129252 | 8770 | 2714 | 6497 | 1488 | 1335 | 0 | 226 | 686 | 1170 | 10272 |  | Unknown
WO-22270-74-A-1 | 590 | 7780 | 792 | 164 | 5337 | 547 | 3318136 | 29 | 12 | 1022 | 1013 | 2321 | 0 | 14702 | 26535 | 124794 | 399227 | 122767 | 598123 | 1856010 | 511395 | 45379 | 18839817 | 4076075 | 4366331 | 1752307 | 8820378 | 7115 | 0 | 809 | 56226 | 891 | 18 | 9670 | 0 | 582455 | 0 | 713 | 8115 | 0 | 124582 | 89636 | 3377 | 38117 | 1108 | 1705 | 0 | 350 | 0 | 0 | 22051 |  | Unknown
WO-22270-75-A-1 | 0 | 7901 | 473 | 110 | 8118 | 1096 | 3335552 | 19 | 78 | 1874 | 752 | 3844 | 0 | 13655 | 36234 | 150869 | 428711 | 189479 | 743176 | 2339942 | 754436 | 83220 | 21679778 | 5001727 | 5477571 | 2099428 | 10411323 | 8457 | 942 | 653 | 80966 | 1765 | 71 | 10830 | 0 | 320347 | 0 | 868 | 18721 | 0 | 131126 | 72539 | 4666 | 118686 | 1504 | 2180 | 0 | 215 | 0 | 908 | 48403 |  | Unknown
WO-22270-76-A-1 | 0 | 8354 | 403 | 0 | 4052 | 364 | 2096551 | 14 | 6 | 713 | 654 | 3248 | 0 | 16895 | 14709 | 87946 | 264520 | 61407 | 414464 | 1167594 | 264233 | 26837 | 11850992 | 2540321 | 2991368 | 1075264 | 5806209 | 4541 | 386 | 0 | 28247 | 1862 | 67 | 8073 | 0 | 220081 | 0 | 427 | 0 | 0 | 128838 | 7309 | 2690 | 19413 | 4611 | 892 | 0 | 103 | 0 | 0 | 26763 |  | Unknown
WO-22270-77-A-1 | 801 | 8146 | 314 | 248 | 7231 | 0 | 4154030 | 24 | 167 | 700 | 814 | 5068 | 0 | 19500 | 40741 | 145771 | 651043 | 193017 | 724842 | 2397560 | 763299 | 86615 | 22956505 | 5293461 | 5720353 | 2155191 | 10925884 | 9382 | 651 | 3841203 | 89528 | 1931 | 375 | 7308 | 0 | 2110610 | 0 | 973 | 16217 | 3827 | 140541 | 9976 | 4459 | 14858 | 6607 | 1837 | 0 | 148 | 4583 | 0 | 15653 |  | Unknown
WO-22270-78-A-1 | 0 | 8391 | 682 | 462 | 5509 | 1313 | 3699256 | 0 | 0 | 1887 | 371 | 3362 | 0 | 16060 | 42786 | 147140 | 504829 | 213636 | 775811 | 2544211 | 831775 | 104731 | 23688616 | 5532173 | 5776537 | 2165997 | 11367131 | 11120 | 0 | 0 | 97909 | 2346 | 60 | 9851 | 0 | 420719 | 0 | 1171 | 16951 | 4851 | 125913 | 17021 | 3043 | 18039 | 6593 | 1537 | 0 | 704 | 0 | 0 | 12977 |  | Unknown
WO-22270-79-A-1 | 1091 | 11316 | 665 | 132 | 6789 | 625 | 3746846 | 17 | 0 | 1218 | 718 | 3357 | 0 | 5471 | 39816 | 149632 | 491960 | 204186 | 756442 | 2503476 | 816834 | 99216 | 23328471 | 5429378 | 5728536 | 2258374 | 11034958 | 11439 | 0 | 0 | 92966 | 1911 | 108 | 10667 | 0 | 340294 | 0 | 906 | 7548 | 5042 | 136957 | 8018 | 1593 | 22072 | 4521 | 1875 | 0 | 216 | 0 | 0 | 17641 |  | Unknown
WO-22270-80-A-1 | 0 | 11070 | 508 | 424 | 4389 | 431 | 3896630 | 0 | 112 | 1913 | 1429 | 2717 | 0 | 3215 | 44513 | 155299 | 540683 | 235939 | 784201 | 2601534 | 876904 | 112711 | 24003414 | 5637572 | 5798240 | 2245614 | 11573539 | 11307 | 1454 | 0 | 103221 | 1263 | 63 | 10909 | 0 | 390842 | 0 | 939 | 20313 | 5827 | 129938 | 38564 | 3864 | 43172 | 7328 | 1345 | 0 | 117 | 0 | 0 | 24660 |  | Unknown
AB-220715-01 LFB-H 5 | 420 | 111982 | 196278 | 305790 | 55658364 | 414 | 2745957 | 1383392 | 4858328 | 429462 | 727330 | 2406752 | 1700250 | 2191323 | 109154 | 197025 | 286216 | 707518 | 1746852 | 6345696 | 2509330 | 543401 | 20292643 | 5351752 | 11001705 | 2265980 | 13828390 | 50507 | 754 | 743 | 307686 | 1930004 | 0 | 2177492 | 45638 | 19004852 | 0 | 612 | 6059930 | 2455116 | 157348 | 12777293 | 2535358 | 10454486 | 31475 | 5454048 | 534838 | 1550301 | 0 | 2106 | 23217998 |  | Unknown
AB-220715-01 LRB 9_INJ001 | 0 | 6898 | 878 | 676 | 90524 | 707 | 3232848 | 426 | 6204 | 1565 | 1063 | 2813 | 17436 | 17785 | 131630 | 359261 | 430506 | 838718 | 2252011 | 7418345 | 2983127 | 607329 | 29943930 | 6482352 | 15068280 | 2921618 | 18439011 | 69981 | 901 | 0 | 351091 | 2161 | 7 | 11932 | 0 | 42210 | 0 | 123 | 0 | 5141 | 145415 | 61688 | 4445 | 102768 | 4143 | 14637 | 5730 | 742 | 0 | 0 | 39976 |  | Unknown
AB-220715-01 LRB 9 | 0 | 6187 | 967 | 198 | 22611 | 995 | 2854269 | 152 | 1117 | 2415 | 1182 | 1426 | 16617 | 4877 | 127754 | 331799 | 400328 | 809751 | 1976682 | 6810433 | 2903771 | 556892 | 26409809 | 5926337 | 13921341 | 2677379 | 17044426 | 48600 | 1424 | 2477 | 345555 | 1205 | 103 | 11073 | 0 | 20834 | 0 | 22 | 4847 | 4203 | 99478 | 66876 | 4691 | 79033 | 4612 | 3444 | 5033 | 656 | 0 | 0 | 38388 |  | Unknown
AB-220715-01 LD 4 | 787 | 9929 | 492 | 203 | 9337 | 0 | 3371119 | 228 | 1319 | 1955 | 2210 | 2638 | 0 | 17149 | 43730 | 138563 | 430349 | 178794 | 672550 | 2294548 | 716808 | 83682 | 21008614 | 4560525 | 5339548 | 2022776 | 10202374 | 8799 | 0 | 0 | 82518 | 1696 | 0 | 9291 | 0 | 344332 | 0 | 961 | 23514 | 5730 | 124757 | 41744 | 1581 | 76384 | 6456 | 1927 | 0 | 218 | 0 | 1608 | 25873 |  | Unknown
AB-220715-01 CCC-H 5 | 1172 | 214490 | 344224 | 548232 | 99461190 | 527 | 2936305 | 2782037 | 8952549 | 862627 | 1454306 | 4263034 | 2963542 | 4047002 | 110818 | 180671 | 270352 | 727228 | 1486649 | 5634422 | 2693442 | 567500 | 18565659 | 5306374 | 10341770 | 2060162 | 12768315 | 62264 | 514 | 0 | 335321 | 3291555 | 47 | 4489163 | 98549 | 33619431 | 0 | 621 | 10292609 | 4937353 | 223687 | 22650407 | 4460795 | 18357758 | 65069 | 10758132 | 1083022 | 2988149 | 0 | 740 | 39887512 |  | Unknown
AB-220715-01 LFSM-H 4 | 0 | 39296 | 112519 | 334592 | 48480791 | 583 | 3426306 | 360623 | 2165387 | 116974 | 198809 | 1348886 | 665615 | 1727331 | 38088 | 112861 | 343129 | 197978 | 709155 | 2391526 | 796480 | 88530 | 18544192 | 4643841 | 5148850 | 1943838 | 9525681 | 12178 | 0 | 0 | 92993 | 687604 | 197 | 2743178 | 0 | 17521458 | 0 | 2089 | 2249845 | 733926 | 154873 | 8739401 | 2652123 | 4758056 | 15459 | 4735905 | 83428 | 522033 | 0 | 1167 | 22429132 |  | Unknown
WO-22270-81-A-1 | 0 | 8305 | 641 | 199 | 24432 | 459 | 3111289 | 34 | 1315 | 1405 | 1047 | 3504 | 0 | 23967 | 28323 | 121074 | 366521 | 123221 | 585225 | 1890612 | 501260 | 50803 | 18293922 | 4205561 | 4423039 | 1690896 | 8639728 | 6348 | 1245 | 1020 | 56564 | 3770 | 358 | 11874 | 0 | 302441 | 0 | 670 | 14503 | 6331 | 129422 | 12073 | 3130 | 27237 | 7777 | 3288 | 0 | 367 | 0 | 0 | 16563 |  | Unknown
WO-22270-82-A-1 | 0 | 9428 | 1841 | 1510 | 16569 | 876 | 5036626 | 0 | 1321 | 2071 | 800 | 4478 | 0 | 3291 | 40997 | 177948 | 1082853 | 197317 | 816846 | 2710178 | 800165 | 89713 | 24731305 | 6124641 | 6277768 | 2372708 | 12553357 | 10423 | 551 | 890 | 89715 | 1529 | 30 | 11445 | 0 | 389467 | 0 | 1855 | 38908 | 9913 | 140968 | 13121 | 5010 | 60602 | 6780 | 3339 | 0 | 416 | 0 | 0 | 16749 |  | Unknown
WO-22270-83-A-1 | 0 | 12084 | 1356 | 350 | 13535 | 1860 | 3904656 | 0 | 835 | 1717 | 883 | 4139 | 0 | 16992 | 45208 | 161946 | 513506 | 205795 | 789113 | 2669831 | 844460 | 95023 | 24460084 | 5757290 | 6020994 | 2323868 | 11784212 | 10377 | 1001 | 0 | 94527 | 2042 | 154 | 9702 | 0 | 380924 | 0 | 1758 | 0 | 12967 | 136701 | 17280 | 6342 | 17751 | 3247 | 2554 | 0 | 518 | 0 | 0 | 10618 |  | Unknown
WO-22270-84-A-1 | 0 | 8232 | 420 | 447 | 11714 | 1058 | 3662467 | 0 | 233 | 0 | 2285 | 3281 | 0 | 3946 | 39531 | 155791 | 502287 | 181820 | 750904 | 2546463 | 772173 | 90614 | 23008855 | 5423355 | 5720144 | 2220971 | 11441580 | 10508 | 0 | 0 | 91522 | 723 | 0 | 11169 | 0 | 250306 | 0 | 1229 | 0 | 35871 | 143166 | 10404 | 4790 | 41011 | 1320 | 3532 | 0 | 181 | 0 | 0 | 12801 |  | Unknown
WO-22270-85-A-1 | 0 | 8213 | 389 | 389 | 10466 | 609 | 3742549 | 7 | 621 | 1144 | 888 | 7525 | 0 | 20926 | 39565 | 150225 | 566827 | 203653 | 758643 | 2643643 | 780033 | 101864 | 24050284 | 5436740 | 5866912 | 2293834 | 11662114 | 10741 | 949 | 2955 | 91493 | 1202 | 278 | 10502 | 0 | 349758 | 0 | 1872 | 25406 | 28135 | 133755 | 5274 | 4864 | 43425 | 6310 | 2734 | 0 | 441 | 1339 | 0 | 14726 |  | Unknown
WO-22270-86-A-1 | 775 | 9313 | 895 | 534 | 10473 | 809 | 5056926 | 0 | 261 | 1240 | 731 | 4685 | 0 | 4299 | 40135 | 170124 | 1026824 | 189036 | 766801 | 2512472 | 757330 | 82981 | 23675098 | 5729521 | 6104626 | 2335104 | 11998524 | 10626 | 1253 | 563 | 86556 | 1870 | 275 | 11785 | 0 | 344562 | 0 | 1857 | 40915 | 27535 | 142258 | 11327 | 5469 | 92389 | 1085 | 1459 | 0 | 134 | 0 | 0 | 13852 |  | Unknown
WO-22270-87-A-1 | 975 | 9242 | 266 | 72 | 8556 | 780 | 5251589 | 26 | 428 | 0 | 0 | 3652 | 0 | 15370 | 41923 | 166079 | 996984 | 202143 | 753415 | 2590736 | 795368 | 89935 | 22623919 | 5981619 | 6274669 | 2416998 | 12290455 | 10305 | 0 | 0 | 86474 | 2193 | 76 | 10567 | 0 | 380055 | 0 | 1528 | 24468 | 108027 | 143887 | 15496 | 4433 | 11044 | 4720 | 3067 | 0 | 568 | 0 | 2251 | 14997 |  | Unknown
WO-22270-88-A-1 | 0 | 10272 | 945 | 181 | 6947 | 500 | 4014703 | 10 | 177 | 1501 | 760 | 3098 | 0 | 18592 | 33766 | 141890 | 508439 | 181172 | 741640 | 2504749 | 709023 | 80438 | 23474682 | 5391922 | 5546647 | 2202924 | 11225127 | 9324 | 804 | 25220582 | 83236 | 3233 | 20 | 11768 | 0 | 453097 | 0 | 1699 | 0 | 97764 | 144953 | 10936 | 4039 | 14225 | 6776 | 3066 | 0 | 360 | 5138 | 0 | 10300 |  | Unknown
WO-22270-89-A-1 | 573 | 8969 | 952 | 189 | 4256 | 580 | 3134655 | 29 | 7 | 759 | 643 | 3878 | 0 | 16214 | 23877 | 116216 | 385850 | 113973 | 574467 | 1852675 | 420888 | 46694 | 18307018 | 4044882 | 4370898 | 1660729 | 8539778 | 7368 | 276 | 0 | 47149 | 834 | 202 | 9326 | 0 | 306446 | 0 | 1543 | 27394 | 68153 | 133721 | 20029 | 1669 | 24498 | 6108 | 1633 | 0 | 158 | 0 | 0 | 31980 |  | Unknown
WO-22270-90-A-1 | 908 | 7161 | 687 | 261 | 4525 | 453 | 3091452 | 0 | 145 | 903 | 580 | 2379 | 0 | 3320 | 27620 | 115994 | 365302 | 130434 | 587176 | 1982942 | 524788 | 55299 | 18276354 | 3819111 | 4568482 | 1514821 | 7750832 | 7240 | 572 | 1226 | 59977 | 1359 | 197 | 10187 | 0 | 277744 | 0 | 1340 | 16901 | 84809 | 132586 | 0 | 3162 | 0 | 5156 | 2013 | 0 | 301 | 0 | 0 | 10142 |  | Unknown
AB-220715-01 LRB 10_INJ001 | 0 | 7674 | 825 | 320 | 7950 | 644 | 2869361 | 0 | 187 | 1804 | 1589 | 1256 | 0 | 4648 | 124925 | 325823 | 364032 | 746001 | 1996278 | 7035897 | 2761186 | 519819 | 26542230 | 5577756 | 14133799 | 2386365 | 16210796 | 51251 | 902 | 658 | 294329 | 2325 | 19 | 9080 | 0 | 24094 | 0 | 0 | 0 | 0 | 124659 | 65741 | 3691 | 73607 | 4204 | 1872 | 5750 | 886 | 0 | 0 | 45985 |  | Unknown
AB-220715-01 LRB 10 | 1363 | 8723 | 778 | 19 | 7780 | 1508 | 2141288 | 22 | 98 | 1595 | 483 | 1500 | 0 | 5824 | 85019 | 247579 | 326823 | 504579 | 1417805 | 5090765 | 1855802 | 375550 | 20692741 | 4535551 | 10464064 | 2159872 | 13478997 | 18497 | 1095 | 734 | 210863 | 2457 | 0 | 8742 | 0 | 23909 | 0 | 0 | 0 | 3141 | 78968 | 47147 | 4414 | 59437 | 1486 | 1286 | 4009 | 190 | 0 | 0 | 26474 |  | Unknown
WO-22270-91-A-1 | 0 | 7002 | 219 | 239 | 8069 | 0 | 3812237 | 0 | 473 | 0 | 553 | 3280 | 0 | 4881 | 41354 | 155800 | 468885 | 199719 | 746753 | 2566212 | 748555 | 85774 | 23113498 | 5284956 | 5809917 | 2198104 | 11308729 | 11147 | 0 | 1326 | 84669 | 1477 | 106 | 10886 | 0 | 296888 | 0 | 1646 | 13086 | 316158 | 121358 | 5055 | 3385 | 65642 | 2083 | 1491 | 0 | 332 | 0 | 0 | 16417 |  | Unknown
WO-22270-92-A-1 | 781 | 12486 | 626 | 215 | 7929 | 1041 | 4747567 | 28 | 84 | 1461 | 974 | 3192 | 0 | 5649 | 38339 | 150126 | 840361 | 211782 | 746938 | 2698378 | 832541 | 96188 | 24576698 | 5678325 | 5952482 | 2382566 | 12500813 | 11620 | 0 | 729 | 89135 | 2199 | 144 | 10830 | 0 | 417804 | 0 | 1988 | 45749 | 468269 | 138780 | 13864 | 3607 | 39305 | 2097 | 1848 | 0 | 347 | 0 | 0 | 19522 |  | Unknown
WO-22270-93-A-1 | 0 | 9196 | 648 | 392 | 7428 | 1534 | 4163346 | 251 | 40 | 1035 | 2250 | 2990 | 0 | 16394 | 45781 | 164928 | 533461 | 224370 | 805398 | 2774581 | 863242 | 103146 | 24694734 | 5666683 | 6257337 | 2352554 | 12289407 | 12938 | 1408 | 670471 | 99668 | 1354 | 183 | 11335 | 0 | 475445 | 0 | 1213 | 36010 | 391620 | 146214 | 14555 | 1058 | 69532 | 5926 | 2302 | 0 | 270 | 0 | 1692 | 35299 |  | Unknown
WO-22270-94-A-1 | 0 | 7849 | 2148 | 413 | 6136 | 975 | 5311255 | 0 | 68 | 1340 | 891 | 2734 | 0 | 3458 | 35484 | 173837 | 1271845 | 169459 | 715334 | 2450242 | 666687 | 78552 | 22499349 | 5566900 | 5846242 | 2195247 | 12149788 | 10572 | 0 | 22544708 | 80065 | 1060 | 80 | 10624 | 0 | 372922 | 0 | 1100 | 34928 | 1634577 | 131701 | 16992 | 619 | 108505 | 6438 | 2115 | 870 | 463 | 4841 | 0 | 35962 |  | Unknown
WO-22270-95-A-1 | 1093 | 9120 | 362 | 60 | 5304 | 965 | 3276199 | 7 | 122 | 1057 | 1064 | 4188 | 0 | 3386 | 34155 | 129496 | 424027 | 141093 | 616178 | 2109594 | 565736 | 63022 | 20794691 | 4462463 | 4906830 | 1862901 | 9458542 | 7724 | 0 | 0 | 64144 | 1331 | 0 | 10489 | 0 | 331029 | 0 | 1891 | 10709 | 800072 | 112670 | 8579 | 625 | 20322 | 1245 | 1098 | 0 | 134 | 0 | 0 | 9938 |  | Unknown
WO-22270-96-A-1 | 0 | 7304 | 394 | 361 | 6595 | 0 | 3634137 | 12 | 11 | 0 | 432 | 4036 | 0 | 6001 | 42375 | 145047 | 462509 | 211412 | 748224 | 2733224 | 813660 | 100463 | 23826961 | 5443547 | 5844232 | 2230558 | 11541020 | 12916 | 771 | 32251827 | 97388 | 1216 | 37 | 11443 | 0 | 413690 | 0 | 1438 | 11665 | 1666284 | 129368 | 7982 | 6074 | 51197 | 7520 | 2501 | 0 | 25 | 7000 | 0 | 12556 |  | Unknown
WO-22270-97-A-1 | 0 | 7703 | 1100 | 93 | 6423 | 905 | 3550060 | 0 | 346 | 1210 | 605 | 3355 | 0 | 3907 | 39361 | 146435 | 429484 | 210196 | 743371 | 2621815 | 843725 | 107670 | 23791887 | 5256730 | 5748496 | 2214819 | 11492106 | 12260 | 0 | 0 | 92828 | 1490 | 151 | 9562 | 0 | 386211 | 0 | 1103 | 0 | 5893 | 139773 | 16496 | 2946 | 19195 | 7907 | 1503 | 0 | 445 | 3407 | 0 | 14296 |  | Unknown
WO-22270-98-A-1 | 0 | 7374 | 937 | 429 | 6318 | 1070 | 3337603 | 0 | 8 | 2142 | 532 | 2140 | 0 | 2917 | 37518 | 131347 | 411539 | 215239 | 750296 | 2577667 | 805018 | 106408 | 23127370 | 5209000 | 5652528 | 2104134 | 11383458 | 10973 | 1200 | 32172240 | 90782 | 0 | 310 | 10383 | 0 | 466992 | 0 | 1613 | 0 | 4660 | 122663 | 9495 | 4229 | 8883 | 6621 | 1673 | 0 | 398 | 8015 | 0 | 13964 |  | Unknown
WO-22270-99-A-1 | 0 | 7009 | 1051 | 393 | 2769 | 532 | 2020279 | 0 | 383 | 639 | 439 | 2034 | 0 | 12759 | 17146 | 79828 | 256072 | 66716 | 412706 | 1239030 | 287145 | 32303 | 12534084 | 2466790 | 3190425 | 984422 | 5970633 | 6270 | 1106 | 417 | 30747 | 540 | 288 | 6705 | 0 | 244377 | 0 | 761 | 0 | 0 | 153933 | 21247 | 2136 | 0 | 6889 | 684 | 0 | 0 | 4505 | 0 | 7031 |  | Unknown
WO-22270-100-A-1 | 0 | 9741 | 225 | 42 | 4286 | 879 | 2945899 | 0 | 0 | 0 | 1128 | 1706 | 0 | 2278 | 27446 | 120132 | 333347 | 122929 | 565199 | 1890414 | 505037 | 57395 | 17880141 | 3864311 | 4410993 | 1587015 | 8549173 | 6952 | 1729 | 0 | 54327 | 2393 | 0 | 9058 | 0 | 345252 | 0 | 647 | 6357 | 2931 | 140567 | 7895 | 3131 | 29320 | 8353 | 987 | 0 | 165 | 0 | 0 | 24362 |  | Unknown
AB-220715-01 LFB-H 6 | 758 | 93352 | 177684 | 267333 | 52254895 | 713 | 2920160 | 1177742 | 4405881 | 320835 | 581243 | 2258951 | 1424820 | 2101300 | 94289 | 184156 | 268701 | 590622 | 1526333 | 5483197 | 2239385 | 454298 | 21637634 | 5221880 | 10675315 | 2205900 | 14485110 | 39912 | 568 | 3330155 | 265343 | 1604932 | 128 | 2120976 | 30479 | 19135615 | 0 | 115 | 5001643 | 2039056 | 178607 | 12232426 | 2556969 | 9646803 | 32291 | 5269778 | 412715 | 1244965 | 5234 | 0 | 23725077 |  | Unknown
AB-220715-01 LRB 11_INJ001 | 0 | 6558 | 1200 | 262 | 71967 | 505 | 2992995 | 344 | 5532 | 2907 | 987 | 2742 | 0 | 23425 | 125522 | 323862 | 401357 | 880847 | 2134600 | 7838529 | 3189079 | 56292 | 29425554 | 6370089 | 15551575 | 2808866 | 19371096 | 76916 | 575 | 37322337 | 377121 | 4827 | 282 | 11294 | 0 | 42416 | 0 | 4215 | 0 | 0 | 144120 | 48396 | 3445 | 94678 | 7401 | 11879 | 3351 | 683 | 7618 | 41498 | 36454 |  | Unknown
AB-220715-01 LRB 11 | 1697 | 8369 | 940 | 292 | 20516 | 342 | 2972327 | 0 | 727 | 1964 | 511 | 2164 | 0 | 17159 | 137342 | 317843 | 379481 | 853202 | 2004282 | 7611678 | 3153962 | 726873 | 27909548 | 5997042 | 15029356 | 2767757 | 18225009 | 53079 | 0 | 32040963 | 347007 | 2148 | 116 | 11166 | 0 | 24449 | 0 | 1486 | 0 | 4407 | 109124 | 46611 | 1374 | 84364 | 5010 | 4832 | 3453 | 362 | 6976 | 0 | 41219 |  | Unknown
AB-220715-01 LD 5 | 0 | 8281 | 874 | 42 | 6063 | 479 | 1986878 | 9 | 205 | 459 | 722 | 1870 | 0 | 16381 | 15506 | 90056 | 215599 | 66288 | 416502 | 1251945 | 270600 | 32221 | 12616466 | 2376266 | 3149240 | 1090966 | 6088699 | 4897 | 0 | 0 | 30142 | 1093 | 141 | 6619 | 0 | 262286 | 0 | 419 | 0 | 2738 | 120360 | 25155 | 3509 | 44862 | 6470 | 1045 | 0 | 0 | 3523 | 1552 | 21553 |  | Unknown
AB-220715-01 CCC-H 6 | 0 | 196917 | 308770 | 513602 | 96291946 | 0 | 2898925 | 2555889 | 8104449 | 709682 | 1282249 | 4087929 | 2657766 | 3921642 | 103456 | 163210 | 262392 | 618538 | 1392965 | 5330652 | 2363626 | 514750 | 18987306 | 5039727 | 10085369 | 2118416 | 13396215 | 52567 | 811 | 4566785 | 286156 | 3141215 | 347 | 4525629 | 87716 | 33730016 | 4903 | 916 | 9781006 | 4153258 | 251355 | 22660272 | 4480858 | 17705680 | 81763 | 9719129 | 995560 | 2449170 | 7012 | 0 | 41422884 |  | Unknown
AB-220715-01 LFSM-H 5 | 0 | 22399 | 78727 | 232164 | 34753341 | 1060 | 2425621 | 185907 | 1329522 | 60744 | 108395 | 850590 | 462375 | 1184609 | 20425 | 80095 | 234791 | 107413 | 510075 | 1644704 | 440231 | 51458 | 14363915 | 3269054 | 3657926 | 1348555 | 7062576 | 6486 | 0 | 18145210 | 49652 | 496955 | 107 | 1932644 | 0 | 13360699 | 0 | 1930001 | 1531504 | 385577 | 153142 | 6259048 | 1993524 | 3379075 | 35937 | 3320598 | 45943 | 331309 | 5236 | 1755 | 19208271 |  | Unknown
WO-22270-101-A-1 | 0 | 7780 | 1093 | 17 | 27790 | 536 | 3835357 | 121 | 2583 | 2655 | 618 | 1900 | 0 | 7610 | 41175 | 148820 | 439884 | 214921 | 839518 | 2897469 | 875364 | 108862 | 26517096 | 5629952 | 6443524 | 2482525 | 13267462 | 12883 | 0 | 35893482 | 94777 | 1086 | 612 | 12543 | 0 | 478153 | 0 | 1915 | 5694 | 9359 | 129595 | 11487 | 7569 | 13752 | 5261 | 4485 | 0 | 90 | 7319 | 0 | 20653 |  | Unknown
WO-22270-102-A-1 | 1302 | 5468 | 1173 | 48 | 15037 | 900 | 3311374 | 77 | 203 | 1290 | 1377 | 1201 | 0 | 17781 | 39713 | 135655 | 411209 | 206879 | 723586 | 2616679 | 821723 | 102235 | 23878042 | 5166919 | 5790764 | 2132395 | 11273981 | 11977 | 0 | 25279006 | 88647 | 1141 | 37 | 13065 | 0 | 430960 | 0 | 1461 | 0 | 8274 | 138971 | 13257 | 6669 | 39453 | 6840 | 3128 | 0 | 646 | 8080 | 0 | 15372 |  | Unknown
WO-22270-103-A-1 | 0 | 17675 | 411 | 16 | 17273 | 426 | 4474384 | 17 | 572 | 1557 | 910 | 3028 | 0 | 19696 | 52570 | 210505 | 635785 | 258852 | 1140323 | 3590592 | 1025468 | 122453 | 32941178 | 6517029 | 9321004 | 3157958 | 18248804 | 11421 | 0 | 32517906 | 109029 | 2597 | 126 | 11278 | 0 | 738232 | 0 | 2598 | 0 | 14244 | 105429 | 49312 | 5477 | 46365 | 3644 | 2183 | 0 | 296 | 7297 | 1166 | 15304 |  | Unknown
WO-22270-104-A-1 | 0 | 8195 | 963 | 33 | 13699 | 264 | 3153089 | 23 | 457 | 994 | 815 | 2394 | 0 | 4205 | 34321 | 119995 | 387220 | 146485 | 627835 | 2098333 | 594156 | 68013 | 20829860 | 4519697 | 4923165 | 1894019 | 9773223 | 7143 | 1092 | 26242891 | 68480 | 756 | 0 | 9908 | 0 | 368749 | 0 | 1487 | 0 | 21439 | 135476 | 5438 | 3532 | 8042 | 7071 | 2866 | 0 | 390 | 5697 | 983 | 10976 |  | Unknown
WO-22270-105-A-1 | 0 | 17011 | 346 | 208 | 8404 | 1922 | 4766902 | 10 | 291 | 1066 | 622 | 2374 | 0 | 4921 | 33101 | 154351 | 1055037 | 149845 | 667558 | 2284706 | 594750 | 66095 | 20703722 | 4675678 | 5264972 | 1880352 | 10364697 | 7358 | 0 | 27092787 | 65824 | 704 | 11 | 8593 | 0 | 443481 | 0 | 576 | 0 | 17952 | 136349 | 25151 | 3600 | 8182 | 4802 | 2643 | 0 | 166 | 4673 | 0 | 9214 |  | Unknown
WO-22270-106-A-1 | 0 | 13257 | 2867 | 1176 | 16756 | 2007 | 5150078 | 34 | 494 | 0 | 470 | 2591 | 0 | 20023 | 35669 | 164080 | 1051736 | 169669 | 693100 | 2492907 | 757298 | 1671 | 23012653 | 5140291 | 5855602 | 2110216 | 11728390 | 9382 | 186976 | 28027075 | 83047 | 2063 | 252 | 10414 | 0 | 412716 | 0 | 1404 | 12995 | 25480 | 147826 | 19912 | 6687 | 6750 | 7557 | 2991 | 3305 | 111 | 3315 | 12213 | 10647 |  | Unknown
WO-22270-107-A-1 | 503 | 8028 | 500 | 114 | 7274 | 892 | 5234659 | 7 | 87 | 1766 | 1876 | 2178 | 0 | 16985 | 42447 | 167909 | 1022189 | 221668 | 776790 | 2807323 | 868187 | 117322 | 24780276 | 5387118 | 6231741 | 2249895 | 12348570 | 12559 | 0 | 31670741 | 96248 | 2055 | 0 | 11158 | 0 | 393938 | 0 | 1758 | 15532 | 105019 | 125386 | 14272 | 4852 | 7556 | 4448 | 2782 | 0 | 62 | 4739 | 0 | 9548 |  | Unknown
WO-22270-108-A-1 | 1749 | 11423 | 901 | 1261 | 5788 | 746 | 4990634 | 10 | 132 | 2182 | 739 | 2000 | 0 | 5400 | 39607 | 172729 | 1279879 | 184261 | 738575 | 2679015 | 789601 | 102087 | 22216407 | 5568319 | 6080203 | 2179844 | 12446477 | 10918 | 0 | 29316635 | 86995 | 0 | 69 | 8935 | 0 | 361029 | 0 | 1566 | 19250 | 91402 | 129729 | 13852 | 1881 | 8002 | 5237 | 3020 | 0 | 166 | 5626 | 0 | 12194 |  | Unknown
WO-22270-109-A-1 | 596 | 9235 | 565 | 136 | 6666 | 1139 | 4391587 | 28 | 319 | 1376 | 1197 | 1840 | 0 | 3990 | 39024 | 146405 | 772287 | 176680 | 691625 | 2452194 | 714474 | 84760 | 23125269 | 4826802 | 5583680 | 2009241 | 11210476 | 9062 | 0 | 26090607 | 83904 | 1019 | 229 | 9574 | 0 | 433820 | 0 | 1279 | 0 | 123250 | 124799 | 55680 | 4812 | 8920 | 5824 | 1775 | 0 | 185 | 6523 | 0 | 12767 |  | Unknown
WO-22270-110-A-1 | 0 | 9725 | 697 | 29 | 5268 | 812 | 3766804 | 51 | 132 | 1680 | 237 | 2323 | 0 | 4289 | 37678 | 131146 | 604494 | 167364 | 639255 | 2422685 | 717074 | 89531 | 22851904 | 4833379 | 5485504 | 2044642 | 10807877 | 9122 | 1591 | 28846263 | 76826 | 1425 | 400 | 12332 | 0 | 396788 | 0 | 1368 | 0 | 120661 | 130326 | 50700 | 3235 | 30208 | 4335 | 1275 | 0 | 458 | 5800 | 1145 | 12901 |  | Unknown
AB-220715-01 LRB 12_INJ001 | 486 | 9746 | 610 | 319 | 8604 | 787 | 2783402 | 77 | 40 | 2183 | 804 | 2673 | 0 | 16940 | 104101 | 282971 | 367593 | 648203 | 1718292 | 6322451 | 2598668 | 40953 | 27455193 | 5781197 | 13146927 | 2596459 | 17583313 | 55209 | 1124 | 34292524 | 289548 | 1754 | 40 | 10418 | 0 | 17547 | 0 | 369 | 0 | 0 | 136448 | 7426 | 5456 | 14392 | 8259 | 2823 | 0 | 134 | 6755 | 11049 | 10183 |  | Unknown
AB-220715-01 LRB 12 | 0 | 10332 | 395 | 212 | 9696 | 420 | 2494929 | 242 | 103 | 1590 | 1744 | 2208 | 0 | 6887 | 105914 | 270566 | 341357 | 653605 | 1613514 | 6247041 | 2474200 | 514414 | 24969885 | 5143456 | 12028132 | 2430888 | 15855749 | 24568 | 582 | 31796499 | 272721 | 1482 | 502 | 9195 | 0 | 17519 | 0 | 677 | 0 | 0 | 96961 | 5700 | 4142 | 13572 | 4227 | 1702 | 2049 | 45 | 7824 | 1122 | 10192 |  | Unknown
WO-22270-111-A-1 | 826 | 7144 | 382 | 31 | 8150 | 1530 | 3227456 | 0 | 107 | 834 | 728 | 2494 | 0 | 2783 | 30058 | 121368 | 406464 | 159741 | 607635 | 2185624 | 629552 | 71022 | 20700911 | 4402708 | 4971085 | 1801407 | 9759566 | 8953 | 653 | 30426687 | 72226 | 1221 | 51 | 9791 | 0 | 323831 | 0 | 708 | 0 | 214478 | 102581 | 11546 | 3104 | 11205 | 7150 | 1090 | 0 | 648 | 6126 | 0 | 10521 |  | Unknown
WO-22270-112-A-1 | 855 | 11443 | 1212 | 265 | 5052 | 1092 | 4799181 | 0 | 80 | 1706 | 917 | 1860 | 0 | 6653 | 35402 | 145063 | 1023248 | 168350 | 659295 | 2444290 | 724965 | 86105 | 22186783 | 5029725 | 5706817 | 2029749 | 11568769 | 8481 | 0 | 26945968 | 82708 | 2022 | 356 | 10316 | 0 | 359615 | 0 | 1381 | 0 | 247544 | 133925 | 31902 | 4175 | 9679 | 6479 | 1682 | 0 | 57 | 4926 | 0 | 11627 |  | Unknown
WO-22270-113-A-1 | 888 | 8638 | 812 | 0 | 12296 | 639 | 3115066 | 0 | 14 | 1234 | 1095 | 2019 | 0 | 22242 | 42574 | 141736 | 373101 | 187370 | 671160 | 2541811 | 828045 | 88615 | 21945106 | 4602464 | 5569755 | 1882586 | 10515390 | 10972 | 0 | 31623082 | 88396 | 746 | 0 | 9017 | 0 | 340596 | 0 | 1140 | 3195 | 369729 | 120342 | 6166 | 5220 | 18946 | 5345 | 1302 | 0 | 168 | 7099 | 1879 | 10675 |  | Unknown
WO-22270-114-A-1 | 1733 | 8496 | 725 | 182 | 8207 | 1193 | 3725906 | 0 | 16 | 0 | 664 | 1934 | 0 | 18438 | 26781 | 118616 | 660687 | 133810 | 588942 | 1973235 | 509911 | 4079 | 20575282 | 4140382 | 4792766 | 1760651 | 9550675 | 7648 | 0 | 28763999 | 58578 | 1870 | 28 | 10289 | 0 | 386865 | 0 | 1091 | 7104 | 793736 | 122090 | 14087 | 4076 | 27204 | 3982 | 1342 | 0 | 214 | 5251 | 6787 | 13248 |  | Unknown
WO-22270-115-A-1 | 0 | 9971 | 1667 | 352 | 7114 | 614 | 4214793 | 0 | 88 | 2335 | 795 | 2707 | 0 | 6824 | 41002 | 153119 | 931989 | 186150 | 684544 | 2544702 | 760625 | 70308 | 21189412 | 5084282 | 6071450 | 2120464 | 11883155 | 10171 | 0 | 28615700 | 86584 | 3279 | 85 | 9351 | 0 | 420953 | 0 | 775 | 0 | 804229 | 130067 | 134021 | 2518 | 53993 | 7007 | 1680 | 0 | 393 | 4618 | 0 | 14157 |  | Unknown
WO-22270-116-A-1 | 370 | 13573 | 2629 | 574 | 10181 | 1447 | 5118525 | 37 | 58 | 0 | 2244 | 3800 | 0 | 4800 | 36663 | 142183 | 1077206 | 203753 | 705787 | 2678440 | 764118 | 5131 | 23720312 | 5224879 | 5794834 | 2244457 | 12331419 | 10969 | 158199 | 30307098 | 83753 | 1713 | 0 | 9218 | 0 | 384390 | 0 | 1318 | 0 | 872699 | 142850 | 13338 | 4697 | 35997 | 8773 | 3451 | 0 | 184 | 4594 | 14547 | 20536 |  | Unknown
WO-22270-117-A-1 | 0 | 6647 | 815 | 17 | 6447 | 853 | 2969305 | 0 | 0 | 1754 | 1385 | 3053 | 0 | 3559 | 37591 | 124895 | 312576 | 179288 | 633138 | 2334350 | 724122 | 65549 | 18305780 | 4357596 | 5166087 | 1775753 | 9898041 | 9589 | 671 | 29085631 | 77324 | 1934 | 56 | 9905 | 0 | 2358530 | 0 | 1825 | 3755 | 4947 | 147306 | 11813 | 2821 | 20667 | 6984 | 2667 | 0 | 482 | 6538 | 1828 | 18040 |  | Unknown
WO-22270-118-A-1 | 1004 | 9648 | 1868 | 107 | 5639 | 0 | 4617915 | 42 | 100 | 0 | 717 | 3359 | 0 | 3273 | 36708 | 131000 | 852029 | 172905 | 660948 | 2484436 | 695421 | 0 | 22323639 | 4580744 | 5271343 | 1922212 | 10991257 | 9560 | 0 | 29292081 | 75347 | 1074 | 0 | 8403 | 0 | 279353 | 0 | 608 | 0 | 4750 | 134924 | 8997 | 5079 | 9643 | 7538 | 1675 | 0 | 280 | 7569 | 16117 | 14987 |  | Unknown
WO-22270-119-A-1 | 794 | 10156 | 511 | 441 | 11946 | 8467 | 3300312 | 11 | 78 | 0 | 978 | 3769 | 0 | 20063 | 25942 | 102191 | 467353 | 130696 | 579807 | 2031972 | 526616 | 0 | 18653659 | 3685317 | 4337529 | 1501090 | 8736161 | 6734 | 345 | 26572196 | 57455 | 703 | 0 | 9167 | 0 | 340759 | 0 | 994 | 0 | 2934 | 130554 | 19007 | 3004 | 30699 | 6715 | 2097 | 4432 | 382 | 5669 | 17897 | 11313 |  | Unknown
WO-22270-120-A-1 | 0 | 13425 | 464 | 91 | 4221 | 1119 | 3566552 | 42 | 0 | 0 | 559 | 2342 | 0 | 4412 | 29187 | 124746 | 475819 | 164834 | 625207 | 2233921 | 700236 | 0 | 19981202 | 4411901 | 5097827 | 1830436 | 10252495 | 9573 | 0 | 28645941 | 71437 | 902 | 97 | 10452 | 0 | 374246 | 0 | 904 | 0 | 3697 | 132565 | 32982 | 4884 | 38221 | 6052 | 1370 | 0 | 216 | 7732 | 13548 | 16714 |  | Unknown
AB-220715-01 LFB-H 7 | 978 | 93569 | 157739 | 217365 | 43323859 | 28268 | 2467414 | 1147400 | 4025961 | 340943 | 595780 | 2084872 | 1321371 | 1980230 | 83936 | 162087 | 237570 | 588835 | 1348651 | 5292971 | 2320535 | 9920 | 19325253 | 4954955 | 10225391 | 2121381 | 14186538 | 52308 | 0 | 25294292 | 271714 | 1529659 | 244963 | 1975252 | 43462 | 17289254 | 0 | 2076662 | 4915900 | 2124906 | 176266 | 12446934 | 2077126 | 9278982 | 37898 | 4972102 | 457899 | 1184998 | 6762 | 36946697 | 22834740 |  | Unknown
AB-220715-01 LRB 13_INJ001 | 0 | 7883 | 262 | 236 | 60667 | 1145 | 2710589 | 654 | 4389 | 2242 | 856 | 4312 | 0 | 10045 | 108854 | 261738 | 331664 | 677789 | 1653649 | 6162095 | 2577659 | 9296 | 25459165 | 5377769 | 12666893 | 2487994 | 16585527 | 51800 | 0 | 34450912 | 298463 | 1951 | 45 | 12011 | 0 | 38896 | 0 | 3080 | 0 | 5229 | 141037 | 15298 | 7838 | 17956 | 7136 | 11246 | 0 | 393 | 7257 | 52089 | 41995 |  | Unknown
AB-220715-01 LRB 13 | 0 | 8739 | 626 | 0 | 16337 | 662 | 1995875 | 291 | 1662 | 1468 | 1292 | 882 | 0 | 4832 | 85307 | 223761 | 277068 | 602438 | 1417471 | 5592209 | 2317303 | 87616 | 19935037 | 4161700 | 10266887 | 1896730 | 13284930 | 18202 | 82355 | 27098095 | 246703 | 1157 | 18 | 8156 | 0 | 16657 | 0 | 426 | 0 | 2774 | 80629 | 6713 | 4101 | 15535 | 3461 | 4517 | 1171 | 474 | 5304 | 18300 | 15883 |  | Unknown
AB-220715-01 LD 6 | 0 | 7780 | 1540 | 887 | 12794 | 3420775 | 3272663 | 20 | 352 | 0 | 1000 | 1598 | 0 | 2676 | 25592 | 102465 | 459732 | 121477 | 520854 | 1889293 | 527612 | 1208 | 18592197 | 3523377 | 4320638 | 1486179 | 8631675 | 6918 | 33089542 | 26544965 | 56706 | 1197 | 571 | 8339 | 0 | 320260 | 0 | 596 | 19181 | 3223 | 152423 | 22188 | 6443 | 30719 | 9102 | 2031 | 0 | 545 | 4664 | 19691 | 12960 |  | Unknown
AB-220715-01 CCC-H 7 | 0 | 192409 | 273859 | 392149 | 84780714 | 208258 | 2422180 | 2368855 | 7566830 | 728256 | 1190381 | 3794292 | 2440505 | 3685206 | 98591 | 143374 | 210135 | 599962 | 1272681 | 5150143 | 2451488 | 13192 | 16990268 | 4860963 | 9574526 | 1962335 | 13051808 | 57256 | 24852574 | 22383436 | 285666 | 2927674 | 699532 | 3800535 | 93615 | 29662858 | 0 | 3540329 | 9415086 | 4391961 | 240664 | 22511073 | 3598328 | 17477003 | 78024 | 9355863 | 990913 | 2402540 | 5193 | 69204424 | 38991038 |  | Unknown
AB-220715-01 LFSM-H 6 | 0 | 29203 | 88546 | 363425 | 40294448 | 2174 | 3446732 | 265530 | 1552368 | 91104 | 151097 | 990717 | 498624 | 1446391 | 28785 | 88168 | 362227 | 143272 | 547779 | 2057003 | 623203 | 1340 | 17823352 | 3875227 | 4488239 | 1676334 | 9111106 | 7883 | 0 | 23802449 | 67886 | 625660 | 283616 | 2699715 | 0 | 16790034 | 0 | 2213197 | 1964343 | 562460 | 177466 | 7982366 | 2345309 | 4230287 | 42973 | 3905607 | 69205 | 420087 | 4955 | 34499758 | 21692791 |  | Unknown
WO-22270-121-A-1 | 0 | 10243 | 527 | 125 | 18186 | 0 | 3346866 | 38 | 1550 | 1010 | 680 | 3253 | 0 | 8165 | 39530 | 130133 | 439426 | 224847 | 663452 | 2607042 | 872626 | 6491 | 18940059 | 4276007 | 5454144 | 1834652 | 10044035 | 12286 | 1281 | 28037079 | 94550 | 1226 | 374 | 9850 | 0 | 310527 | 0 | 1568 | 0 | 8873 | 142331 | 7412 | 6196 | 0 | 5746 | 3804 | 1306 | 55 | 4516 | 4736 | 18785 |  | Unknown
WO-22270-122-A-1 | 0 | 8578 | 988 | 17 | 8019 | 395 | 1962344 | 10 | 1166 | 0 | 894 | 1998 | 0 | 2232 | 17755 | 80274 | 221875 | 76465 | 432530 | 1402745 | 341178 | 0 | 13693187 | 2599912 | 3497743 | 1091594 | 6324605 | 5276 | 10858558 | 22223242 | 37659 | 847 | 192 | 7371 | 0 | 234144 | 0 | 1279 | 0 | 5618 | 123086 | 13440 | 7040 | 15085 | 5220 | 1511 | 0 | 128 | 5410 | 20522 | 11932 |  | Unknown
WO-22270-123-A-1 | 0 | 8318 | 2228 | 517 | 10452 | 614616 | 4742762 | 0 | 536 | 0 | 465 | 4764 | 0 | 26394 | 32527 | 128336 | 1031737 | 187488 | 585518 | 2343953 | 716815 | 0 | 21031285 | 4403054 | 5333152 | 1912900 | 11064215 | 9628 | 25733920 | 27745914 | 73169 | 1309 | 456 | 9267 | 0 | 381921 | 0 | 633 | 0 | 9823 | 130187 | 10571 | 6036 | 6673 | 8039 | 3599 | 0 | 94 | 3786 | 22525 | 13481 | D,Q | Unknown
WO-22270-124-A-1 | 0 | 12360 | 1735 | 243 | 8863 | 3418 | 4502979 | 0 | 468 | 0 | 774 | 2052 | 0 | 5754 | 38426 | 143232 | 847099 | 194237 | 698798 | 2698175 | 796284 | 3584 | 20479034 | 4615263 | 5699289 | 1990011 | 11270126 | 11548 | 0 | 29015502 | 89594 | 968 | 0 | 8455 | 0 | 415043 | 0 | 1152 | 0 | 31180 | 152101 | 17305 | 6374 | 81684 | 10025 | 1872 | 0 | 357 | 4677 | 20952 | 12617 |  | Unknown
WO-22270-125-A-1 | 1225 | 13423 | 7145 | 532 | 20382 | 4291 | 5032037 | 14 | 146 | 0 | 1854 | 3524 | 0 | 2895 | 39321 | 281848 | 1297419 | 188027 | 675536 | 2529167 | 780730 | 3257 | 19296311 | 5132345 | 6037692 | 1941943 | 10993672 | 12097 | 5478569 | 26748912 | 76656 | 2872 | 420 | 8501 | 0 | 429275 | 0 | 933 | 0 | 27801 | 120977 | 33743 | 7287 | 32679 | 6943 | 3805 | 1930 | 136 | 4899 | 26663 | 29774 |  | Unknown
WO-22270-126-A-1 | 0 | 6507 | 2567 | 81 | 8040 | 661 | 3806806 | 0 | 53 | 0 | 1051 | 3399 | 0 | 2943 | 26233 | 122073 | 636822 | 117206 | 531254 | 1943876 | 528099 | 751 | 18499001 | 3614134 | 4521694 | 1525797 | 8799705 | 7859 | 982148 | 27128926 | 54211 | 4482 | 642 | 6918 | 0 | 324594 | 0 | 790 | 0 | 20550 | 127694 | 25164 | 5251 | 37246 | 7510 | 1829 | 0 | 348 | 6551 | 21066 | 12203 |  | Unknown
WO-22270-127-A-1 | 0 | 5302 | 1034 | 458 | 4998 | 1442 | 2750708 | 0 | 438 | 0 | 763 | 1819 | 0 | 2359 | 30992 | 113753 | 304044 | 145717 | 589442 | 2298372 | 625802 | 3620 | 17297313 | 3645695 | 4825909 | 1437646 | 8553778 | 10110 | 584 | 27287211 | 70416 | 1775 | 59 | 10012 | 0 | 298617 | 0 | 981 | 0 | 81746 | 116200 | 9339 | 6266 | 55508 | 7689 | 1341 | 0 | 81 | 5379 | 19944 | 9109 |  | Unknown
WO-22270-128-A-1 | 0 | 6631 | 1055 | 100 | 6814 | 378 | 3121770 | 0 | 33 | 0 | 1831 | 3002 | 0 | 8992 | 28550 | 111707 | 393039 | 145476 | 602592 | 2268452 | 626511 | 2659 | 19982121 | 3898364 | 4834988 | 1634387 | 9499258 | 8183 | 0 | 28858281 | 63632 | 1919 | 58 | 8495 | 0 | 308845 | 0 | 846 | 0 | 98720 | 128830 | 11881 | 4655 | 29249 | 7461 | 1709 | 0 | 430 | 7588 | 18071 | 11308 |  | Unknown
WO-22270-129-A-1 | 0 | 9793 | 2037 | 106 | 5890 | 891 | 4363666 | 0 | 105 | 867 | 463 | 1091 | 0 | 4164 | 35759 | 129362 | 827472 | 187501 | 625168 | 2399759 | 745632 | 23632 | 23228761 | 4588869 | 5481662 | 1991819 | 11240230 | 9689 | 0 | 30378953 | 77786 | 2421 | 63 | 11657 | 0 | 431246 | 0 | 846 | 31601 | 133282 | 141067 | 30092 | 816 | 62780 | 7855 | 1653 | 0 | 238 | 4870 | 0 | 11223 |  | Unknown
WO-22270-130-A-1 | 0 | 14532 | 2338 | 495 | 0 | 4459 | 3950125 | 0 | 0 | 0 | 565 | 1880 | 0 | 1705 | 18842 | 114440 | 764223 | 91354 | 470343 | 1717563 | 416689 | 1154 | 17402469 | 3322756 | 4352324 | 1412516 | 8791448 | 7022 | 0 | 24236317 | 50714 | 1636 | 40 | 8402 | 0 | 460842 | 0 | 892 | 21686 | 77218 | 138307 | 50982 | 3982 | 430240 | 9130 | 1401 | 0 | 248 | 6006 | 16405 | 12223 |  | Unknown
AB-220715-01 LRB 14_INJ001 | 0 | 7677 | 1702 | 101 | 5562 | 468 | 2482971 | 35 | 136 | 1840 | 1215 | 2326 | 0 | 10254 | 88950 | 234898 | 322193 | 605013 | 1479328 | 5808207 | 2390547 | 5671 | 24648982 | 4996868 | 11851395 | 2218538 | 15909323 | 49023 | 37095111 | 32739726 | 265948 | 1289 | 379 | 8951 | 0 | 18502 | 0 | 338 | 3158 | 4604 | 143390 | 5670 | 3621 | 17286 | 9071 | 1895 | 0 | 120 | 5147 | 19384 | 15288 |  | Unknown
AB-220715-01 LRB 14 | 421 | 7296 | 1730 | 84 | 10567 | 1837976 | 1845814 | 0 | 97 | 1476 | 784 | 2455 | 0 | 8467 | 64644 | 199445 | 249140 | 439049 | 1164324 | 4588548 | 1665274 | 0 | 19358553 | 3749485 | 9277931 | 1765251 | 12816098 | 17287 | 30284513 | 25021746 | 184050 | 1485 | 171 | 8083 | 0 | 16077 | 0 | 464 | 0 | 0 | 83227 | 5603 | 3490 | 12316 | 3280 | 1671 | 0 | 315 | 5124 | 17433 | 8296 | D,Q | Unknown
WO-22270-131-A-1 | 485 | 40760 | 1080 | 346 | 4312 | 1487 | 3345422 | 0 | 85 | 0 | 574 | 2375 | 0 | 1047 | 23513 | 101404 | 531293 | 123781 | 550218 | 2033265 | 517798 | 0 | 17528152 | 3573179 | 4487682 | 1494266 | 9084723 | 7442 | 31167389 | 24998725 | 56755 | 1881 | 33 | 9544 | 0 | 273071 | 0 | 1137 | 8153 | 350150 | 110637 | 22620 | 3988 | 24909 | 6627 | 1570 | 609 | 129 | 5109 | 17801 | 11775 | D,Q | Unknown
WO-22270-132-A-1 | 0 | 11594 | 2742 | 1215 | 5351 | 4213661 | 5291859 | 0 | 0 | 0 | 1264 | 1618 | 0 | 0 | 34739 | 173513 | 1297582 | 180112 | 653695 | 2637207 | 766261 | 0 | 21597146 | 4614516 | 5903689 | 1980270 | 11965642 | 10844 | 31592473 | 27560248 | 82526 | 1637 | 207 | 9130 | 0 | 348548 | 0 | 732 | 0 | 410810 | 127795 | 25998 | 925 | 95039 | 6827 | 1057 | 0 | 67 | 3264 | 16765 | 14274 | D,Q | Unknown
WO-22270-133-A-1 | 0 | 8044 | 1438 | 747 | 4184 | 2328089 | 4036385 | 0 | 37 | 0 | 668 | 4609 | 0 | 23418 | 24781 | 109917 | 772092 | 120387 | 538697 | 2093334 | 537745 | 589 | 18901668 | 3804378 | 4592238 | 1634970 | 9484061 | 7134 | 30951840 | 26388231 | 53059 | 5017 | 28 | 10086 | 0 | 389545 | 0 | 775 | 13163 | 213496 | 136200 | 82659 | 757 | 54975 | 9798 | 1184 | 0 | 447 | 5119 | 16608 | 14105 |  | Unknown
WO-22270-134-A-1 | 634 | 22076 | 2585 | 1543 | 3774 | 2350021 | 4630474 | 34 | 264 | 0 | 1372 | 3914 | 0 | 7484 | 29312 | 121983 | 978151 | 126845 | 566135 | 2118182 | 503867 | 815 | 19266677 | 3860438 | 5122490 | 1689109 | 9798107 | 8036 | 28298022 | 25782986 | 62531 | 5307 | 0 | 7402 | 0 | 428910 | 0 | 347 | 50138 | 5370496 | 108394 | 138525 | 3852 | 45655 | 6028 | 2157 | 7752 | 136 | 5626 | 12767 | 27645 |  | Unknown
WO-22270-135-A-1 | 0 | 10932 | 962 | 352 | 4920 | 3978846 | 3779774 | 0 | 229 | 0 | 417 | 2636 | 0 | 19849 | 21293 | 104846 | 646508 | 99331 | 494603 | 1792899 | 420742 | 728 | 18105377 | 3172844 | 4251316 | 1445158 | 8965474 | 6475 | 29960619 | 25520988 | 49385 | 3615 | 786 | 7124 | 0 | 300841 | 0 | 364 | 11290 | 445424 | 127293 | 18616 | 5370 | 6418 | 6970 | 1038 | 0 | 118 | 4349 | 14622 | 9140 |  | Unknown
WO-22270-136-A-1 | 25040 | 159225 | 4582 | 871 | 10941 | 6469830 | 6284887 | 0 | 144 | 1455 | 894 | 1514 | 0 | 3679 | 21297 | 122537 | 1290263 | 121399 | 516680 | 2079913 | 591592 | 0 | 22119472 | 4962161 | 5855521 | 2165432 | 13194586 | 3981 | 34135354 | 26642322 | 55484 | 2440 | 679 | 8280 | 0 | 410301 | 126530 | 1115 | 34152 | 144584 | 135676 | 31567 | 1222 | 44053 | 4545 | 1311 | 5620 | 0 | 360776 | 488809 | 15690 |  | Unknown
WO-22270-137-A-1 | 10686 | 37878 | 484 | 64 | 6043 | 2293590 | 1738739 | 0 | 0 | 5182 | 718 | 2710 | 0 | 2740 | 28521 | 99835 | 222982 | 136130 | 477072 | 2235862 | 659323 | 600 | 14943000 | 2940610 | 4941408 | 1297714 | 9263209 | 6665 | 8797599 | 12713582 | 66307 | 1126 | 647 | 2164 | 0 | 194761 | 46031 | 440 | 0 | 2307 | 27028 | 8562 | 422 | 18969 | 1976 | 1168 | 0 | 63 | 831424 | 402768 | 14273 |  | Unknown
WO-22270-138-A-1 | 2198 | 80680 | 952 | 263 | 3486 | 2776532 | 2726384 | 0 | 0 | 6052 | 487 | 2449 | 0 | 2247 | 20592 | 87499 | 475667 | 101071 | 416637 | 1801389 | 520661 | 596 | 13390151 | 3202624 | 4391642 | 1403595 | 9369027 | 3857 | 4500446 | 8289063 | 53890 | 4874 | 240 | 608 | 0 | 131743 | 0 | 279 | 15538 | 0 | 3308 | 5231 | 1541 | 5138 | 933 | 1269 | 4614 | 333 | 14032 | 10631 | 5081 |  | Unknown
WO-22270-139-A-1 | 4065 | 93691 | 780 | 57 | 2539 | 2607462 | 2055113 | 0 | 18 | 3132 | 295 | 3929 | 0 | 1518 | 12308 | 66938 | 254224 | 53525 | 350065 | 1293108 | 323209 | 1096 | 14070998 | 2331221 | 3271597 | 1022183 | 6537586 | 3358 | 25647005 | 22168034 | 30142 | 1513 | 75 | 5418 | 0 | 311132 | 17330 | 284 | 0 | 0 | 13302 | 8700 | 325 | 5471 | 1687 | 547 | 3800 | 28 | 27797 | 22856 | 9140 |  | Unknown
WO-22270-140-A-1 | 5175 | 115120 | 2149 | 228 | 3229 | 2810226 | 2695366 | 31 | 0 | 3967 | 649 | 1862 | 0 | 2735 | 25012 | 99852 | 462866 | 104719 | 463804 | 1782722 | 495760 | 670 | 17201517 | 3206086 | 4299628 | 1339615 | 8340094 | 6284 | 29305932 | 23883446 | 53287 | 2465 | 121 | 7305 | 0 | 303248 | 13875 | 257 | 0 | 0 | 53824 | 11949 | 1068 | 6763 | 1750 | 436 | 0 | 183 | 30156 | 23463 | 24759 |  | Unknown
AB-220715-01 LFB-H 8 | 41960077 | 82761 | 133600 | 238536 | 45947449 | 1867084 | 2428242 | 967091 | 3234682 | 290850 | 522117 | 1897952 | 1117295 | 1827269 | 84195 | 141629 | 255958 | 523805 | 1162800 | 4783389 | 2042640 | 0 | 20685732 | 4351996 | 8931373 | 2019840 | 12567876 | 52700 | 34422581 | 25511094 | 240908 | 1353206 | 1520412 | 1884393 | 46394 | 17740796 | 31817093 | 2082675 | 4399065 | 1849319 | 117100 | 11684793 | 2334818 | 8448635 | 10761 | 4367669 | 404436 | 1002477 | 27019269 | 32079761 | 21801379 |  | Unknown
AB-220715-01 LRB 15_INJ001 | 505691 | 47643 | 979 | 1837 | 237603 | 2389937 | 2663934 | 2683 | 13403 | 2249 | 1764 | 8172 | 4872 | 13351 | 97640 | 243868 | 360597 | 524588 | 1528681 | 5858095 | 2322692 | 0 | 25839013 | 4906561 | 12555336 | 2366627 | 15948317 | 33463 | 37704073 | 33275338 | 231321 | 5099 | 10872 | 10502 | 0 | 153971 | 444900 | 19790 | 30517 | 10325 | 106091 | 56189 | 14402 | 66864 | 7404 | 33092 | 0 | 3634 | 191832 | 344056 | 273704 | C | Unknown
AB-220715-01 LRB 15 | 240320 | 52125 | 1290 | 302 | 28497 | 2340833 | 2477558 | 0 | 723 | 2518 | 987 | 1391 | 0 | 8617 | 95789 | 234891 | 344095 | 525317 | 1521114 | 5984993 | 2179518 | 0 | 26677815 | 4618474 | 11929136 | 2262813 | 15116895 | 23381 | 36337608 | 32216398 | 251665 | 4047 | 2380 | 7605 | 0 | 37245 | 302707 | 7321 | 0 | 3144 | 101504 | 9024 | 1776 | 33914 | 4912 | 7959 | 0 | 231 | 168123 | 300571 | 99026 | D,Q | Unknown
AB-220715-01 LD 7 | 61927 | 128092 | 468 | 216 | 8271 | 2418492 | 2304726 | 11 | 211 | 2414 | 523 | 1267 | 0 | 4090 | 20498 | 90500 | 441457 | 77348 | 442798 | 1586895 | 370515 | 2199 | 14886295 | 2544183 | 3675123 | 1199582 | 7328787 | 6424 | 26427335 | 21641575 | 45449 | 3716 | 768 | 5148 | 0 | 331890 | 128981 | 1667 | 0 | 0 | 52477 | 13544 | 2031 | 22048 | 1045 | 2019 | 14443 | 141 | 107814 | 159126 | 24505 |  | Unknown
AB-220715-01 CCC-H 8 | 71551754 | 163649 | 238936 | 420933 | 80292227 | 1473347 | 2186247 | 1889168 | 5703583 | 603501 | 974031 | 3258977 | 2116612 | 3204232 | 84399 | 121635 | 206957 | 489858 | 1056800 | 4246032 | 2077923 | 0 | 15691752 | 3501116 | 7967798 | 1667999 | 11021968 | 41276 | 28511691 | 19334403 | 237285 | 2355443 | 2505872 | 3395912 | 69395 | 27614293 | 49644592 | 3085030 | 8041984 | 3749471 | 81433 | 18747354 | 3809028 | 14080172 | 7415 | 6959968 | 771694 | 1862004 | 44127122 | 55259551 | 34201767 |  | Unknown
AB-220715-01 LFSM-H 7 | 37105095 | 139840 | 87836 | 344174 | 33535600 | 2004697 | 2618416 | 221558 | 1324591 | 87769 | 127874 | 954060 | 475136 | 1251619 | 29557 | 86215 | 373138 | 123238 | 494490 | 1961354 | 621772 | 0 | 15804340 | 3113149 | 4360705 | 1414281 | 8084466 | 8678 | 27019135 | 21492600 | 60200 | 589238 | 1789709 | 2042724 | 0 | 14956285 | 27153692 | 1786605 | 1857875 | 557035 | 34595 | 7131542 | 1967561 | 4027523 | 1711 | 3201341 | 62369 | 370014 | 17490655 | 25555270 | 19703675 |  | Unknown
WO-22270-141-A-1 | 575929 | 84001 | 1121 | 1537 | 85465 | 2579865 | 2489714 | 552 | 3539 | 16313 | 889 | 3773 | 0 | 5574 | 35756 | 122088 | 457927 | 145861 | 573003 | 2303238 | 683065 | 699 | 19782291 | 3631663 | 4969822 | 1717122 | 9793481 | 8651 | 32647379 | 26500673 | 71295 | 1989 | 11584 | 8340 | 0 | 280697 | 719559 | 24313 | 7759 | 3053 | 40801 | 24917 | 12133 | 22293 | 944 | 10752 | 28208 | 1687 | 390343 | 716273 | 240809 |  | Unknown
WO-22270-142-A-1 | 219943 | 74990 | 1195 | 295 | 32220 | 2568325 | 2482730 | 0 | 583 | 14348 | 719 | 2843 | 0 | 3797 | 30746 | 113243 | 433527 | 125794 | 502274 | 2037922 | 546595 | 0 | 19080449 | 3402434 | 4796678 | 1508916 | 9330821 | 6102 | 32489542 | 27207567 | 63843 | 1541 | 2400 | 7089 | 0 | 215670 | 530483 | 5330 | 0 | 0 | 58831 | 10243 | 5836 | 19308 | 895 | 4160 | 487 | 139 | 433071 | 934919 | 61093 | D,Q | Unknown
WO-22270-143-A-1 | 36101 | 110233 | 878 | 117 | 28103 | 1961016 | 1827300 | 0 | 479 | 5707 | 556 | 4391 | 0 | 2301 | 15423 | 87021 | 300838 | 61997 | 362414 | 1301505 | 297470 | 853 | 14059759 | 2311155 | 3388263 | 1086537 | 6503170 | 5425 | 25268723 | 20352708 | 31973 | 1938 | 602 | 4663 | 0 | 223461 | 118065 | 650 | 0 | 0 | 19964 | 10710 | 7495 | 5156 | 745 | 2185 | 1922 | 89 | 145691 | 229957 | 11740 |  | Unknown
WO-22270-144-A-1 | 50150 | 130273 | 2711 | 239 | 9371 | 3326905 | 3257402 | 19 | 440 | 47864 | 574 | 2300 | 0 | 4609 | 31948 | 133975 | 786429 | 145314 | 573430 | 2245959 | 686914 | 0 | 20363449 | 3862700 | 5437062 | 1655493 | 10268993 | 10473 | 29628833 | 26788780 | 70659 | 6624 | 795 | 7820 | 0 | 289769 | 200162 | 1394 | 50315 | 0 | 22252 | 27127 | 7572 | 38297 | 845 | 2489 | 0 | 362 | 482097 | 718507 | 18652 |  | Unknown
WO-22270-145-A-1 | 14646 | 147215 | 1273 | 126 | 15453 | 3202739 | 2718872 | 12 | 87 | 45090 | 165 | 1629 | 0 | 2070 | 29870 | 117916 | 496409 | 143294 | 547070 | 2173396 | 643791 | 562 | 21179565 | 3903414 | 4886714 | 1600079 | 9773329 | 10870 | 32218106 | 26419173 | 77164 | 2952 | 442 | 7175 | 0 | 243683 | 62290 | 688 | 21613 | 0 | 24955 | 12445 | 6230 | 24354 | 872 | 2015 | 28201 | 0 | 190291 | 255661 | 15720 |  | Unknown
WO-22270-146-A-1 | 12313 | 71512 | 352 | 29 | 6999 | 2443134 | 2130717 | 0 | 10 | 29507 | 208 | 2446 | 0 | 2461 | 24527 | 92427 | 372452 | 90916 | 433524 | 1722900 | 460155 | 822 | 16305276 | 2881326 | 4034522 | 1225828 | 7903729 | 8906 | 24455743 | 24486686 | 48029 | 5230 | 477 | 6231 | 0 | 155044 | 42607 | 608 | 0 | 0 | 14092 | 12418 | 5948 | 31822 | 504 | 1313 | 33629 | 56 | 190866 | 206992 | 13336 |  | Unknown
WO-22270-147-A-1 | 10509 | 80758 | 694 | 121 | 5021 | 1770695 | 1576669 | 0 | 0 | 58147 | 401 | 1692 | 0 | 5035 | 13461 | 77591 | 270031 | 47744 | 339378 | 1133748 | 246120 | 1267 | 11583296 | 2265564 | 3225824 | 1042155 | 6775621 | 5007 | 7112602 | 9955180 | 27133 | 3999 | 315 | 2703 | 0 | 192780 | 22789 | 454 | 0 | 0 | 1728 | 7539 | 6376 | 36620 | 812 | 1835 | 0 | 83 | 87073 | 117758 | 9943 |  | Unknown
WO-22270-148-A-1 | 13344 | 191064 | 1426 | 870 | 5624 | 3343703 | 3170322 | 0 | 64 | 157426 | 629 | 2882 | 0 | 9240 | 36067 | 127360 | 782452 | 165268 | 569266 | 2418578 | 774954 | 0 | 21295071 | 4031105 | 5704854 | 1659164 | 10865694 | 9707 | 30180524 | 26972536 | 78227 | 2424 | 1534 | 6181 | 0 | 215138 | 41573 | 434 | 31305 | 7606 | 33030 | 14250 | 4970 | 116007 | 835 | 2099 | 0 | 11 | 231521 | 224666 | 15863 | D,Q | Unknown
WO-22270-149-A-1 | 10026 | 110124 | 1233 | 155 | 6379 | 2541509 | 2806006 | 0 | 111 | 150437 | 415 | 2401 | 0 | 3012 | 32501 | 118281 | 422855 | 169894 | 594572 | 2415749 | 737088 | 1191 | 20994098 | 3793380 | 5450588 | 1681935 | 10149306 | 9394 | 34033054 | 28283258 | 72788 | 2591 | 0 | 9045 | 0 | 389076 | 33901 | 822 | 0 | 8006 | 85493 | 19585 | 4112 | 42824 | 1570 | 1230 | 21330 | 175 | 265915 | 180891 | 12164 |  | Unknown
WO-22270-150-A-1 | 9078 | 100230 | 400 | 183 | 5143 | 2488460 | 2323430 | 0 | 18 | 114556 | 574 | 2723 | 0 | 2008 | 28659 | 102217 | 399978 | 123721 | 499943 | 2141597 | 588086 | 602 | 18687443 | 3432329 | 4809651 | 1482777 | 8941046 | 8914 | 32135818 | 27064942 | 64735 | 2303 | 42 | 7619 | 0 | 252746 | 29455 | 623 | 0 | 0 | 73821 | 19612 | 3650 | 6208 | 3045 | 2301 | 0 | 30 | 178534 | 164917 | 15186 |  | Unknown
AB-220715-01 LRB 16_INJ001 | 12725 | 70519 | 628 | 450 | 11012 | 2071304 | 2230968 | 11 | 132 | 3328 | 456 | 1139 | 0 | 4618 | 80865 | 201897 | 317824 | 412519 | 1165125 | 4796231 | 1839962 | 1907 | 24724590 | 4145597 | 10435240 | 2040059 | 14346854 | 31362 | 36647315 | 29843384 | 192514 | 3440 | 95 | 6540 | 0 | 27172 | 48869 | 1133 | 0 | 0 | 107690 | 20526 | 1915 | 28522 | 10151 | 2319 | 4272 | 75 | 258269 | 229153 | 19935 |  | Unknown
AB-220715-01 LRB 16 | 14788 | 58229 | 955 | 589 | 7048 | 1999097 | 2140762 | 0 | 99 | 2890 | 908 | 388 | 0 | 3389 | 80058 | 200163 | 298191 | 398216 | 1207377 | 4825536 | 1841664 | 0 | 23758425 | 4080586 | 10238552 | 2094631 | 13706642 | 28700 | 33160647 | 29602926 | 183473 | 1183 | 658 | 6412 | 0 | 25924 | 42249 | 611 | 0 | 0 | 79166 | 8413 | 2097 | 28349 | 6368 | 1333 | 0 | 149 | 425162 | 291585 | 42717 | D,Q | Unknown
WO-22270-151-A-1 | 7370 | 75554 | 697 | 109 | 8172 | 2042724 | 2119261 | 0 | 0 | 375117 | 290 | 1369 | 0 | 1687 | 27389 | 104109 | 343248 | 115917 | 501542 | 1904176 | 565315 | 0 | 17809073 | 3099270 | 4398704 | 1396366 | 8451782 | 6017 | 31128854 | 24697936 | 57239 | 773 | 158 | 6233 | 0 | 184634 | 25573 | 411 | 0 | 2336 | 82564 | 57405 | 3071 | 7151 | 1970 | 1150 | 0 | 0 | 222136 | 126347 | 14480 | D,Q | Unknown
WO-22270-152-A-1 | 8657 | 81871 | 1563 | 448 | 5309 | 2336261 | 2127264 | 0 | 72 | 441864 | 743 | 1456 | 0 | 14896 | 30326 | 111869 | 380712 | 147061 | 535569 | 2171600 | 658354 | 0 | 19950926 | 3349826 | 4858332 | 1549125 | 9492805 | 9329 | 30673222 | 26557310 | 63729 | 730 | 218 | 6194 | 0 | 229472 | 30181 | 390 | 0 | 3039 | 67866 | 11739 | 3163 | 10813 | 2016 | 1744 | 0 | 0 | 131567 | 64111 | 19700 | D,Q | Unknown
WO-22270-153-A-1 | 7298 | 60621 | 393 | 24 | 5492 | 1701362 | 1564749 | 0 | 0 | 139674 | 367 | 1108 | 0 | 4432 | 13014 | 67719 | 289275 | 43772 | 309984 | 1052335 | 226704 | 907 | 12714269 | 2069753 | 2952713 | 1051780 | 6427292 | 4068 | 8207649 | 14212626 | 22075 | 622 | 163 | 1654 | 0 | 184796 | 18893 | 95 | 0 | 1900 | 4817 | 11665 | 2199 | 12071 | 591 | 915 | 0 | 0 | 89022 | 51438 | 33975 |  | Unknown
WO-22270-154-A-1 | 12664 | 71973 | 260 | 254 | 5691 | 2261118 | 2299737 | 0 | 165 | 898816 | 306 | 853 | 0 | 4798 | 33043 | 114107 | 379144 | 130265 | 527595 | 2077924 | 633498 | 476 | 18892054 | 3274452 | 4786316 | 1516944 | 9308754 | 7703 | 30639980 | 25007841 | 65835 | 757 | 196 | 6746 | 0 | 355975 | 28161 | 517 | 0 | 5154 | 61549 | 219849 | 1100 | 11180 | 1746 | 1138 | 0 | 95 | 159251 | 66842 | 24528 |  | Unknown
WO-22270-155-A-1 | 8327 | 86666 | 1184 | 216 | 6589 | 2757769 | 3129704 | 0 | 109 | 1940827 | 540 | 1937 | 0 | 4566 | 47830 | 231690 | 624293 | 212998 | 1124693 | 3992764 | 1072043 | 0 | 29629816 | 4792469 | 10742136 | 2483686 | 17632076 | 9634 | 31468489 | 31276463 | 92967 | 1012 | 400 | 6309 | 0 | 483522 | 21344 | 223 | 0 | 11617 | 9572 | 28329 | 1943 | 19536 | 1380 | 1559 | 0 | 262 | 58735 | 38304 | 14125 | D,Q | Unknown
WO-22270-156-A-1 | 5650 | 126718 | 780 | 25 | 13526 | 3039166 | 3437986 | 0 | 0 | 865942 | 776 | 1254 | 0 | 5525 | 29471 | 109388 | 788804 | 113207 | 477127 | 1908518 | 545367 | 1945 | 19334524 | 3479976 | 4447850 | 1419354 | 9283976 | 6618 | 30123706 | 25117097 | 48029 | 3510 | 545 | 5543 | 0 | 298517 | 20506 | 36 | 0 | 5468 | 37581 | 13930 | 3376 | 26027 | 1114 | 1448 | 0 | 44 | 151890 | 54341 | 39135 |  | Unknown
WO-22270-157-A-1 | 6313 | 60042 | 309 | 68 | 3363 | 2440098 | 2535751 | 0 | 0 | 7802 | 733 | 1546 | 0 | 4511 | 31705 | 116444 | 390497 | 127295 | 544228 | 2158416 | 593474 | 777 | 19537435 | 3281876 | 5196545 | 1504434 | 8946909 | 7115 | 33235078 | 27039566 | 65854 | 1090 | 255 | 7453 | 0 | 335444 | 21675 | 273 | 25756 | 6657 | 51130 | 177514 | 2018 | 40313 | 565 | 1544 | 0 | 104 | 94032 | 40485 | 40286 |  | Unknown
WO-22270-158-A-1 | 7691 | 122801 | 850 | 498 | 5292 | 2940924 | 2731400 | 0 | 0 | 3226 | 449 | 3011 | 0 | 14211 | 28416 | 112130 | 719429 | 124087 | 496865 | 2017716 | 611183 | 0 | 18063766 | 3312252 | 4676702 | 1445449 | 9566553 | 7034 | 20541693 | 21948470 | 59343 | 2268 | 280 | 5101 | 0 | 166705 | 21515 | 326 | 0 | 1278 | 8449 | 35506 | 1138 | 18772 | 4616 | 874 | 0 | 46 | 131663 | 44244 | 13867 | D,Q | Unknown
WO-22270-159-A-1 | 7627 | 109970 | 862 | 387 | 7261 | 2739398 | 2961268 | 9 | 56 | 3675 | 743 | 2118 | 0 | 5084 | 30087 | 138831 | 495058 | 115004 | 701554 | 2369286 | 591800 | 0 | 25609527 | 4010470 | 7154057 | 2006936 | 13331964 | 5910 | 29263325 | 30028896 | 63918 | 910 | 539 | 5514 | 0 | 402999 | 21513 | 388 | 0 | 0 | 5823 | 29008 | 1661 | 42393 | 569 | 1176 | 0 | 20 | 48453 | 28922 | 15111 |  | Unknown
WO-22270-160-A-1 | 7049 | 75945 | 1596 | 45 | 5199 | 1775674 | 1506043 | 0 | 0 | 1129 | 321 | 1921 | 0 | 14375 | 20030 | 81480 | 266946 | 72027 | 366427 | 1480547 | 378715 | 0 | 12729966 | 2417975 | 3403835 | 1030482 | 6645223 | 5944 | 10546442 | 14907093 | 37585 | 526 | 505 | 2321 | 0 | 181425 | 14552 | 731 | 0 | 0 | 4494 | 11769 | 2329 | 4423 | 691 | 642 | 0 | 0 | 99724 | 26202 | 22332 | D,Q | Unknown
AB-220715-01 LFB-H 9 | 30319282 | 78129 | 122040 | 219772 | 37325000 | 1408186 | 1858966 | 767079 | 2546116 | 251436 | 414913 | 1634635 | 908452 | 1443127 | 79993 | 134892 | 212725 | 437559 | 1000400 | 4073788 | 1841312 | 3694 | 17287027 | 3262654 | 8235689 | 1663172 | 11264146 | 24022 | 25085889 | 20351319 | 211075 | 1242675 | 1259489 | 1434965 | 25157 | 16294504 | 21535008 | 1468070 | 3928616 | 1709479 | 6960 | 10564742 | 1853562 | 7730878 | 790 | 3230802 | 329348 | 791170 | 6392993 | 15621679 | 18182280 |  | Unknown
AB-220715-01 LRB 17_INJ001 | 514617 | 88811 | 1307 | 2207 | 262345 | 2037892 | 2143829 | 2177 | 15255 | 1873 | 1629 | 9907 | 4870 | 11071 | 86018 | 205239 | 306042 | 440566 | 1283922 | 5236737 | 1982393 | 0 | 24594162 | 4027139 | 10618169 | 2018735 | 14436278 | 30471 | 34299972 | 31065468 | 206198 | 5388 | 12909 | 11665 | 0 | 154766 | 413021 | 26092 | 22428 | 10521 | 107546 | 64440 | 20926 | 70871 | 5258 | 32178 | 30257 | 3376 | 167236 | 219854 | 340948 |  | Unknown
AB-220715-01 LRB 17 | 134814 | 91508 | 783 | 559 | 20980 | 2161197 | 2162359 | 58 | 637 | 5154 | 419 | 1435 | 0 | 4821 | 87418 | 209558 | 297338 | 482924 | 1362703 | 5235691 | 1894534 | 2331 | 23348606 | 4310166 | 10012992 | 2164904 | 14326661 | 26613 | 34862689 | 30359741 | 195002 | 3387 | 1198 | 7496 | 0 | 29550 | 186663 | 3877 | 0 | 3875 | 100185 | 9421 | 2079 | 29019 | 5411 | 3957 | 0 | 304 | 165314 | 138999 | 50245 |  | Unknown
AB-220715-01 LD 8 | 41436 | 68227 | 624 | 271 | 8508 | 2072463 | 2144285 | 12 | 275 | 2244 | 300 | 1491 | 0 | 4278 | 20853 | 82944 | 297844 | 78067 | 392269 | 1462805 | 349396 | 0 | 13275971 | 2557496 | 3429914 | 1034197 | 6126156 | 6367 | 26544928 | 19344985 | 38257 | 1813 | 696 | 5871 | 0 | 198115 | 96624 | 1115 | 0 | 3465 | 84806 | 11517 | 2084 | 53943 | 2745 | 2282 | 8465 | 65 | 207477 | 149077 | 21195 |  | Unknown
AB-220715-01 CCC-H 9 | 70864251 | 166242 | 247337 | 398246 | 78987400 | 1458367 | 2208073 | 1931914 | 5848146 | 599342 | 1030682 | 3278063 | 2064707 | 3122077 | 82629 | 125467 | 201430 | 510230 | 1072569 | 4332367 | 1971418 | 0 | 16032009 | 3554542 | 8055316 | 1695888 | 10457366 | 51690 | 27716630 | 18446645 | 245512 | 2541952 | 2551914 | 3397584 | 82766 | 27329670 | 51739431 | 3240393 | 8253523 | 3679669 | 160843 | 18375133 | 3695689 | 14687515 | 24595 | 7147165 | 870112 | 1817444 | 47642794 | 54798454 | 32583449 |  | Unknown
AB-220715-01 LFSM-H 8 | 35421212 | 71756 | 71941 | 245142 | 31933295 | 1757891 | 2180831 | 159622 | 1170507 | 68901 | 102562 | 816895 | 420590 | 1156594 | 24239 | 75736 | 262702 | 97378 | 462711 | 1659169 | 479856 | 0 | 14810303 | 3050140 | 3756214 | 1294954 | 7389918 | 7833 | 25417726 | 19974113 | 48644 | 508961 | 1485139 | 1670928 | 0 | 13004847 | 25959675 | 1764465 | 1554075 | 420789 | 70091 | 6609977 | 1767581 | 3579067 | 5895 | 3149538 | 45697 | 309971 | 22019946 | 26459234 | 17938028 |  | Unknown
WO-22270-161-A-1 | 454888 | 51350 | 742 | 337 | 38099 | 1986392 | 2019447 | 22 | 1467 | 4113 | 712 | 2380 | 0 | 4573 | 25280 | 90301 | 330552 | 96858 | 421825 | 1505230 | 412731 | 981 | 16116943 | 2994895 | 3674705 | 1294928 | 7362581 | 4753 | 28572131 | 21517596 | 49045 | 1948 | 6605 | 6752 | 0 | 182171 | 696407 | 15416 | 0 | 1787 | 65756 | 31372 | 8530 | 6047 | 1850 | 5073 | 2682 | 637 | 619865 | 881519 | 153728 |  | Unknown
WO-22270-162-A-1 | 233216 | 79567 | 507 | 248 | 23189 | 2331250 | 2109942 | 0 | 611 | 5724 | 474 | 1714 | 0 | 3584 | 27638 | 92255 | 362497 | 114259 | 500477 | 1830214 | 487440 | 0 | 16928868 | 3135963 | 4109872 | 1468628 | 8525979 | 5357 | 28589800 | 25150482 | 54481 | 898 | 1993 | 8149 | 0 | 288812 | 478690 | 5407 | 21493 | 0 | 52940 | 11432 | 5562 | 13179 | 1753 | 2541 | 0 | 169 | 529921 | 729200 | 54215 | D,Q | Unknown
WO-22270-163-A-1 | 61314 | 52091 | 116 | 263 | 10208 | 1672142 | 1536856 | 0 | 346 | 3097 | 316 | 2213 | 0 | 3966 | 15786 | 79985 | 229617 | 54525 | 345906 | 1149852 | 273405 | 1388 | 12857812 | 2325081 | 2898261 | 1032507 | 5966876 | 3741 | 24504190 | 19001604 | 30667 | 689 | 727 | 5579 | 0 | 156096 | 167409 | 1306 | 0 | 0 | 60066 | 9584 | 5538 | 18172 | 1560 | 2799 | 0 | 115 | 373145 | 436123 | 19697 |  | Unknown
WO-22270-164-A-1 | 52296 | 87690 | 671 | 8 | 10025 | 2626230 | 2813938 | 0 | 418 | 20883 | 736 | 1119 | 0 | 18398 | 33841 | 119909 | 425434 | 145041 | 564007 | 2236290 | 666467 | 0 | 19781444 | 3816615 | 5072580 | 1585522 | 9497668 | 8580 | 32264373 | 26196559 | 75186 | 689 | 106 | 9871 | 0 | 265157 | 174526 | 468 | 0 | 1834 | 77452 | 41781 | 4479 | 36720 | 1481 | 2522 | 0 | 39 | 352723 | 339437 | 16091 | D,Q | Unknown
WO-22270-165-A-1 | 28070 | 74718 | 390 | 103 | 10876 | 2160422 | 2037933 | 0 | 166 | 18987 | 857 | 2307 | 0 | 2947 | 30179 | 114002 | 361304 | 131824 | 505746 | 2046580 | 556366 | 0 | 17130650 | 3127881 | 4585177 | 1487528 | 8954291 | 4607 | 28547043 | 23654570 | 67796 | 2325 | 265 | 4805 | 0 | 242310 | 89464 | 1126 | 0 | 0 | 71158 | 10220 | 5113 | 8295 | 2082 | 2917 | 0 | 369 | 283246 | 319686 | 19570 | D,Q | Unknown
WO-22270-166-A-1 | 22860 | 89312 | 1406 | 186 | 5704 | 2725002 | 2604755 | 0 | 212 | 18950 | 491 | 1704 | 0 | 3956 | 35152 | 116542 | 571070 | 132070 | 571861 | 2081678 | 591614 | 1444 | 20176247 | 3617881 | 4721342 | 1578553 | 9307376 | 6403 | 32476254 | 25849404 | 64993 | 953 | 711 | 7417 | 0 | 343588 | 74555 | 697 | 0 | 0 | 77237 | 7977 | 4505 | 9444 | 2314 | 2032 | 0 | 72 | 262154 | 256537 | 13327 |  | Unknown
WO-22270-167-A-1 | 14732 | 98836 | 1176 | 430 | 5648 | 2798596 | 2800951 | 0 | 0 | 37554 | 536 | 1657 | 0 | 3200 | 19690 | 104460 | 614421 | 74108 | 416676 | 1539979 | 344472 | 0 | 16788719 | 3056745 | 3611791 | 1214396 | 8053607 | 3531 | 27618267 | 22126715 | 38149 | 2155 | 1559 | 5383 | 0 | 396046 | 50441 | 346 | 0 | 0 | 20702 | 33127 | 2486 | 24058 | 1177 | 1071 | 11539 | 82 | 217016 | 185585 | 12013 |  | Unknown
WO-22270-168-A-1 | 19277 | 111019 | 962 | 279 | 8070 | 3425960 | 3555450 | 0 | 50 | 57210 | 470 | 1655 | 0 | 15800 | 27777 | 138586 | 903088 | 133297 | 509001 | 1951951 | 576787 | 0 | 18385424 | 3786516 | 4653025 | 1540673 | 9125726 | 7327 | 28966846 | 24199721 | 62015 | 2885 | 634 | 8290 | 0 | 223342 | 50765 | 782 | 28099 | 0 | 50359 | 11998 | 3522 | 17749 | 1004 | 1973 | 0 | 357 | 317569 | 285695 | 15495 | D,Q | Unknown
WO-22270-169-A-1 | 15747 | 129191 | 2221 | 541 | 6454 | 3251637 | 3030011 | 14 | 51 | 44090 | 471 | 1792 | 0 | 5296 | 17178 | 99263 | 935816 | 72228 | 387787 | 1305458 | 318093 | 912 | 14113607 | 2790302 | 3405037 | 1161526 | 7248648 | 5124 | 25512240 | 19715255 | 31504 | 3667 | 587 | 7767 | 0 | 222308 | 37634 | 1231 | 0 | 1894 | 76201 | 18854 | 4666 | 21597 | 1326 | 2038 | 0 | 88 | 288974 | 223936 | 12841 |  | Unknown
WO-22270-170-A-1 | 12562 | 99301 | 1438 | 274 | 5780 | 3311183 | 3116038 | 0 | 88 | 79811 | 454 | 1986 | 0 | 15434 | 26498 | 116824 | 819148 | 107664 | 501441 | 1835729 | 491147 | 0 | 17762304 | 3443669 | 4397285 | 1480671 | 8815748 | 8010 | 28478833 | 24417699 | 53943 | 3513 | 540 | 5492 | 0 | 233304 | 34502 | 456 | 0 | 0 | 61507 | 8628 | 2484 | 5625 | 1362 | 1778 | 0 | 107 | 230905 | 173685 | 11498 | D,Q | Unknown
AB-220715-01 LRB 18_INJ001 | 10350 | 58167 | 609 | 221 | 7751 | 2283502 | 2471648 | 0 | 285 | 4044 | 704 | 930 | 13574 | 4283 | 93252 | 233752 | 356241 | 482731 | 1416177 | 5197232 | 1918687 | 3151 | 24380115 | 4685267 | 11009099 | 2214346 | 15008419 | 29718 | 36131084 | 31226055 | 214031 | 824 | 260 | 8483 | 0 | 34632 | 40527 | 308 | 0 | 4709 | 117657 | 8082 | 2742 | 32618 | 7523 | 1268 | 0 | 279 | 233574 | 177962 | 21549 |  | Unknown
WO-22209-62-A-1B | 8903 | 119216 | 1097 | 388 | 5147 | 2904826 | 2578303 | 0 | 166 | 0 | 262 | 3010 | 0 | 1783 | 19813 | 91270 | 386639 | 81465 | 388031 | 1685855 | 407490 | 1677 | 15793284 | 2772252 | 4145766 | 1129293 | 8313996 | 5468 | 28165428 | 22997577 | 43172 | 2652 | 470 | 7659 | 0 | 8501198 | 19793 | 749 | 0 | 0 | 114727 | 7923 | 2677 | 58318 | 5592 | 656 | 0 | 0 | 110652 | 63242 | 11886 |  | Unknown
AB-220715-01 LRB 18 | 13718 | 51526 | 390 | 419 | 11047 | 2076112 | 2104985 | 0 | 68 | 2420 | 734 | 1316 | 0 | 3732 | 84515 | 210157 | 292534 | 438976 | 1260583 | 4690238 | 1770614 | 0 | 21369570 | 4256085 | 9609305 | 2051760 | 13137166 | 15606 | 32285653 | 28180034 | 200118 | 3829 | 494 | 8272 | 0 | 25981 | 34180 | 248 | 0 | 4176 | 98528 | 7173 | 1710 | 30230 | 6856 | 1778 | 10997 | 28 | 234892 | 142608 | 14113 |  | Unknown
WO-22270-171-A-1 | 10230 | 75049 | 312 | 50 | 4262 | 2335439 | 2198294 | 55 | 219 | 200208 | 337 | 1642 | 0 | 1926 | 38037 | 112197 | 364912 | 146873 | 549928 | 2147506 | 642829 | 984 | 18517951 | 3454792 | 4696955 | 1549442 | 9109177 | 8750 | 30973664 | 25645065 | 67583 | 833 | 447 | 7555 | 0 | 217723 | 25911 | 121 | 0 | 0 | 67908 | 45737 | 2878 | 19811 | 1737 | 1147 | 0 | 147 | 218310 | 132844 | 8593 |  | Unknown
WO-22270-172-A-1 | 9152 | 71905 | 727 | 367 | 4623 | 2228636 | 2240503 | 0 | 77 | 226030 | 278 | 1556 | 0 | 14459 | 34064 | 115792 | 370298 | 140379 | 534822 | 2116280 | 624566 | 0 | 18818518 | 3503523 | 4712013 | 1540644 | 9120655 | 7624 | 30940338 | 23825739 | 69748 | 768 | 431 | 7374 | 0 | 283647 | 26172 | 495 | 0 | 2896 | 78559 | 13092 | 2500 | 17754 | 1533 | 1019 | 2036 | 270 | 206796 | 107433 | 11751 | D,Q | Unknown
WO-22270-173-A-1 | 9452 | 105234 | 1035 | 417 | 6829 | 2978223 | 2855709 | 12 | 213 | 300289 | 765 | 1957 | 0 | 2509 | 34909 | 115272 | 612041 | 152186 | 535398 | 2166115 | 663709 | 1938 | 20252525 | 3712945 | 4820123 | 1539161 | 9704164 | 7022 | 31655033 | 25841038 | 69738 | 1637 | 261 | 7477 | 0 | 263382 | 26324 | 331 | 0 | 2746 | 91204 | 17588 | 3137 | 27570 | 2379 | 1471 | 0 | 744 | 215880 | 99241 | 12594 |  | Unknown
WO-22270-174-A-1 | 7675 | 131266 | 2504 | 230 | 4929 | 3612097 | 3660921 | 0 | 27 | 848586 | 650 | 1347 | 0 | 4336 | 32719 | 249185 | 1160423 | 124965 | 521543 | 2029520 | 519485 | 0 | 17439954 | 3952504 | 4975969 | 1533210 | 9471485 | 5824 | 28412921 | 23723366 | 55322 | 1666 | 202 | 7778 | 0 | 208597 | 22834 | 315 | 26242 | 6442 | 86761 | 77864 | 1382 | 32792 | 2041 | 1367 | 0 | 37 | 158201 | 72391 | 20492 | D,Q | Unknown
WO-22270-175-A-1 | 10729 | 65677 | 450 | 61 | 5202 | 2398714 | 2275071 | 0 | 0 | 662854 | 331 | 2002 | 0 | 4690 | 35768 | 115601 | 384726 | 129497 | 543406 | 2122488 | 594416 | 0 | 18830491 | 3471206 | 4637843 | 1545338 | 9280381 | 6686 | 30332074 | 26278341 | 64679 | 986 | 341 | 6819 | 0 | 174743 | 21806 | 155 | 0 | 2410 | 66510 | 8258 | 2168 | 12266 | 2435 | 733 | 0 | 89 | 102596 | 54781 | 11404 | D,Q | Unknown
WO-22270-176-A-1 | 6273 | 84757 | 713 | 285 | 3574 | 2016722 | 1756199 | 0 | 0 | 477838 | 1079 | 1386 | 0 | 5335 | 17910 | 81447 | 280251 | 64699 | 376408 | 1376487 | 308150 | 0 | 14235878 | 2612632 | 3435999 | 1113208 | 6876162 | 4826 | 27025433 | 21697185 | 33535 | 1546 | 8 | 5326 | 0 | 176403 | 20359 | 460 | 0 | 4228 | 76627 | 10643 | 2512 | 5745 | 2423 | 932 | 3357 | 153 | 98580 | 47231 | 11021 |  | Unknown
WO-22270-177-A-1 | 7539 | 78026 | 671 | 27 | 6115 | 2616987 | 2578011 | 0 | 339 | 7277 | 408 | 1586 | 0 | 3809 | 37809 | 133344 | 401985 | 142206 | 645123 | 2473796 | 688421 | 1487 | 21958673 | 3899051 | 5828726 | 1831898 | 11134728 | 6742 | 34124614 | 29440950 | 73075 | 1309 | 146 | 6572 | 0 | 182687 | 20864 | 266 | 0 | 0 | 63191 | 60919 | 2008 | 8634 | 681 | 1630 | 0 | 34 | 101784 | 41845 | 11364 |  | Unknown
WO-22270-178-A-1 | 8086 | 72943 | 416 | 35 | 3559 | 1901999 | 1671047 | 0 | 0 | 2525 | 262 | 1288 | 0 | 13948 | 18875 | 71500 | 267865 | 62119 | 328193 | 1229164 | 311949 | 0 | 13680992 | 2378705 | 2867754 | 1060428 | 6148355 | 3037 | 25643746 | 19854001 | 31875 | 1074 | 0 | 6980 | 0 | 137333 | 18984 | 83 | 0 | 0 | 60858 | 11636 | 1712 | 20372 | 2081 | 892 | 18950 | 21 | 123552 | 47848 | 16835 |  | Unknown
WO-22270-179-A-1 | 5946 | 58955 | 315 | 60 | 6725 | 2217106 | 2124782 | 0 | 0 | 2837 | 319 | 2367 | 0 | 2372 | 21784 | 93854 | 341678 | 91873 | 449059 | 1618004 | 411207 | 0 | 16153915 | 2963710 | 3815560 | 1270501 | 7580320 | 7275 | 28800320 | 23226701 | 44654 | 692 | 218 | 6047 | 0 | 221329 | 17078 | 234 | 0 | 0 | 65301 | 10328 | 1008 | 5573 | 2126 | 1153 | 0 | 392 | 93064 | 45012 | 31965 | D,Q | Unknown
WO-22270-180-A-1 | 7146 | 75050 | 524 | 189 | 5399 | 2338431 | 2254397 | 0 | 40 | 3259 | 609 | 1351 | 0 | 13118 | 29334 | 113564 | 347165 | 115193 | 509927 | 1975791 | 543370 | 622 | 18072983 | 3246147 | 4484080 | 1482883 | 9321170 | 6772 | 30790723 | 25551444 | 59021 | 1030 | 152 | 6380 | 0 | 282275 | 24502 | 173 | 0 | 2108 | 80175 | 15954 | 1751 | 22248 | 2089 | 853 | 0 | 120 | 104948 | 42488 | 29425 |  | Unknown
WO-22209-64-A-1B | 8640 | 62513 | 313 | 264 | 6614 | 2562909 | 2371990 | 0 | 46 | 0 | 329 | 4717 | 0 | 1259 | 23151 | 92870 | 298614 | 74984 | 380524 | 1739384 | 393360 | 0 | 13756687 | 2729733 | 4102904 | 1135938 | 8234356 | 6654 | 28508518 | 21987613 | 41538 | 1570 | 978 | 5521 | 0 | 36603694 | 24008 | 259 | 0 | 0 | 120200 | 19521 | 653 | 149080 | 7812 | 1063 | 0 | 70 | 123418 | 63514 | 14660 | D,Q | Unknown
AB-220715-01 LFB-H 10 | 39725586 | 94571 | 138584 | 233492 | 42132782 | 1688084 | 2176229 | 942232 | 3161751 | 278478 | 490428 | 1824989 | 1110535 | 1732321 | 94706 | 146815 | 237331 | 502091 | 1142118 | 4594567 | 2024045 | 3002 | 19219979 | 4022973 | 8528689 | 1788550 | 12295535 | 47073 | 30849628 | 23540197 | 231281 | 1365503 | 1451158 | 1664423 | 34438 | 17314369 | 28979289 | 1846690 | 4307415 | 1859407 | 133057 | 10818129 | 2020472 | 8049585 | 12920 | 4049144 | 407238 | 948622 | 26605827 | 32870799 | 21500403 |  | Unknown
AB-220715-01 LRB 19_INJ001 | 504690 | 56071 | 1066 | 1682 | 210154 | 2497754 | 2649689 | 2748 | 14750 | 5189 | 2280 | 7454 | 16066 | 10772 | 110492 | 245240 | 374577 | 523247 | 1433477 | 5673945 | 2083916 | 4661 | 27205917 | 5174139 | 12104147 | 2324384 | 16362010 | 32898 | 39810080 | 34283040 | 228759 | 4750 | 9317 | 11065 | 0 | 128694 | 659581 | 21462 | 20461 | 10905 | 142594 | 55520 | 11051 | 61130 | 7424 | 27880 | 18506 | 4448 | 609600 | 874784 | 242861 |  | Unknown
AB-220715-01 LRB 19 | 108562 | 72779 | 643 | 67 | 23495 | 2054776 | 2181846 | 285 | 2046 | 4041 | 1337 | 3373 | 8642 | 5104 | 90755 | 210782 | 304150 | 455894 | 1226736 | 5085319 | 1989883 | 0 | 23794591 | 4300225 | 10277703 | 1992401 | 13901883 | 18754 | 33400239 | 28685978 | 226224 | 651 | 2083 | 7001 | 0 | 31781 | 208111 | 2668 | 0 | 3730 | 86210 | 13625 | 4884 | 33090 | 4614 | 4115 | 24865 | 307 | 346446 | 346628 | 39985 |  | Unknown
AB-220715-01 LD 9 | 40772 | 66339 | 1069 | 0 | 7654 | 2322582 | 2205038 | 0 | 281 | 2369 | 648 | 1950 | 0 | 16559 | 25466 | 97876 | 338383 | 87728 | 437075 | 1677253 | 427754 | 1164 | 16173618 | 2971139 | 4044534 | 1344566 | 7983247 | 6853 | 19664981 | 19992938 | 44848 | 1144 | 1285 | 4280 | 0 | 211734 | 96771 | 837 | 0 | 0 | 10670 | 16718 | 3692 | 6618 | 1068 | 2658 | 13391 | 64 | 262460 | 233626 | 18612 |  | Unknown
AB-220715-01 CCC-H 10 | 68644616 | 174355 | 224417 | 393470 | 72610027 | 1434971 | 2012785 | 1894444 | 5774867 | 568949 | 995393 | 3052473 | 1979024 | 3133013 | 83396 | 115407 | 213050 | 496252 | 1001249 | 4174848 | 1937687 | 0 | 16156602 | 3656733 | 7871819 | 1533437 | 11186151 | 48911 | 27926845 | 19222363 | 222916 | 2567670 | 2522446 | 3160907 | 79847 | 29026980 | 49987879 | 3229065 | 7990749 | 3592858 | 40365 | 19235370 | 3413907 | 14786797 | 2261 | 7391520 | 847744 | 1797970 | 35996371 | 51167032 | 33466968 |  | Unknown
AB-220715-01 LFSM-H 9 | 27164147 | 76329 | 64968 | 197232 | 26738027 | 1545947 | 1731979 | 145288 | 1106989 | 58714 | 87396 | 696884 | 366311 | 1011650 | 22246 | 70844 | 216694 | 87609 | 399748 | 1477259 | 388732 | 0 | 13002454 | 2605984 | 3360966 | 1117379 | 6758566 | 4407 | 17821948 | 16600286 | 44911 | 448532 | 1366225 | 1398315 | 0 | 11696560 | 17222211 | 1351237 | 1379538 | 370839 | 11142 | 5988743 | 1545673 | 3101264 | 1216 | 2583888 | 36490 | 254195 | 7202326 | 12275479 | 15394190 |  | Unknown
WO-22270-181-A-1 | 558093 | 55529 | 474 | 1054 | 99527 | 2541559 | 2903610 | 524 | 4373 | 6965 | 809 | 3586 | 0 | 4919 | 34583 | 113576 | 376135 | 128274 | 530007 | 2134145 | 637872 | 1240 | 18956485 | 3302136 | 4816138 | 1431554 | 8673978 | 7099 | 31064721 | 26373722 | 66730 | 2295 | 14386 | 9424 | 0 | 503638 | 662380 | 24261 | 17675 | 2471 | 31120 | 57475 | 8142 | 44214 | 675 | 9817 | 0 | 1087 | 561464 | 645039 | 258790 |  | Unknown
WO-22270-182-A-1 | 206978 | 101392 | 1472 | 564 | 17706 | 2969994 | 2675088 | 29 | 912 | 4101 | 648 | 1910 | 0 | 3184 | 20706 | 101785 | 648459 | 77988 | 404719 | 1552925 | 398398 | 996 | 16459906 | 2850392 | 3981976 | 1188611 | 8167641 | 6042 | 28196325 | 22726913 | 43052 | 881 | 3270 | 7786 | 0 | 299920 | 332537 | 5882 | 0 | 0 | 76707 | 16544 | 6883 | 14191 | 4005 | 3296 | 0 | 100 | 481480 | 517108 | 60018 |  | Unknown
WO-22270-183-A-1 | 94721 | 55543 | 725 | 205 | 15674 | 2552126 | 2414103 | 8 | 285 | 6959 | 269 | 2187 | 0 | 4589 | 33425 | 111515 | 393390 | 138842 | 528202 | 2116849 | 608931 | 443 | 19253640 | 3482020 | 4724255 | 1473230 | 9552472 | 8715 | 31010478 | 26567073 | 62339 | 1703 | 798 | 6886 | 0 | 329735 | 201755 | 2311 | 0 | 0 | 76787 | 14249 | 5012 | 34840 | 2205 | 3899 | 6405 | 78 | 349117 | 413846 | 28012 |  | Unknown
WO-22270-184-A-1 | 39634 | 77187 | 2671 | 390 | 9243 | 3045376 | 3107179 | 0 | 408 | 16754 | 588 | 1793 | 0 | 2222 | 33445 | 108132 | 670962 | 134816 | 517852 | 2138816 | 639985 | 0 | 20139215 | 3527137 | 4647348 | 1462737 | 9434538 | 8644 | 31957941 | 25600649 | 70329 | 1463 | 728 | 6225 | 0 | 352874 | 105457 | 1496 | 0 | 0 | 86619 | 66840 | 5080 | 17512 | 1623 | 2663 | 15167 | 18 | 288500 | 312140 | 18848 |  | Unknown
WO-22270-185-A-1 | 26032 | 132927 | 1871 | 1140 | 9694 | 4114102 | 4439424 | 0 | 374 | 27259 | 1325 | 2701 | 0 | 3072 | 44432 | 284415 | 1528768 | 173171 | 741466 | 3012761 | 814638 | 525 | 23150911 | 4995261 | 8056361 | 2206331 | 14661485 | 8066 | 31879476 | 27876175 | 90282 | 6127 | 2843 | 6718 | 0 | 363564 | 77254 | 979 | 18229 | 13524 | 42552 | 22413 | 5335 | 21957 | 865 | 2651 | 30786 | 47 | 208887 | 233683 | 16360 |  | Unknown
WO-22209-50-A-1B | 8393 | 77531 | 565 | 64 | 5433 | 2748932 | 2476931 | 0 | 143 | 0 | 681 | 1770 | 0 | 14281 | 28180 | 96116 | 226583 | 96543 | 438961 | 1940787 | 514912 | 1392 | 13086565 | 3045019 | 4449372 | 1291088 | 9071610 | 6448 | 30817615 | 25281866 | 50146 | 3584 | 996 | 6272 | 0 | 63726070 | 22597 | 160 | 0 | 0 | 112266 | 42454 | 1077 | 676202 | 6269 | 1052 | 0 | 132 | 129924 | 84239 | 18722 |  | Unknown
WO-22270-186-A-1 | 17864 | 94535 | 330 | 464 | 5582 | 2517533 | 2315006 | 0 | 87 | 11289 | 318 | 1277 | 0 | 17415 | 18513 | 83338 | 427778 | 69596 | 358339 | 1340149 | 326720 | 0 | 13512307 | 2417070 | 3342417 | 1068896 | 6822586 | 5478 | 26634056 | 20612123 | 34076 | 1015 | 764 | 5417 | 0 | 183481 | 56030 | 444 | 0 | 2374 | 76422 | 14043 | 4439 | 13588 | 2503 | 2845 | 0 | 177 | 198907 | 183684 | 13584 | D,Q | Unknown
WO-22270-187-A-1 | 17312 | 78141 | 1598 | 273 | 8616 | 3105591 | 3137435 | 0 | 56 | 154304 | 902 | 2595 | 0 | 6979 | 54399 | 152561 | 613865 | 247399 | 678660 | 3081640 | 1155541 | 0 | 23880077 | 4428297 | 6336298 | 1965195 | 12985380 | 20413 | 36863822 | 30522292 | 122710 | 750 | 762 | 7349 | 0 | 467064 | 53312 | 346 | 0 | 4148 | 118115 | 25029 | 4170 | 50651 | 4851 | 2321 | 0 | 117 | 295254 | 245126 | 13431 | D,Q | Unknown
WO-22270-188-A-1 | 25306 | 130219 | 2530 | 592 | 16763 | 4410106 | 3966564 | 96 | 575 | 138516 | 1383 | 2872 | 0 | 3859 | 65085 | 279624 | 1495305 | 269362 | 761956 | 3426436 | 1236410 | 0 | 24832194 | 5412626 | 7920627 | 2152651 | 14474228 | 20853 | 35133627 | 30346587 | 125616 | 2655 | 1076 | 7757 | 0 | 834435 | 50264 | 2114 | 12160 | 5070 | 103190 | 21897 | 4890 | 166862 | 6044 | 3648 | 0 | 465 | 185414 | 162019 | 19642 |  | Unknown
WO-22270-189-A-1 | 13289 | 202250 | 4075 | 338 | 8319 | 4195390 | 4444594 | 80 | 426 | 130511 | 982 | 2598 | 0 | 4657 | 52056 | 185304 | 1211984 | 230129 | 719147 | 3039471 | 1042060 | 0 | 26099177 | 4691945 | 6857499 | 2026103 | 14163833 | 16292 | 35332446 | 30898688 | 120679 | 792 | 745 | 7261 | 0 | 828208 | 39865 | 1245 | 35460 | 6383 | 90873 | 48333 | 3936 | 54022 | 3025 | 1352 | 4266 | 205 | 178957 | 144277 | 20596 | D,Q | Unknown
WO-22270-190-A-1 | 13196 | 129742 | 848 | 480 | 7324 | 3385725 | 3380513 | 0 | 0 | 94096 | 573 | 2816 | 0 | 2611 | 40796 | 132285 | 794217 | 173696 | 555196 | 2360636 | 765946 | 0 | 22583744 | 4021777 | 5309754 | 1732972 | 10823491 | 10518 | 33708414 | 27803523 | 90012 | 3380 | 760 | 7999 | 0 | 770711 | 29169 | 575 | 0 | 2743 | 84349 | 47088 | 2736 | 29831 | 4638 | 808 | 3200 | 41 | 154598 | 119433 | 13868 | D,Q | Unknown
AB-220715-01 LRB 20_INJ001 | 15132 | 56595 | 727 | 285 | 5963 | 2301883 | 2441407 | 145 | 649 | 4696 | 1660 | 949 | 0 | 15497 | 97905 | 235916 | 333220 | 495819 | 1365573 | 5549776 | 2217661 | 4936 | 25699741 | 4284905 | 11680942 | 2111009 | 15617144 | 46383 | 37831393 | 32251869 | 252520 | 1617 | 176 | 7712 | 0 | 26104 | 43746 | 58 | 0 | 0 | 127902 | 10756 | 4332 | 20108 | 6724 | 2033 | 4970 | 13 | 260293 | 195700 | 16641 |  | Unknown
AB-220715-01 LRB 20 | 11362 | 56621 | 1139 | 21 | 6284 | 1664744 | 1727037 | 85 | 0 | 3194 | 726 | 925 | 8515 | 3211 | 84316 | 200638 | 255559 | 425025 | 1047795 | 4725506 | 1842133 | 3885 | 19889324 | 3421158 | 9592561 | 1689166 | 12643900 | 19932 | 29509342 | 24939790 | 204594 | 833 | 674 | 5962 | 0 | 19596 | 30109 | 789 | 0 | 0 | 81536 | 23237 | 3484 | 18105 | 3571 | 1006 | 3765 | 188 | 243674 | 125983 | 16151 |  | Unknown
WO-22270-196-A-1 | 9559 | 74913 | 1193 | 323 | 5913 | 3475175 | 3609015 | 33 | 56 | 1205951 | 558 | 2455 | 0 | 18645 | 56168 | 152660 | 635473 | 240146 | 643972 | 3077886 | 1139282 | 0 | 25598630 | 4480390 | 6522938 | 1933863 | 13141993 | 20721 | 36274402 | 31830044 | 122240 | 1855 | 250 | 9976 | 0 | 908718 | 28550 | 352 | 0 | 8632 | 107972 | 24622 | 2822 | 30238 | 4913 | 1128 | 0 | 0 | 121927 | 56077 | 15345 |  | Unknown
WO-22270-191-A-1 | 11371 | 114386 | 2845 | 240 | 6041 | 4168593 | 4169182 | 0 | 19 | 493174 | 625 | 2253 | 0 | 20085 | 61418 | 197773 | 1066820 | 266195 | 718456 | 3340064 | 1247643 | 0 | 25991201 | 4756345 | 7343352 | 2119930 | 14534261 | 21024 | 36443751 | 32417895 | 126505 | 2741 | 1023 | 9191 | 0 | 932421 | 30458 | 898 | 0 | 4682 | 113135 | 58931 | 2669 | 45380 | 3891 | 1427 | 0 | 39 | 150400 | 96171 | 17359 | D,Q | Unknown
WO-22270-197-A-1 | 8544 | 108285 | 223 | 91 | 5127 | 3060822 | 2907200 | 0 | 0 | 661577 | 451 | 1524 | 0 | 10039 | 41507 | 114366 | 533290 | 140258 | 532702 | 2301723 | 685938 | 655 | 20959273 | 3613718 | 5169965 | 1546636 | 10400087 | 10025 | 32629540 | 27735261 | 71453 | 1900 | 882 | 7637 | 0 | 307083 | 26348 | 653 | 0 | 5812 | 110655 | 38383 | 2551 | 40740 | 5413 | 1710 | 0 | 289 | 111271 | 47637 | 18723 |  | Unknown
WO-22270-192-A-1 | 13755 | 89051 | 1404 | 534 | 8024 | 3255030 | 3194752 | 0 | 0 | 642411 | 479 | 2437 | 0 | 17061 | 62809 | 159453 | 593436 | 303874 | 712567 | 3138719 | 1275267 | 0 | 27126592 | 4732495 | 6666580 | 2093905 | 13334156 | 24153 | 37008033 | 32957844 | 136209 | 1805 | 152 | 7224 | 0 | 905056 | 31473 | 383 | 10172 | 6731 | 121922 | 20139 | 2967 | 52485 | 5302 | 1839 | 0 | 98 | 163229 | 103489 | 17224 | D,Q | Unknown
WO-22270-198-A-1 | 10937 | 70483 | 1127 | 55 | 4930 | 2716667 | 2716111 | 0 | 6 | 4957461 | 3088 | 1487 | 0 | 18341 | 34243 | 132773 | 415753 | 169954 | 588901 | 2632540 | 634144 | 1138 | 22458245 | 3893802 | 5584033 | 1639334 | 11065748 | 11645 | 35158316 | 28517614 | 67161 | 1036 | 306 | 10045 | 0 | 422744 | 24109 | 1170 | 7455 | 10095978 | 88421 | 15621 | 2156 | 16044 | 2019 | 1533 | 4428 | 340 | 114429 | 49156 | 12408 |  | Unknown
WO-22270-193-A-1 | 10275 | 87391 | 516 | 302 | 8211 | 2751426 | 3256161 | 0 | 0 | 443882 | 410 | 1573 | 0 | 4575 | 48029 | 132788 | 537847 | 209770 | 600249 | 2709957 | 956387 | 0 | 22468485 | 4265959 | 5772170 | 1812329 | 11873065 | 16422 | 36378824 | 26626199 | 106735 | 785 | 573 | 6719 | 0 | 845512 | 29150 | 772 | 0 | 3564 | 118873 | 21558 | 1784 | 17370 | 3532 | 1548 | 0 | 0 | 135098 | 89654 | 14078 | D,Q | Unknown
WO-22270-199-A-1 | 7735 | 108890 | 1194 | 234 | 4712 | 2702542 | 2648514 | 0 | 0 | 5251817 | 2362 | 2229 | 0 | 14472 | 33880 | 127379 | 428635 | 162087 | 567704 | 2477762 | 604404 | 0 | 21670250 | 3858263 | 5411755 | 1606146 | 10627487 | 13829 | 32825112 | 26677468 | 64619 | 1133 | 54 | 7189 | 0 | 511934 | 21091 | 130 | 0 | 10165882 | 80677 | 20830 | 2446 | 25640 | 3219 | 1061 | 1309 | 36 | 119372 | 45944 | 12191 | D,Q | Unknown
WO-22209-26-A-1B | 7853 | 108223 | 513 | 80 | 0 | 2693180 | 2227796 | 0 | 0 | 0 | 459 | 2718 | 0 | 1844 | 21111 | 83175 | 298085 | 85097 | 385010 | 1687173 | 424919 | 0 | 16095196 | 2766395 | 3955236 | 1107166 | 8137461 | 8466 | 28553540 | 23616166 | 46123 | 1104 | 696 | 5862 | 0 | 2196621 | 25042 | 609 | 0 | 2057 | 122848 | 23075 | 3036 | 504866 | 6185 | 1120 | 0 | 21 | 147347 | 89679 | 9849 | D,Q | Unknown
WO-22270-200-A-1 | 12249 | 129741 | 1612 | 196 | 5551 | 2822996 | 2538817 | 0 | 14 | 4731481 | 1915 | 1368 | 0 | 6931 | 36247 | 120612 | 404347 | 151874 | 541995 | 2337775 | 599146 | 1032 | 19825367 | 3522961 | 5338460 | 1516460 | 9989664 | 9850 | 33040892 | 29166318 | 61754 | 633 | 565 | 7614 | 0 | 312091 | 19905 | 218 | 0 | 9570492 | 92934 | 36644 | 1801 | 22090 | 3646 | 1366 | 0 | 0 | 106847 | 42409 | 15896 |  | Unknown
WO-22270-194-A-1 | 10482 | 71986 | 2249 | 2424 | 8461 | 3164550 | 3459614 | 0 | 196 | 1176117 | 1013 | 2838 | 0 | 2792 | 67543 | 192429 | 547904 | 299005 | 815770 | 3675290 | 1414189 | 0 | 29615948 | 4979431 | 8333244 | 2263229 | 15929215 | 23998 | 41782086 | 35639961 | 139349 | 863 | 118 | 7965 | 0 | 789105 | 28655 | 476 | 4519 | 7212 | 101307 | 27304 | 1516 | 48620 | 3805 | 1155 | 0 | 119 | 113023 | 61368 | 15283 | D,Q | Unknown
WO-22270-195-A-1 | 8073 | 64268 | 629 | 81 | 4864 | 2492256 | 2449810 | 7 | 151 | 538074 | 409 | 1719 | 0 | 15896 | 43499 | 121157 | 374565 | 172405 | 569531 | 2495069 | 813259 | 1537 | 21702220 | 3889121 | 5337636 | 1629132 | 10417731 | 12327 | 34253435 | 28763560 | 89450 | 1027 | 446 | 7593 | 0 | 430315 | 26220 | 281 | 0 | 5092 | 110642 | 13009 | 2114 | 30408 | 3408 | 1094 | 0 | 0 | 133442 | 63509 | 10859 |  | Unknown
AB-220715-01 LFB-H 11 | 8300103 | 102578 | 151891 | 165795 | 37091052 | 1339382 | 1510707 | 1012549 | 3279981 | 411277 | 544623 | 1747135 | 1155467 | 1564650 | 104259 | 152134 | 177524 | 551945 | 1211979 | 5429895 | 2370944 | 3251 | 12123754 | 2815373 | 10268475 | 1658269 | 12889642 | 59029 | 2439918 | 6507855 | 268344 | 1604874 | 1104598 | 1197644 | 52513 | 10912514 | 1308668 | 441616 | 5147848 | 2236546 | 2805 | 11534321 | 1229182 | 9203621 | 280 | 2869232 | 488593 | 1062719 | 84104 | 142683 | 5596511 |  | Unknown
AB-220715-01 LRB 21 | 19127 | 72520 | 305 | 97 | 22973 | 2018130 | 1993054 | 296 | 2300 | 3767 | 1047 | 1435 | 0 | 18413 | 98501 | 212014 | 300489 | 459063 | 1190488 | 5381324 | 2093452 | 4562 | 22951723 | 4078338 | 10613237 | 1873834 | 14466924 | 22186 | 32214217 | 28860293 | 241550 | 3371 | 528 | 5444 | 0 | 27883 | 27888 | 507 | 0 | 6170 | 95559 | 11141 | 3277 | 30829 | 6430 | 4054 | 14949 | 449 | 90782 | 31055 | 14516 |  | Unknown
AB-220715-01 LD 10 | 11695 | 121897 | 1982 | 518 | 11995 | 3239830 | 3591831 | 8 | 334 | 2325449 | 795 | 1782 | 0 | 17694 | 31041 | 139851 | 727302 | 140679 | 538589 | 2342617 | 627072 | 0 | 21373739 | 3718049 | 5347087 | 1698249 | 11200934 | 8432 | 31359191 | 27209242 | 64412 | 812 | 728 | 6200 | 0 | 276811 | 22220 | 755 | 0 | 6592735 | 103695 | 24217 | 2671 | 27542 | 4237 | 1997 | 0 | 57 | 90434 | 30085 | 13490 |  | Unknown
AB-220715-01 CCC-H 11 | 62552188 | 207935 | 243513 | 389966 | 73359033 | 1506161 | 1945719 | 2119115 | 6083579 | 799752 | 1106896 | 3087642 | 2131782 | 3129270 | 106256 | 120679 | 204008 | 533110 | 1057160 | 5033259 | 2516157 | 4721 | 16211907 | 3906454 | 9109447 | 1602482 | 12641679 | 63416 | 24572230 | 18986827 | 286852 | 2979081 | 2596559 | 3138564 | 97016 | 30991789 | 47238687 | 3043485 | 9369287 | 4468238 | 2595 | 22557576 | 3201908 | 16729028 | 587 | 7766560 | 1025063 | 2082237 | 9435820 | 31931568 | 33784552 |  | Unknown
WO-22209-16-A-1B | 17900 | 121431 | 1659 | 728 | 13230 | 3294712 | 2922934 | 0 | 163 | 19678 | 1213 | 2944 | 0 | 4692 | 17453 | 82440 | 363929 | 108593 | 467768 | 2254105 | 652479 | 0 | 18631992 | 3757811 | 5062016 | 1271826 | 9339887 | 11790 | 30737034 | 26027307 | 78051 | 5214 | 1924 | 8522 | 0 | 3359472 | 42546 | 1142 | 0 | 27354 | 123174 | 36059 | 2662 | 51075 | 6611 | 3959 | 0 | 64 | 102206 | 94166 | 22796 | D,Q | Unknown
AB-220715-01 LFSM-H 10 | 33053985 | 205488 | 78188 | 442677 | 29238898 | 2339082 | 2863278 | 157275 | 1077678 | 1855928 | 96435 | 753422 | 398517 | 1055766 | 23100 | 84798 | 472744 | 97878 | 425695 | 1809408 | 423311 | 0 | 15765296 | 2919915 | 4128931 | 1217569 | 7908652 | 6755 | 26227227 | 20577475 | 51976 | 517654 | 2004048 | 2183224 | 0 | 13828310 | 25060463 | 1565779 | 1647284 | 5026234 | 150556 | 7019115 | 1672745 | 3662555 | 34056 | 2924044 | 47637 | 261838 | 25631714 | 29181725 | 18640318 |  | Unknown
WO-22270-201-A-1 | 312025 | 125584 | 1668 | 830 | 47016 | 2599212 | 2414262 | 27 | 1847 | 4156516 | 1183 | 3393 | 0 | 8864 | 26124 | 104482 | 368656 | 115076 | 479193 | 2042148 | 503451 | 0 | 18531576 | 3044015 | 4412096 | 1387673 | 8828642 | 7649 | 30722272 | 24511707 | 53468 | 1756 | 5019 | 6648 | 0 | 300939 | 414824 | 11902 | 0 | 7634346 | 100381 | 33011 | 8242 | 25639 | 5839 | 7917 | 0 | 186 | 291308 | 401806 | 113185 |  | Unknown
WO-22270-202-A-1 | 101140 | 56850 | 1152 | 238 | 21530 | 2531157 | 2535540 | 10 | 286 | 4121898 | 1114 | 2454 | 0 | 4629 | 26447 | 106538 | 373671 | 110319 | 528127 | 2111930 | 453729 | 418 | 20517245 | 3936619 | 4885475 | 1603564 | 11059800 | 10213 | 34087430 | 29226073 | 49345 | 3798 | 866 | 7285 | 0 | 180415 | 183489 | 2445 | 0 | 7660050 | 98722 | 14872 | 5347 | 9480 | 5078 | 2657 | 0 | 464 | 210371 | 226280 | 29170 |  | Unknown
WO-22270-203-A-1 | 37735 | 55936 | 1008 | 254 | 12645 | 2518716 | 2578022 | 0 | 694 | 5006172 | 1661 | 2040 | 0 | 3532 | 30084 | 115064 | 361149 | 159799 | 587666 | 2376762 | 637910 | 1511 | 20235970 | 4040754 | 5152350 | 1567207 | 11351161 | 11191 | 34267592 | 27898593 | 61876 | 669 | 1012 | 6892 | 0 | 308249 | 108198 | 1231 | 0 | 10376157 | 104773 | 16260 | 3645 | 9158 | 7815 | 2953 | 9618 | 31 | 166601 | 192487 | 13852 |  | Unknown
AB-220715-01 LRB 21_INJ001 | 56952 | 57243 | 1098 | 1070 | 176228 | 2297486 | 2189545 | 3001 | 11701 | 5716 | 1551 | 7340 | 0 | 20760 | 98287 | 256407 | 308502 | 476769 | 1347453 | 5928264 | 2165647 | 5387 | 25827315 | 4537931 | 11944023 | 2116050 | 16532369 | 37697 | 40632069 | 33874755 | 247394 | 6750 | 6376 | 8550 | 0 | 65437 | 39367 | 2617 | 21458 | 17770 | 133636 | 47524 | 5579 | 62002 | 8094 | 16145 | 12642 | 3180 | 92429 | 35941 | 37570 |  | Unknown
WO-22270-204-A-1 | 17832 | 70008 | 763 | 123 | 8223 | 2241633 | 2150724 | 0 | 141 | 2942196 | 934 | 1408 | 0 | 5061 | 15081 | 89635 | 277149 | 72212 | 369779 | 1566810 | 310477 | 0 | 16437401 | 2840280 | 3656935 | 1206656 | 7737605 | 8216 | 30122983 | 23599311 | 40094 | 1453 | 386 | 5573 | 0 | 136873 | 86023 | 598 | 0 | 5164732 | 108136 | 13374 | 2782 | 48956 | 5594 | 2048 | 12677 | 253 | 179928 | 168606 | 12649 |  | Unknown
WO-22270-205-A-1 | 15074 | 108560 | 2050 | 132 | 5614 | 2456603 | 2269103 | 0 | 37 | 4759559 | 1300 | 1962 | 0 | 8896 | 31630 | 112654 | 331253 | 135280 | 524284 | 2318626 | 553318 | 2302 | 19113363 | 3302608 | 5052642 | 1438636 | 9676819 | 9178 | 33850264 | 27441481 | 63613 | 2620 | 834 | 4381 | 0 | 187568 | 44727 | 246 | 0 | 9131872 | 104820 | 21548 | 5072 | 5896 | 5492 | 1611 | 0 | 170 | 153193 | 123758 | 10804 |  | Unknown
WO-22270-206-A-1 | 9450 | 85810 | 2372 | 236 | 4433 | 2487020 | 2282657 | 0 | 68 | 3845542 | 902 | 2281 | 0 | 18253 | 21988 | 100941 | 285885 | 89706 | 405642 | 1939750 | 430233 | 849 | 17110554 | 2645828 | 4363613 | 1161523 | 8297687 | 6202 | 31380877 | 25625571 | 51225 | 4381 | 63 | 6613 | 0 | 312911 | 23821 | 140 | 0 | 6748674 | 127915 | 28084 | 3819 | 20076 | 7971 | 1352 | 0 | 109 | 115861 | 51787 | 11526 |  | Unknown
WO-22270-207-A-1 | 5656 | 70531 | 891 | 112 | 4494 | 2359358 | 2073401 | 0 | 0 | 2851255 | 724 | 1083 | 0 | 3571 | 16339 | 96343 | 336241 | 62211 | 331950 | 1463601 | 288327 | 0 | 14515480 | 2327833 | 3735302 | 981687 | 7687885 | 7888 | 26318461 | 21979217 | 39905 | 2460 | 174 | 5142 | 0 | 336239 | 17927 | 456 | 0 | 4946374 | 123029 | 26850 | 3482 | 35013 | 9012 | 743 | 0 | 39 | 74391 | 36076 | 7504 | D,Q | Unknown
WO-22270-208-A-1 | 8909 | 66061 | 1135 | 0 | 5979 | 2534823 | 2372244 | 0 | 0 | 3047778 | 1633 | 2014 | 0 | 3187 | 30005 | 106441 | 349290 | 121904 | 466090 | 2172200 | 530208 | 0 | 19465500 | 3179151 | 4888982 | 1331510 | 9756775 | 7325 | 33714520 | 27843778 | 63483 | 3201 | 674 | 8227 | 0 | 209115 | 25463 | 285 | 0 | 7210798 | 135077 | 14443 | 3224 | 23029 | 8176 | 1341 | 0 | 226 | 116617 | 43068 | 12722 | D,Q | Unknown
WO-22270-209-A-1 | 8357 | 80740 | 690 | 239 | 9629 | 2888848 | 2587611 | 17 | 18 | 3831700 | 1566 | 2479 | 0 | 52412 | 39740 | 132766 | 374695 | 172114 | 537979 | 2707808 | 673820 | 0 | 22370756 | 3639080 | 5839527 | 1542083 | 11268998 | 11136 | 34597397 | 30503898 | 77582 | 3791 | 594 | 6684 | 0 | 637576 | 27913 | 963 | 0 | 9205464 | 109800 | 28858 | 1957 | 52630 | 6194 | 1164 | 0 | 54 | 96436 | 41347 | 13893 |  | Unknown
WO-22270-210-A-1 | 8490 | 64695 | 720 | 320 | 5265 | 2874013 | 2644704 | 0 | 0 | 4286981 | 1489 | 1589 | 0 | 12614 | 36511 | 124293 | 402229 | 164860 | 571580 | 2783024 | 719747 | 0 | 23175947 | 3795063 | 5975251 | 1600983 | 12077200 | 13982 | 36592000 | 30984253 | 81550 | 1467 | 210 | 4933 | 0 | 634337 | 24013 | 884 | 0 | 9863046 | 130234 | 15470 | 2547 | 24611 | 6962 | 1211 | 0 | 0 | 89340 | 35803 | 13116 | D,Q | Unknown
AB-220715-01 LRB 22_INJ001 | 11291 | 56332 | 675 | 118 | 8959 | 2202590 | 2170194 | 66 | 17 | 42480 | 1639 | 981 | 0 | 14276 | 102358 | 245910 | 332309 | 477532 | 1304177 | 6359957 | 2460376 | 6040 | 27996557 | 4183363 | 12532133 | 2087106 | 17969396 | 44556 | 39671586 | 35426456 | 280454 | 1134 | 286 | 6361 | 0 | 34913 | 32126 | 407 | 0 | 91011 | 135162 | 40008 | 2161 | 31607 | 7918 | 1298 | 5923 | 148 | 116496 | 49318 | 19013 |  | Unknown
AB-220715-01 LRB 22 | 16166 | 45826 | 689 | 231 | 10304 | 1571959 | 1489872 | 0 | 156 | 8101 | 617 | 1418 | 0 | 15527 | 74040 | 188098 | 238767 | 343182 | 922690 | 4543262 | 1697926 | 4263 | 19412130 | 3053423 | 9138883 | 1463968 | 12911343 | 20648 | 29270139 | 25550953 | 179572 | 961 | 166 | 4451 | 0 | 27297 | 27868 | 535 | 17719 | 22152 | 86822 | 26680 | 2418 | 34000 | 6039 | 1372 | 3229 | 315 | 113610 | 40269 | 14669 |  | Unknown
WO-22270-211-A-1 | 11391 | 42656 | 554 | 103 | 5220 | 2677206 | 2601572 | 8 | 32 | 3669083 | 1659 | 1305 | 0 | 12844 | 37653 | 132177 | 357338 | 155620 | 546469 | 2662864 | 688156 | 2203 | 22430735 | 3889188 | 5715933 | 1597403 | 11448530 | 9728 | 35545385 | 31362643 | 71482 | 1547 | 487 | 7008 | 0 | 314661 | 27444 | 251 | 0 | 9031795 | 124718 | 13818 | 1259 | 19279 | 6975 | 2420 | 0 | 115 | 95874 | 30307 | 15591 |  | Unknown
WO-22270-212-A-1 | 7669 | 46684 | 240 | 78 | 5300 | 2486375 | 2419707 | 0 | 73 | 3137236 | 1865 | 2031 | 0 | 43893 | 32765 | 116012 | 358441 | 124086 | 494734 | 2339850 | 604138 | 0 | 21237452 | 3878303 | 5130742 | 1497004 | 10811392 | 11841 | 33654171 | 28488882 | 66935 | 2779 | 237 | 7485 | 0 | 409709 | 26068 | 475 | 0 | 7723499 | 116204 | 16919 | 1989 | 40653 | 7874 | 1238 | 0 | 231 | 88624 | 32020 | 12209 | D,Q | Unknown
WO-22270-213-A-1 | 6285 | 43971 | 1046 | 105 | 5609 | 2770240 | 2843749 | 0 | 0 | 4053827 | 1725 | 1679 | 0 | 19171 | 40341 | 138558 | 427921 | 156739 | 566561 | 2715015 | 718723 | 4317 | 22977323 | 4053729 | 5825835 | 1722576 | 12507360 | 14404 | 37145700 | 30265495 | 85676 | 1073 | 153 | 7546 | 0 | 483397 | 25420 | 199 | 0 | 9537727 | 119441 | 24451 | 1860 | 36698 | 7955 | 967 | 0 | 123 | 84445 | 28264 | 10146 |  | Unknown
WO-22270-214-A-1 | 6522 | 108209 | 2894 | 621 | 4072 | 4539473 | 4771740 | 21 | 0 | 7378664 | 3003 | 2367 | 0 | 6054 | 51702 | 229641 | 1126258 | 225661 | 682200 | 3498357 | 926666 | 2438 | 26389497 | 4734977 | 7941049 | 2028914 | 15341442 | 21435 | 38816731 | 32974346 | 89100 | 5438 | 239 | 6689 | 0 | 1789912 | 25413 | 347 | 0 | 14812187 | 129451 | 96237 | 1324 | 52113 | 6391 | 2232 | 0 | 9 | 84787 | 29131 | 11917 |  | Unknown
WO-22270-215-A-1 | 5402 | 135273 | 2433 | 702 | 3391 | 4176153 | 4060069 | 0 | 36 | 4546002 | 2076 | 1602 | 0 | 3091 | 25747 | 161656 | 1085132 | 99597 | 480468 | 2269269 | 530274 | 0 | 19718899 | 3500936 | 5368983 | 1424552 | 10532768 | 7520 | 34149628 | 27595557 | 56715 | 6454 | 638 | 6280 | 0 | 252462 | 22651 | 284 | 12693 | 8194978 | 133687 | 26662 | 817 | 41157 | 8066 | 1608 | 10478 | 8 | 86045 | 32263 | 11696 |  | Unknown
WO-22270-216-A-1 | 6195 | 73877 | 529 | 206 | 5304 | 2726022 | 2530439 | 0 | 0 | 4057707 | 2204 | 1228 | 0 | 3059 | 35796 | 114146 | 372582 | 145052 | 545555 | 2624182 | 695961 | 1882 | 21788958 | 3462874 | 5124682 | 1591366 | 11166318 | 11818 | 36370572 | 30596928 | 85626 | 3975 | 52 | 7587 | 0 | 807030 | 21239 | 305 | 0 | 11129646 | 119383 | 40456 | 1485 | 53662 | 5025 | 746 | 19663 | 0 | 75444 | 26490 | 19377 |  | Unknown
WO-22270-217-A-1 | 4961 | 61103 | 436 | 469 | 4827 | 2764257 | 2458533 | 0 | 0 | 3736563 | 1297 | 2111 | 0 | 14497 | 32291 | 126558 | 370237 | 128228 | 515257 | 2459430 | 628020 | 0 | 20589884 | 3213941 | 5442785 | 1379006 | 10309220 | 9822 | 34627258 | 30356599 | 77522 | 0 | 106 | 6725 | 0 | 449416 | 19824 | 70 | 0 | 10349730 | 107665 | 31383 | 2050 | 47461 | 5187 | 703 | 5640 | 182 | 79191 | 23647 | 15123 |  | Unknown
WO-22270-218-A-1 | 5976 | 54756 | 486 | 224 | 3844 | 2262345 | 2236572 | 0 | 19 | 1694263 | 1077 | 1615 | 0 | 11841 | 44056 | 141010 | 331739 | 167267 | 596399 | 2791465 | 832272 | 0 | 22102196 | 3572074 | 6062329 | 1592824 | 10781459 | 12976 | 35667119 | 28942155 | 103059 | 628 | 482 | 7118 | 0 | 270234 | 25330 | 330 | 0 | 4535970 | 128451 | 19757 | 1294 | 12592 | 7771 | 2053 | 0 | 75 | 73645 | 23161 | 11889 | D,Q | Unknown
WO-22270-219-A-1 | 5598 | 101909 | 2236 | 359 | 4349 | 2728844 | 3087515 | 43 | 0 | 1997080 | 1026 | 700 | 0 | 4046 | 45844 | 162319 | 450121 | 200937 | 635989 | 3227442 | 985580 | 0 | 25175785 | 3672166 | 6853538 | 1608310 | 12721918 | 14625 | 36859882 | 31183282 | 109178 | 1707 | 36 | 6307 | 0 | 164155 | 25429 | 195 | 0 | 5330895 | 116486 | 17854 | 1239 | 26057 | 8005 | 1216 | 0 | 0 | 78119 | 23486 | 13433 |  | Unknown
WO-22270-220-A-1 | 5652 | 110541 | 930 | 126 | 3611 | 3049007 | 2627370 | 0 | 0 | 1504121 | 516 | 1610 | 0 | 3812 | 43123 | 132933 | 458514 | 159008 | 548267 | 2684486 | 794512 | 1937 | 21250524 | 3493487 | 5814175 | 1490656 | 11275577 | 11184 | 33822810 | 30032673 | 90086 | 3141 | 141 | 7163 | 0 | 226547 | 21441 | 292 | 0 | 4422984 | 120749 | 13532 | 1156 | 10021 | 6094 | 1006 | 10682 | 116 | 69977 | 20350 | 12322 |  | Unknown
AB-220715-01 LFB-H 12 | 36381095 | 98671 | 148393 | 216699 | 39316537 | 1772012 | 2105501 | 829757 | 3168897 | 361754 | 451190 | 1659065 | 1074141 | 1563996 | 106636 | 155772 | 232840 | 463823 | 1177803 | 5712844 | 2498642 | 12454 | 21976360 | 4028401 | 11098232 | 1741780 | 15111322 | 50643 | 35157116 | 26995358 | 262976 | 1642505 | 1404723 | 1609855 | 36228 | 19048452 | 31989001 | 1682250 | 5278846 | 2166880 | 174172 | 13076749 | 1951907 | 9744951 | 34687 | 3899485 | 418087 | 899203 | 30475713 | 36043861 | 23477223 |  | Unknown
AB-220715-01 LRB 23_INJ001 | 402647 | 54320 | 693 | 2106 | 193873 | 2098694 | 2145973 | 3324 | 14301 | 3888 | 1877 | 7688 | 4589 | 11674 | 102683 | 256403 | 330224 | 476218 | 1315732 | 6234700 | 2491489 | 7660 | 27014804 | 4192262 | 12901297 | 2037344 | 17624904 | 49716 | 39662450 | 33305233 | 267466 | 8009 | 9942 | 7963 | 0 | 134202 | 477765 | 18222 | 26506 | 15596 | 130912 | 66260 | 10061 | 59737 | 5168 | 27729 | 5708 | 4107 | 398280 | 468384 | 215321 |  | Unknown
AB-220715-01 LRB 23 | 93460 | 53457 | 426 | 0 | 24551 | 1659891 | 1571594 | 446 | 1618 | 10433 | 618 | 813 | 0 | 16281 | 104603 | 213783 | 254963 | 453195 | 1061182 | 5575634 | 2186411 | 11598 | 20842471 | 3299747 | 10513831 | 1617515 | 14217666 | 16664 | 30332378 | 27464005 | 236105 | 750 | 2016 | 6253 | 0 | 26992 | 162845 | 3314 | 24680 | 22345 | 80100 | 11760 | 3647 | 24824 | 5061 | 4881 | 3494 | 748 | 257419 | 227917 | 38219 |  | Unknown
AB-220715-01 LD 11 | 46827 | 54489 | 1121 | 147 | 7820 | 2691807 | 2966040 | 24 | 700 | 1736444 | 665 | 2014 | 12575 | 2778 | 48689 | 154958 | 452702 | 192560 | 610756 | 3095658 | 968559 | 12544 | 23851746 | 3831322 | 6647475 | 1667184 | 12323242 | 13637 | 35827354 | 31028782 | 108695 | 1150 | 1405 | 6992 | 0 | 154288 | 105028 | 1200 | 19240 | 5047543 | 119340 | 17365 | 3123 | 34040 | 4356 | 3328 | 0 | 0 | 158137 | 155769 | 23642 |  | Unknown
AB-220715-01 CCC-H 12 | 67132820 | 194827 | 240422 | 377617 | 68418576 | 1564101 | 1987061 | 1717726 | 5829225 | 738343 | 906823 | 2905206 | 2004732 | 2871604 | 92603 | 132950 | 201369 | 464159 | 1021654 | 5283209 | 2381097 | 16789 | 18249143 | 3864453 | 9563423 | 1520143 | 13445640 | 52383 | 31317683 | 22353940 | 267099 | 3125450 | 2581997 | 3154617 | 78787 | 32677781 | 55545293 | 2935265 | 9809087 | 4370039 | 209188 | 23575348 | 3129772 | 17434277 | 63595 | 7608544 | 885608 | 1776807 | 55762356 | 64199347 | 39515268 |  | Unknown
AB-220715-01 LFSM-H 11 | 33877302 | 92892 | 94439 | 275479 | 32206137 | 2047813 | 2503716 | 303454 | 1592427 | 1772733 | 152159 | 943257 | 485040 | 1163110 | 38308 | 101426 | 283991 | 155733 | 533729 | 2570384 | 794392 | 2503 | 18420698 | 2928486 | 5431109 | 1303000 | 10005229 | 10404 | 31492016 | 24496569 | 90253 | 747005 | 1786691 | 1954672 | 0 | 16774095 | 30204012 | 1667092 | 2424442 | 5038635 | 149980 | 8743862 | 1840911 | 4975936 | 27292 | 3101917 | 100473 | 386227 | 28960602 | 33131987 | 22199581 |  | Unknown
WO-22270-221-A-1 | 327231 | 48422 | 1134 | 309 | 20959 | 2204623 | 2213568 | 56 | 1022 | 1910886 | 1185 | 1762 | 0 | 10425 | 49986 | 146188 | 310052 | 197478 | 643498 | 2999326 | 945716 | 3690 | 21968010 | 3757084 | 6540357 | 1709687 | 11936740 | 7905 | 37118808 | 29994751 | 108003 | 2276 | 2974 | 6606 | 0 | 163902 | 539891 | 9007 | 0 | 4898854 | 121602 | 15234 | 2078 | 12803 | 6726 | 6235 | 2762 | 202 | 549754 | 739675 | 100230 |  | Unknown
WO-22270-222-A-1 | 117573 | 54135 | 666 | 345 | 13671 | 2048660 | 1886591 | 13 | 892 | 1500100 | 914 | 2177 | 0 | 4755 | 38128 | 134996 | 252600 | 153941 | 547460 | 2586570 | 802223 | 0 | 19031010 | 3300496 | 5659326 | 1394304 | 10703960 | 10046 | 33625122 | 27343427 | 81361 | 1356 | 1495 | 5568 | 0 | 118063 | 256473 | 2769 | 0 | 4088731 | 115643 | 21939 | 2578 | 46159 | 6538 | 3596 | 2565 | 246 | 374812 | 451662 | 32889 | D,Q | Unknown
WO-22270-223-A-1 | 56357 | 44046 | 854 | 159 | 12160 | 2306835 | 2242020 | 66 | 301 | 1690883 | 491 | 1706 | 0 | 21117 | 51354 | 145398 | 295444 | 195432 | 601342 | 2940838 | 956808 | 7540 | 21877461 | 3756828 | 6492293 | 1625551 | 11896518 | 14152 | 36519222 | 30167351 | 105189 | 0 | 1174 | 6563 | 0 | 101586 | 162486 | 1431 | 0 | 4804981 | 123881 | 25650 | 5251 | 38151 | 7315 | 2949 | 19131 | 211 | 340521 | 363697 | 18826 |  | Unknown
WO-22270-224-A-1 | 28983 | 51410 | 250 | 84 | 6567 | 2233609 | 2103128 | 0 | 520 | 1802416 | 773 | 1647 | 0 | 2872 | 45899 | 148616 | 289846 | 197599 | 617437 | 3042827 | 912321 | 2671 | 20876279 | 3569848 | 6240995 | 1546054 | 11117897 | 14602 | 36850187 | 29761044 | 106397 | 3302 | 909 | 6525 | 0 | 214795 | 110638 | 995 | 0 | 4933483 | 117826 | 24928 | 3819 | 23863 | 7666 | 2768 | 9313 | 120 | 305987 | 318724 | 14545 |  | Unknown
WO-22270-225-A-1 | 22627 | 48756 | 721 | 171 | 8893 | 2584551 | 2494007 | 0 | 80 | 1727911 | 495 | 1529 | 0 | 3217 | 47010 | 158089 | 390749 | 178774 | 581314 | 2928703 | 939533 | 2848 | 22968960 | 3562636 | 6278527 | 1586849 | 11814640 | 10087 | 36259506 | 30970398 | 97477 | 4454 | 255 | 7930 | 0 | 194548 | 82943 | 420 | 0 | 4785112 | 118832 | 19043 | 3608 | 41925 | 6742 | 1783 | 14436 | 694 | 274154 | 272325 | 16941 |  | Unknown
WO-22270-226-A-1 | 15222 | 110917 | 630 | 272 | 5599 | 2417391 | 2608656 | 0 | 163 | 1949239 | 1286 | 895 | 0 | 17495 | 48364 | 142178 | 316378 | 206238 | 607152 | 3028326 | 982839 | 1728 | 21528776 | 3614861 | 6307692 | 1582656 | 11399572 | 15159 | 34828417 | 30742043 | 105271 | 1415 | 358 | 6222 | 0 | 156407 | 56994 | 623 | 0 | 5239862 | 116938 | 22761 | 2827 | 32118 | 5852 | 1692 | 19175 | 243 | 243343 | 233486 | 11284 |  | Unknown
AB-220715-01 LRB 24_INJ001 | 15379 | 44414 | 565 | 0 | 7990 | 2225709 | 2290815 | 96 | 102 | 16337 | 787 | 2672 | 12140 | 3083 | 110469 | 267599 | 338576 | 491306 | 1362996 | 6449184 | 2594506 | 0 | 27937892 | 4208453 | 13271086 | 2089775 | 18250216 | 52145 | 40894237 | 35837291 | 262966 | 1558 | 631 | 7656 | 0 | 23395 | 53067 | 348 | 0 | 46009 | 131271 | 8731 | 3142 | 22634 | 5314 | 1689 | 16204 | 257 | 288790 | 235985 | 14950 | D,Q | Unknown
AB-220715-01 LRB 24 | 10088 | 38928 | 369 | 0 | 6529 | 1683999 | 1610896 | 204 | 146 | 4986 | 589 | 573 | 9220 | 15961 | 67176 | 195250 | 275126 | 281158 | 993605 | 4386362 | 1527220 | 9287 | 20882693 | 3206341 | 9800208 | 1581659 | 13529513 | 12666 | 30643917 | 26524847 | 159622 | 2005 | 435 | 7826 | 0 | 20739 | 38331 | 450 | 0 | 10622 | 96302 | 8013 | 3249 | 18008 | 6804 | 1511 | 4152 | 40 | 314319 | 199915 | 12003 |  | Unknown
AB-220715-01 CCC-H 13 | 67004759 | 191770 | 252607 | 388877 | 71610730 | 1548163 | 2075909 | 1807160 | 5949231 | 717339 | 948636 | 3012853 | 1965897 | 2876286 | 99354 | 129878 | 201832 | 469969 | 1024841 | 5578768 | 2465554 | 14631 | 18796639 | 3741444 | 10213630 | 1505355 | 13836837 | 49744 | 31652205 | 23151359 | 288393 | 3226640 | 2646721 | 3216405 | 73797 | 34114036 | 58666982 | 2870342 | 10294949 | 4400492 | 205025 | 24009307 | 3277793 | 18721523 | 50506 | 7452157 | 924004 | 1833859 | 55550190 | 66674518 | 41008050 |  | Unknown
AB-220715-01 QCS 2 | 39795487 | 128633 | 179751 | 268643 | 44786197 | 1883348 | 2188869 | 17882 | 63492 | 474535 | 609328 | 1938473 | 1327902 | 1748193 | 121508 | 175128 | 266751 | 604229 | 1333423 | 7234142 | 3117107 | 20796 | 25271347 | 4232125 | 12800782 | 1829471 | 16710888 | 62716 | 34739142 | 29282534 | 360443 | 2205882 | 39271 | 38025 | 49099 | 22625606 | 35249778 | 55636 | 7078152 | 2865614 | 146614 | 15721824 | 48501 | 11856062 | 22248 | 99151 | 655909 | 1223248 | 1308224 | 1735494 | 27986677 |  | Unknown