## Data Flags Legend
High internal standard recovery, corresponding analyte in sample. Reviewed and not disqualifying | IS1 | Internal standard recovery was above the method acceptance limits of 150%. Target analytes were detected in the sample. Impact on results is not expected based on analyst review. See comments. | Analyst and validator concurrence required to release data. If this is not met, reanalysis and/or repreparation is required.
High internal standard recovery, no analyte in sample | IS2 | Internal standard recovery was above the method acceptance limits. No target analytes were detected in the sample. | Analyst and validator concurrence required to release data. If this is not met, reanalysis and/or repreparation is required.
Low internal standard recovery, corresponding analyte in sample. Reviewed and not disqualifying | IS3 | Internal standard recovery was below the method acceptance limit of 50%. Target analytes were detected in the sample. Impact on results is not expected based on analyst review. See comments. | Analyst and validator concurrence required to release data. If this is not met, reanalysis and/or repreparation is required.
Low internal standard recovery, no analyte in sample. | IS4 | Internal standard recovery was below the method acceptance limits. No target analytes were detected in the sample. | Analyst and validator concurrence required to release data. If this is not met, reanalysis and/or repreparation is required.
Internal standards outside limit, disqualifying. | IS5 | Internal standard recovery was outside of method limits of 50-150%, and upon review is expected to impact data quality. Data not useable. [Margie would rather we said things like 'data biased low', but my impression is you prefer a go/no-go for data, right?] | Sample re-analysis required; If failure persists, re-prepare sample and re-analyze.
Laboratory duplicate precision failure | D | RPD/RSD for duplicate analysis exceeded method acceptance. | Sample re-analysis required; If failure persists, further troubleshooting is required to improve precision and all samples analyzed since previous passing duplicate must be re-analyzed.
Laboratory Reagent Blank (LRB) failure | B | Result is >50% of MRL. | Flag LRB failure and sample results within 10-fold of the LRB concentration as biased high.
Continuing Calibration Check (CCC) failure | V1 | Continuing calibration check (CCC) recovery was outside the method acceptance limits (70-130%). | Samples analyzed since previous passing CCC must be re-analyzed. If failure persists, re-calibration is required. Upon recalibration, all other QC must be re-analyzed as well.
Quality Control Sample (QCS) failure | V2 | Quality control sample (QCS) recovery was outside of the method acceptance limits (70-130%). | Re-analysis of QCS required; if failure persists, re-preparation of QCS is required and re-calibration may be necessary. Upon recalibration, all other QC must be re-analyzed as well. No sample results may be reported if QCS failure occurs in batch.
Inter-Batch Consistency check (IBC) failure | V3 | The inter-batch consistency check (IBC) sample was outside of acceptance criteria (80-120%). This suggests a systematic bias for the analytical batch relative to prior sample batches. | Compare peak areas of calibration standards to those of previous batches to determine if source of failure is the calibration curve or the IBC itself. For IBC failure, re-preparation and/or re-analysis of IBC is required. For calibration curve failure, re-preparation and/or re-analysis of calibration curve is required. Upon recalibration, all other QC must be re-analyzed as well. No sample results may be reported if IBC failure occurs in batch.
High Laboratory Fortified Blank (LFB) recovery | L1 | LFB results >150% for LFB-Low or >130% for LFB-High. Affected sample results biased high. | Re-analysis of LFB required; If failure persists, re-prepare LFB and re-analyze. Re-preparation of spiking solution (PDS/SDS) may also be necessary. Affected samples, i.e. all samples analyzed since previous passing LFB, must be flagged.
Low Laboratory Fortified Blank (LFB) recovery | L2 | LFB results <50% for LFB-Low or <70% for LFB-High. Affected sample results biased low. | Re-analysis of LFB required; If failure persists, re-prepare LFB and re-analyze. Re-preparation of spiking solution (PDS/SDS) may also be necessary. Affected samples, i.e. all samples analyzed since previous passing LFB, must be flagged.
Laboratory Fortified Sample Matrix (LFSM) failure | M1 | LFSM recovery outside | Absent LFB issues:
|  | 70-130% for LFSM-High or | Re-analysis of LFSM is required; Passing LFB confirms proper preparation and analysis so no further troubleshooting is necessary. If failure persists, samples analyzed since previous passing LFSM of same type (-Low/-High) must be flagged.
50-150% for LFSM-Low. Sample results may be biased proportionally due to matrix interferences.
Coordinating LFSM/LFB failures:
|  | Note: Simultaneous failure of LFB and LFSM indicates systematic bias in the preparation or analysis processes. | Re-preparation and re-analysis of LFB/LFSM required. Re-preparation of spiking solution (PDS/SDS) may also be necessary. If failure persists, samples analyzed since previous passing set of LFSM and LFB must be flagged.
LFSM interference | M2 | Spike recovery unable to be calculated due to other factors, e.g. spiking concentration too low. | Re-preparation and re-analysis of LFSM required. If failure persists, samples analyzed since previous passing LFSM of same type (-Low/-High) must be flagged. Analyst must describe in comments.
Upper calibration limit exceeded | UL | Sample results are above the upper calibration limit. | Sample must be re-analyzed at a higher dilution factor to obtain results within calibration range. Adjacent samples must be inspected for signs of carryover (see flag “C”).
Lower calibration limit exceeded for diluted sample | LL | Sample dilution factor is greater than “2” but results are <MRL. Sample was therefore over-diluted. | Sample must be re-analyzed at a lower dilution factor, undiluted if necessary, to obtain results within calibration range.
Carryover/Contamination observed | C | Carryover and/or contamination observed in sample. Sample results are suspected to be biased high. | Sample re-analysis is required.

## Template
Batch Name: | [5 empty] | SA: | Area: | SAH: | AreaH: | Last Acquisition Date: |  |  | Sample Count: |  | Analytes | PFBA | L-PFBS | HFPO-DA | PFHxA | L-PFHxS | PFOA | PFNA | L-PFOS | PFDA | HCal Std |  | Sheet1_1 | [15 empty] | Verifier Count | Validator Count | MPFBA | M3PFBS | MPHFPO-DA | M5PFHxA | M3PFHxS | M8PFOA | M9PFNA | M8PFOS | M6PFDA |  | Type | [7 empty] | D (LD) | K (LRB) | M (Cal) | CCC/QCS | S (LFSM) | Z (LFB) | IBC | <MRL | <MRL OVRD
2023-01-18 | [5 empty] | Sheet2 | Sheet4 | 8 | 8 | 2023-01-31 07:17 |  |  | 365 |  | MRL> | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 5 | 0 |  | Sheet3 | [15 empty] | 0 | 0 | 365082.5 | 740479.5 | 1224004.5 | 6800994.5 | 3376990.5 | 27604846.5 | 31889867.5 | 3960175.5 | 32794787 |  | U |  |  |  |  | Quality Control Maximum Recovery %: |  |  | 30% | 50% | 30% | 30% | 30% | 30% | 20% | 50% | ON
[6 empty] | Measured Concentration, ng/L | [9 empty] | Concentration, ng/L, Corrected for Sample Dilution | [10 empty] | Sheet4 |  |  | Flags | [8 empty] | Re-Analysis Flag |  |  | Verifier/Validator Initials |  |  | Area (Internal Standard) | [17 empty] | Measured Concentration, ng/L | [9 empty] | Quality Control Checks | [18 empty] | These fields are editable with password: 'fixit'
# | Time Stamp: | Sample ID: | # | Type | QC? | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | Dilution Factor: | C-PFBA | C-PFBS | C-HFPO-DA | C-PFHxA | C-PFHxS | C-PFOA | C-PFNA | C-PFOS | C-PFDA | True Value | True Value 2 | True Value 3 | Sample Type | Flags (38) | F-PFBA | F-PFBS | F-HFPO-DA | F-PFHxA | F-PFHxS | F-PFOA | F-PFNA | F-PFOS | F-PFDA |  | List for drop down flags | Re-Analysis Flags | Sample Comments | Verifier Comments | Validator Comments | MPFBA | MPFBS | MPHFPO-DA | MPFHxA | MPFHxS | MPFOA | MPFNA | MPFOS | MPFDA | IS Comments | M&Q | K&Z | D | S | # | QC Samples | Type | Dilution Factor: | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | True Value | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | PFBA | PFBS | GenX | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | Pairing
1 | 2023-01-18 10:32 | AB-230117-02 Cal Blank 1 | 0 | M | QC | 0 | #N/A | 0 | 0.566 | #N/A | 2.166 | 2.445 | #N/A | 2.511 | 1 | 0 | #N/A | 0 | 0.566 | #N/A | 2.166 | 2.445 | #N/A | 2.511 | Blank 1 | Blank 1 | NA | Calib Std | No TrueV! | IS1 | [9 empty] | IS1 | R1 |  |  |  | 632991 | 791404 | 999100 | 6168538 | 3461600 | 27779178 | 30410864 | 4273192 | 30301110 |  | 1 |  |  |  | 1 | AB-230117-02 Cal Blank 1 | M | 1 | 0 | #N/A | 0 | 0.566 | #N/A | 2.166 | 2.445 | #N/A | 2.511 | Blank 1
2 | 2023-01-18 11:00 | AB-230117-02 Cal Std 1 1 | 1 | M | QC | 0 | #N/A | 1.078 | 1.415 | #N/A | 1.437 | 1.368 | #N/A | 1.325 | 1 | 0 | #N/A | 1.078 | 1.415 | #N/A | 1.437 | 1.368 | #N/A | 1.325 | Std 1 1 | Std 1 1 | NA | Calib Std | No TrueV! | [10 empty] | IS2 | R2 |  |  |  | 374515 | 474992 | 1412490 | 7254807 | 3721715 | 31153664 | 35852585 | 4544352 | 40439150 |  | 2 |  |  |  | 2 | AB-230117-02 Cal Std 1 1 | M | 1 | 0 | #N/A | 1.078 | 1.415 | #N/A | 1.437 | 1.368 | #N/A | 1.325 | Std 1 1
3 | 2023-01-18 11:27 | AB-230117-02 Cal Std 2 1 | 2 | M | QC | 0 | #N/A | 2.003 | 2.095 | #N/A | 2.071 | 2.05 | #N/A | 2.016 | 1 | 0 | #N/A | 2.003 | 2.095 | #N/A | 2.071 | 2.05 | #N/A | 2.016 | Std 2 1 | Std 2 1 | NA | Calib Std | No TrueV! | [10 empty] | IS3 | R3 |  |  |  | 334728 | 914458 | 1300088 | 6943986 | 3812773 | 31781766 | 35137464 | 4444601 | 39003692 |  | 3 |  |  |  | 3 | AB-230117-02 Cal Std 2 1 | M | 1 | 0 | #N/A | 2.003 | 2.095 | #N/A | 2.071 | 2.05 | #N/A | 2.016 | Std 2 1
4 | 2023-01-18 11:54 | AB-230117-02 Cal Std 3 1 | 3 | M | QC | 5.762 | #N/A | 4.412 | 3.86 | #N/A | 3.96 | 3.843 | #N/A | 4.043 | 1 | 5.762 | #N/A | 4.412 | 3.86 | #N/A | 3.96 | 3.843 | #N/A | 4.043 | Std 3 1 | Std 3 1 | NA | Calib Std | No TrueV! | [10 empty] | IS4 |  | Within 50-150% limit |  |  | 433302 | 964959 | 1271882 | 7322397 | 3747583 | 27679054 | 34506802 | 4174088 | 41799297 |  | 4 |  |  |  | 4 | AB-230117-02 Cal Std 3 1 | M | 1 | 5.762 | #N/A | 4.412 | 3.86 | #N/A | 3.96 | 3.843 | #N/A | 4.043 | Std 3 1 | [20 empty] | Within 50-150% tolerance < MRL
5 | 2023-01-18 12:21 | AB-230117-02 Cal Std 4 1 | 4 | M | QC | 11.646 | #N/A | 9.51 | 9.126 | #N/A | 9.194 | 9.265 | #N/A | 9.394 | 1 | 11.646 | #N/A | 9.51 | 9.126 | #N/A | 9.194 | 9.265 | #N/A | 9.394 | Std 4 1 | Std 4 1 | NA | Calib Std | No TrueV! | [10 empty] | IS5 |  |  |  |  | 389333 | 775858 | 1286073 | 6862671 | 3751819 | 31867700 | 37211667 | 4570958 | 42979199 |  | 5 |  |  |  | 5 | AB-230117-02 Cal Std 4 1 | M | 1 | 11.646 | #N/A | 9.51 | 9.126 | #N/A | 9.194 | 9.265 | #N/A | 9.394 | Std 4 1
6 | 2023-01-18 12:48 | AB-230117-02 Cal Std 5 1 | 5 | M | QC | 18.831 | #N/A | 18.606 | 17.853 | #N/A | 18.158 | 18.63 | #N/A | 18.415 | 1 | 18.831 | #N/A | 18.606 | 17.853 | #N/A | 18.158 | 18.63 | #N/A | 18.415 | Std 5 1 | Std 5 1 | NA | Calib Std | No TrueV! | [10 empty] | D |  |  |  |  | 384293 | 536479 | 1283661 | 6974866 | 3716010 | 27530639 | 33993178 | 4177449 | 41768495 |  | 6 |  |  |  | 6 | AB-230117-02 Cal Std 5 1 | M | 1 | 18.831 | #N/A | 18.606 | 17.853 | #N/A | 18.158 | 18.63 | #N/A | 18.415 | Std 5 1
7 | 2023-01-18 13:15 | AB-230117-02 Cal Std 6 1 | 6 | M | QC | 37.308 | #N/A | 36.592 | 35.14 | #N/A | 36.553 | 37.996 | #N/A | 37.389 | 1 | 37.308 | #N/A | 36.592 | 35.14 | #N/A | 36.553 | 37.996 | #N/A | 37.389 | Std 6 1 | Std 6 1 | NA | Calib Std | No TrueV! | [10 empty] | B |  |  |  |  | 411914 | 715733 | 1184376 | 6780447 | 3596018 | 30647728 | 32905329 | 3900993 | 38998738 |  | 7 |  |  |  | 7 | AB-230117-02 Cal Std 6 1 | M | 1 | 37.308 | #N/A | 36.592 | 35.14 | #N/A | 36.553 | 37.996 | #N/A | 37.389 | Std 6 1
8 | 2023-01-18 13:42 | AB-230117-02 Cal Std 7 1 | 7 | M | QC | 95.665 | #N/A | 95.727 | 93.802 | #N/A | 95.968 | 96.468 | #N/A | 97.43 | 1 | 95.665 | #N/A | 95.727 | 93.802 | #N/A | 95.968 | 96.468 | #N/A | 97.43 | Std 7 1 | Std 7 1 | NA | Calib Std | No TrueV! | [10 empty] | V1 |  |  |  |  | 266916 | 737784 | 1006545 | 5695371 | 3292381 | 27053869 | 32561100 | 4019358 | 33433459 |  | 8 |  |  |  | 8 | AB-230117-02 Cal Std 7 1 | M | 1 | 95.665 | #N/A | 95.727 | 93.802 | #N/A | 95.968 | 96.468 | #N/A | 97.43 | Std 7 1
9 | 2023-01-18 14:09 | AB-230117-02 Cal Std 8 1 | 8 | M | QC | 163.759 | #N/A | 197.427 | 191.961 | #N/A | 187.35 | 189.259 | #N/A | 193.114 | 1 | 163.759 | #N/A | 197.427 | 191.961 | #N/A | 187.35 | 189.259 | #N/A | 193.114 | Std 8 1 | Std 8 1 | NA | Calib Std | No TrueV! | [10 empty] | V2 |  |  |  |  | 355650 | 943092 | 1298685 | 7138944 | 3257272 | 27803507 | 31218635 | 3648036 | 32156115 |  | 9 |  |  |  | 9 | AB-230117-02 Cal Std 8 1 | M | 1 | 163.759 | #N/A | 197.427 | 191.961 | #N/A | 187.35 | 189.259 | #N/A | 193.114 | Std 8 1
10 | 2023-01-18 14:36 | AB-230117-02 Cal Std 9 1 | 9 | M | QC | 328.948 | #N/A | 402.773 | 384.687 | #N/A | 378.816 | 386.336 | #N/A | 385.495 | 1 | 328.948 | #N/A | 402.773 | 384.687 | #N/A | 378.816 | 386.336 | #N/A | 385.495 | Std 9 1 | Std 9 1 | NA | Calib Std | No TrueV! | [10 empty] | V3 |  |  |  |  | 307127 | 823849 | 1263633 | 6821542 | 2689261 | 23663020 | 25709605 | 2894577 | 26016671 |  | 10 |  |  |  | 10 | AB-230117-02 Cal Std 9 1 | M | 1 | 328.948 | #N/A | 402.773 | 384.687 | #N/A | 378.816 | 386.336 | #N/A | 385.495 | Std 9 1
11 | 2023-01-18 15:03 | AB-230117-02 Cal Std 10 1 | 10 | M | QC | 951.413 | #N/A | 1029.861 | 1001.157 | #N/A | 980.955 | 982.021 | #N/A | 972.285 | 1 | 951.413 | #N/A | 1029.861 | 1001.157 | #N/A | 980.955 | 982.021 | #N/A | 972.285 | Std 10 1 | Std 10 1 | NA | Calib Std | No TrueV! | [7 empty] | IS3 | IS3 |  | L1 |  |  |  |  | 355318 | 735204 | 1166714 | 5884437 | 1848502 | 16194891 | 16620897 | 1890722 | 16098522 |  | 11 |  |  |  | 11 | AB-230117-02 Cal Std 10 1 | M | 1 | 951.413 | #N/A | 1029.861 | 1001.157 | #N/A | 980.955 | 982.021 | #N/A | 972.285 | Std 10 1
12 | 2023-01-18 15:30 | AB-230117-02 Cal Std 11 1 | 11 | M | QC | 1822.989 | #N/A | 2030.625 | 1991.195 | #N/A | 1905.749 | 1944.508 | #N/A | 1949.845 | 1 | 1822.989 | #N/A | 2030.625 | 1991.195 | #N/A | 1905.749 | 1944.508 | #N/A | 1949.845 | Std 11 1 | Std 11 1 | NA | Calib Std | No TrueV! |  |  |  |  | IS3 | IS3 | IS3 | IS3 | IS3 |  | L2 |  |  |  |  | 399743 | 743175 | 930811 | 4663536 | 1182015 | 10669119 | 10205299 | 1117397 | 9992543 |  | 12 |  |  |  | 12 | AB-230117-02 Cal Std 11 1 | M | 1 | 1822.989 | #N/A | 2030.625 | 1991.195 | #N/A | 1905.749 | 1944.508 | #N/A | 1949.845 | Std 11 1
13 | 2023-01-18 15:57 | AB-230117-02 Cal Std 12 1 | 12 | M | QC | 4173.586 | #N/A | 4196.925 | 4184.365 | #N/A | 4022.699 | 3980.463 | #N/A | 4002.829 | 1 | 4173.586 | #N/A | 4196.925 | 4184.365 | #N/A | 4022.699 | 3980.463 | #N/A | 4002.829 | Std 12 1 | Std 12 1 | NA | Calib Std | No TrueV! |  |  |  | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  | M1 |  |  |  |  | 344196 | 624962 | 650094 | 3204749 | 801246 | 6088887 | 5450072 | 717869 | 5309212 |  | 13 |  |  |  | 13 | AB-230117-02 Cal Std 12 1 | M | 1 | 4173.586 | #N/A | 4196.925 | 4184.365 | #N/A | 4022.699 | 3980.463 | #N/A | 4002.829 | Std 12 1
14 | 2023-01-18 16:24 | AB-230117-02 Cal Std 13 1 | 13 | M | QC | 10164.091 | #N/A | 9751.461 | 9860.345 | #N/A | 10134.089 | 10124.793 | #N/A | 10103.42 | 1 | 10164.091 | #N/A | 9751.461 | 9860.345 | #N/A | 10134.089 | 10124.793 | #N/A | 10103.42 | Std 13 1 | Std 13 1 | NA | Calib Std | No TrueV! |  |  | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  | M2 |  |  |  |  | 307430 | 398687 | 337049 | 1609916 | 462298 | 2671403 | 2633403 | 409631 | 2537051 |  | 14 |  |  |  | 14 | AB-230117-02 Cal Std 13 1 | M | 1 | 10164.091 | #N/A | 9751.461 | 9860.345 | #N/A | 10134.089 | 10124.793 | #N/A | 10103.42 | Std 13 1
15 | 2023-01-18 17:46 | AB-230117-02 LRB 1 | 18 | K | QC | 0 | #N/A | 1.125 | 0 | #N/A | 3.537 | 4.565 | #N/A | 4.651 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [5 empty] | b | b | b | b | b |  | UL |  |  |  |  | 377315 | 881418 | 1295992 | 7132049 | 3663879 | 30802350 | 35806446 | 4205808 | 37221036 |  |  | 15 |  |  | 17 | AB-230117-02 QCS 1 | Q | 1 | 215.739 | #N/A | 219.599 | 212.607 | #N/A | 211.384 | 210.554 | #N/A | 213.968 | 1 | 21574% |  | 21960% | 21261% |  | 21138% | 21055% |  | 21397% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL
16 | 2023-01-18 18:13 | AB-230117-02 LFB-H 1 | 44 | Z | QC | 183.485 | #N/A | 178.384 | 194.206 | #N/A | 198.792 | 204.981 | #N/A | 191.802 | 1 | 183.485 | #N/A | 178.384 | 194.206 | #N/A | 198.792 | 204.981 | #N/A | 191.802 | H 1 | H 1 | NA | Blank Spike | No TrueV! | [10 empty] | LL |  |  |  |  | 344493 | 625103 | 1214377 | 6609916 | 3203708 | 27298231 | 29701401 | 3380199 | 31292462 |  |  | 16 |  |  | 18 | AB-230117-02 IBC 1 | Q | 1 | 169.148 | #N/A | 189.323 | 190.496 | #N/A | 191.07 | 196.495 | #N/A | 195.489 | 1 | 16915% |  | 18932% | 19050% |  | 19107% | 19650% |  | 19549% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL
17 | 2023-01-18 18:40 | AB-230117-02 QCS 1 | 72 | Q | QC | 215.739 | #N/A | 219.599 | 212.607 | #N/A | 211.384 | 210.554 | #N/A | 213.968 | 1 | 215.739 | #N/A | 219.599 | 212.607 | #N/A | 211.384 | 210.554 | #N/A | 213.968 | 1 | NA | 1 | QCS | [11 empty] | C |  |  |  |  | 368604 | 761275 | 1279834 | 6912932 | 3056412 | 26340525 | 27591477 | 2988900 | 34623193 |  | 17 |  |  |  | 19 | AB-230117-02 CCC-H 1 | Q | 1 | 331.373 | #N/A | 392.93 | 384.696 | #N/A | 381.982 | 379.732 | #N/A | 378.211 | 2 | 16569% |  | 19646% | 19235% |  | 19099% | 18987% |  | 18911% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL
18 | 2023-01-18 19:07 | AB-230117-02 IBC 1 | 17 | Q | QC | 169.148 | #N/A | 189.323 | 190.496 | #N/A | 191.07 | 196.495 | #N/A | 195.489 | 1 | 169.148 | #N/A | 189.323 | 190.496 | #N/A | 191.07 | 196.495 | #N/A | 195.489 | 1 | NA | 1 | IBC | [16 empty] | 343074 | 694321 | 1265968 | 6831085 | 3145626 | 26669533 | 30178374 | 3525478 | 33994344 |  | 18 |  |  |  | 45 | AB-230117-02 CCC-H 2 | Q | 1 | 342.723 | #N/A | 403.865 | 379.785 | #N/A | 381.16 | 387.516 | #N/A | 380.815 | 2 | 17136% |  | 20193% | 18989% |  | 19058% | 19376% |  | 19041% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL
19 | 2023-01-18 19:34 | AB-230117-02 CCC-H 1 | 42 | Q | QC | 331.373 | #N/A | 392.93 | 384.696 | #N/A | 381.982 | 379.732 | #N/A | 378.211 | 1 | 331.373 | #N/A | 392.93 | 384.696 | #N/A | 381.982 | 379.732 | #N/A | 378.211 | 2 | 02 | NA | CCC | [16 empty] | 360810 | 782284 | 1247833 | 6532921 | 2601336 | 22963695 | 23069723 | 2474341 | 24555753 |  | 19 |  |  |  | 73 | AB-230117-02 CCC-H 3 | Q | 1 | 366.428 | #N/A | 400.863 | 377.097 | #N/A | 377.585 | 383.968 | #N/A | 393.376 | 2 | 18321% |  | 20043% | 18855% |  | 18879% | 19198% |  | 19669% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL
20 | 2023-01-18 20:55 | WO-23003-01-A-1 @10 | 70 | U |  | 429.802 | #N/A | 0 | 1.148 | #N/A | 1.127 | 1.291 | #N/A | 1.464 | 1 | 429.802 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS3 | IS4 | [25 empty] | 110 | AB-230117-02 CCC-H 4 | Q | 1 | 333.561 | #N/A | 394.235 | 375.413 | #N/A | 375.67 | 378.57 | #N/A | 378.791 | 2 | 16678% |  | 19712% | 18771% |  | 18784% | 18928% |  | 18940% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL
21 | 2023-01-18 21:22 | WO-23003-02-A-1 | 15 | U |  | 458.755 | #N/A | 2.86 | 0 | #N/A | 0.97 | 1.012 | #N/A | 1.402 | 1 | 458.755 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS3 | IS4 | [11 empty] | 39407 | 259309 | 447622 | 2457350 | 2189835 | 19370109 | 26032384 | 4064659 | 31419146 | [5 empty] | 138 | AB-230117-02 CCC-H 5 | Q | 1 | 309.052 | #N/A | 375.935 | 355.016 | #N/A | 353.128 | 357.216 | #N/A | 355.469 | 2 | 15453% |  | 18797% | 17751% |  | 17656% | 17861% |  | 17773% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL
22 | 2023-01-18 21:49 | WO-23003-03-A-1 | 43 | U |  | 463.337 | #N/A | 0 | 0.902 | #N/A | 0.843 | 0.908 | #N/A | 1.12 | 1 | 463.337 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | [14 empty] | 96340 | 435784 | 662710 | 3443525 | 2930056 | 24292510 | 28062070 | 4041821 | 28375547 | [5 empty] | 166 | AB-230117-02 CCC-H 6 | Q | 1 | 304.261 | #N/A | 352.766 | 339.3 | #N/A | 333.08 | 341.007 | #N/A | 341.563 | 2 | 15213% |  | 17638% | 16965% |  | 16654% | 17050% |  | 17078% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL
23 | 2023-01-18 22:16 | WO-23003-04-A-1 | 71 | U |  | 434.867 | #N/A | 0 | 0 | #N/A | 0.757 | 0.772 | #N/A | 0.822 | 1 | 434.867 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | IS4 | [10 empty] | 8462 | 170395 | 231614 | 1372283 | 1635815 | 16332245 | 26011891 | 3964193 | 34514886 | [5 empty] | 194 | AB-230117-02 CCC-H 7 | Q | 1 | 307.839 | #N/A | 372.196 | 355.298 | #N/A | 349.262 | 350.888 | #N/A | 358.688 | 2 | 15392% |  | 18610% | 17765% |  | 17463% | 17544% |  | 17934% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL
24 | 2023-01-18 22:43 | WO-23003-05-A-1 | 45 | U |  | 511.067 | #N/A | 0 | 0 | #N/A | 0.675 | 0.744 | #N/A | 1.093 | 1 | 511.067 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | IS4 | [10 empty] | 16128 | 207211 | 279961 | 1572217 | 1528397 | 15172803 | 24593314 | 3859187 | 33872567 | [5 empty] | 228 | AB-230117-02 CCC-H 8 | Q | 1 | 306.068 | #N/A | 374.743 | 347.778 | #N/A | 345.655 | 353.314 | #N/A | 346.394 | 2 | 15303% |  | 18737% | 17389% |  | 17283% | 17666% |  | 17320% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL
25 | 2023-01-18 23:10 | WO-23003-06-A-1 | 73 | U |  | 462.818 | #N/A | 0 | 0 | #N/A | 0.657 | 0.771 | #N/A | 0.769 | 1 | 462.818 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 10215 | 120550 | 259336 | 1375153 | 1284902 | 13088314 | 22472696 | 3521650 | 32033462 | [5 empty] | 256 | AB-230117-02 CCC-H 9 | Q | 1 | 302.616 | #N/A | 351.796 | 350.18 | #N/A | 337.277 | 341.66 | #N/A | 345.726 | 2 | 15131% |  | 17590% | 17509% |  | 16864% | 17083% |  | 17286% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL
26 | 2023-01-18 23:37 | WO-23003-07-A-1 | 29 | U |  | 0 | #N/A | 0 | 0 | #N/A | 1.008 | 1.8 | #N/A | 3898.943 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 3898.943 | NA | NA | NA | Unknown | [9 empty] | IS4 | [6 empty] | 330783 | 816905 | 1141879 | 5638969 | 3383417 | 29457912 | 33917331 | 3844275 | 11240896 | [5 empty] | 284 | AB-230117-02 CCC-H 10 | Q | 1 | 301.419 | #N/A | 356.398 | 350.601 | #N/A | 341.141 | 348.019 | #N/A | 342.023 | 2 | 15071% |  | 17820% | 17530% |  | 17057% | 17401% |  | 17101% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL
27 | 2023-01-19 00:05 | WO-23003-08-A-1 | 30 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.97 | 1.8 | #N/A | 4309.18 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4309.18 | NA | NA | NA | Unknown | [9 empty] | IS4 | [6 empty] | 342805 | 763057 | 1182503 | 6072207 | 3440710 | 29295926 | 33862932 | 3787128 | 10905532 | [5 empty] | 312 | AB-230117-02 CCC-H 11 | Q | 1 | 309.141 | #N/A | 366.619 | 355.509 | #N/A | 355.982 | 358.493 | #N/A | 360.917 | 2 | 15457% |  | 18331% | 17775% |  | 17799% | 17925% |  | 18046% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL
28 | 2023-01-19 00:32 | WO-23003-09-A-1 | 41 | U |  | 0 | #N/A | 0 | 0 | #N/A | 1.051 | 1.666 | #N/A | 3931.95 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 3931.95 | NA | NA | NA | Unknown | [9 empty] | IS4 | [6 empty] | 285021 | 605447 | 1180678 | 4196726 | 3398277 | 29795410 | 34190517 | 3893301 | 11603574 | [5 empty] | 326 | AB-230117-02 CCC-H 12 | Q | 1 | 343.096 | #N/A | 391.777 | 383.581 | #N/A | 375.367 | 373.921 | #N/A | 372.448 | 2 | 17155% |  | 19589% | 19179% |  | 18768% | 18696% |  | 18622% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL |  |  | Carryover from prior sample
29 | 2023-01-19 00:59 | WO-23003-10-A-1 | 56 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.95 | 1.794 | #N/A | 4432.787 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4432.787 | NA | NA | NA | Unknown | [9 empty] | IS4 | [6 empty] | 307870 | 677764 | 1147085 | 6539975 | 3351815 | 29682382 | 34149551 | 3876799 | 10743050 | [5 empty] | 327 | AB-230117-02 QCS 2 | Q | 1 | 194.121 | #N/A | 224.431 | 212.761 | #N/A | 218.131 | 232.69 | #N/A | 220.165 | 2 | 9706% |  | 11222% | 10638% |  | 10907% | 11634% |  | 11008% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL
30 | 2023-01-19 01:26 | LRB2-Inj1 | 57 | K | QC | 0 | #N/A | 0 | 0 | #N/A | 0.46 | 0.414 | #N/A | 204.132 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 204.132 | NA | NA | NA | Blank | [9 empty] | b | [6 empty] | 404157 | 709896 | 1239312 | 6426584 | 3550835 | 30052213 | 34535528 | 4082623 | 38121486 |  |  | 30 |  |  | 363 | QCS-2-Reinject | Q | 1 | 174.903 | #N/A | 211.41 | 200.785 | #N/A | 189.964 | 202.109 | #N/A | 205.073 | 2 | 8745% |  | 10570% | 10039% |  | 9498% | 10105% |  | 10254% | FAIL |  | FAIL | FAIL |  | FAIL | FAIL |  | FAIL
31 | 2023-01-19 01:53 | AB-230117-02 LRB 2 | 68 | K | QC | 0 | #N/A | 0 | 0.364 | #N/A | 0.466 | 0.349 | #N/A | 7.368 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 7.368 | NA | NA | NA | Blank | [9 empty] | b |  |  |  | Only detects within 10-fold are carryover and flagged |  |  | 287597 | 579252 | 1250979 | 6983145 | 3186044 | 26557316 | 32251309 | 3692244 | 38345817 |  |  | 31 |  |  | 365 | CCC-H-12-reinject | Q | 1 | 305.952 | #N/A | 364.623 | 349.782 | #N/A | 351.819 | 354.686 | #N/A | 353.248 | #VALUE!
32 | 2023-01-19 02:20 | WO-23003-11-A-1 | 69 | U |  | 0 | #N/A | 0.293 | 0 | #N/A | 0.861 | 1.711 | #N/A | 4146.818 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4146.818 | NA | NA | NA | Unknown | [9 empty] | IS4 | [6 empty] | 295229 | 706630 | 1151204 | 6441159 | 3316637 | 28043019 | 31943220 | 3669322 | 10981126
33 | 2023-01-19 02:47 | WO-23003-12-A-1 | 74 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.96 | 1.709 | #N/A | 4137.805 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4137.805 | NA | NA | NA | Unknown | [9 empty] | IS4 | [6 empty] | 286308 | 626626 | 1048456 | 5594863 | 3308669 | 28505799 | 32764869 | 3817697 | 11063233
34 | 2023-01-19 03:14 | WO-23003-13-A-1 | 85 | U |  | 0 | #N/A | 0.791 | 0 | #N/A | 0.961 | 1.759 | #N/A | 4255.588 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4255.588 | NA | NA | NA | Unknown | [9 empty] | IS4 | [6 empty] | 343592 | 671594 | 1136431 | 6350247 | 3307942 | 28549100 | 33171763 | 3731629 | 10872151
35 | 2023-01-19 03:41 | WO-23003-14-A-1 | 86 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.529 | 0.504 | #N/A | 193.554 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 193.554 | NA | NA | NA | Unknown |  | IS3 | IS4 | [6 empty] | C | R1 | [5 empty] | 133833 | 328067 | 642898 | 3410265 | 2563857 | 22823195 | 29279951 | 4013959 | 28911726
36 | 2023-01-19 04:08 | WO-23003-15-A-1 | 14 | U |  | 0 | #N/A | 0 | 0.817 | #N/A | 0.502 | 0.488 | #N/A | 7.75 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 7.75 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 |  |  |  |  | C | R1 | [5 empty] | 145471 | 260209 | 424411 | 1935784 | 2694573 | 23280241 | 28846629 | 4034252 | 35402869
37 | 2023-01-19 04:35 | WO-23003-16-A-1 | 16 | U |  | 0 | #N/A | 0.678 | 0.767 | #N/A | 0.465 | 0.505 | #N/A | 4.083 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | [13 empty] | 137413 | 355383 | 667889 | 3583419 | 2613035 | 23060882 | 30758028 | 3988486 | 36763017
38 | 2023-01-19 05:02 | WO-23003-17-A-1 | 19 | U |  | 0 | #N/A | 0.851 | 0.834 | #N/A | 0.579 | 0.469 | #N/A | 2.358 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 127206 | 255377 | 606825 | 3321449 | 2527701 | 22483101 | 29735618 | 3934598 | 36557273
39 | 2023-01-19 05:29 | WO-23003-18-A-1 | 20 | U |  | 0 | #N/A | 0.314 | 0.925 | #N/A | 0.69 | 0.508 | #N/A | 1.905 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 121050 | 314019 | 593985 | 3151576 | 2335472 | 21654419 | 29123978 | 3888248 | 33091451
40 | 2023-01-19 05:57 | WO-23003-19-A-1 | 21 | U |  | 0 | #N/A | 0.589 | 0 | #N/A | 0.681 | 0.569 | #N/A | 1.605 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 143465 | 342972 | 398253 | 2525868 | 2232321 | 20665717 | 23736871 | 3207515 | 33150001
41 | 2023-01-19 06:24 | WO-23003-20-A-1 | 22 | U |  | 0 | #N/A | 0.342 | 0 | #N/A | 0.524 | 0.531 | #N/A | 1.392 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  |  |  |  | IS4 | [11 empty] | 326795 | 675868 | 717575 | 3341284 | 2914086 | 25587257 | 20617946 | 2077513 | 37950471
42 | 2023-01-19 06:51 | AB-230117-02 LRB 3 | 23 | K | QC | 0 | #N/A | 0 | 0 | #N/A | 0.452 | 0.37 | #N/A | 0.983 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [8 empty] | b | [7 empty] | 329839 | 608685 | 1180544 | 6519835 | 3316109 | 28054916 | 32548149 | 3830699 | 39060283 |  |  | 42
43 | 2023-01-19 07:18 | AB-230117-02 LD 1 | 24 | D | QC | 574.418 | #N/A | 0 | 1.066 | #N/A | 0.424 | 0.632 | #N/A | 1.343 | 1 | 574.418 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Duplicate |  | IS3 | [14 empty] | 92502 | 532312 | 969578 | 5676603 | 3556895 | 27273438 | 27937670 | 4080404 | 32960813 |  |  |  | 43 |  | LRB & LFB Samples |  |  |  | Measured Concentration, ng/L | [9 empty] | Quality Control Checks
44 | 2023-01-19 07:45 | AB-230117-02 LFB-H 2 | 25 | Z | QC | 168.178 | #N/A | 196.167 | 190.739 | #N/A | 188.953 | 194.708 | #N/A | 191.306 | 1 | 168.178 | #N/A | 196.167 | 190.739 | #N/A | 188.953 | 194.708 | #N/A | 191.306 | H 2 | H 2 | NA | Blank Spike | No TrueV! | [12 empty] | Normal, just highlighted due to LRB carrryover |  |  | 332746 | 786185 | 1141890 | 6097995 | 2698290 | 23586649 | 27131872 | 3069555 | 31196287 |  |  | 44 |  |  | # | QC Samples | Type | Dilution Factor: | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | True Value | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA
45 | 2023-01-19 08:12 | AB-230117-02 CCC-H 2 | 26 | Q | QC | 342.723 | #N/A | 403.865 | 379.785 | #N/A | 381.16 | 387.516 | #N/A | 380.815 | 1 | 342.723 | #N/A | 403.865 | 379.785 | #N/A | 381.16 | 387.516 | #N/A | 380.815 | 2 | 02 | NA | CCC | [16 empty] | 329585 | 687121 | 1089784 | 5777486 | 2292144 | 20239051 | 22720073 | 2438598 | 25113502 |  | 45
46 | 2023-01-19 08:39 | AB-230117-02 LFSM-H 1 @5 | 27 | S | QC | 570.426 | #N/A | 197.796 | 197.561 | #N/A | 202.268 | 209.365 | #N/A | 212.68 | 1 | 570.426 | #N/A | 197.796 | 197.561 | #N/A | 202.268 | 209.365 | #N/A | 212.68 | H 1 @5 | NA | H 1 @5 | Sample Spike | No TrueV! | IS3 | IS3 | IS3 | IS3 | [24 empty] | 46 | 15 | AB-230117-02 LRB 1 | K | 1 | 0 | #N/A | 1.125 | 0 | #N/A | 3.537 | 4.565 | #N/A | 4.651 | NA | [14 empty] | FAIL | FAIL |  | FAIL |  |  | Carryover
47 | 2023-01-19 09:33 | WO-23003-21-A-1 | 28 | U |  | 0 | #N/A | 1.109 | 1.055 | #N/A | 1.898 | 2.538 | #N/A | 3.141 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown | [16 empty] | 280253 | 697557 | 1114067 | 6276252 | 3199409 | 26868769 | 21316953 | 2063733 | 36473897 | [5 empty] | 16 | AB-230117-02 LFB-H 1 | Z | 1 | 183.485 | #N/A | 178.384 | 194.206 | #N/A | 198.792 | 204.981 | #N/A | 191.802 | H 1
48 | 2023-01-19 10:00 | WO-23003-22-A-1 | 31 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.665 | 0.601 | #N/A | 0.812 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown | [16 empty] | 330967 | 542265 | 1142107 | 6347969 | 3311045 | 27368674 | 22130353 | 2181366 | 39365969 | [5 empty] | 30 | LRB2-Inj1 | K | 1 | 0 | #N/A | 0 | 0 | #N/A | 0.46 | 0.414 | #N/A | 204.132 | NA | [17 empty] | FAIL |  |  | Carryover
49 | 2023-01-19 10:27 | WO-23003-23-A-1 | 32 | U |  | 0 | #N/A | 0 | 0.597 | #N/A | 0.544 | 0.484 | #N/A | 0.718 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown | [16 empty] | 283492 | 771498 | 1078946 | 5864427 | 3325718 | 27119147 | 21230023 | 2083832 | 38399615 | [5 empty] | 31 | AB-230117-02 LRB 2 | K | 1 | 0 | #N/A | 0 | 0.364 | #N/A | 0.466 | 0.349 | #N/A | 7.368 | NA | [17 empty] | FAIL |  |  | Carryover
50 | 2023-01-19 10:54 | WO-23003-24-A-1 | 33 | U |  | 0 | #N/A | 0.512 | 0.905 | #N/A | 0.863 | 0.727 | #N/A | 0.913 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown | [16 empty] | 324427 | 614004 | 1056532 | 5411334 | 3195824 | 27217484 | 21992191 | 2133264 | 39298517 | [5 empty] | 42 | AB-230117-02 LRB 3 | K | 1 | 0 | #N/A | 0 | 0 | #N/A | 0.452 | 0.37 | #N/A | 0.983 | NA | [20 empty] | Carryover
51 | 2023-01-19 11:21 | WO-23003-25-A-1 | 34 | U |  | 0 | #N/A | 0.818 | 0.512 | #N/A | 0.515 | 0.436 | #N/A | 0.578 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown | [16 empty] | 329385 | 666989 | 1115481 | 5929044 | 3109176 | 27087305 | 21199897 | 2030032 | 38685040 | [5 empty] | 44 | AB-230117-02 LFB-H 2 | Z | 1 | 168.178 | #N/A | 196.167 | 190.739 | #N/A | 188.953 | 194.708 | #N/A | 191.306 | H 2 | [20 empty] | Within limits except for blank detection (163 / 100 = ~80)
52 | 2023-01-19 11:49 | WO-23003-26-A-1 | 35 | U |  | 0 | #N/A | 1.004 | 0 | #N/A | 0.349 | 0.392 | #N/A | 0.489 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  | C | [7 empty] | 66085 | 284220 | 428529 | 2503721 | 1958657 | 19606438 | 27032188 | 3615594 | 35019415 | [5 empty] | 57 | LRB4-inj1 | K | 1 | 0 | #N/A | 0 | 0.299 | #N/A | 0.368 | 0.282 | #N/A | 0.334 | NA
53 | 2023-01-19 12:16 | WO-23003-27-A-1 | 36 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.424 | 0.363 | #N/A | 0.461 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS4 | IS4 | [11 empty] | 85744 | 387147 | 567057 | 3127946 | 2251039 | 20932697 | 26742961 | 3558184 | 30702952 | [5 empty] | 58 | AB-230117-02 LRB 4 | K | 1 | 0 | #N/A | 0.341 | 0 | #N/A | 0.32 | 0.302 | #N/A | 0.359 | NA
54 | 2023-01-19 12:43 | WO-23003-28-A-1 | 37 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.354 | 0.397 | #N/A | 0.444 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 96008 | 253068 | 537621 | 2979528 | 2098475 | 20636117 | 28505227 | 3857353 | 35387265 | [5 empty] | 69 | LRB5-inj1 | K | 1 | 0 | #N/A | 0 | 0 | #N/A | 0.429 | 0.372 | #N/A | 0.473 | NA
55 | 2023-01-19 13:10 | WO-23003-29-A-1 | 38 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.28 | 0.324 | #N/A | 0.406 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 85066 | 369494 | 567838 | 3070656 | 2128875 | 21057908 | 29615862 | 3906672 | 37248270 | [5 empty] | 70 | AB-230117-02 LRB 5 | K | 1 | 0 | #N/A | 0 | 0 | #N/A | 0.37 | 0.288 | #N/A | 0.334 | NA
56 | 2023-01-19 13:37 | WO-23003-30-A-1 | 39 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.295 | 0.37 | #N/A | 0.484 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 70563 | 341668 | 501333 | 2835254 | 1990198 | 19285062 | 26277998 | 3740955 | 31784162 | [5 empty] | 72 | AB-230117-02 LFB-H 3 | Z | 1 | 169.371 | #N/A | 198.181 | 191.991 | #N/A | 190.049 | 189.526 | #N/A | 186.355 | H 3 | [20 empty] | Carryover
57 | 2023-01-19 14:04 | LRB4-inj1 | 40 | K | QC | 0 | #N/A | 0 | 0.299 | #N/A | 0.368 | 0.282 | #N/A | 0.334 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 331054 | 783700 | 1167946 | 6208058 | 3160405 | 27216054 | 33682350 | 3910568 | 40439399 |  |  | 57 |  |  | 86 | LRB6-inj1 | K | 1 | 0 | #N/A | 0 | 0.352 | #N/A | 0.362 | 0.221 | #N/A | 0.343 | NA | [20 empty] | Rinse injection
58 | 2023-01-19 14:31 | AB-230117-02 LRB 4 | 46 | K | QC | 0 | #N/A | 0.341 | 0 | #N/A | 0.32 | 0.302 | #N/A | 0.359 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 255481 | 813269 | 727112 | 4316238 | 3118516 | 27292063 | 32347881 | 3707401 | 39118334 |  |  | 58 |  |  | 87 | AB-230117-02 LRB 6 | K | 1 | 0 | #N/A | 0 | 0 | #N/A | 0.308 | 0.24 | #N/A | 0.346 | NA
59 | 2023-01-19 14:58 | WO-23003-31-A-1 | 47 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.702 | 0.548 | #N/A | 0.955 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | IS4 | [10 empty] | 55125 | 303279 | 321171 | 1478468 | 1609450 | 17080853 | 26018884 | 3446208 | 34027840 | [5 empty] | 99 | LRB-7-inj2 | K | 1 | 0 | #N/A | 0 | 0 | #N/A | 0.294 | 0.334 | #N/A | 153.845 | NA | [17 empty] | FAIL
60 | 2023-01-19 15:25 | WO-23003-32-A-1 | 48 | U |  | 0 | #N/A | 2.118 | 0 | #N/A | 0.394 | 0.346 | #N/A | 0.385 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | IS4 | [10 empty] | 40994 | 286371 | 308123 | 1450216 | 1665290 | 17533167 | 26624290 | 3615937 | 34280801 | [5 empty] | 100 | LRB-7-inj3 | K | 1 | 157.537 | #N/A | 1.41 | 0 | #N/A | 0.295 | 0.324 | #N/A | 4.078 | NA | [9 empty] | FAIL | [7 empty] | FAIL |  |  | Carryover
61 | 2023-01-19 15:52 | WO-23003-33-A-1 | 49 | U |  | 0 | #N/A | 0 | 1.126 | #N/A | 0.355 | 0.488 | #N/A | 0.674 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 91680 | 258808 | 298250 | 2886147 | 1698006 | 17870923 | 27006693 | 3737407 | 35053863 | [5 empty] | 102 | LFB-H-3-Inj2 | Z | 1 | 174.08 | #N/A | 196.02 | 191.147 | #N/A | 193.469 | 195.058 | #N/A | 197.419 | H | [20 empty] | Same as above
62 | 2023-01-19 16:19 | WO-23003-34-A-1 | 50 | U |  | 0 | #N/A | 0 | 0.991 | #N/A | 0.484 | 0.521 | #N/A | 0.663 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 80214 | 364625 | 512736 | 2851048 | 2128378 | 19697415 | 28660834 | 3886488 | 35959334 | [5 empty] | 106 | LRB7-inj1 | K | 1 | 0 | #N/A | 0 | 0 | #N/A | 0.326 | 0.308 | #N/A | 27.187 | NA | [17 empty] | FAIL |  |  | Carryover
63 | 2023-01-19 16:46 | WO-23003-35-A-1 | 51 | U |  | 0 | #N/A | 0 | 1.091 | #N/A | 0.491 | 0.514 | #N/A | 0.71 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 82771 | 327495 | 522159 | 2912434 | 2086886 | 19886142 | 27913901 | 3905300 | 34775503 | [5 empty] | 107 | AB-230117-02 LRB 7 | K | 1 | 0 | #N/A | 0 | 0 | #N/A | 14.3 | 4.824 | #N/A | 77.188 | NA | [14 empty] | FAIL | FAIL |  | FAIL |  |  | Carryover
64 | 2023-01-19 17:14 | WO-23003-36-A-1 | 52 | U |  | 0 | #N/A | 0.99 | 0.885 | #N/A | 0.516 | 0.587 | #N/A | 0.763 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 89538 | 316645 | 531139 | 2949618 | 2130115 | 20278508 | 28228487 | 3891160 | 35097046 | [5 empty] | 109 | AB-230117-02 LFB-H 4 | Z | 1 | 160.162 | #N/A | 199.569 | 187.965 | #N/A | 186.073 | 195.326 | #N/A | 192.196 | H 4 | [20 empty] | Same as above
65 | 2023-01-19 17:41 | WO-23003-37-A-1 | 53 | U |  | 0 | #N/A | 0 | 0.965 | #N/A | 0.344 | 0.473 | #N/A | 0.577 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 74620 | 310697 | 539847 | 2863062 | 2033498 | 20464234 | 28772017 | 3927748 | 35420039 | [5 empty] | 122 | LRB-8-inj1 | K | 1 | 0 | #N/A | 0.346 | 0 | #N/A | 0.355 | 0.28 | #N/A | 0.783 | NA
66 | 2023-01-19 18:08 | WO-23003-38-A-1 | 54 | U |  | 0 | #N/A | 0 | 0.903 | #N/A | 0.433 | 0.435 | #N/A | 0.639 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS3 | IS4 | [11 empty] | 69658 | 391407 | 561070 | 3071048 | 2210565 | 20761933 | 28472937 | 3900660 | 36217472 | [5 empty] | 123 | AB-230117-02 LRB 8 | K | 1 | 0 | #N/A | 0 | 0 | #N/A | 0.409 | 0.277 | #N/A | 0.257 | NA
67 | 2023-01-19 18:35 | WO-23003-39-A-1 | 55 | U |  | 126.713 | #N/A | 0.513 | 0 | #N/A | 0.402 | 0.302 | #N/A | 0.309 | 1 | 126.713 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | IS4 | [10 empty] | 34161 | 240789 | 370251 | 2020304 | 1597783 | 16282203 | 26067846 | 3623185 | 34449016 | [5 empty] | 134 | AB-230117-02 LRB 9 | K | 1 | 0 | #N/A | 0 | 0.794 | #N/A | 0.552 | 0.229 | #N/A | 0.312 | NA
68 | 2023-01-19 19:02 | WO-23003-40-A-1 | 58 | U |  | 134.042 | #N/A | 1.336 | 0 | #N/A | 0.33 | 0.443 | #N/A | 0.518 | 1 | 134.042 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 74297 | 345027 | 501418 | 2778990 | 1961904 | 19481074 | 27583554 | 3727989 | 35276162 | [5 empty] | 137 | AB-230117-02 LFB-H 5 | Z | 1 | 158.258 | #N/A | 179.705 | 179.093 | #N/A | 179.201 | 181.423 | #N/A | 180.317 | H 5
69 | 2023-01-19 19:29 | LRB5-inj1 | 59 | K | QC | 0 | #N/A | 0 | 0 | #N/A | 0.429 | 0.372 | #N/A | 0.473 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 333530 | 790561 | 1158540 | 6538012 | 3170322 | 27224949 | 33056389 | 3827007 | 39665652 |  |  | 69 |  |  | 150 | LRB-10-inj1 | K | 1 | 0 | #N/A | 0 | 0.79 | #N/A | 0.515 | 0.261 | #N/A | 0.283 | NA
70 | 2023-01-19 19:56 | AB-230117-02 LRB 5 | 60 | K | QC | 0 | #N/A | 0 | 0 | #N/A | 0.37 | 0.288 | #N/A | 0.334 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 284660 | 617862 | 1129294 | 5896299 | 2983135 | 26102836 | 30826818 | 3543112 | 37000585 |  |  | 70 |  |  | 151 | AB-230117-02 LRB 10 | K | 1 | 0 | #N/A | 0.079 | 0.848 | #N/A | 0.571 | 0.264 | #N/A | 0.321 | NA
71 | 2023-01-19 20:23 | AB-230117-02 LD 2 | 61 | D | QC | 0 | #N/A | 0 | 0 | #N/A | 0.553 | 0 | #N/A | 0.39 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Duplicate | [7 empty] | IS4 | IS4 | [7 empty] | 314101 | 757165 | 1139792 | 6420086 | 3240989 | 25510276 | 9308318 | 897677 | 36730462 |  |  |  | 71 |  | 162 | AB-230117-02 LRB 11 | K | 1 | 0 | #N/A | 0.345 | 0.756 | #N/A | 0.55 | 0.267 | #N/A | 0.264 | NA
72 | 2023-01-19 20:50 | AB-230117-02 LFB-H 3 | 62 | Z | QC | 169.371 | #N/A | 198.181 | 191.991 | #N/A | 190.049 | 189.526 | #N/A | 186.355 | 1 | 169.371 | #N/A | 198.181 | 191.991 | #N/A | 190.049 | 189.526 | #N/A | 186.355 | H 3 | H 3 | NA | Blank Spike | No TrueV! | [7 empty] | C | [7 empty] | 336965 | 728784 | 1098817 | 6134568 | 2651172 | 23679978 | 25011867 | 2536330 | 31838461 |  |  | 72 |  |  | 165 | AB-230117-02 LFB-H 6 | Z | 1 | 158.298 | #N/A | 179.368 | 175.354 | #N/A | 173.631 | 176.742 | #N/A | 176.054 | H 6
73 | 2023-01-19 21:17 | AB-230117-02 CCC-H 3 | 63 | Q | QC | 366.428 | #N/A | 400.863 | 377.097 | #N/A | 377.585 | 383.968 | #N/A | 393.376 | 1 | 366.428 | #N/A | 400.863 | 377.097 | #N/A | 377.585 | 383.968 | #N/A | 393.376 | 2 | 02 | NA | CCC | [16 empty] | 355150 | 681707 | 747444 | 3462214 | 2126517 | 20023795 | 22761914 | 2428572 | 25012779 |  | 73 |  |  |  | 178 | LRB-12-inj1 | K | 1 | 0 | #N/A | 0 | 0.799 | #N/A | 0.648 | 0.302 | #N/A | 24.998 | NA | [17 empty] | FAIL |  |  | Carryover
74 | 2023-01-19 21:44 | AB-230117-02 LFSM-H 2 | 64 | S | QC | 169.438 | #N/A | 198.319 | 190.263 | #N/A | 196.626 | 202.465 | #N/A | 206.37 | 1 | 169.438 | #N/A | 198.319 | 190.263 | #N/A | 196.626 | 202.465 | #N/A | 206.37 | H 2 | NA | H 2 | Sample Spike | No TrueV! | [6 empty] | IS3 | IS3 | [7 empty] | 365992 | 703762 | 1111637 | 6218454 | 2642276 | 21889536 | 8742054 | 871340 | 27198676 |  |  |  |  | 74 | 179 | AB-230117-02 LRB 12 | K | 1 | 0 | #N/A | 0 | 0.758 | #N/A | 0.534 | 0.278 | #N/A | 1.003 | NA
75 | 2023-01-19 22:11 | Rinse5 | 65 | U |  | 0 | #N/A | 0 | 39.805 | #N/A | 160.825 | 169.964 | #N/A | 175.37 | 1 | <MRL | #N/A | <MRL | 39.805 | #N/A | 160.825 | 169.964 | #N/A | 175.37 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 | [6 empty] | 12456 | 16713 | 23414 | 117238 | 145690 | 1308521 | 1786131 | 219481 | 2805805 | [5 empty] | 190 | AB-230117-02 LRB 13 | K | 1 | 0 | #N/A | 0.263 | 0.748 | #N/A | 0.62 | 0.374 | #N/A | 69.087 | NA | [17 empty] | FAIL |  |  | Carryover
76 | 2023-01-19 22:38 | WO-23003-41-A-1 | 66 | U |  | 150.848 | #N/A | 0.582 | 0 | #N/A | 0.665 | 0.764 | #N/A | 0.677 | 1 | 150.848 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 |  |  |  | C |  | R1 |  |  | Rerun due to carryover |  |  | 56781 | 227095 | 380418 | 2619758 | 1921842 | 19253007 | 27213616 | 3646302 | 34946438 | [5 empty] | 193 | AB-230117-02 LFB-H 7 | Z | 1 | 157.151 | #N/A | 181.009 | 182.079 | #N/A | 177.55 | 177.827 | #N/A | 303.883 | H 7
77 | 2023-01-19 23:06 | WO-23003-42-A-1 | 67 | U |  | 136.634 | #N/A | 0 | 0 | #N/A | 0.454 | 0.466 | #N/A | 0.442 | 1 | 136.634 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 44191 | 230489 | 458326 | 2180358 | 1909302 | 18875295 | 27039117 | 3761177 | 35680754 | [5 empty] | 201 | AB-230117-02 LRB 12 20230123102918 | K | 1 | 0 | #N/A | 0 | 0 | #N/A | 49.489 | 26.569 | #N/A | 943.114 | NA | [14 empty] | FAIL | FAIL |  | FAIL |  |  | Rinse injection
78 | 2023-01-19 23:33 | WO-23003-43-A-1 | 75 | U |  | 111.294 | #N/A | 0 | 0 | #N/A | 0.354 | 0.416 | #N/A | 0.4 | 1 | 111.294 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 60196 | 251081 | 380678 | 2502751 | 1777769 | 17773173 | 23914508 | 3622146 | 31060291 | [5 empty] | 202 | AB-230117-02 LRB 13 20230123105625 | K | 1 | 0 | #N/A | 0.169 | 0.805 | #N/A | 0.61 | 0.281 | #N/A | 3.33 | NA | [17 empty] | FAIL |  |  | Rinse injection
79 | 2023-01-20 | WO-23003-44-A-1 | 76 | U |  | 109.959 | #N/A | 1.537 | 0 | #N/A | 0.355 | 0.413 | #N/A | 0.359 | 1 | 109.959 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 74464 | 220833 | 369180 | 1667776 | 1963128 | 19266918 | 28279691 | 3978394 | 34820074 | [5 empty] | 211 | LRB-14-inj1 | K | 1 | 0 | #N/A | 0 | 0.799 | #N/A | 0.486 | 0.247 | #N/A | 0.68 | NA | [20 empty] | Carryover
80 | 2023-01-20 00:27 | WO-23003-45-A-1 | 77 | U |  | 126.129 | #N/A | 0 | 0 | #N/A | 0.401 | 0.366 | #N/A | 0.368 | 1 | 126.129 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 83660 | 249583 | 365309 | 1785092 | 1938531 | 19560443 | 27438800 | 3725257 | 34854193 | [5 empty] | 212 | AB-230117-02 LRB 14 | K | 1 | 0 | #N/A | 0 | 0.844 | #N/A | 0.577 | 0.274 | #N/A | 0.664 | NA
81 | 2023-01-20 00:54 | WO-23003-46-A-1 | 78 | U |  | 117.077 | #N/A | 0 | 0 | #N/A | 0.328 | 0.427 | #N/A | 0.45 | 1 | 117.077 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | IS4 | [10 empty] | 43989 | 285681 | 349917 | 2429265 | 1376805 | 15617674 | 25851385 | 3768171 | 33672580 | [5 empty] | 224 | LRB-15-inj1 | K | 1 | 0 | #N/A | 0.374 | 0.91 | #N/A | 0.603 | 0.31 | #N/A | 1.417 | NA
82 | 2023-01-20 01:21 | WO-23003-47-A-1 | 79 | U |  | 268.234 | #N/A | 1.131 | 0 | #N/A | 0.429 | 0.355 | #N/A | 0.318 | 1 | 268.234 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 9210 | 145000 | 272427 | 1573853 | 1394350 | 13784367 | 21632529 | 3462368 | 27574645 | [5 empty] | 225 | AB-230117-02 LRB 15 | K | 1 | 0 | #N/A | 0.658 | 0.9 | #N/A | 0.525 | 0.302 | #N/A | 1.286 | NA
83 | 2023-01-20 01:48 | WO-23003-48-A-1 | 80 | U |  | 253.554 | #N/A | 0.421 | 0 | #N/A | 0.491 | 0.363 | #N/A | 0.334 | 1 | 253.554 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 66177 | 211177 | 508707 | 2722086 | 1847476 | 18954867 | 27459965 | 3623848 | 35375815 | [5 empty] | 227 | AB-230117-02 LFB-H 8 | Z | 1 | 149.916 | #N/A | 178.569 | 174.977 | #N/A | 174.263 | 174.528 | #N/A | 172.643 | H 8
84 | 2023-01-20 02:15 | WO-23003-49-A-1 | 81 | U |  | 250.815 | #N/A | 0 | 0 | #N/A | 0.807 | 0.381 | #N/A | 0.323 | 1 | 250.815 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 69942 | 262748 | 508233 | 2492440 | 2007731 | 19278484 | 27906762 | 3894264 | 35076715 | [5 empty] | 240 | LRB-16-inj1 | K | 1 | 0 | #N/A | 0.811 | 0.741 | #N/A | 0.516 | 0.273 | #N/A | 0.375 | NA
85 | 2023-01-20 02:42 | WO-23003-50-A-1 | 82 | U |  | 214.775 | #N/A | 2.007 | 0 | #N/A | 0 | 0.388 | #N/A | 0.264 | 1 | 214.775 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 10422 | 99474 | 129790 | 1119243 | 976903 | 11365017 | 21242150 | 3133749 | 31442073 | [5 empty] | 241 | AB-230117-02 LRB 16 | K | 1 | 0 | #N/A | 0 | 0.749 | #N/A | 0.617 | 0.266 | #N/A | 0.431 | NA
86 | 2023-01-20 03:09 | LRB6-inj1 | 83 | K | QC | 0 | #N/A | 0 | 0.352 | #N/A | 0.362 | 0.221 | #N/A | 0.343 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 258969 | 491110 | 1091249 | 6487216 | 3338896 | 28081546 | 33107061 | 3831279 | 34146175 |  |  | 86 |  |  | 252 | LRB-17-inj1 | K | 1 | 0 | #N/A | 0 | 0.786 | #N/A | 0.588 | 0.256 | #N/A | 0.322 | NA | [20 empty] | See QAQC Overflow page for additional LRBs/LFBs
87 | 2023-01-20 03:36 | AB-230117-02 LRB 6 | 84 | K | QC | 0 | #N/A | 0 | 0 | #N/A | 0.308 | 0.24 | #N/A | 0.346 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 288321 | 526049 | 764375 | 5772719 | 2312291 | 23284126 | 29128442 | 3361080 | 36778822 |  |  | 87
88 | 2023-01-20 04:04 | WO-23003-51-A-1 | 87 | U |  | 498.597 | #N/A | 0 | 0 | #N/A | 0.375 | 0.324 | #N/A | 0.295 | 1 | 498.597 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | IS4 | [10 empty] | 40187 | 214580 | 241504 | 1164454 | 1641215 | 17883015 | 27104782 | 3685757 | 35034929 | [5 empty] | LD Samples |  |  |  | Measured Concentration, ng/L | [9 empty] | Quality Control Checks
89 | 2023-01-20 04:31 | WO-23003-52-A-1 | 88 | U |  | 413.609 | #N/A | 1.442 | 0 | #N/A | 0.295 | 0.342 | #N/A | 0.235 | 1 | 413.609 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 86481 | 259214 | 283266 | 2021364 | 1771089 | 18689352 | 25366446 | 3788554 | 32019259 | [5 empty] | # | QC Samples | Type | Dilution Factor: | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | True Value | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA
90 | 2023-01-20 04:58 | WO-23003-53-A-1 | 89 | U |  | 419.335 | #N/A | 0 | 0 | #N/A | 0.355 | 0.362 | #N/A | 0.428 | 1 | 419.335 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 91104 | 265947 | 600538 | 3350863 | 2580651 | 23159173 | 30016345 | 4157094 | 36530098
91 | 2023-01-20 05:25 | WO-23003-54-A-1 | 90 | U |  | 446.172 | #N/A | 0 | 0 | #N/A | 0.466 | 0.507 | #N/A | 0.617 | 1 | 446.172 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS3 | IS4 | [11 empty] | 48294 | 203735 | 472734 | 2518745 | 1747890 | 17113731 | 26555731 | 3725653 | 34325965
92 | 2023-01-20 05:52 | WO-23003-55-A-1 | 91 | U |  | 456.386 | #N/A | 0 | 0.584 | #N/A | 0.523 | 0.541 | #N/A | 1.323 | 1 | 456.386 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS3 | IS4 | [11 empty] | 53430 | 239124 | 566303 | 3250452 | 2842774 | 24470312 | 29039247 | 3959883 | 33577065 | [5 empty] | 43 | AB-230117-02 LD 1 | D | 1 | 574.418 | #N/A | 0 | 1.066 | #N/A | 0.424 | 0.632 | #N/A | 1.343 | NA | [18 empty] | FAIL
93 | 2023-01-20 06:19 | WO-23003-56-A-1 | 92 | U |  | 433.565 | #N/A | 0 | 0 | #N/A | 0.523 | 0.563 | #N/A | 0.576 | 1 | 433.565 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS3 | IS4 | IS4 | [10 empty] | 53390 | 205437 | 423742 | 2287421 | 1620637 | 16203550 | 24908344 | 3684997 | 31568213
94 | 2023-01-20 06:46 | WO-23003-57-A-1 | 101 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.337 | 0.365 | #N/A | 4.897 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 44556 | 169211 | 455739 | 2450418 | 1718804 | 16898084 | 25858895 | 3724831 | 34045107 | [5 empty] | 71 | AB-230117-02 LD 2 | D | 1 | 0 | #N/A | 0 | 0 | #N/A | 0.553 | 0 | #N/A | 0.39 | NA | [18 empty] | FAIL |  | Original sample was at 1000-fold dilution, so this isn't calculating correctly.
95 | 2023-01-20 10:01 | WO-23003-26-A-1B | 100 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.288 | 0.417 | #N/A | 0.466 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 |  |  |  |  | R1 |  |  | Analyzed again below to verify if this was 'real', repeated pattern |  |  | 42089 | 221861 | 386664 | 2061350 | 1686468 | 15830622 | 22726578 | 3128580 | 23123741
96 | 2023-01-20 10:28 | WO-23003-58-A-1B | 99 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.492 | 1.257 | #N/A | 61.125 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 61.125 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [5 empty] | R1 |  |  | Analyzed again below to verify if this was 'real', repeated pattern |  |  | 89357 | 318973 | 502433 | 2343061 | 1823727 | 17636975 | 24650996 | 3520584 | 31407933 | [5 empty] | 108 | AB-230117-02 LD 3 | D | 1 | 140.88 | #N/A | 0 | 0 | #N/A | 0.387 | 0.41 | #N/A | 1.737 | NA | [18 empty] | FAIL |  | Carryover
97 | 2023-01-20 10:55 | WO-23003-59-A-1B | 98 | U |  | 0 | #N/A | 0.234 | 0 | #N/A | 0.616 | 0.839 | #N/A | 632.89 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 632.89 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [5 empty] | R1 |  |  | Analyzed again below to verify if this was 'real', repeated pattern |  |  | 69366 | 243432 | 527354 | 2796228 | 2211458 | 20857792 | 27380388 | 3757713 | 23508883
98 | 2023-01-20 11:23 | WO-23003-60-A-1B | 97 | U |  | 0 | #N/A | 0 | 0 | #N/A | 1.587 | 2.809 | #N/A | 4183.607 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4183.607 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | IS4 |  |  |  | IS4 | R1 |  |  | Analyzed again below to verify if this was 'real', repeated pattern |  |  | 9744 | 111850 | 283109 | 1522380 | 1565684 | 15651949 | 24657370 | 3340961 | 8884746 | [5 empty] | 135 | LD4-inj1 | D | 1 | 0 | #N/A | 1.695 | 0 | #N/A | 0.397 | 0.342 | #N/A | 3.644 | NA | [18 empty] | FAIL |  | Carryover
99 | 2023-01-20 11:50 | LRB-7-inj2 | 96 | K | QC | 0 | #N/A | 0 | 0 | #N/A | 0.294 | 0.334 | #N/A | 153.845 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 153.845 | NA | NA | NA | Blank | [9 empty] | b | [6 empty] | 281986 | 443639 | 1175727 | 5967825 | 3248161 | 27653250 | 28828488 | 3378218 | 34617381 |  |  | 99
100 | 2023-01-20 12:17 | LRB-7-inj3 | 95 | K | QC | 157.537 | #N/A | 1.41 | 0 | #N/A | 0.295 | 0.324 | #N/A | 4.078 | 1 | 157.537 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | Original sample had carryover, not a duplicate failure |  |  | 50728 | 224057 | 489201 | 2671578 | 1984692 | 18860380 | 26903633 | 3754570 | 35105550 |  |  | 100 |  |  | 136 | AB-230117-02 LD 4 | D | 1 | 0 | #N/A | 1.281 | 0 | #N/A | 0.502 | 0.314 | #N/A | 3.87 | NA | [18 empty] | FAIL
101 | 2023-01-20 12:44 | WO-23003-14-A-1B | 94 | U |  | 0 | #N/A | 0 | 0.614 | #N/A | 0.47 | 0.383 | #N/A | 1.981 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [5 empty] | R1 |  |  | Reprepared and run below. |  |  | 98056 | 191303 | 586884 | 3228103 | 2367076 | 20850445 | 26497504 | 3522701 | 31673484
102 | 2023-01-20 13:11 | LFB-H-3-Inj2 | 198 | Z | QC | 174.08 | #N/A | 196.02 | 191.147 | #N/A | 193.469 | 195.058 | #N/A | 197.419 | 1 | 174.08 | #N/A | 196.02 | 191.147 | #N/A | 193.469 | 195.058 | #N/A | 197.419 | H | H | NA | Blank Spike | No TrueV! | [15 empty] | 263890 | 416290 | 1113353 | 5737650 | 2467529 | 18087861 | 22319355 | 2470966 | 26770409 |  |  | 102 |  |  | 163 | LD-5-inj1 | D | 1 | 0 | #N/A | 0 | 0.802 | #N/A | 0.657 | 0.477 | #N/A | 82.9 | NA | [18 empty] | FAIL
103 | 2023-01-20 14:05 | WO-23003-58-A-1 | 197 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.547 | 1.322 | #N/A | 60.711 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 60.711 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b |  |  |  | Subsequently reprepared entirely, but showed no improvement. Probably contamination of samples |  |  | 43031 | 191825 | 419100 | 2336952 | 1777737 | 16830488 | 24137122 | 3252455 | 27314869
104 | 2023-01-20 14:32 | WO-23003-59-A-1 | 196 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.656 | 0.846 | #N/A | 660.136 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 660.136 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | Subsequently reprepared entirely, but showed no improvement. Probably contamination of samples |  |  | 35406 | 267109 | 404907 | 2062473 | 2051000 | 18992116 | 25958296 | 3414709 | 20147347 | [5 empty] | 164 | AB-230117-02 LD 5 | D | 1 | 0 | #N/A | 0 | 0.77 | #N/A | 0.645 | 0.438 | #N/A | 88.093 | NA | [18 empty] | FAIL |  | Dup failure. Rerun below
105 | 2023-01-20 15:00 | WO-23003-60-A-1 | 195 | U |  | 0 | #N/A | 0 | 0 | #N/A | 1.794 | 2.954 | #N/A | 4256.541 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4256.541 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | IS4 |  |  |  | IS4 |  |  |  | Subsequently reprepared entirely, but showed no improvement. Probably contamination of samples |  |  | 3676 | 99288 | 229596 | 1307468 | 1576405 | 15734824 | 21786661 | 2805663 | 9685815
106 | 2023-01-20 15:27 | LRB7-inj1 | 194 | K | QC | 0 | #N/A | 0 | 0 | #N/A | 0.326 | 0.308 | #N/A | 27.187 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 27.187 | NA | NA | NA | Blank | [9 empty] | b | [6 empty] | 240978 | 501100 | 1135717 | 6075374 | 1875354 | 26555056 | 30394983 | 3619684 | 34977386 |  |  | 106 |  |  | 191 | LD6-inj1 | D | 1 | 0 | #N/A | 0 | 0.883 | #N/A | 1.362 | 2.893 | #N/A | 3922.707 | NA | [18 empty] | FAIL |  | Reprepared below
107 | 2023-01-20 15:54 | AB-230117-02 LRB 7 | 193 | K | QC | 0 | #N/A | 0 | 0 | #N/A | 14.3 | 4.824 | #N/A | 77.188 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 14.3 | <MRL | #N/A | 77.188 | NA | NA | NA | Blank |  | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | b | [6 empty] | 5123 | 6978 | 12758 | 78411 | 78334 | 663256 | 1096848 | 154930 | 1533140 |  |  | 107
108 | 2023-01-20 16:21 | AB-230117-02 LD 3 | 192 | D | QC | 140.88 | #N/A | 0 | 0 | #N/A | 0.387 | 0.41 | #N/A | 1.737 | 1 | 140.88 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Duplicate |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | Original sample had carryover, not a duplicate failure |  |  | 33217 | 195206 | 398970 | 2303037 | 1797806 | 17701996 | 24386723 | 3312750 | 28505998 |  |  |  | 108 |  | 192 | AB-230117-02 LD 6 | D | 1 | 0 | #N/A | 0 | 0.865 | #N/A | 1.374 | 3.079 | #N/A | 4142.057 | NA | [18 empty] | FAIL
109 | 2023-01-20 16:48 | AB-230117-02 LFB-H 4 | 191 | Z | QC | 160.162 | #N/A | 199.569 | 187.965 | #N/A | 186.073 | 195.326 | #N/A | 192.196 | 1 | 160.162 | #N/A | 199.569 | 187.965 | #N/A | 186.073 | 195.326 | #N/A | 192.196 | H 4 | H 4 | NA | Blank Spike | No TrueV! | [15 empty] | 238833 | 443262 | 1181283 | 6098605 | 2745742 | 23171554 | 25512501 | 2951750 | 28066312 |  |  | 109
110 | 2023-01-20 17:15 | AB-230117-02 CCC-H 4 | 190 | Q | QC | 333.561 | #N/A | 394.235 | 375.413 | #N/A | 375.67 | 378.57 | #N/A | 378.791 | 1 | 333.561 | #N/A | 394.235 | 375.413 | #N/A | 375.67 | 378.57 | #N/A | 378.791 | 2 | 02 | NA | CCC | [16 empty] | 245197 | 463618 | 1173296 | 5962507 | 2194755 | 20043337 | 21827783 | 2555118 | 22953023 |  | 110 |  |  |  | 226 | AB-230117-02 LD 7 | D | 1 | 0 | #N/A | 0 | 1.76 | #N/A | 0.417 | 0.336 | #N/A | 1.197 | NA | [18 empty] | FAIL
111 | 2023-01-20 17:42 | AB-230117-02 LFSM-H 3 | 189 | S | QC | 334.468 | #N/A | 193.406 | 190.483 | #N/A | 191.067 | 198.421 | #N/A | 203.756 | 1 | 334.468 | #N/A | 193.406 | 190.483 | #N/A | 191.067 | 198.421 | #N/A | 203.756 | H 3 | NA | H 3 | Sample Spike | No TrueV! | IS3 | IS3 | IS3 | IS3 | [11 empty] | 34135 | 210395 | 468608 | 2497078 | 1830236 | 17167760 | 22029527 | 3212403 | 25558148 |  |  |  |  | 111
112 | 2023-01-20 18:36 | WO-23003-61-A-1 | 188 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.576 | 0.672 | #N/A | 4.3 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 39650 | 184116 | 473310 | 2553636 | 1857673 | 18165095 | 25807274 | 3818796 | 31165442 | [5 empty] | 254 | AB-230117-02 LD 8 | D | 1 | 0 | #N/A | 0 | 0 | #N/A | 0.414 | 0.38 | #N/A | 0.49 | NA | [18 empty] | FAIL
113 | 2023-01-20 19:03 | WO-23003-62-A-1 | 187 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.494 | 0.45 | #N/A | 5.986 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 5.986 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b | [6 empty] | 58574 | 244470 | 533107 | 2937539 | 2105867 | 20501940 | 28012206 | 4020798 | 33327527
114 | 2023-01-20 19:30 | WO-23003-63-A-1 | 186 | U |  | 0 | #N/A | 0.45 | 0 | #N/A | 0.39 | 0.432 | #N/A | 9.272 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 9.272 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b | [6 empty] | 47697 | 245238 | 501018 | 2596358 | 1889404 | 19240240 | 27207511 | 4026837 | 33426157 | [5 empty] | 282 | AB-230117-02 LD 9 | D | 1 | 0 | #N/A | 0.632 | 0 | #N/A | 0.32 | 0.269 | #N/A | 0.236 | NA | [18 empty] | FAIL
115 | 2023-01-20 19:57 | WO-23003-64-A-1 | 185 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.38 | 0.475 | #N/A | 12.026 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 12.026 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b | [6 empty] | 34007 | 201717 | 423399 | 2264569 | 1729750 | 16931737 | 25722526 | 3776287 | 31846879
116 | 2023-01-20 20:24 | WO-23003-65-A-1 | 184 | U |  | 0 | #N/A | 0 | 0.25 | #N/A | 0.317 | 0.34 | #N/A | 8.885 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 8.885 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b | [6 empty] | 39898 | 226076 | 449024 | 2397431 | 1923390 | 18142797 | 25240218 | 3902899 | 30651184 | [5 empty] | 310 | AB-230117-02 LD 10 | D | 1 | 0 | #N/A | 1.057 | 3.438 | #N/A | 10.794 | 0 | #N/A | 0.273 | NA | [18 empty] | FAIL
117 | 2023-01-20 20:51 | WO-23003-66-A-1 | 183 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.338 | 0.344 | #N/A | 5.647 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 5.647 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b | [6 empty] | 83558 | 301071 | 575404 | 3053264 | 2196066 | 21091039 | 26677106 | 4144308 | 29837440
118 | 2023-01-20 21:19 | WO-23003-67-A-1 | 182 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.362 | 0.338 | #N/A | 5.711 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 5.711 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b | [6 empty] | 72354 | 315898 | 568400 | 2982161 | 2198727 | 20463965 | 27985442 | 4021916 | 33824164 | [5 empty] | 349 | 217295-d1000-LD-reprep | D | 1000 | 0 | #N/A | 0.177 | 0.308 | #N/A | 0.299 | 0.28 | #N/A | 1.087 | NA | [18 empty] | FAIL
119 | 2023-01-20 21:46 | WO-23003-68-A-1 | 181 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.382 | 0.389 | #N/A | 10.214 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 10.214 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b | [6 empty] | 72278 | 313311 | 540784 | 2893781 | 2117441 | 19475658 | 24328483 | 3818392 | 29097804
120 | 2023-01-20 22:13 | WO-23003-69-A-1 | 180 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.524 | 0.312 | #N/A | 9.563 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 9.563 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b | [6 empty] | 68747 | 323597 | 505162 | 2797238 | 2013235 | 19684253 | 27044723 | 3994878 | 33335946 | [5 empty] | 350 | 217335-d2-LD-reprep | D | 2 | 430.994 | #N/A | 0 | 0 | #N/A | 0.574 | 0.328 | #N/A | 1.008 | NA | [18 empty] | FAIL |  | See QAQC Overflow page
121 | 2023-01-20 22:40 | WO-23003-70-A-1 | 179 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.365 | 0.408 | #N/A | 14.244 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 14.244 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b | [6 empty] | 64949 | 325383 | 523354 | 2863392 | 2085017 | 19678201 | 27121773 | 3802834 | 33314099
122 | 2023-01-20 23:07 | LRB-8-inj1 | 178 | K | QC | 0 | #N/A | 0.346 | 0 | #N/A | 0.355 | 0.28 | #N/A | 0.783 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 344404 | 734076 | 1155057 | 6387645 | 3310690 | 28461204 | 32962886 | 4128230 | 38150621 |  |  | 122 |  |  | 351 | 217375-d2-LD-reprep | D | 2 | 0 | #N/A | 0 | 0.931 | #N/A | 0.848 | 0.367 | #N/A | 28.506 | NA | [18 empty] | FAIL |  | See QAQC Overflow page
123 | 2023-01-20 23:34 | AB-230117-02 LRB 8 | 177 | K | QC | 0 | #N/A | 0 | 0 | #N/A | 0.409 | 0.277 | #N/A | 0.257 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 346903 | 736670 | 1219569 | 6418043 | 3110107 | 27135711 | 31335940 | 3791022 | 35232704 |  |  | 123
124 | 2023-01-21 00:01 | WO-23003-71-A-1 | 176 | U |  | 0 | #N/A | 0.602 | 0 | #N/A | 0.33 | 0.328 | #N/A | 14.541 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 14.541 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 86431 | 342097 | 521621 | 2904367 | 2109933 | 20358612 | 27434685 | 3924923 | 32070510 | [5 empty] | 352 | 217395-d4-LD-reprep | D | 4 | 0 | #N/A | 0 | 1.523 | #N/A | 8.776 | 8.406 | #N/A | 2607.572 | NA | [18 empty] | FAIL |  | See QAQC Overflow page
125 | 2023-01-21 00:28 | WO-23003-72-A-1 | 175 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.515 | 0.315 | #N/A | 19.333 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 19.333 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 67634 | 316272 | 494171 | 2682500 | 2063857 | 18645532 | 23551651 | 3652557 | 28298413
126 | 2023-01-21 00:55 | WO-23003-73-A-1 | 174 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.532 | 0.365 | #N/A | 16.405 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 16.405 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 68866 | 323190 | 484900 | 2800246 | 1979874 | 19586057 | 27116037 | 3936148 | 32476129 | [5 empty] | 353 | 217359-d2-LD | D | 2 | 0 | #N/A | 0 | 0 | #N/A | 0.968 | 2.774 | #N/A | 8748.882 | NA | [18 empty] | FAIL
127 | 2023-01-21 01:22 | WO-23003-74-A-1 | 173 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.311 | 0.357 | #N/A | 33.419 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 33.419 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 94132 | 363087 | 555461 | 3016217 | 2248390 | 20917351 | 28137426 | 3985448 | 33163671
128 | 2023-01-21 01:49 | WO-23003-75-A-1 | 172 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.403 | 0.51 | #N/A | 46.694 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 46.694 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 70217 | 346097 | 513406 | 2759560 | 2037652 | 19069664 | 25054025 | 3906238 | 28726809 | [5 empty] | 354 | 217415-d2-LD-reprep | D | 2 | 0 | #N/A | 0.57 | 0 | #N/A | 0.48 | 0.339 | #N/A | 10.557 | NA | [18 empty] | FAIL
129 | 2023-01-21 02:16 | WO-23003-76-A-1 | 171 | U |  | 201.051 | #N/A | 0 | 0 | #N/A | 0.32 | 0.35 | #N/A | 1.751 | 1 | 201.051 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 82684 | 354298 | 553255 | 3045691 | 2199167 | 21155629 | 28576413 | 4064344 | 35596451
130 | 2023-01-21 02:43 | WO-23003-77-A-1 | 170 | U |  | 212.264 | #N/A | 0.689 | 0 | #N/A | 0.35 | 0.385 | #N/A | 0.393 | 1 | 212.264 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 98954 | 367850 | 593294 | 3165726 | 2404844 | 22688275 | 30051573 | 4163590 | 36194142
131 | 2023-01-21 03:11 | WO-23003-78-A-1 | 169 | U |  | 343.62 | #N/A | 0 | 0 | #N/A | 0.336 | 0.301 | #N/A | 0.283 | 1 | 343.62 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 87338 | 354479 | 566963 | 3047045 | 2259321 | 20974092 | 28502611 | 4070156 | 33969588 | [39 empty] | See QAQC Overflow page
132 | 2023-01-21 03:38 | WO-23003-79-A-1 | 168 | U |  | 341.696 | #N/A | 0 | 0 | #N/A | 0.404 | 0.384 | #N/A | 0.401 | 1 | 341.696 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 75986 | 352862 | 527901 | 2837187 | 2187823 | 20666124 | 27348729 | 4033920 | 32949161
133 | 2023-01-21 04:05 | WO-23003-80-A-1 | 167 | U |  | 354.309 | #N/A | 0 | 0 | #N/A | 0.352 | 0.326 | #N/A | 0.25 | 1 | 354.309 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 43395 | 273658 | 423898 | 2327100 | 1872359 | 18043143 | 26625541 | 3979224 | 33311920 | [5 empty] | LFSM Samples |  |  |  | Measured Concentration, ng/L | [9 empty] | Quality Control Checks
134 | 2023-01-21 04:32 | AB-230117-02 LRB 9 | 166 | K | QC | 0 | #N/A | 0 | 0.794 | #N/A | 0.552 | 0.229 | #N/A | 0.312 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 386958 | 837006 | 1298730 | 7073193 | 3396911 | 30238560 | 34633235 | 4029941 | 41477591 |  |  | 134 |  |  | # | QC Samples | Type | Dilution Factor: | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | True Value | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA
135 | 2023-01-21 04:59 | LD4-inj1 | 165 | D | QC | 0 | #N/A | 1.695 | 0 | #N/A | 0.397 | 0.342 | #N/A | 3.644 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Duplicate |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 37954 | 246215 | 380776 | 2132653 | 1778462 | 16990993 | 24656530 | 3622521 | 31087726 |  |  |  | 135
136 | 2023-01-21 05:26 | AB-230117-02 LD 4 | 164 | D | QC | 0 | #N/A | 1.281 | 0 | #N/A | 0.502 | 0.314 | #N/A | 3.87 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Duplicate |  | IS3 | IS4 | IS4 | IS4 | IS4 | [10 empty] | 20917 | 196260 | 312969 | 1760888 | 1650355 | 15897027 | 23339370 | 3415913 | 29138161 |  |  |  | 136
137 | 2023-01-21 05:53 | AB-230117-02 LFB-H 5 | 163 | Z | QC | 158.258 | #N/A | 179.705 | 179.093 | #N/A | 179.201 | 181.423 | #N/A | 180.317 | 1 | 158.258 | #N/A | 179.705 | 179.093 | #N/A | 179.201 | 181.423 | #N/A | 180.317 | H 5 | H 5 | NA | Blank Spike | No TrueV! | [15 empty] | 399255 | 830324 | 1328738 | 6832898 | 2845482 | 25545598 | 28676707 | 3116264 | 32549183 |  |  | 137 |  |  | 46 | AB-230117-02 LFSM-H 1 @5 | S | 1 | 570.426 | #N/A | 197.796 | 197.561 | #N/A | 202.268 | 209.365 | #N/A | 212.68 | H 1 @5 | [18 empty] | FAIL
138 | 2023-01-21 06:20 | AB-230117-02 CCC-H 5 | 162 | Q | QC | 309.052 | #N/A | 375.935 | 355.016 | #N/A | 353.128 | 357.216 | #N/A | 355.469 | 1 | 309.052 | #N/A | 375.935 | 355.016 | #N/A | 353.128 | 357.216 | #N/A | 355.469 | 2 | 02 | NA | CCC | [16 empty] | 442018 | 857556 | 1310115 | 6803754 | 2362328 | 21883588 | 23985052 | 2634839 | 26265725 |  | 138
139 | 2023-01-21 06:47 | AB-230117-02 LFSM-H 4 | 161 | S | QC | 189.256 | #N/A | 200.054 | 184.951 | #N/A | 197.378 | 204.634 | #N/A | 213.294 | 1 | 189.256 | #N/A | 200.054 | 184.951 | #N/A | 197.378 | 204.634 | #N/A | 213.294 | H 4 | NA | H 4 | Sample Spike | No TrueV! | IS3 | IS3 | IS3 | IS3 | [11 empty] | 87624 | 355626 | 549664 | 3032232 | 2072653 | 18871381 | 23253872 | 3159809 | 26066773 |  |  |  |  | 139 | 74 | AB-230117-02 LFSM-H 2 | S | 1 | 169.438 | #N/A | 198.319 | 190.263 | #N/A | 196.626 | 202.465 | #N/A | 206.37 | H 2 | [18 empty] | FAIL |  | See QAQC Overflow page
140 | 2023-01-21 07:41 | WO-23003-81-A-1 | 160 | U |  | 319.278 | #N/A | 0 | 0 | #N/A | 0.541 | 0.62 | #N/A | 0.574 | 1 | 319.278 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 93811 | 360637 | 559477 | 3165164 | 2247196 | 21767220 | 29882331 | 4168274 | 37402474
141 | 2023-01-21 08:08 | WO-23003-82-A-1 | 159 | U |  | 602.933 | #N/A | 2.353 | 0 | #N/A | 0.828 | 0.705 | #N/A | 0.63 | 1 | 602.933 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | R2 |  |  | Low internal standard in this sample, looks like around half. |  |  | 54739 | 200694 | 314867 | 1697629 | 1288854 | 12461196 | 17761291 | 2465134 | 21307863 | [5 empty] | 111 | AB-230117-02 LFSM-H 3 | S | 1 | 334.468 | #N/A | 193.406 | 190.483 | #N/A | 191.067 | 198.421 | #N/A | 203.756 | H 3 | [18 empty] | FAIL
142 | 2023-01-21 08:36 | WO-23003-83-A-1 | 158 | U |  | 311.52 | #N/A | 0 | 0 | #N/A | 0.375 | 0.424 | #N/A | 0.398 | 1 | 311.52 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 88731 | 355826 | 550934 | 2958244 | 2218956 | 20622987 | 29319967 | 4111584 | 35975599
143 | 2023-01-21 09:03 | WO-23003-84-A-1 | 157 | U |  | 405.654 | #N/A | 0 | 0 | #N/A | 0.327 | 0.405 | #N/A | 0.229 | 1 | 405.654 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS4 | IS4 | [11 empty] | 97362 | 394791 | 559854 | 3076516 | 2350619 | 21416800 | 26361109 | 4017463 | 31050434 | [5 empty] | 139 | AB-230117-02 LFSM-H 4 | S | 1 | 189.256 | #N/A | 200.054 | 184.951 | #N/A | 197.378 | 204.634 | #N/A | 213.294 | H 4 | [18 empty] | FAIL
144 | 2023-01-21 09:30 | WO-23003-85-A-1 | 156 | U |  | 419.227 | #N/A | 0 | 0 | #N/A | 0.31 | 0.395 | #N/A | 0.326 | 1 | 419.227 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS4 | IS4 | [11 empty] | 95808 | 375658 | 567741 | 3211135 | 2413172 | 22201556 | 28031744 | 4087810 | 32886770
145 | 2023-01-21 09:57 | WO-23003-86-A-1 | 155 | U |  | 433.77 | #N/A | 0.871 | 0 | #N/A | 0.35 | 0.431 | #N/A | 0.246 | 1 | 433.77 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 45454 | 285832 | 431507 | 2547631 | 2010931 | 19859541 | 28078988 | 4037136 | 34794228 | [5 empty] | 167 | AB-230117-02 LFSM-H 5 | S | 1 | 182.099 | #N/A | 178.008 | 176.535 | #N/A | 182.066 | 186.755 | #N/A | 223.297 | H 5 | [18 empty] | FAIL
146 | 2023-01-21 10:24 | WO-23003-87-A-1 | 154 | U |  | 406.54 | #N/A | 0 | 0 | #N/A | 0.359 | 0.368 | #N/A | 0.219 | 1 | 406.54 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 30259 | 220206 | 359063 | 2003607 | 1712563 | 16981543 | 25806511 | 3880549 | 33573338
147 | 2023-01-21 10:51 | WO-23003-88-A-1 | 153 | U |  | 382.709 | #N/A | 0 | 0 | #N/A | 0.287 | 0.342 | #N/A | 0.223 | 1 | 382.709 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS4 | IS4 | [11 empty] | 98078 | 390901 | 573552 | 3178984 | 2367506 | 21706626 | 29585549 | 4184832 | 36457724 | [5 empty] | 195 | AB-230117-02 LFSM-H 6 | S | 1 | 160.732 | #N/A | 171.837 | 176.045 | #N/A | 184.547 | 191.689 | #N/A | 2843.453 | H 6 | [18 empty] | FAIL
148 | 2023-01-21 11:18 | WO-23003-89-A-1 | 152 | U |  | 415.919 | #N/A | 0 | 0 | #N/A | 0.284 | 0.291 | #N/A | 0.192 | 1 | 415.919 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 82000 | 350298 | 557271 | 2981804 | 2267966 | 21841201 | 29903150 | 4240476 | 36900501
149 | 2023-01-21 11:45 | WO-23003-90-A-1 | 151 | U |  | 365.485 | #N/A | 0 | 0 | #N/A | 0.344 | 0.321 | #N/A | 0.371 | 1 | 365.485 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 69073 | 309238 | 489232 | 2768161 | 2158117 | 20567756 | 28199051 | 4188114 | 33684003 | [5 empty] | 229 | AB-230117-02 LFSM-H 7 | S | 1 | 188.826 | #N/A | 175.013 | 173.568 | #N/A | 182.658 | 188.523 | #N/A | 188.842 | H 7 | [18 empty] | FAIL
150 | 2023-01-21 12:12 | LRB-10-inj1 | 150 | K | QC | 0 | #N/A | 0 | 0.79 | #N/A | 0.515 | 0.261 | #N/A | 0.283 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 390942 | 854107 | 1318397 | 7266040 | 3525650 | 30607554 | 35895538 | 4033211 | 43561553 |  |  | 150
151 | 2023-01-21 12:39 | AB-230117-02 LRB 10 | 149 | K | QC | 0 | #N/A | 0.079 | 0.848 | #N/A | 0.571 | 0.264 | #N/A | 0.321 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 393815 | 821316 | 1316014 | 7168889 | 3334539 | 28633631 | 33684803 | 3766065 | 40242698 |  |  | 151 |  |  | 257 | AB-230117-02 LFSM-H 8 | S | 1 | 177.244 | #N/A | 183.28 | 170.633 | #N/A | 178.895 | 185.896 | #N/A | 191.129 | H 8 | [18 empty] | FAIL
152 | 2023-01-21 13:06 | WO-23003-91-A-1 | 148 | U |  | 385.432 | #N/A | 0 | 0 | #N/A | 0.38 | 0.317 | #N/A | 0.408 | 1 | 385.432 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 48649 | 289060 | 458185 | 2551234 | 2073458 | 20445424 | 28354687 | 4097531 | 34928638
153 | 2023-01-21 13:33 | WO-23003-92-A-1 | 147 | U |  | 385.868 | #N/A | 0 | 0 | #N/A | 0.312 | 0.343 | #N/A | 0.174 | 1 | 385.868 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | [14 empty] | 111912 | 415217 | 631015 | 3722523 | 2653892 | 21748898 | 26672539 | 4256224 | 33403984 | [5 empty] | 285 | AB-230117-02 LFSM-H 9 | S | 1 | 168.758 | #N/A | 184.196 | 179.695 | #N/A | 182.596 | 190.975 | #N/A | 189.486 | H 9 | [18 empty] | FAIL
154 | 2023-01-21 14:00 | WO-23003-93-A-1 | 146 | U |  | 377.744 | #N/A | 0 | 0 | #N/A | 0.277 | 0.391 | #N/A | 0.324 | 1 | 377.744 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 62051 | 337582 | 531128 | 2874110 | 2216247 | 21726871 | 29449513 | 4303332 | 35550674
155 | 2023-01-21 14:27 | WO-23003-94-A-1 | 145 | U |  | 396.942 | #N/A | 0 | 0 | #N/A | 0.27 | 0.352 | #N/A | 0.307 | 1 | 396.942 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS4 | IS4 | [11 empty] | 90503 | 400040 | 587438 | 3204689 | 2432299 | 23272982 | 32934546 | 4484651 | 40658864 | [5 empty] | 313 | AB-230117-02 LFSM-H 10 | S | 1 | 194.12 | #N/A | 183.533 | 188.607 | #N/A | 190.703 | 192.309 | #N/A | 194.699 | H 10 | [18 empty] | FAIL
156 | 2023-01-21 14:54 | WO-23003-95-A-1 | 144 | U |  | 396.948 | #N/A | 0 | 0 | #N/A | 0.24 | 0.271 | #N/A | 0.195 | 1 | 396.948 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS4 | IS4 | [11 empty] | 87444 | 414155 | 591866 | 3290070 | 2438136 | 23651394 | 33004661 | 4621082 | 41590416
157 | 2023-01-21 15:22 | WO-23003-96-A-1 | 143 | U |  | 396.134 | #N/A | 0 | 0 | #N/A | 0.191 | 0.291 | #N/A | 0.168 | 1 | 396.134 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 86582 | 369718 | 570404 | 3137650 | 2335994 | 22640971 | 30472169 | 4407111 | 38779746 | [5 empty] | 355 | 217296-d1000-LFSM-200-reprep | S | 1000 | 166.141 | #N/A | 199.356 | 189.823 | #N/A | 180.753 | 184.164 | #N/A | 190.796 | 200 | [18 empty] | FAIL |  | See QAQC Overflow page
158 | 2023-01-21 15:49 | WO-23003-97-A-1 | 142 | U |  | 382.008 | #N/A | 0.489 | 0 | #N/A | 0.314 | 0.355 | #N/A | 0.572 | 1 | 382.008 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS4 | IS4 | [11 empty] | 85366 | 388074 | 528734 | 3111777 | 2284322 | 21661622 | 28777567 | 4408049 | 35663824
159 | 2023-01-21 16:16 | WO-23003-98-A-1 | 141 | U |  | 403.126 | #N/A | 0 | 0 | #N/A | 0.294 | 0.335 | #N/A | 0.211 | 1 | 403.126 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS4 | IS4 | [11 empty] | 90738 | 373117 | 576514 | 3129094 | 2241053 | 22766224 | 31947734 | 4420815 | 40201443 | [5 empty] | 356 | 217336-d2-LFSM-200-reprep | S | 2 | 606.641 | #N/A | 185.585 | 184.866 | #N/A | 175.24 | 187.674 | #N/A | 194.566 | 200 | [18 empty] | FAIL |  | See QAQC Overflow page
160 | 2023-01-21 16:43 | WO-23003-99-A-1 | 140 | U |  | 447.416 | #N/A | 0 | 0 | #N/A | 0.262 | 0.309 | #N/A | 0.186 | 1 | 447.416 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 41441 | 296170 | 435231 | 2419840 | 1936346 | 19214059 | 28842466 | 4178947 | 38353517
161 | 2023-01-21 17:10 | WO-23003-100-A-1 | 139 | U |  | 410.536 | #N/A | 0.321 | 0 | #N/A | 0.224 | 0.31 | #N/A | 0.224 | 1 | 410.536 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 75533 | 344729 | 539778 | 2963581 | 2109992 | 21267105 | 31068711 | 4328504 | 39237589 | [5 empty] | 357 | 217376-d2-LFSM-200-reprep | S | 2 | 196.844 | #N/A | 173.032 | 173.811 | #N/A | 166.325 | 177.327 | #N/A | 266.746 | 200 | [18 empty] | FAIL |  | See QAQC Overflow page
162 | 2023-01-21 17:37 | AB-230117-02 LRB 11 | 138 | K | QC | 0 | #N/A | 0.345 | 0.756 | #N/A | 0.55 | 0.267 | #N/A | 0.264 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 366041 | 852826 | 1260144 | 6981552 | 3472176 | 30665008 | 35547285 | 4082172 | 43350177 |  |  | 162
163 | 2023-01-21 18:04 | LD-5-inj1 | 137 | D | QC | 0 | #N/A | 0 | 0.802 | #N/A | 0.657 | 0.477 | #N/A | 82.9 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 82.9 | NA | NA | NA | Duplicate |  | IS3 |  | IS4 | IS4 |  |  |  |  | D |  |  |  | Duplicate failure. Reprepared below and passed. |  |  | 82474 | 380794 | 584499 | 3291411 | 2654414 | 25128152 | 32939070 | 4291168 | 38110657 |  |  |  | 163 |  | 358 | 217396-d200-LFSM-200 | S | 200 | 153.37 | #N/A | 175.395 | 175.651 | #N/A | 168.361 | 178.6 | #N/A | 4040.399 | 200 | [18 empty] | FAIL |  | See QAQC Overflow page
164 | 2023-01-21 18:31 | AB-230117-02 LD 5 | 136 | D | QC | 0 | #N/A | 0 | 0.77 | #N/A | 0.645 | 0.438 | #N/A | 88.093 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 88.093 | NA | NA | NA | Duplicate |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | D |  |  |  | Duplicate failure. Reprepared below and passed. |  |  | 69999 | 333710 | 543923 | 3003137 | 2467575 | 23295803 | 30311159 | 3998089 | 34578186 |  |  |  | 164
165 | 2023-01-21 18:58 | AB-230117-02 LFB-H 6 | 135 | Z | QC | 158.298 | #N/A | 179.368 | 175.354 | #N/A | 173.631 | 176.742 | #N/A | 176.054 | 1 | 158.298 | #N/A | 179.368 | 175.354 | #N/A | 173.631 | 176.742 | #N/A | 176.054 | H 6 | H 6 | NA | Blank Spike | No TrueV! | [15 empty] | 363560 | 822027 | 1249401 | 6712510 | 2775932 | 25321860 | 28650436 | 3062121 | 33569510 |  |  | 165 |  |  | 359 | 217360-d2-LFSM-200 | S | 2 | 182.595 | #N/A | 187.722 | 186.158 | #N/A | 185.148 | 185.543 | #N/A | 384.137 | 200 | [18 empty] | FAIL |  | See QAQC Overflow page
166 | 2023-01-21 19:25 | AB-230117-02 CCC-H 6 | 134 | Q | QC | 304.261 | #N/A | 352.766 | 339.3 | #N/A | 333.08 | 341.007 | #N/A | 341.563 | 1 | 304.261 | #N/A | 352.766 | 339.3 | #N/A | 333.08 | 341.007 | #N/A | 341.563 | 2 | 02 | NA | CCC | [16 empty] | 379586 | 810488 | 1219465 | 6477511 | 2380427 | 22052524 | 24084407 | 2540641 | 27124276 |  | 166
167 | 2023-01-21 19:52 | AB-230117-02 LFSM-H 5 | 133 | S | QC | 182.099 | #N/A | 178.008 | 176.535 | #N/A | 182.066 | 186.755 | #N/A | 223.297 | 1 | 182.099 | #N/A | 178.008 | 176.535 | #N/A | 182.066 | 186.755 | #N/A | 223.297 | H 5 | NA | H 5 | Sample Spike | No TrueV! | IS3 | [14 empty] | 89635 | 455406 | 727272 | 4003771 | 2828511 | 24178040 | 26412031 | 3427596 | 28982045 |  |  |  |  | 167 | 360 | 217416-d2-LFSM-200-reprep | S | 2 | 199.788 | #N/A | 183.534 | 184.672 | #N/A | 186.591 | 187.492 | #N/A | 195.064 | 200 | [18 empty] | FAIL |  | See QAQC Overflow page
168 | 2023-01-21 20:47 | WO-23003-101-A-1 | 132 | U |  | 0 | #N/A | 0 | 0.912 | #N/A | 0.618 | 0.649 | #N/A | 28.256 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 28.256 | NA | NA | NA | Unknown |  | IS3 |  | IS4 | IS4 | [11 empty] | 71332 | 376523 | 558194 | 3044917 | 2309340 | 22389687 | 31868378 | 4216694 | 38768749
169 | 2023-01-21 21:14 | WO-23003-102-A-1 | 131 | U |  | 0 | #N/A | 0.788 | 1.032 | #N/A | 0.732 | 0.666 | #N/A | 64.801 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 64.801 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 60077 | 274722 | 415107 | 2251411 | 1056082 | 10581129 | 18015535 | 2695615 | 27775775
170 | 2023-01-21 21:41 | WO-23003-103-A-1 | 130 | U |  | 0 | #N/A | 0 | 0.975 | #N/A | 0.702 | 0.525 | #N/A | 66.78 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 66.78 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | IS4 | [10 empty] | 62960 | 308437 | 448814 | 2466184 | 1552503 | 15504438 | 26318356 | 3727708 | 35768513
171 | 2023-01-21 22:08 | WO-23003-104-A-1 | 129 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.713 | 0.618 | #N/A | 128.717 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 128.717 | NA | NA | NA | Unknown |  | IS3 | [14 empty] | 119110 | 447858 | 697535 | 3711085 | 2776678 | 26677085 | 34076493 | 4577412 | 37655788
172 | 2023-01-21 22:35 | WO-23003-105-A-1 | 128 | U |  | 0 | #N/A | 0.4 | 0 | #N/A | 0.819 | 0.695 | #N/A | 128.03 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 128.03 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 44312 | 275513 | 432451 | 2374984 | 1697435 | 17543190 | 27261056 | 3832481 | 34342726
173 | 2023-01-21 23:02 | WO-23003-106-A-1 | 127 | U |  | 0 | #N/A | 0.23 | 0 | #N/A | 0.742 | 0.515 | #N/A | 137.351 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 137.351 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 61625 | 280040 | 421655 | 2339066 | 1306281 | 13398584 | 22560018 | 3342309 | 30720915
174 | 2023-01-21 23:29 | WO-23003-107-A-1 | 126 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.668 | 0.464 | #N/A | 141.587 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 141.587 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 79995 | 336779 | 474385 | 2685215 | 1848875 | 18844557 | 28511479 | 3963746 | 34808537
175 | 2023-01-21 23:56 | WO-23003-108-A-1 | 125 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.504 | 0.484 | #N/A | 117.876 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 117.876 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 40644 | 243079 | 371179 | 2020116 | 1222131 | 12475249 | 22459853 | 3400362 | 31394130
176 | 2023-01-22 00:23 | WO-23003-109-A-1 | 124 | U |  | 0 | #N/A | 0 | 0.53 | #N/A | 0.575 | 0.506 | #N/A | 137.434 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 137.434 | NA | NA | NA | Unknown |  | IS3 |  | IS4 | IS4 | [11 empty] | 80475 | 370341 | 559441 | 3179144 | 2499747 | 24204764 | 33272051 | 4340752 | 37906152
177 | 2023-01-22 00:50 | WO-23003-110-A-1 | 123 | U |  | 0 | #N/A | 0 | 1.761 | #N/A | 3.639 | 2.107 | #N/A | 666.017 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 666.017 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 85601 | 358982 | 543250 | 3090429 | 2298698 | 23004724 | 31172627 | 4257264 | 25868952
178 | 2023-01-22 01:17 | LRB-12-inj1 | 122 | K | QC | 0 | #N/A | 0 | 0.799 | #N/A | 0.648 | 0.302 | #N/A | 24.998 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 24.998 | NA | NA | NA | Blank | [9 empty] | b | [6 empty] | 368063 | 837112 | 1263690 | 6972498 | 3497300 | 31493553 | 37572020 | 4072598 | 43582855 |  |  | 178
179 | 2023-01-22 01:44 | AB-230117-02 LRB 12 | 121 | K | QC | 0 | #N/A | 0 | 0.758 | #N/A | 0.534 | 0.278 | #N/A | 1.003 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 362305 | 846668 | 1260290 | 7028879 | 3298787 | 29379328 | 34883484 | 3753277 | 41966592 |  |  | 179
180 | 2023-01-22 02:12 | WO-23003-111-A-1 | 120 | U |  | 0 | #N/A | 0 | 1.35 | #N/A | 4.279 | 2.348 | #N/A | 740.764 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 740.764 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [5 empty] | R1 |  |  | Electively rerun below. |  |  | 67943 | 328678 | 508806 | 2876144 | 2170294 | 21485097 | 31212515 | 4174581 | 25777832
181 | 2023-01-22 02:39 | WO-23003-112-A-1 | 119 | U |  | 0 | #N/A | 0 | 1.139 | #N/A | 2.811 | 1.711 | #N/A | 536.317 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 536.317 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b | [6 empty] | 74317 | 364283 | 528078 | 2902462 | 2237851 | 22209911 | 30748381 | 4242494 | 27511818
182 | 2023-01-22 03:06 | WO-23003-113-A-1 | 118 | U |  | 0 | #N/A | 0 | 0.765 | #N/A | 2.136 | 1.514 | #N/A | 641.876 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 641.876 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b | [6 empty] | 70741 | 353182 | 516099 | 2967227 | 2249019 | 22864066 | 31830157 | 4327492 | 27159198
183 | 2023-01-22 03:33 | WO-23003-114-A-1 | 117 | U |  | 0 | #N/A | 0 | 1.029 | #N/A | 2.034 | 1.569 | #N/A | 622.53 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 622.53 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b | [6 empty] | 74455 | 347830 | 529838 | 3000690 | 2270014 | 22992151 | 31365991 | 4320860 | 27456944
184 | 2023-01-22 04:00 | WO-23003-115-A-1 | 116 | U |  | 0 | #N/A | 0.437 | 1.365 | #N/A | 2.371 | 1.895 | #N/A | 714.013 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 714.013 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 88662 | 354215 | 550322 | 3003752 | 2318518 | 23193579 | 30846423 | 4320818 | 25004515
185 | 2023-01-22 04:27 | WO-23003-116-A-1 | 115 | U |  | 0 | #N/A | 0 | 1.71 | #N/A | 11.677 | 8.8 | #N/A | 2477.13 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 11.677 | 8.8 | #N/A | 2477.13 | NA | NA | NA | Unknown |  | IS3 | [7 empty] | IS4 | [6 empty] | 140200 | 415987 | 654291 | 3620587 | 2579445 | 24439633 | 33143414 | 3932868 | 15028164
186 | 2023-01-22 04:54 | WO-23003-117-A-1 | 114 | U |  | 0 | #N/A | 0.376 | 1.608 | #N/A | 10.943 | 8.243 | #N/A | 2511.685 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 10.943 | 8.243 | #N/A | 2511.685 | NA | NA | NA | Unknown |  | IS3 | [7 empty] | IS4 | [6 empty] | 137545 | 434364 | 656751 | 3649594 | 2638588 | 25776455 | 34091479 | 4130614 | 15207107
187 | 2023-01-22 05:21 | WO-23003-118-A-1 | 113 | U |  | 0 | #N/A | 0 | 1.751 | #N/A | 12.301 | 9.413 | #N/A | 2857.187 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 12.301 | 9.413 | #N/A | 2857.187 | NA | NA | NA | Unknown |  | IS3 | [7 empty] | IS4 | [6 empty] | 157920 | 455736 | 702467 | 3944874 | 2874973 | 27047440 | 34571765 | 4072631 | 13732261
188 | 2023-01-22 05:48 | WO-23003-119-A-1 | 112 | U |  | 0 | #N/A | 0 | 1.049 | #N/A | 3.789 | 2.851 | #N/A | 1085.45 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 1085.45 | NA | NA | NA | Unknown |  | IS3 | [8 empty] | R1 |  |  | Electively rerun below |  |  | 152503 | 470814 | 705098 | 3887057 | 2850510 | 27604625 | 34960984 | 4353409 | 22798424
189 | 2023-01-22 06:15 | WO-23003-120-A-1 | 111 | U |  | 0 | #N/A | 0.615 | 1.294 | #N/A | 4.447 | 3.724 | #N/A | 1329.015 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 1329.015 | NA | NA | NA | Unknown |  | IS3 | [8 empty] | R1 |  |  | Electively rerun below |  |  | 139198 | 434712 | 680859 | 3851120 | 2644378 | 26321859 | 34541785 | 4314884 | 21611490
190 | 2023-01-22 06:42 | AB-230117-02 LRB 13 | 110 | K | QC | 0 | #N/A | 0.263 | 0.748 | #N/A | 0.62 | 0.374 | #N/A | 69.087 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 69.087 | NA | NA | NA | Blank | [9 empty] | b | [6 empty] | 327227 | 808357 | 1190410 | 6512473 | 3308258 | 30286976 | 36636843 | 3878390 | 42552876 |  |  | 190
191 | 2023-01-22 07:10 | LD6-inj1 | 109 | D | QC | 0 | #N/A | 0 | 0.883 | #N/A | 1.362 | 2.893 | #N/A | 3922.707 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 3922.707 | NA | NA | NA | Duplicate | [9 empty] | IS4 | [6 empty] | 346314 | 816555 | 1177883 | 6657933 | 3345917 | 31273925 | 37076510 | 3868396 | 11938961 |  |  |  | 191
192 | 2023-01-22 07:37 | AB-230117-02 LD 6 | 108 | D | QC | 0 | #N/A | 0 | 0.865 | #N/A | 1.374 | 3.079 | #N/A | 4142.057 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4142.057 | NA | NA | NA | Duplicate | [9 empty] | IS4 | [6 empty] | 340517 | 796687 | 1197901 | 6648637 | 3246144 | 29696384 | 34722557 | 3770630 | 11439467 |  |  |  | 192
193 | 2023-01-22 08:04 | AB-230117-02 LFB-H 7 | 107 | Z | QC | 157.151 | #N/A | 181.009 | 182.079 | #N/A | 177.55 | 177.827 | #N/A | 303.883 | 1 | 157.151 | #N/A | 181.009 | 182.079 | #N/A | 177.55 | 177.827 | #N/A | 303.883 | H 7 | H 7 | NA | Blank Spike | No TrueV! | [15 empty] | 356495 | 795193 | 1194388 | 6414093 | 2802953 | 26765152 | 30536985 | 3134189 | 32994516 |  |  | 193
194 | 2023-01-22 08:31 | AB-230117-02 CCC-H 7 | 106 | Q | QC | 307.839 | #N/A | 372.196 | 355.298 | #N/A | 349.262 | 350.888 | #N/A | 358.688 | 1 | 307.839 | #N/A | 372.196 | 355.298 | #N/A | 349.262 | 350.888 | #N/A | 358.688 | 2 | 02 | NA | CCC | [16 empty] | 353402 | 749296 | 1158056 | 6157485 | 2371745 | 23190512 | 26586794 | 2628619 | 29369183 |  | 194
195 | 2023-01-22 08:58 | AB-230117-02 LFSM-H 6 | 105 | S | QC | 160.732 | #N/A | 171.837 | 176.045 | #N/A | 184.547 | 191.689 | #N/A | 2843.453 | 1 | 160.732 | #N/A | 171.837 | 176.045 | #N/A | 184.547 | 191.689 | #N/A | 2843.453 | H 6 | NA | H 6 | Sample Spike | No TrueV! | IS3 | [7 empty] | IS3 | [6 empty] | 142739 | 413834 | 634016 | 3534758 | 2373337 | 22728676 | 28306027 | 3280289 | 13096971 |  |  |  |  | 195
196 | 2023-01-22 09:52 | WO-23003-121-A-1 | 104 | U |  | 0 | #N/A | 0 | 1.724 | #N/A | 8.728 | 8.418 | #N/A | 2598.122 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 8.728 | 8.418 | #N/A | 2598.122 | NA | NA | NA | Unknown |  | IS3 | [7 empty] | IS4 | R1 |  |  | Electively rerun below |  |  | 164724 | 477397 | 701356 | 3912310 | 2724072 | 26466602 | 33793745 | 4032880 | 14897412
197 | 2023-01-22 10:19 | WO-23003-122-A-1 | 103 | U |  | 0 | #N/A | 0 | 0.942 | #N/A | 1.505 | 3.298 | #N/A | 4052.461 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4052.461 | NA | NA | NA | Unknown | [9 empty] | IS4 | [6 empty] | 317514 | 799705 | 1172222 | 6513227 | 3260898 | 30676932 | 35822718 | 3824035 | 11794822
198 | 2023-01-22 10:46 | WO-23003-123-A-1 | 102 | U |  | 0 | #N/A | 0 | 0.825 | #N/A | 1.448 | 3.799 | #N/A | 4114.272 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4114.272 | NA | NA | NA | Unknown | [9 empty] | IS4 | [6 empty] | 341203 | 787922 | 1191848 | 6652997 | 3447890 | 31781255 | 36705423 | 3914703 | 11760718
199 | 2023-01-22 11:13 | WO-23003-124-A-1 | 219 | U |  | 0 | #N/A | 0 | 0.721 | #N/A | 1.325 | 2.549 | #N/A | 4254.003 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4254.003 | NA | NA | NA | Unknown | [9 empty] | IS4 | [6 empty] | 344771 | 782216 | 1147141 | 6545217 | 3413219 | 31540850 | 37306194 | 4028406 | 11648092
200 | 2023-01-23 10:08 | Rinse12 | 218 | U |  | 0 | #N/A | 40.266 | 0 | #N/A | 11.109 | 6.109 | #N/A | 3907 | 1 | <MRL | #N/A | 40.266 | <MRL | #N/A | 11.109 | 6.109 | #N/A | 3907 | NA | NA | NA | Unknown |  | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 |  |  |  | These are conditioning rinses, not true QC samples |  |  | 12916 | 6697 | 18818 | 100284 | 124329 | 1287286 | 2143667 | 265578 | 2431795
201 | 2023-01-23 10:35 | AB-230117-02 LRB 12 20230123102918 | 217 | K | QC | 0 | #N/A | 0 | 0 | #N/A | 49.489 | 26.569 | #N/A | 943.114 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 49.489 | 26.569 | #N/A | 943.114 | NA | NA | NA | Blank |  | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 |  |  |  | These are conditioning rinses, not true QC samples |  |  | 7694 | 7032 | 9823 | 50717 | 22189 | 176137 | 226469 | 24467 | 273976 |  |  | 201
202 | 2023-01-23 11:02 | AB-230117-02 LRB 13 20230123105625 | 216 | K | QC | 0 | #N/A | 0.169 | 0.805 | #N/A | 0.61 | 0.281 | #N/A | 3.33 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [9 empty] | b | [6 empty] | 477597 | 899201 | 1377450 | 7254703 | 3126958 | 27657980 | 32002182 | 3550476 | 37983588 |  |  | 202
203 | 2023-01-23 11:29 | WO-23003-111-A-1B | 215 | U |  | 0 | #N/A | 0 | 1.749 | #N/A | 4.534 | 2.374 | #N/A | 738.645 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 738.645 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 52232 | 285769 | 445432 | 2481005 | 2043807 | 19681184 | 27636877 | 3769791 | 24339869
204 | 2023-01-23 11:56 | WO-23003-121-A-1B | 214 | U |  | 0 | #N/A | 0.285 | 1.486 | #N/A | 8.541 | 7.928 | #N/A | 2538.308 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 8.541 | 7.928 | #N/A | 2538.308 | NA | NA | NA | Unknown | [9 empty] | IS4 | R1 |  |  | Reprepared and rerun below |  |  | 191988 | 507402 | 745087 | 4054371 | 2728083 | 24419468 | 30006898 | 3747443 | 13605851
205 | 2023-01-23 12:24 | WO-23003-125-A-1 | 213 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.512 | 0.413 | #N/A | 46.972 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 46.972 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 |  |  |  |  | C | R1 |  |  | Rerun due to carryover concern |  |  | 32552 | 256722 | 367291 | 2125778 | 1690241 | 15625123 | 22372578 | 3598788 | 28262592
206 | 2023-01-23 12:51 | WO-23003-126-A-1 | 212 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0 | 0.412 | #N/A | 4.716 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  | R1 |  |  | Rerun due to carryover concern |  |  | 30509 | 184292 | 292620 | 1591971 | 846387 | 8130357 | 12752173 | 1976140 | 19273236
207 | 2023-01-23 13:18 | WO-23003-127-A-1 | 211 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.451 | 0.375 | #N/A | 1.887 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | [9 empty] | 79591 | 289169 | 446521 | 2451430 | 1117778 | 11281829 | 19103573 | 2945807 | 28731300
208 | 2023-01-23 13:45 | WO-23003-128-A-1 | 210 | U |  | 0 | #N/A | 0.672 | 0 | #N/A | 0.583 | 0.347 | #N/A | 1.569 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [8 empty] | 45411 | 237811 | 350118 | 1982972 | 934843 | 9105334 | 14623735 | 2260065 | 23495238
209 | 2023-01-23 14:12 | WO-23003-129-A-1 | 209 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.451 | 0.345 | #N/A | 1.157 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | [9 empty] | 60467 | 277829 | 408787 | 2274633 | 1148923 | 11565885 | 18874520 | 3174631 | 28034179
210 | 2023-01-23 14:39 | WO-23003-130-A-1 | 208 | U |  | 0 | #N/A | 0.482 | 0 | #N/A | 0.481 | 0.296 | #N/A | 0.778 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 37183 | 263797 | 392233 | 2212515 | 1830424 | 17821794 | 25163894 | 3863138 | 32804320
211 | 2023-01-23 15:06 | LRB-14-inj1 | 207 | K | QC | 0 | #N/A | 0 | 0.799 | #N/A | 0.486 | 0.247 | #N/A | 0.68 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 523742 | 970248 | 1423067 | 7939528 | 3645393 | 31459571 | 36416657 | 4127495 | 44684916 |  |  | 211
212 | 2023-01-23 15:33 | AB-230117-02 LRB 14 | 206 | K | QC | 0 | #N/A | 0 | 0.844 | #N/A | 0.577 | 0.274 | #N/A | 0.664 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 515255 | 984231 | 1427915 | 7990089 | 3160520 | 28640931 | 33465281 | 3665930 | 40930026 |  |  | 212
213 | 2023-01-23 16:00 | WO-23003-131-A-1 | 205 | U |  | 0 | #N/A | 0.329 | 0 | #N/A | 0.366 | 0.306 | #N/A | 0.748 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 94356 | 347480 | 530790 | 2844506 | 1809424 | 18207144 | 26549325 | 3956493 | 34747393
214 | 2023-01-23 16:27 | WO-23003-132-A-1 | 204 | U |  | 0 | #N/A | 0 | 0 | #N/A | 1.113 | 1.47 | #N/A | 0.77 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | [10 empty] | 31432 | 237812 | 362798 | 2055522 | 1478371 | 15305524 | 24372583 | 3707534 | 33892710
215 | 2023-01-23 16:54 | WO-23003-133-A-1 | 203 | U |  | 0 | #N/A | 1.203 | 0 | #N/A | 0.367 | 0.295 | #N/A | 0.523 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | [9 empty] | 61954 | 274414 | 406838 | 2242924 | 1214269 | 12101965 | 19143175 | 3200615 | 27027843
216 | 2023-01-23 17:21 | WO-23003-134-A-1 | 202 | U |  | 0 | #N/A | 0.198 | 0 | #N/A | 0.371 | 0.261 | #N/A | 0.414 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 98231 | 365967 | 543832 | 3030173 | 2043914 | 19418592 | 26244060 | 4282981 | 34656130
217 | 2023-01-23 17:49 | WO-23003-135-A-1 | 201 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.412 | 0.333 | #N/A | 0.443 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | [10 empty] | 103340 | 351604 | 508344 | 2814769 | 1568439 | 14745212 | 22759139 | 3590842 | 30085287
218 | 2023-01-23 18:16 | WO-23003-136-A-1 | 200 | U |  | 0 | #N/A | 1.066 | 0 | #N/A | 0.251 | 0.251 | #N/A | 0.437 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | [9 empty] | 10602 | 156718 | 255224 | 1406570 | 1386247 | 13518015 | 22565939 | 3699614 | 30705787
219 | 2023-01-23 18:43 | WO-23003-137-A-1 | 326 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.319 | 0.258 | #N/A | 0.386 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | [14 empty] | 132259 | 456434 | 629946 | 3548669 | 2499664 | 22937395 | 29179604 | 4280646 | 35718603
220 | 2023-01-23 19:10 | WO-23003-138-A-1 | 325 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.209 | 0.284 | #N/A | 0.368 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS3 | IS4 | [11 empty] | 96973 | 379201 | 583191 | 3349023 | 2401875 | 22679273 | 28595325 | 4161902 | 35416341
221 | 2023-01-27 08:27 | WO-23003-139-A-1 | 324 | U |  | 0 | #N/A | 3.112 | 0 | #N/A | 0.695 | 0.57 | #N/A | 133.922 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 133.922 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 |  |  |  |  | C | R1 |  |  | Analyzed after instrument leak, some carryover observed and sample rerun |  |  | 75063 | 200900 | 501825 | 2565629 | 1893044 | 18606048 | 26342001 | 3544326 | 31886772
222 | 2023-01-27 08:54 | WO-23003-139-A-1B | 323 | U |  | 0 | #N/A | 2.591 | 0 | #N/A | 0.378 | 0.345 | #N/A | 6.248 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 6.248 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 |  |  |  |  | C | [6 empty] | 47777 | 247705 | 393619 | 2194892 | 1842386 | 18277425 | 27124738 | 3553271 | 35046837
223 | 2023-01-27 09:21 | WO-23003-140-A-1 | 322 | U |  | 0 | #N/A | 0.914 | 0 | #N/A | 0.357 | 0.338 | #N/A | 2.01 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 50264 | 269934 | 386869 | 2241844 | 1776710 | 18299776 | 27809792 | 3732851 | 36342071
224 | 2023-01-27 09:48 | LRB-15-inj1 | 321 | K | QC | 0 | #N/A | 0.374 | 0.91 | #N/A | 0.603 | 0.31 | #N/A | 1.417 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 353297 | 776399 | 1122844 | 6363611 | 3163272 | 28246415 | 33904567 | 3750497 | 42072343 |  |  | 224
225 | 2023-01-27 10:15 | AB-230117-02 LRB 15 | 320 | K | QC | 0 | #N/A | 0.658 | 0.9 | #N/A | 0.525 | 0.302 | #N/A | 1.286 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 302709 | 725349 | 1054039 | 5947417 | 2627401 | 23444026 | 29548719 | 3163740 | 35095576 |  |  | 225
226 | 2023-01-27 10:42 | AB-230117-02 LD 7 | 319 | D | QC | 0 | #N/A | 0 | 1.76 | #N/A | 0.417 | 0.336 | #N/A | 1.197 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Duplicate |  | IS3 | IS4 | IS3 | IS4 | IS4 | [10 empty] | 17781 | 171470 | 258314 | 1583020 | 1522140 | 14327453 | 20794686 | 3377280 | 30179486 |  |  |  | 226
227 | 2023-01-27 11:09 | AB-230117-02 LFB-H 8 | 318 | Z | QC | 149.916 | #N/A | 178.569 | 174.977 | #N/A | 174.263 | 174.528 | #N/A | 172.643 | 1 | 149.916 | #N/A | 178.569 | 174.977 | #N/A | 174.263 | 174.528 | #N/A | 172.643 | H 8 | H 8 | NA | Blank Spike | No TrueV! | [15 empty] | 373163 | 766022 | 1103741 | 6265642 | 2647643 | 24434056 | 29190429 | 3035404 | 35223208 |  |  | 227
228 | 2023-01-27 11:36 | AB-230117-02 CCC-H 8 | 317 | Q | QC | 306.068 | #N/A | 374.743 | 347.778 | #N/A | 345.655 | 353.314 | #N/A | 346.394 | 1 | 306.068 | #N/A | 374.743 | 347.778 | #N/A | 345.655 | 353.314 | #N/A | 346.394 | 2 | 02 | NA | CCC | [16 empty] | 380122 | 793697 | 1095340 | 6025403 | 2157964 | 21121392 | 24578455 | 2383178 | 28531073 |  | 228
229 | 2023-01-27 12:04 | AB-230117-02 LFSM-H 7 | 316 | S | QC | 188.826 | #N/A | 175.013 | 173.568 | #N/A | 182.658 | 188.523 | #N/A | 188.842 | 1 | 188.826 | #N/A | 175.013 | 173.568 | #N/A | 182.658 | 188.523 | #N/A | 188.842 | H 7 | NA | H 7 | Sample Spike | No TrueV! | IS3 | IS3 | IS3 | IS3 | IS3 | [7 empty] | PFDA okay vs second injection, as is PFBS |  |  | 29375 | 219237 | 326061 | 1886454 | 1618655 | 15199777 | 20777930 | 3058586 | 24540149 |  |  |  |  | 229
230 | 2023-01-27 12:58 | WO-23003-141-A-1 | 315 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.824 | 0.724 | #N/A | 0.925 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | [9 empty] | 12317 | 149370 | 225755 | 1333669 | 1267875 | 13563890 | 22187134 | 3333604 | 29934648
231 | 2023-01-27 13:25 | WO-23003-142-A-1 | 314 | U |  | 0 | #N/A | 0 | 1.467 | #N/A | 0.453 | 0.486 | #N/A | 0.92 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 43086 | 269680 | 393507 | 2269744 | 1970255 | 18317393 | 23460827 | 3757708 | 33259350
232 | 2023-01-27 13:52 | WO-23003-143-A-1 | 313 | U |  | 0 | #N/A | 1.526 | 0 | #N/A | 0.43 | 0.363 | #N/A | 0.671 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 54078 | 295991 | 411477 | 2412260 | 1912046 | 19532034 | 29438075 | 3855879 | 40160262
233 | 2023-01-27 14:19 | WO-23003-144-A-1 | 312 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.442 | 0.337 | #N/A | 0.465 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 56921 | 297135 | 430289 | 2550021 | 1800955 | 19453429 | 30179345 | 4008874 | 40148666
234 | 2023-01-27 14:46 | WO-23003-145-A-1 | 311 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.406 | 0.332 | #N/A | 0.419 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 72268 | 317013 | 458870 | 2684127 | 2031906 | 20234548 | 27672780 | 3890902 | 34635365
235 | 2023-01-27 15:13 | WO-23003-146-A-1 | 310 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.535 | 0.324 | #N/A | 0.408 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 61502 | 295815 | 443958 | 2652987 | 2056685 | 20046846 | 25700237 | 3906282 | 32156990
236 | 2023-01-27 15:40 | WO-23003-147-A-1 | 309 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.467 | 0.329 | #N/A | 0.366 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 49981 | 273863 | 429523 | 2453173 | 1906985 | 19525177 | 27348352 | 3832871 | 35745303
237 | 2023-01-27 16:07 | WO-23003-148-A-1 | 308 | U |  | 0 | #N/A | 1.373 | 0 | #N/A | 0.388 | 0.358 | #N/A | 0.359 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 42946 | 242650 | 376745 | 2224399 | 1794748 | 18475522 | 26373254 | 3756816 | 33859088
238 | 2023-01-27 16:34 | WO-23003-149-A-1 | 307 | U |  | 0 | #N/A | 1.108 | 0 | #N/A | 0.237 | 0.346 | #N/A | 0.365 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 33126 | 233992 | 340903 | 2007245 | 1703351 | 17551789 | 26469640 | 3793251 | 33375145
239 | 2023-01-27 17:01 | WO-23003-150-A-1 | 306 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.429 | 0.324 | #N/A | 0.351 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | [10 empty] | 31830 | 226839 | 344916 | 2039599 | 1624270 | 17346607 | 27896754 | 3851752 | 37504622
240 | 2023-01-27 17:28 | LRB-16-inj1 | 305 | K | QC | 0 | #N/A | 0.811 | 0.741 | #N/A | 0.516 | 0.273 | #N/A | 0.375 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 356839 | 785930 | 1105469 | 6366458 | 3153054 | 28695723 | 34424377 | 3752352 | 43038656 |  |  | 240
241 | 2023-01-27 17:56 | AB-230117-02 LRB 16 | 304 | K | QC | 0 | #N/A | 0 | 0.749 | #N/A | 0.617 | 0.266 | #N/A | 0.431 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 319065 | 823464 | 1158028 | 6549995 | 2838192 | 26693963 | 31526459 | 3350159 | 38439401 |  |  | 241
242 | 2023-01-27 18:23 | WO-23003-151-A-1 | 303 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.572 | 0.323 | #N/A | 1.037 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | [10 empty] | 64985 | 328737 | 439616 | 2494156 | 1307811 | 13904521 | 23442267 | 3304266 | 34276127
243 | 2023-01-27 18:50 | WO-23003-152-A-1 | 302 | U |  | 0 | #N/A | 1.181 | 0 | #N/A | 0.4 | 0.266 | #N/A | 0.27 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | [9 empty] | 35316 | 235400 | 337422 | 1952820 | 1233668 | 13012060 | 22355921 | 3345514 | 33483385
244 | 2023-01-27 19:17 | WO-23003-153-A-1 | 301 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.36 | 0.347 | #N/A | 0.404 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [8 empty] | 45525 | 256724 | 358758 | 2031301 | 969843 | 9743480 | 15843986 | 2604457 | 25804341
245 | 2023-01-27 19:44 | WO-23003-154-A-1 | 300 | U |  | 0 | #N/A | 1.097 | 0 | #N/A | 0.404 | 1.635 | #N/A | 0.293 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | [10 empty] | 63488 | 310399 | 471730 | 2742671 | 1614797 | 16827970 | 24393253 | 3772794 | 33684806
246 | 2023-01-27 20:11 | WO-23003-155-A-1 | 299 | U |  | 0 | #N/A | 1.192 | 0 | #N/A | 0.326 | 0.374 | #N/A | 0.263 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 64013 | 311653 | 488040 | 2785854 | 1783363 | 17934940 | 26608371 | 3803417 | 36089654
247 | 2023-01-27 20:38 | WO-23003-156-A-1 | 298 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.357 | 0.26 | #N/A | 0.25 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS3 | IS4 | [11 empty] | 82072 | 377701 | 554647 | 3178765 | 2226147 | 19713169 | 24511254 | 3762476 | 31947500
248 | 2023-01-27 21:05 | WO-23003-157-A-1 | 297 | U |  | 0 | #N/A | 0.409 | 0 | #N/A | 0.38 | 0.286 | #N/A | 0.251 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | [9 empty] | 15191 | 165777 | 255885 | 1513836 | 1257699 | 12335038 | 19264744 | 3244749 | 27568727
249 | 2023-01-27 21:32 | WO-23003-158-A-1 | 296 | U |  | 0 | #N/A | 0.863 | 0 | #N/A | 0.308 | 0.267 | #N/A | 1.766 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS3 | IS4 | [11 empty] | 92849 | 386892 | 556363 | 3121432 | 1987225 | 19224864 | 26801827 | 3959551 | 36461695
250 | 2023-01-27 21:59 | WO-23003-159-A-1 | 295 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.293 | 0.325 | #N/A | 0.343 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | [10 empty] | 38069 | 260875 | 410528 | 2368282 | 1670059 | 17583008 | 25763344 | 3888346 | 35695339
251 | 2023-01-27 22:26 | WO-23003-160-A-1 | 294 | U |  | 0 | #N/A | 0 | 1.797 | #N/A | 0.361 | 0.313 | #N/A | 0.487 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 54321 | 282179 | 451646 | 2716039 | 1772186 | 15260217 | 20522506 | 3469802 | 30127225
252 | 2023-01-27 22:53 | LRB-17-inj1 | 293 | K | QC | 0 | #N/A | 0 | 0.786 | #N/A | 0.588 | 0.256 | #N/A | 0.322 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 484708 | 1050190 | 1439985 | 7902759 | 3379366 | 29850546 | 32933569 | 4215275 | 40057113 |  |  | 252
253 | 2023-01-27 23:20 | AB-230117-02 LRB 17 | 292 | K | QC | 0 | #N/A | 0 | 0.889 | #N/A | 0.626 | 0.23 | #N/A | 0.331 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 472207 | 1036892 | 1408963 | 8001582 | 3197912 | 28252225 | 32070431 | 3737957 | 38781650 |  |  | 253
254 | 2023-01-27 23:48 | AB-230117-02 LD 8 | 291 | D | QC | 0 | #N/A | 0 | 0 | #N/A | 0.414 | 0.38 | #N/A | 0.49 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Duplicate |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 41236 | 244540 | 387823 | 2240979 | 1713812 | 16793031 | 23159551 | 3731604 | 31064596 |  |  |  | 254
255 | 2023-01-28 00:15 | AB-230117-02 LFB-H 9 | 290 | Z | QC | 160.935 | #N/A | 183.084 | 181.434 | #N/A | 178.355 | 179.694 | #N/A | 177.434 | 1 | 160.935 | #N/A | 183.084 | 181.434 | #N/A | 178.355 | 179.694 | #N/A | 177.434 | H 9 | H 9 | NA | Blank Spike | No TrueV! |  | IS1 | [13 empty] | 546872 | 1168510 | 1664664 | 8495625 | 2731349 | 24747417 | 26967994 | 3098356 | 32099934 |  |  | 255
256 | 2023-01-28 00:42 | AB-230117-02 CCC-H 9 | 289 | Q | QC | 302.616 | #N/A | 351.796 | 350.18 | #N/A | 337.277 | 341.66 | #N/A | 345.726 | 1 | 302.616 | #N/A | 351.796 | 350.18 | #N/A | 337.277 | 341.66 | #N/A | 345.726 | 2 | 02 | NA | CCC | [16 empty] | 502223 | 997179 | 1384656 | 7007399 | 2169590 | 19631655 | 21703755 | 2417403 | 24904792 |  | 256
257 | 2023-01-28 01:09 | AB-230117-02 LFSM-H 8 | 288 | S | QC | 177.244 | #N/A | 183.28 | 170.633 | #N/A | 178.895 | 185.896 | #N/A | 191.129 | 1 | 177.244 | #N/A | 183.28 | 170.633 | #N/A | 178.895 | 185.896 | #N/A | 191.129 | H 8 | NA | H 8 | Sample Spike | No TrueV! | IS3 | IS3 | IS3 | IS3 | [11 empty] | 67281 | 314744 | 497745 | 2890177 | 1896632 | 15740897 | 18646299 | 2939093 | 21881749 |  |  |  |  | 257
258 | 2023-01-28 02:03 | WO-23003-161-A-1 | 287 | U |  | 0 | #N/A | 0.794 | 0 | #N/A | 0.552 | 0.561 | #N/A | 0.674 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 69693 | 320634 | 541326 | 3033931 | 2091007 | 18538399 | 23473716 | 3656138 | 30116754
259 | 2023-01-28 02:30 | WO-23003-162-A-1 | 286 | U |  | 0 | #N/A | 0.801 | 0 | #N/A | 0.454 | 0.462 | #N/A | 0.632 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 53537 | 295461 | 465424 | 2617183 | 1993288 | 18864513 | 24824767 | 3892407 | 32557192
260 | 2023-01-28 02:57 | WO-23003-163-A-1 | 285 | U |  | 0 | #N/A | 1.814 | 0 | #N/A | 0.449 | 0.346 | #N/A | 0.298 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | [9 empty] | 16334 | 168104 | 245372 | 1435424 | 1319432 | 13691119 | 21426804 | 3292263 | 30799128
261 | 2023-01-28 03:24 | WO-23003-164-A-1 | 284 | U |  | 0 | #N/A | 0.989 | 0 | #N/A | 0.485 | 0.339 | #N/A | 0.303 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | [10 empty] | 33077 | 222523 | 361499 | 2117720 | 1648049 | 15546125 | 20805999 | 3467978 | 28647728
262 | 2023-01-28 03:51 | WO-23003-165-A-1 | 283 | U |  | 0 | #N/A | 1.082 | 0.789 | #N/A | 0.724 | 0.482 | #N/A | 0.497 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 69809 | 284546 | 520027 | 2977046 | 2077889 | 18802474 | 23474981 | 3702482 | 31033433
263 | 2023-01-28 04:18 | WO-23003-166-A-1 | 282 | U |  | 0 | #N/A | 0 | 1.19 | #N/A | 0.392 | 0.504 | #N/A | 0.628 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | [13 empty] | 136720 | 362297 | 681771 | 3810986 | 2539693 | 21890359 | 25407205 | 3901813 | 31878795
264 | 2023-01-28 04:45 | WO-23003-167-A-1 | 281 | U |  | 0 | #N/A | 0 | 1.668 | #N/A | 0.488 | 0.566 | #N/A | 0.71 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 66326 | 283386 | 515902 | 2920180 | 2079503 | 17747007 | 22371425 | 3630087 | 29415023
265 | 2023-01-28 05:12 | WO-23003-168-A-1 | 280 | U |  | 0 | #N/A | 0.745 | 0.889 | #N/A | 0.503 | 0.536 | #N/A | 0.645 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | [13 empty] | 107661 | 338226 | 612076 | 3560271 | 2389505 | 19229820 | 22941001 | 3746978 | 30011092
266 | 2023-01-28 05:40 | WO-23003-169-A-1 | 279 | U |  | 0 | #N/A | 0.498 | 0 | #N/A | 0.365 | 0.271 | #N/A | 0.257 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 12005 | 151875 | 230313 | 1260188 | 1289295 | 13380648 | 22045140 | 3482955 | 31308682
267 | 2023-01-28 06:07 | WO-23003-170-A-1 | 278 | U |  | 0 | #N/A | 0.735 | 0 | #N/A | 0.311 | 0.245 | #N/A | 0.221 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 48476 | 321024 | 441216 | 2560727 | 1918476 | 18979309 | 25869636 | 3701497 | 35655967
268 | 2023-01-28 06:34 | LRB-18-inj1 | 277 | K | QC | 0 | #N/A | 0 | 0.716 | #N/A | 0.512 | 0.231 | #N/A | 0.287 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 465299 | 999773 | 1407079 | 7666312 | 3231707 | 28983106 | 31732010 | 3593951 | 40339726 |  |  | 268
269 | 2023-01-28 07:01 | AB-230117-02 LRB 18 | 276 | K | QC | 0 | #N/A | 0 | 0.712 | #N/A | 0.506 | 0.239 | #N/A | 0.309 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 480642 | 1016943 | 1380572 | 7600592 | 3250951 | 27888958 | 30466100 | 3514841 | 38512113 |  |  | 269
270 | 2023-01-28 07:28 | WO-23003-171-A-1 | 275 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.386 | 0.273 | #N/A | 0.273 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS4 | IS4 | [11 empty] | 97618 | 417810 | 558713 | 3174860 | 2132079 | 20015689 | 24823141 | 3679650 | 32821029
271 | 2023-01-28 07:55 | WO-23003-172-A-1 | 274 | U |  | 0 | #N/A | 0.751 | 0 | #N/A | 0.17 | 0.29 | #N/A | 0.282 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS4 | IS4 | [11 empty] | 61777 | 372477 | 493198 | 2828001 | 2023323 | 19258794 | 25363940 | 3705559 | 33821093
272 | 2023-01-28 08:22 | WO-23003-173-A-1 | 273 | U |  | 0 | #N/A | 0.325 | 0 | #N/A | 0.354 | 0.305 | #N/A | 0.267 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 14126 | 167549 | 258711 | 1406071 | 1318094 | 13489980 | 21850161 | 3306529 | 33014680
273 | 2023-01-28 08:49 | WO-23003-174-A-1 | 272 | U |  | 0 | #N/A | 0.845 | 0 | #N/A | 0.338 | 0.271 | #N/A | 0.229 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 59001 | 351613 | 468510 | 2663445 | 1943904 | 18862711 | 25850470 | 3765508 | 35919221
274 | 2023-01-28 09:16 | WO-23003-175-A-1 | 271 | U |  | 0 | #N/A | 0.341 | 0 | #N/A | 0.339 | 0.286 | #N/A | 0.19 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 74350 | 369722 | 511880 | 2959853 | 1999689 | 19082738 | 25711792 | 3626223 | 35315191
275 | 2023-01-28 09:43 | WO-23003-176-A-1 | 270 | U |  | 0 | #N/A | 1.229 | 0 | #N/A | 0.374 | 0.249 | #N/A | 0.183 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | IS4 | [10 empty] | 62970 | 334777 | 475018 | 2623095 | 1633243 | 16601625 | 24917770 | 3523101 | 34980436
276 | 2023-01-28 10:10 | WO-23003-177-A-1 | 269 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.378 | 0.285 | #N/A | 0.246 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 47934 | 304495 | 412953 | 2401343 | 1708034 | 16951043 | 24435144 | 3653857 | 35158130
277 | 2023-01-28 10:38 | WO-23003-178-A-1 | 268 | U |  | 0 | #N/A | 0.198 | 0 | #N/A | 0.166 | 0.258 | #N/A | 0.206 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 49961 | 287900 | 378797 | 2095557 | 1080273 | 10682658 | 17221198 | 2808552 | 26034693
278 | 2023-01-28 11:05 | WO-23003-179-A-1 | 267 | U |  | 0 | #N/A | 0.544 | 0 | #N/A | 0.376 | 0 | #N/A | 0.249 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 38951 | 235710 | 306279 | 1727550 | 797813 | 7635678 | 11761625 | 1761842 | 19752521
279 | 2023-01-28 11:32 | WO-23003-180-A-1 | 266 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.341 | 0.282 | #N/A | 0.228 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [9 empty] | 41414 | 266808 | 351343 | 2117294 | 1270011 | 12516657 | 18800561 | 3073816 | 26608435
280 | 2023-01-28 11:59 | LRB-19-inj1 | 265 | K | QC | 0 | #N/A | 0.179 | 0.906 | #N/A | 0.519 | 0.249 | #N/A | 0.281 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 462613 | 975650 | 1333367 | 7363776 | 3402141 | 28767903 | 32008275 | 3416155 | 40849895 |  |  | 280
281 | 2023-01-28 12:26 | AB-230117-02 LRB 19 | 264 | K | QC | 0 | #N/A | 0 | 0.827 | #N/A | 0.512 | 0.277 | #N/A | 0.287 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 475072 | 979270 | 1364602 | 7657476 | 3210571 | 27607743 | 31221498 | 3510414 | 39049072 |  |  | 281
282 | 2023-01-28 12:53 | AB-230117-02 LD 9 | 263 | D | QC | 0 | #N/A | 0.632 | 0 | #N/A | 0.32 | 0.269 | #N/A | 0.236 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Duplicate |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 49183 | 317353 | 423819 | 2457924 | 1834832 | 18049984 | 25240310 | 3726813 | 33893145 |  |  |  | 282
283 | 2023-01-28 13:20 | AB-230117-02 LFB-H 10 | 262 | Z | QC | 151.291 | #N/A | 180.078 | 170.015 | #N/A | 171.506 | 172.512 | #N/A | 173.244 | 1 | 151.291 | #N/A | 180.078 | 170.015 | #N/A | 171.506 | 172.512 | #N/A | 173.244 | H 10 | H 10 | NA | Blank Spike | No TrueV! | [15 empty] | 474655 | 985252 | 1319412 | 7015382 | 2433931 | 22259117 | 24459732 | 2675897 | 29407136 |  |  | 283
284 | 2023-01-28 13:47 | AB-230117-02 CCC-H 10 | 261 | Q | QC | 301.419 | #N/A | 356.398 | 350.601 | #N/A | 341.141 | 348.019 | #N/A | 342.023 | 1 | 301.419 | #N/A | 356.398 | 350.601 | #N/A | 341.141 | 348.019 | #N/A | 342.023 | 2 | 02 | NA | CCC | [16 empty] | 493383 | 952125 | 1279916 | 6632334 | 2100146 | 18727431 | 20658508 | 2348444 | 23743526 |  | 284
285 | 2023-01-28 14:14 | AB-230117-02 LFSM-H 9 | 260 | S | QC | 168.758 | #N/A | 184.196 | 179.695 | #N/A | 182.596 | 190.975 | #N/A | 189.486 | 1 | 168.758 | #N/A | 184.196 | 179.695 | #N/A | 182.596 | 190.975 | #N/A | 189.486 | H 9 | NA | H 9 | Sample Spike | No TrueV! | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | [9 empty] | 14244 | 165991 | 242322 | 1389125 | 1255344 | 12329974 | 18646924 | 2649459 | 25494037 |  |  |  |  | 285
286 | 2023-01-28 15:08 | WO-23003-181-A-1 | 259 | U |  | 0 | #N/A | 0.867 | 0 | #N/A | 0.611 | 0.524 | #N/A | 0.526 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS4 | IS4 | [11 empty] | 79367 | 387068 | 520554 | 2959952 | 2055317 | 19786204 | 26881852 | 3623051 | 35900244
287 | 2023-01-28 15:35 | WO-23003-182-A-1 | 258 | U |  | 0 | #N/A | 0.34 | 0 | #N/A | 0.464 | 0.391 | #N/A | 0.316 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS4 | IS4 | [11 empty] | 73792 | 376204 | 498288 | 2889862 | 2053163 | 19655372 | 25915073 | 3699468 | 35678436
288 | 2023-01-28 16:03 | WO-23003-183-A-1 | 257 | U |  | 0 | #N/A | 0.59 | 0 | #N/A | 0.39 | 0.295 | #N/A | 0.325 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS3 | IS4 | [11 empty] | 82736 | 379271 | 520642 | 2915834 | 1920242 | 19124889 | 26049925 | 3641152 | 36206525
289 | 2023-01-28 16:30 | WO-23003-184-A-1 | 256 | U |  | 0 | #N/A | 0.63 | 0 | #N/A | 0.413 | 0.302 | #N/A | 0.249 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS3 | IS4 | [11 empty] | 76146 | 403204 | 537967 | 3094685 | 2344458 | 21936335 | 27768915 | 3691711 | 35514802
290 | 2023-01-28 16:57 | WO-23003-185-A-1 | 255 | U |  | 0 | #N/A | 0.51 | 0.431 | #N/A | 0.377 | 0.368 | #N/A | 0.241 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS3 | IS4 | [11 empty] | 77847 | 380600 | 506374 | 2904735 | 1990488 | 19844099 | 26407794 | 3640898 | 35863933
291 | 2023-01-28 17:24 | WO-23003-186-A-1 | 254 | U |  | 0 | #N/A | 0.239 | 0 | #N/A | 0.37 | 0.258 | #N/A | 0.224 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 74934 | 361916 | 529083 | 2949424 | 2032877 | 19807118 | 25912912 | 3634040 | 35994281
292 | 2023-01-28 17:51 | WO-23003-187-A-1 | 253 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.27 | 0.235 | #N/A | 0.23 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 67936 | 355746 | 474748 | 2762067 | 1924367 | 18966931 | 26285689 | 3637597 | 36518213
293 | 2023-01-28 18:18 | WO-23003-188-A-1 | 252 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.342 | 0.293 | #N/A | 0.203 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 69449 | 351928 | 484668 | 2742557 | 1994692 | 19244736 | 26529526 | 3607841 | 36804811
294 | 2023-01-28 18:45 | WO-23003-189-A-1 | 251 | U |  | 0 | #N/A | 0.778 | 1.662 | #N/A | 0.995 | 0.474 | #N/A | 0.38 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS3 | IS4 | [11 empty] | 85580 | 382234 | 512194 | 2905790 | 2040648 | 18794074 | 21142994 | 2918972 | 34075321
295 | 2023-01-28 19:12 | WO-23003-190-A-1 | 250 | U |  | 0 | #N/A | 0 | 1.55 | #N/A | 1.102 | 0.329 | #N/A | 0.385 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | [10 empty] | 39881 | 293861 | 394071 | 2251107 | 1684141 | 17055709 | 21148930 | 2896586 | 33587846
296 | 2023-01-28 19:39 | LRB-20-inj1 | 249 | K | QC | 0 | #N/A | 0 | 0.763 | #N/A | 0.521 | 0.244 | #N/A | 0.307 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [8 empty] | b | [7 empty] | 439978 | 957130 | 1357430 | 7502454 | 3272710 | 28880873 | 32546918 | 3544377 | 41119870 |  |  | 296
297 | 2023-01-28 20:06 | AB-230117-02 LRB 20 | 248 | K | QC | 0 | #N/A | 0.168 | 0.841 | #N/A | 0.497 | 0.242 | #N/A | 0.282 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [16 empty] | 437246 | 968936 | 1374949 | 7324266 | 3149172 | 28240850 | 31973298 | 3481649 | 39648533 |  |  | 297
298 | 2023-01-28 20:33 | WO-23003-191-A-1 | 247 | U |  | 0 | #N/A | 1.379 | 1.531 | #N/A | 0.847 | 0.292 | #N/A | 0.168 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS3 | [9 empty] | 15693 | 182248 | 258299 | 1507161 | 1279172 | 13583557 | 19573895 | 2775798 | 31907415
299 | 2023-01-28 21:01 | WO-23003-192-A-1 | 246 | U |  | 0 | #N/A | 0 | 1.124 | #N/A | 0.889 | 0.327 | #N/A | 0.175 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 67594 | 358394 | 491857 | 2784636 | 1938047 | 18266146 | 22132711 | 2910447 | 35628495
300 | 2023-01-28 21:28 | WO-23003-193-A-1 | 245 | U |  | 0 | #N/A | 0.206 | 0.838 | #N/A | 0.633 | 0.258 | #N/A | 0.163 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 64886 | 353054 | 481457 | 2720983 | 1916812 | 18691242 | 22666720 | 2918698 | 36116897
301 | 2023-01-28 21:55 | WO-23003-194-A-1 | 244 | U |  | 0 | #N/A | 1.246 | 0.838 | #N/A | 0.905 | 0.277 | #N/A | 0.209 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS3 | IS4 | [11 empty] | 81652 | 383637 | 509360 | 2968645 | 2008907 | 19273079 | 21045555 | 2755911 | 35050320
302 | 2023-01-28 22:22 | WO-23003-195-A-1 | 243 | U |  | 0 | #N/A | 0.301 | 1.117 | #N/A | 0.858 | 0 | #N/A | 0.201 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 64274 | 345615 | 483473 | 2680389 | 1913407 | 18682395 | 21631474 | 2779978 | 35999627
303 | 2023-01-28 22:49 | WO-23003-196-A-1 | 242 | U |  | 0 | #N/A | 1.277 | 1.056 | #N/A | 0.806 | 0.256 | #N/A | 0.188 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | [10 empty] | 35397 | 250865 | 346963 | 2018624 | 1581333 | 15600340 | 20086012 | 2732450 | 34051812
304 | 2023-01-28 23:16 | WO-23003-197-A-1 | 241 | U |  | 0 | #N/A | 1.222 | 3.823 | #N/A | 13.769 | 0 | #N/A | 0.24 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 13.769 | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 |  | IS3 | IS3 | IS4 | [7 empty] | 68455 | 356339 | 480511 | 2828714 | 1753471 | 11764041 | 6191372 | 818531 | 33394941
305 | 2023-01-28 23:43 | WO-23003-198-A-1 | 240 | U |  | 0 | #N/A | 0.694 | 4.452 | #N/A | 14.717 | 0.435 | #N/A | 0.238 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 14.717 | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | [7 empty] | 22061 | 200546 | 280313 | 1671541 | 1312846 | 9768776 | 5348640 | 821874 | 27215197
306 | 2023-01-29 00:10 | WO-23003-199-A-1 | 239 | U |  | 0 | #N/A | 0.607 | 3.086 | #N/A | 6.684 | 0.239 | #N/A | 0.248 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 6.684 | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | R2 |  |  | This one looks like it has twice the internal standard amount, cutting concentrations in half. I'll rerun it this weekend. |  |  | 103257 | 401948 | 569404 | 3228719 | 2049307 | 16534708 | 10973119 | 1399210 | 33610321
307 | 2023-01-29 00:37 | WO-23003-200-A-1 | 238 | U |  | 0 | #N/A | 0 | 4.582 | #N/A | 13.529 | 0 | #N/A | 0.206 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 13.529 | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 |  |  | IS3 | IS4 | [7 empty] | 75617 | 366120 | 489517 | 2804146 | 1842376 | 14015330 | 7062188 | 896167 | 34422132
308 | 2023-01-29 01:04 | LRB-21-inj1 | 237 | K | QC | 0 | #N/A | 0 | 0.929 | #N/A | 0.653 | 0.233 | #N/A | 0.291 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [8 empty] | b | [7 empty] | 437337 | 904832 | 1252977 | 7134991 | 3249198 | 28451576 | 29645595 | 3086215 | 42486830 |  |  | 308
309 | 2023-01-29 01:31 | AB-230117-02 LRB 21 | 236 | K | QC | 0 | #N/A | 0.064 | 0.722 | #N/A | 0.609 | 0.237 | #N/A | 0.325 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [8 empty] | b | [7 empty] | 354881 | 869335 | 1185440 | 6373522 | 2577534 | 23046320 | 26813713 | 2926896 | 32716061 |  |  | 309
310 | 2023-01-29 01:58 | AB-230117-02 LD 10 | 235 | D | QC | 0 | #N/A | 1.057 | 3.438 | #N/A | 10.794 | 0 | #N/A | 0.273 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 10.794 | <MRL | #N/A | <MRL | NA | NA | NA | Duplicate |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | [7 empty] | 11231 | 153826 | 219738 | 1269400 | 1171678 | 9817354 | 5515946 | 825097 | 26862747 |  |  |  | 309
311 | 2023-01-29 02:26 | AB-230117-02 LFB-H 11 | 234 | Z | QC | 163.487 | #N/A | 195.443 | 186.635 | #N/A | 183.897 | 185.609 | #N/A | 186.414 | 1 | 163.487 | #N/A | 195.443 | 186.635 | #N/A | 183.897 | 185.609 | #N/A | 186.414 | H 11 | H 11 | NA | Blank Spike | No TrueV! | [15 empty] | 409354 | 839595 | 1116067 | 6137861 | 2303348 | 20602763 | 21417114 | 2379547 | 26941243 |  |  | 311
312 | 2023-01-29 02:53 | AB-230117-02 CCC-H 11 | 233 | Q | QC | 309.141 | #N/A | 366.619 | 355.509 | #N/A | 355.982 | 358.493 | #N/A | 360.917 | 1 | 309.141 | #N/A | 366.619 | 355.509 | #N/A | 355.982 | 358.493 | #N/A | 360.917 | 2 | 02 | NA | CCC | [16 empty] | 460334 | 906117 | 1221504 | 6258063 | 2011273 | 17894682 | 19829498 | 2175740 | 22917705 |  | 312
313 | 2023-01-29 03:20 | AB-230117-02 LFSM-H 10 | 232 | S | QC | 194.12 | #N/A | 183.533 | 188.607 | #N/A | 190.703 | 192.309 | #N/A | 194.699 | 1 | 194.12 | #N/A | 183.533 | 188.607 | #N/A | 190.703 | 192.309 | #N/A | 194.699 | H 10 | NA | H 10 | Sample Spike | No TrueV! | IS3 |  | IS3 | IS3 |  | IS3 | IS3 | IS3 | [7 empty] | 88543 | 383532 | 525551 | 2898810 | 1727810 | 12232708 | 5712771 | 831195 | 22408400 |  |  |  |  | 313
314 | 2023-01-29 04:14 | WO-23003-201-A-1 | 231 | U |  | 0 | #N/A | 1.429 | 3.72 | #N/A | 11.225 | 0.809 | #N/A | 0.604 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 11.225 | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 |  | IS3 | IS3 | IS4 | [7 empty] | 65972 | 344747 | 454091 | 2646422 | 1756825 | 12833795 | 6282357 | 824947 | 30882941
315 | 2023-01-29 04:41 | WO-23003-202-A-1 | 230 | U |  | 0 | #N/A | 0 | 4.087 | #N/A | 11.115 | 0 | #N/A | 0.397 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 11.115 | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 |  | IS3 | IS3 | IS4 | [7 empty] | 80081 | 348026 | 481883 | 2742275 | 1808625 | 13140881 | 6111114 | 854523 | 31378239
316 | 2023-01-29 05:08 | WO-23003-203-A-1 | 229 | U |  | 0 | #N/A | 1.134 | 6.695 | #N/A | 39.631 | 0 | #N/A | 0.358 | 1 | <MRL | #N/A | <MRL | 6.695 | #N/A | 39.631 | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS3 | IS3 | IS3 | UL |  | R1 | [5 empty] | 36289 | 247272 | 361261 | 2082824 | 1473104 | 5989206 | 2065247 | 363290 | 24116360
317 | 2023-01-29 05:35 | WO-23003-204-A-1 | 228 | U |  | 0 | #N/A | 0.521 | 5.247 | #N/A | 29.505 | 0 | #N/A | 0.31 | 1 | <MRL | #N/A | <MRL | 5.247 | #N/A | 29.505 | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | IS3 | IS3 | IS3 | UL |  | R1 | [5 empty] | 44987 | 289048 | 400465 | 2327833 | 1553806 | 6360974 | 2250114 | 409960 | 24559170
318 | 2023-01-29 06:02 | WO-23003-206-A-1 | 227 | U |  | 0 | #N/A | 0.479 | 5.258 | #N/A | 35.231 | 0 | #N/A | 0.308 | 1 | <MRL | #N/A | <MRL | 5.258 | #N/A | 35.231 | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  |  |  |  | IS3 | IS3 | UL |  | R1 | [5 empty] | 157626 | 503077 | 674686 | 3682234 | 2217176 | 9117678 | 2571564 | 396341 | 31615078
319 | 2023-01-29 06:29 | WO-23003-207-A-1 | 226 | U |  | 0 | #N/A | 0 | 4.559 | #N/A | 26.636 | 0 | #N/A | 0.258 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 26.636 | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  |  |  |  | IS3 | IS3 | UL |  | R1 | [5 empty] | 133800 | 452265 | 613027 | 3493807 | 2228845 | 10652995 | 3472443 | 526947 | 33979020
320 | 2023-01-29 06:56 | WO-23003-208-A-1 | 225 | U |  | 0 | #N/A | 0.392 | 4.923 | #N/A | 24.868 | 0 | #N/A | 0.26 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 24.868 | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  |  |  |  | IS3 | IS3 | UL |  | R1 | [5 empty] | 137549 | 463464 | 641950 | 3484935 | 2194316 | 11019572 | 3671026 | 534977 | 34058358
321 | 2023-01-29 07:23 | WO-23003-209-A-1 | 224 | U |  | 0 | #N/A | 0.264 | 4.472 | #N/A | 25.391 | 0 | #N/A | 0.31 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 25.391 | <MRL | #N/A | <MRL | NA | NA | NA | Unknown | [6 empty] | IS3 | IS3 | UL |  | R1 | [5 empty] | 189791 | 511993 | 695284 | 3913333 | 2358730 | 11487087 | 3580538 | 490738 | 34154426
322 | 2023-01-29 07:50 | WO-23003-210-A-1 | 223 | U |  | 0 | #N/A | 0.152 | 0.762 | #N/A | 1.497 | 0 | #N/A | 0.289 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown | [7 empty] | IS4 | IS4 | [7 empty] | 417165 | 883799 | 1243565 | 6958152 | 3225027 | 25194154 | 6993795 | 765246 | 37852841
323 | 2023-01-29 08:18 | WO-23003-211-A-1 | 222 | U |  | 0 | #N/A | 0 | 0.916 | #N/A | 1.074 | 0 | #N/A | 0.265 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown | [7 empty] | IS4 | IS4 | [7 empty] | 453256 | 925945 | 1336035 | 7069050 | 3336733 | 26980041 | 7497363 | 851737 | 38808812
324 | 2023-01-29 08:45 | AB-230117-02 LRB 22 | 221 | K | QC | 0 | #N/A | 0.161 | 0.756 | #N/A | 0.574 | 0.186 | #N/A | 0.267 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank | [8 empty] | b | [7 empty] | 462212 | 956316 | 1317309 | 7312818 | 3315056 | 28520880 | 27771783 | 2760987 | 41046510 |  |  | 324
325 | 2023-01-29 09:12 | WO-23003-212-A-1 | 220 | U |  | 0 | #N/A | 0.224 | 0.803 | #N/A | 1.148 | 0 | #N/A | 0.284 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown | [7 empty] | IS4 | IS4 | [7 empty] | 454164 | 937794 | 1295538 | 7037727 | 3311931 | 26368237 | 7639842 | 818293 | 37972771
326 | 2023-01-29 09:39 | AB-230117-02 CCC-H 12 | 199 | Q | QC | 343.096 | #N/A | 391.777 | 383.581 | #N/A | 375.367 | 373.921 | #N/A | 372.448 | 1 | 343.096 | #N/A | 391.777 | 383.581 | #N/A | 375.367 | 373.921 | #N/A | 372.448 | 2 | 02 | NA | CCC | [8 empty] | C |  |  |  |  | Carryover in CCC, not a cal failure. QCS Okay |  |  | 397349 | 796744 | 1095778 | 5657886 | 1940837 | 17254104 | 17173406 | 1870799 | 22029282 |  | 326
327 | 2023-01-29 10:06 | AB-230117-02 QCS 2 | 329 | Q | QC | 194.121 | #N/A | 224.431 | 212.761 | #N/A | 218.131 | 232.69 | #N/A | 220.165 | 1 | 194.121 | #N/A | 224.431 | 212.761 | #N/A | 218.131 | 232.69 | #N/A | 220.165 | 2 | NA | 2 | QCS | [16 empty] | 401060 | 802983 | 1130478 | 6179804 | 2333913 | 20001675 | 22790999 | 2572196 | 29133809 |  | 327
328 | 2023-01-30 11:15 | LRB-22-reprep | 334 | K | QC | 0 | #N/A | 0 | 0.435 | #N/A | 2.334 | 0.439 | #N/A | 0.417 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank |  | IS1 | [6 empty] | b | [7 empty] | 587046 | 1038861 | 1501253 | 7867451 | 2908422 | 26885937 | 28422575 | 3077562 | 35010831 |  |  | 329
329 | 2023-01-30 11:42 | LRB-22-reprep-inj2 | 333 | K | QC | 0 | #N/A | 0.212 | 0.494 | #N/A | 2.24 | 0.435 | #N/A | 0.404 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank |  | IS1 | [6 empty] | b | [7 empty] | 572685 | 1013485 | 1356743 | 7210561 | 2515105 | 23303576 | 25322946 | 2794056 | 30900382 |  |  | 331
330 | 2023-01-30 12:09 | WO-23003-203-A-1B | 332 | U |  | 0 | #N/A | 0.231 | 0.868 | #N/A | 4.491 | 0.326 | #N/A | 0.378 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown | [8 empty] | IS4 | [7 empty] | 500203 | 965643 | 1342376 | 7261125 | 3062842 | 25456947 | 17509722 | 1662963 | 38456781
331 | 2023-01-30 12:36 | WO-23003-204-A-1B | 331 | U |  | 0 | #N/A | 0.313 | 1.758 | #N/A | 9.339 | 0 | #N/A | 0.315 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 9.339 | <MRL | #N/A | <MRL | NA | NA | NA | Unknown | [7 empty] | IS3 | IS4 | [7 empty] | 317187 | 668503 | 918572 | 5203349 | 2797550 | 19003081 | 9018587 | 973200 | 36292475
332 | 2023-01-30 13:03 | WO-23003-205-A-1B | 330 | U |  | 0 | #N/A | 0.35 | 1.411 | #N/A | 7.698 | 0 | #N/A | 0.343 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 7.698 | <MRL | #N/A | <MRL | NA | NA | NA | Unknown | [7 empty] | IS3 | IS4 | [7 empty] | 295207 | 646357 | 865937 | 4933200 | 2759228 | 18925067 | 9179316 | 1005913 | 35260483
333 | 2023-01-30 13:30 | WO-23003-206-A-1B | 328 | U |  | 0 | #N/A | 0.356 | 0.794 | #N/A | 4.89 | 0.362 | #N/A | 0.347 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown | [7 empty] | IS4 | IS4 | [7 empty] | 495795 | 982724 | 1330125 | 7396823 | 3234755 | 25605917 | 14106716 | 1336272 | 38969513
334 | 2023-01-30 13:58 | WO-23003-207-A-1B | 327 | U |  | 0 | #N/A | 0 | 0.815 | #N/A | 3.98 | 0.366 | #N/A | 0.34 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown | [8 empty] | IS4 | [7 empty] | 465880 | 955947 | 1261198 | 7203777 | 3170272 | 26394783 | 17579970 | 1638880 | 39965297
335 | 2023-01-30 14:25 | WO-23003-208-A-1B | 335 | U |  | 0 | #N/A | 0.718 | 0.753 | #N/A | 3.878 | 0.294 | #N/A | 0.316 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown | [8 empty] | IS4 | [7 empty] | 449514 | 936372 | 1264936 | 7138306 | 3246724 | 26469318 | 18251157 | 1709606 | 39806308
336 | 2023-01-30 15:56 | 217288-d4-reprep | 336 | U |  | 0 | #N/A | 0.593 | 0.768 | #N/A | 0.739 | 0.407 | #N/A | 0.577 | 4 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | [14 empty] | 157464 | 393326 | 650768 | 3761331 | 2579520 | 24106931 | 27874542 | 3658609 | 33642603
337 | 2023-01-30 16:23 | 217289-d4-reprep | 348 | U |  | 0 | #N/A | 0 | 0.642 | #N/A | 0.655 | 0.343 | #N/A | 0.368 | 4 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | [14 empty] | 168253 | 421121 | 695262 | 3904277 | 2727214 | 25204572 | 28154739 | 3967468 | 37427258
338 | 2023-01-30 16:50 | 217315-d2-reprep | 354 | U |  | 137.147 | #N/A | 0 | 0 | #N/A | 0.537 | 0.321 | #N/A | 0.228 | 2 | 274.294 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 52733 | 315346 | 398883 | 2492778 | 1848991 | 19174601 | 25701650 | 3730780 | 35013426
339 | 2023-01-30 17:17 | 217356-d2-reprep | 337 | U |  | 0 | #N/A | 0.299 | 0.477 | #N/A | 0.546 | 0.326 | #N/A | 4.268 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 64045 | 362507 | 488862 | 2734706 | 2077732 | 21287780 | 26439273 | 3725669 | 35078474
340 | 2023-01-30 17:44 | 217357-d2-reprep | 349 | U |  | 0 | #N/A | 0.835 | 0 | #N/A | 0.607 | 0.929 | #N/A | 48.327 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 96.654 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 | [11 empty] | 54807 | 337621 | 453180 | 2671781 | 1990548 | 19919468 | 26011235 | 3765206 | 33193268
341 | 2023-01-30 18:12 | 217358-d2-reprep | 355 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.733 | 1.195 | #N/A | 2575.277 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 5150.554 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | IS4 | [6 empty] | 66583 | 358746 | 489764 | 2800127 | 1989970 | 20363549 | 26309173 | 3693069 | 11191331
342 | 2023-01-30 18:39 | 217359-d2-reprep | 338 | U |  | 0 | #N/A | 0 | 0 | #N/A | 1.217 | 3.011 | #N/A | 9903.264 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 19806.528 | NA | NA | NA | Unknown |  | IS3 |  | IS4 | IS4 |  |  |  |  | IS4 | [6 empty] | 65076 | 371808 | 490242 | 2777479 | 1898308 | 19783128 | 27355959 | 3700322 | 5216648
343 | 2023-01-30 19:06 | 217360-d2-reprep | 339 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.642 | 0.345 | #N/A | 378.052 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 756.104 | NA | NA | NA | Unknown |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | C | R1 |  |  | Electively rerun, but this has carrover, original sample did not. Retain original results |  |  | 36267 | 289713 | 381007 | 2278169 | 1758120 | 17853311 | 23347612 | 3663221 | 24984103
344 | 2023-01-30 19:33 | 217393-d4-reprep | 340 | U |  | 0 | #N/A | 0 | 0.775 | #N/A | 3.747 | 2.949 | #N/A | 1061.939 | 4 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4247.756 | NA | NA | NA | Unknown |  | IS3 | [14 empty] | 167116 | 543343 | 714264 | 4115075 | 2761566 | 25726568 | 29222054 | 3815020 | 18289200
345 | 2023-01-30 20:00 | 217394-d4-reprep | 352 | U |  | 0 | #N/A | 0.511 | 1.218 | #N/A | 4.579 | 4.034 | #N/A | 1414.578 | 4 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 5658.312 | NA | NA | NA | Unknown |  | IS3 | [14 empty] | 164144 | 526280 | 686028 | 3908243 | 2660302 | 26064711 | 30325157 | 3844497 | 16523189
346 | 2023-01-30 21:21 | 217395-d4-reprep | 341 | U |  | 0 | #N/A | 0.132 | 1.534 | #N/A | 9.032 | 8.354 | #N/A | 2714.735 | 4 | <MRL | #N/A | <MRL | <MRL | #N/A | 36.128 | 33.416 | #N/A | 10858.94 | NA | NA | NA | Unknown |  | IS3 | [7 empty] | IS4 | [6 empty] | 167429 | 528791 | 707509 | 4108574 | 2784250 | 26148752 | 30556450 | 3838100 | 11213632
347 | 2023-01-30 22:15 | 217399-d2-reprep | 358 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.679 | 0.312 | #N/A | 4.311 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 60835 | 368700 | 445606 | 2773002 | 1942365 | 18007646 | 21440348 | 3282239 | 27407650
348 | 2023-01-30 22:42 | 217400-d2-reprep | 342 | U |  | 0 | #N/A | 0 | 0 | #N/A | 0.707 | 0.321 | #N/A | 1.945 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Unknown |  | IS3 |  | IS3 | IS4 | [11 empty] | 57369 | 387634 | 543005 | 3286410 | 2705573 | 23929613 | 24811018 | 3662110 | 31039101
349 | 2023-01-30 23:09 | 217295-d1000-LD-reprep | 350 | D | QC | 0 | #N/A | 0.177 | 0.308 | #N/A | 0.299 | 0.28 | #N/A | 1.087 | 1000 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Duplicate |  | IS1 | IS2 | [13 empty] | 641965 | 1318561 | 1798305 | 9513170 | 3518786 | 30393908 | 21408323 | 1986992 | 42625232 |  |  |  | 351
350 | 2023-01-30 23:36 | 217335-d2-LD-reprep | 356 | D | QC | 430.994 | #N/A | 0 | 0 | #N/A | 0.574 | 0.328 | #N/A | 1.008 | 2 | 861.988 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Duplicate |  | IS3 | IS3 | IS4 | IS4 | [11 empty] | 39079 | 305429 | 398711 | 2367644 | 1888781 | 19509902 | 25665879 | 3588274 | 33019176 |  |  |  | 352
351 | 2023-01-31 00:04 | 217375-d2-LD-reprep | 343 | D | QC | 0 | #N/A | 0 | 0.931 | #N/A | 0.848 | 0.367 | #N/A | 28.506 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 57.012 | NA | NA | NA | Duplicate |  | IS3 | IS4 | IS4 | IS4 | IS4 | [10 empty] | 32562 | 252033 | 353343 | 2050127 | 1396354 | 15040827 | 21663026 | 3283182 | 30356496 |  |  |  | 353
352 | 2023-01-31 00:31 | 217395-d4-LD-reprep | 344 | D | QC | 0 | #N/A | 0 | 1.523 | #N/A | 8.776 | 8.406 | #N/A | 2607.572 | 4 | <MRL | #N/A | <MRL | <MRL | #N/A | 35.104 | 33.624 | #N/A | 10430.288 | NA | NA | NA | Duplicate |  | IS3 | [7 empty] | IS4 | [6 empty] | 162917 | 546714 | 686795 | 3962948 | 2715983 | 26488218 | 30315119 | 3864173 | 11471986 |  |  |  | 354
353 | 2023-01-31 01:25 | 217359-d2-LD | 351 | D | QC | 0 | #N/A | 0 | 0 | #N/A | 0.968 | 2.774 | #N/A | 8748.882 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 17497.764 | NA | NA | NA | Duplicate |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | IS4 | [6 empty] | 52260 | 310530 | 423000 | 2495623 | 1734908 | 17329609 | 23961101 | 3248796 | 5267128 |  |  |  | 355
354 | 2023-01-31 02:19 | 217415-d2-LD-reprep | 345 | D | QC | 0 | #N/A | 0.57 | 0 | #N/A | 0.48 | 0.339 | #N/A | 10.557 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 21.114 | NA | NA | NA | Duplicate |  | IS3 | IS4 | IS3 | IS4 | [11 empty] | 37833 | 291605 | 374812 | 2263857 | 1720166 | 16258879 | 20903479 | 3296585 | 26651054 |  |  |  | 356
355 | 2023-01-31 02:46 | 217296-d1000-LFSM-200-reprep | 357 | S | QC | 166.141 | #N/A | 199.356 | 189.823 | #N/A | 180.753 | 184.164 | #N/A | 190.796 | 1000 | 166141 | #N/A | 199356 | 189823 | #N/A | 180753 | 184164 | #N/A | 190796 | 200 | NA | 200 | Sample Spike |  | IS1 | IS1 | [5 empty] | IS3 | [7 empty] | 556150 | 1113075 | 1449487 | 7752940 | 2384164 | 21920330 | 16594272 | 1563970 | 28143191 |  |  |  |  | 357
356 | 2023-01-31 03:13 | 217336-d2-LFSM-200-reprep | 346 | S | QC | 606.641 | #N/A | 185.585 | 184.866 | #N/A | 175.24 | 187.674 | #N/A | 194.566 | 2 | 1213.282 | #N/A | 371.17 | 369.732 | #N/A | 350.48 | 375.348 | #N/A | 389.132 | 200 | NA | 200 | Sample Spike |  | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | [9 empty] | 9796 | 135649 | 201239 | 1164473 | 1150272 | 11821768 | 16744072 | 2616570 | 21729644 |  |  |  |  | 358
357 | 2023-01-31 03:40 | 217376-d2-LFSM-200-reprep | 347 | S | QC | 196.844 | #N/A | 173.032 | 173.811 | #N/A | 166.325 | 177.327 | #N/A | 266.746 | 2 | 393.688 | #N/A | 346.064 | 347.622 | #N/A | 332.65 | 354.654 | #N/A | 533.492 | 200 | NA | 200 | Sample Spike |  | IS3 |  | IS3 | IS3 | [11 empty] | 47681 | 384978 | 513604 | 2947540 | 2245190 | 20267034 | 22022785 | 3180995 | 25146296 |  |  |  |  | 359
358 | 2023-01-31 04:07 | 217396-d200-LFSM-200 | 353 | S | QC | 153.37 | #N/A | 175.395 | 175.651 | #N/A | 168.361 | 178.6 | #N/A | 4040.399 | 200 | 30674 | #N/A | 35079 | 35130.2 | #N/A | 33672.2 | 35720 | #N/A | 808079.8 | 200 | NA | 200 | Sample Spike |  | IS1 | IS1 | [6 empty] | IS3 | [6 empty] | 574760 | 1163052 | 1554299 | 8146487 | 2649285 | 23886278 | 25767847 | 2716368 | 8403033 |  |  |  |  | 360
359 | 2023-01-31 04:34 | 217360-d2-LFSM-200 | 359 | S | QC | 182.595 | #N/A | 187.722 | 186.158 | #N/A | 185.148 | 185.543 | #N/A | 384.137 | 2 | 365.19 | #N/A | 375.444 | 372.316 | #N/A | 370.296 | 371.086 | #N/A | 768.274 | 200 | NA | 200 | Sample Spike |  | IS3 | IS3 | IS3 | IS3 | IS3 |  |  |  | C | [6 empty] | 41412 | 299353 | 387900 | 2189551 | 1637116 | 15306684 | 20095160 | 2838369 | 21867262 |  |  |  |  | 361
360 | 2023-01-31 05:01 | 217416-d2-LFSM-200-reprep | 364 | S | QC | 199.788 | #N/A | 183.534 | 184.672 | #N/A | 186.591 | 187.492 | #N/A | 195.064 | 2 | 399.576 | #N/A | 367.068 | 369.344 | #N/A | 373.182 | 374.984 | #N/A | 390.128 | 200 | NA | 200 | Sample Spike |  | IS3 | IS3 | IS3 | IS3 |  | IS3 | IS3 | [8 empty] | 38456 | 301510 | 389577 | 2212125 | 1743168 | 13558401 | 15469898 | 2693800 | 20335293 |  |  |  |  | 362
361 | 2023-01-30 11:15 | LRB-22-reprep | 363 | K | QC | 0 | #N/A | 0 | 0.435 | #N/A | 2.334 | 0.439 | #N/A | 0.417 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank |  | IS1 | IS2 |  |  | b | b | b | b | b | [6 empty] | 587046 | 1038861 | 1501253 | 7867451 | 2908422 | 26885937 | 28422575 | 3077562 | 35010831 |  |  | 329
362 | 2023-01-30 11:42 | LRB-22-reprep-inj2 | 362 | K | QC | 0 | #N/A | 0.212 | 0.494 | #N/A | 2.24 | 0.435 | #N/A | 0.404 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | NA | NA | NA | Blank |  | IS1 | IS2 | [13 empty] | 572685 | 1013485 | 1356743 | 7210561 | 2515105 | 23303576 | 25322946 | 2794056 | 30900382 |  |  | 331
363 | 2023-01-31 06:23 | QCS-2-Reinject | 361 | Q | QC | 174.903 | #N/A | 211.41 | 200.785 | #N/A | 189.964 | 202.109 | #N/A | 205.073 | 1 | 174.903 | #N/A | 211.41 | 200.785 | #N/A | 189.964 | 202.109 | #N/A | 205.073 | 2 | NA | 2 | QCS |  | IS1 | [14 empty] | 579051 | 1055790 | 1405041 | 7455815 | 2267959 | 21990461 | 23190728 | 2463527 | 26824835 |  | 363
364 | 2023-01-31 06:50 | LFB-H-11-reinject | 360 | Z | QC | 194.643 | #N/A | 225.808 | 219.041 | #N/A | 205.408 | 224.483 | #N/A | 213.202 | 1 | 194.643 | #N/A | 225.808 | 219.041 | #N/A | 205.408 | 224.483 | #N/A | 213.202 | H | H | NA | Blank Spike | No TrueV! | IS1 | [14 empty] | 590402 | 1096484 | 1449455 | 7447833 | 2263951 | 22092585 | 23966959 | 2412550 | 28409833 |  |  | 364
365 | 2023-01-31 07:17 | CCC-H-12-reinject | 93 | Q | QC | 305.952 | #N/A | 364.623 | 349.782 | #N/A | 351.819 | 354.686 | #N/A | 353.248 | 1 | 305.952 | #N/A | 364.623 | 349.782 | #N/A | 351.819 | 354.686 | #N/A | 353.248 | #VALUE! | #VALUE! | NA | CCC | No TrueV! | [15 empty] | 382165 | 865786 | 1120831 | 5985086 | 1929331 | 17713375 | 20159351 | 2071302 | 23410432 |  | 365
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA
[26 empty] | NA | NA

## QAQC Overflow
LRB-11 | K | 1 | 0 | 0 | 0.345 | 0.756 | 0.727 | 0.55 | 0.267 | 0.017 | 0.264 | NA
LFB-200-6 | Z | 1 | 158.298 | 149.25 | 179.368 | 175.354 | 148.382 | 173.631 | 176.742 | 151.608 | 176.054 | 200 | 79.15% | 74.62% | 89.51% | 87.3% | 73.83% | 86.54% | 88.24% | 75.8% | 87.89%
LRB-17 | K | QC | 0 | 0 | 0 | 0.786 | 0.767 | 0.588 | 0.256 | -0.151 | 0.322 | NA
LRB-17-inj2 | K | QC | 0 | 0 | 0 | 0.889 | 0.833 | 0.626 | 0.23 | -0.144 | 0.331 | NA
LFB-200-9 | Z | QC | 160.935 | 171.888 | 183.084 | 181.434 | 155.518 | 178.355 | 179.694 | 152.895 | 177.434 | 200 | 80.47% | 85.94% | 91.54% | 90.72% | 77.76% | 89.18% | 89.85% | 76.45% | 88.72%
LRB-19 | K | QC | 0 | 0 | 0.179 | 0.906 | 0.723 | 0.519 | 0.249 | 0.66 | 0.281 | NA
LRB-19-inj2 | K | QC | 0 | 0 | 0 | 0.827 | 0.753 | 0.512 | 0.277 | -0.2 | 0.287 | NA
LFB-200-10 | Z | QC | 151.291 | 158.45 | 180.078 | 170.015 | 149.734 | 171.506 | 172.512 | 146.83 | 173.244 | 200 | 75.65% | 79.22% | 90.04% | 85.01% | 74.87% | 85.75% | 86.26% | 73.42% | 86.62%
LRB-20 | K | QC | 0 | 0 | 0 | 0.763 | 0.972 | 0.521 | 0.244 | 12.246 | 0.307 | NA | [9 empty] | Carryover
LRB-20-inj2 | K | QC | 0 | 0 | 0.168 | 0.841 | 0.761 | 0.497 | 0.242 | 0.331 | 0.282 | NA
LRB-21 | K | QC | 0 | 0 | 0 | 0.929 | 2.409 | 0.653 | 0.233 | 238.564 | 0.291 | NA | [9 empty] | Carryover
LRB-21-inj2 | K | QC | 0 | 1.048 | 0.064 | 0.722 | 0.868 | 0.609 | 0.237 | 8.959 | 0.325 | NA | [9 empty] | Carryover
LFB-200-11 | Z | QC | 163.487 | 176.561 | 195.443 | 186.635 | 164.505 | 183.897 | 185.609 | 441.806 | 186.414 | 200 | 81.74% | 88.28% | 97.72% | 93.32% | 82.25% | 91.95% | 92.8% | 220.9% | 93.21% | Carryover. Not an LFB failure
LRB-22 | K | QC | 0 | 0 | 0.161 | 0.756 | 1.076 | 0.574 | 0.186 | 429.568 | 0.267 | NA | [9 empty] | Carryover
LRB-22-reprep | K | QC | 0 | 0 | 0 | 0.435 | 0.85 | 2.334 | 0.439 | 3.365 | 0.417 | NA | [9 empty] | Carryover
LRB-22-reprep-inj2 | K | QC | 0 | 0 | 0.212 | 0.494 | 0.799 | 2.24 | 0.435 | 3.082 | 0.404 | NA | [9 empty] | Carryover
LRB-22-postrun | K | QC | 0 | 0 | 0 | 0.383 | 3.048 | 3.603 | 5.822 | 6.458 | 9.672 | NA | [9 empty] | Carryover
LRB-22-inj2-postrun | K | QC | 0 | 0 | 0.24 | 0.315 | 0.899 | 0.764 | 0.537 | 0.384 | 1.636 | NA
LFB-200-11-postrun | Z | QC | 194.643 | 199.836 | 225.808 | 219.041 | 195.122 | 205.408 | 224.483 | 194.608 | 213.202 | 200 | 97.32% | 99.92% | 112.9% | 109.5% | 97.56% | 102.7% | 112.2% | 97.3% | 106.6%
217295-d1000 | U |  | 0 | 0 | 1.109 | 1.055 | 2.776 | 1.898 | 2.538 | 1416.32 | 3.141
217295-d1000-reprep-LD | D | 1000 | 0 | 0 | 0.177 | 0.308 | 1.598 | 0.299 | 0.28 | 1222.979 | 1.087 | NA | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 14.65% | <MRL
217335-d2 | U |  | 433.77 | 2.233 | 0.871 | 0 | 0.75 | 0.35 | 0.431 | 0.071 | 0.246
217335-d2-reprep-LD | D | 2 | 430.994 | 0 | 0 | 0 | 0.725 | 0.574 | 0.328 | 53.394 | 1.008 | NA | 0.642% | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 199.5% | <MRL | Carryover in LD. Not a precision failure
217375-d2 | U |  | 0 | 0 | 0 | 0.912 | 0.928 | 0.618 | 0.649 | 0.365 | 28.256
217375-d2-reprep-LD | D | 2 | 0 | 0 | 0 | 0.931 | 0.772 | 0.848 | 0.367 | 1.407 | 28.506 | NA | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 0.8809%
217395-d4-reprep | U | 4 | 0 | 0 | 0.132 | 1.534 | 0.747 | 9.032 | 8.354 | 0.241 | 2714.735 | NA
217395-d4-reprep-LD | D | 4 | 0 | 0 | 0 | 1.523 | 0.717 | 8.776 | 8.406 | 0.313 | 2607.572 | NA | <MRL | <MRL | <MRL | <MRL | <MRL | 2.875% | 0.6205% | <MRL | 4.027%
217359-d2-reprep | U | 2 | 0 | 0 | 0 | 0 | 0.815 | 1.217 | 3.011 | 0.671 | 9903.264 | NA
217359-d2-reprep-LD | D | 2 | 0 | 0 | 0 | 0 | 0.759 | 0.968 | 2.774 | 0.25 | 8748.882 | NA | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 12.38%
217425-d2 | U |  | 0 | 113.991 | 0 | 0 | 0 | 0.824 | 0.724 | 0.287 | 0.925
217425-d2-reprep-LD | D | 2 | 0 | 121.83 | 0.57 | 0 | 0 | 0.48 | 0.339 | 0.179 | 10.557 | NA | <MRL | 6.648% | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 167.8% | Carryover in LD. Not a precision failure
217296-d1000 | U |  | 0 | 0 | 0 | 0 | 1.996 | 0.665 | 0.601 | 1386.347 | 0.812
217296-d1000-reprep-LFSM-200 | S | 1000 | 166.141 | 171.849 | 199.356 | 189.823 | 164.478 | 180.753 | 184.164 | 1611.097 | 190.796 | 200 | 83.07% | 85.92% | 99.68% | 94.91% | 81.24% | 90.04% | 91.78% | 112.4% | 94.99%
217336-d2 | U |  | 406.54 | 1.572 | 0 | 0 | 0.749 | 0.359 | 0.368 | 0.201 | 0.219
217336-d2-reprep-LFSM-200 | S | 2 | 606.641 | 175.682 | 185.585 | 184.866 | 149.657 | 175.24 | 187.674 | 211.498 | 194.566 | 200 | 100.1% | 87.05% | 92.79% | 92.43% | 74.45% | 87.44% | 93.65% | 105.6% | 97.17%
217376-d2 | U |  | 0 | 0 | 0.788 | 1.032 | 0 | 0.732 | 0.666 | 0.352 | 64.801
217376-d2-reprep-LFSM-200 | S | 2 | 196.844 | 158.711 | 173.032 | 173.811 | 146.373 | 166.325 | 177.327 | 153.34 | 266.746 | 200 | 98.42% | 79.36% | 86.12% | 86.39% | 73.19% | 82.8% | 88.33% | 76.49% | 101%
217396-d200 | U |  | 0 | 0 | 0 | 0.942 | 0.815 | 1.505 | 3.298 | 0.037 | 4052.461
217396-d200-reprep-LFSM-200 | S | 200 | 153.37 | 162.554 | 175.395 | 175.651 | 144.704 | 168.361 | 178.6 | 154.142 | 4040.399 | 200 | 76.69% | 81.28% | 87.7% | 87.35% | 71.94% | 83.43% | 87.65% | 77.05% | -6.031% | Low spike vs ambient PFAS. Not a matrix failure
217360-d2-reprep | U | 2 | 0 | 0 | 0 | 0 | 0.8 | 0.642 | 0.345 | 0.63 | 378.052 | NA
217360-d2-reprep-LFSM-200 | S | 2 | 182.595 | 168.197 | 187.722 | 186.158 | 144.38 | 185.148 | 185.543 | 160.884 | 384.137 | 200 | 91.3% | 84.1% | 93.86% | 93.08% | 71.79% | 92.25% | 92.6% | 80.13% | 3.042% | Carryover in the sample and LD; not a matrix failure but showing need to rinse.
217416-d2 | U |  | 0 | 151.9 | 0 | 1.467 | 0 | 0.453 | 0.486 | 0.024 | 0.92
217416-d2-reprep-LFSM-200 | S | 2 | 199.788 | 312.093 | 183.534 | 184.672 | 152.563 | 186.591 | 187.492 | 161.847 | 195.064 | 200 | 99.89% | 80.1% | 91.77% | 91.6% | 76.28% | 93.07% | 93.5% | 80.91% | 97.07%

## Report Template
Type
U
|  |  |  | Concentration (Raw) | [9 empty] | Concentration | [9 empty] | Flags | [15 empty] | Raw Concentrations, ng/L | [9 empty] | Concentration, ng/L, Corrected for Sample Dilution | [9 empty] | Flags
# | Time Stamp: | Sample ID: | Type | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | Dilution Factor: | C-PFBA | C-PFBS | C-HFPO-DA | C-PFHxA | C-PFHxS | C-PFOA | C-PFNA | C-PFOS | C-PFDA | Flags | F-PFBA | F-PFBS | F-HFPO-DA | F-PFHxA | F-PFHxS | F-PFOA | F-PFNA | F-PFOS | F-PFDA | Re-Analysis Flag | Sample Comments | IS Comments | # | Time Stamp: | Sample ID: | Type | PFBA | PFBS | HFPO-DA | PFHxA | PFHxS | PFOA | PFNA | PFOS | PFDA | Dilution Factor: | C-PFBA | C-PFBS | C-HFPO-DA | C-PFHxA | C-PFHxS | C-PFOA | C-PFNA | C-PFOS | C-PFDA | Flags | F-PFBA | F-PFBS | F-HFPO-DA | F-PFHxA | F-PFHxS | F-PFOA | F-PFNA | F-PFOS | F-PFDA | Re-Analysis Flag | Sample Comments | IS Comments | Final Notes
1 | 2023-01-18 10:32 | AB-230117-02 Cal Blank 1 | M | 0 | #N/A | 0 | 0.566 | #N/A | 2.166 | 2.445 | #N/A | 2.511 | 1 | 0 | #N/A | 0 | 0.566 | #N/A | 2.166 | 2.445 | #N/A | 2.511 | No TrueV! | IS1 | [11 empty] | 20 | 2023-01-18 20:55 | 217275-d2 | U | 429.802 | 11.936 | 0 | 1.148 | 1.272 | 1.127 | 1.291 | 1.714 | 1.464 | 2 | 859.604 | 23.872 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS3 | IS4
2 | 2023-01-18 11:00 | AB-230117-02 Cal Std 1 1 | M | 0 | #N/A | 1.078 | 1.415 | #N/A | 1.437 | 1.368 | #N/A | 1.325 | 1 | 0 | #N/A | 1.078 | 1.415 | #N/A | 1.437 | 1.368 | #N/A | 1.325 | No TrueV! | [12 empty] | 21 | 2023-01-18 21:22 | 217276-d2 | U | 458.755 | 12.332 | 2.86 | 0 | 1.232 | 0.97 | 1.012 | 1.579 | 1.402 | 2 | 917.51 | 24.664 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS3 | IS4
3 | 2023-01-18 11:27 | AB-230117-02 Cal Std 2 1 | M | 0 | #N/A | 2.003 | 2.095 | #N/A | 2.071 | 2.05 | #N/A | 2.016 | 1 | 0 | #N/A | 2.003 | 2.095 | #N/A | 2.071 | 2.05 | #N/A | 2.016 | No TrueV! | [12 empty] | 22 | 2023-01-18 21:49 | 217277-d2 | U | 463.337 | 11.056 | 0 | 0.902 | 1.223 | 0.843 | 0.908 | 1.271 | 1.12 | 2 | 926.674 | 22.112 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3
4 | 2023-01-18 11:54 | AB-230117-02 Cal Std 3 1 | M | 5.762 | #N/A | 4.412 | 3.86 | #N/A | 3.96 | 3.843 | #N/A | 4.043 | 1 | 5.762 | #N/A | 4.412 | 3.86 | #N/A | 3.96 | 3.843 | #N/A | 4.043 | No TrueV! | [10 empty] | Within 50-150% limit |  | 23 | 2023-01-18 22:16 | 217278-d2 | U | 434.867 | 0 | 0 | 0 | 0 | 0.757 | 0.772 | 1.588 | 0.822 | 2 | 869.734 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4
5 | 2023-01-18 12:21 | AB-230117-02 Cal Std 4 1 | M | 11.646 | #N/A | 9.51 | 9.126 | #N/A | 9.194 | 9.265 | #N/A | 9.394 | 1 | 11.646 | #N/A | 9.51 | 9.126 | #N/A | 9.194 | 9.265 | #N/A | 9.394 | No TrueV! | [12 empty] | 24 | 2023-01-18 22:43 | 217279-d2 | U | 511.067 | 0 | 0 | 0 | 0 | 0.675 | 0.744 | 1.348 | 1.093 | 2 | 1022.134 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4
6 | 2023-01-18 12:48 | AB-230117-02 Cal Std 5 1 | M | 18.831 | #N/A | 18.606 | 17.853 | #N/A | 18.158 | 18.63 | #N/A | 18.415 | 1 | 18.831 | #N/A | 18.606 | 17.853 | #N/A | 18.158 | 18.63 | #N/A | 18.415 | No TrueV! | [12 empty] | 25 | 2023-01-18 23:10 | 217280-d2 | U | 462.818 | 0 | 0 | 0 | 0 | 0.657 | 0.771 | 8.176 | 0.769 | 2 | 925.636 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 16.352 | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4 | IS4
7 | 2023-01-18 13:15 | AB-230117-02 Cal Std 6 1 | M | 37.308 | #N/A | 36.592 | 35.14 | #N/A | 36.553 | 37.996 | #N/A | 37.389 | 1 | 37.308 | #N/A | 36.592 | 35.14 | #N/A | 36.553 | 37.996 | #N/A | 37.389 | No TrueV! | [12 empty] | 26 | 2023-01-18 23:37 | 217281-d200 | U | 0 | 0 | 0 | 0 | 0.934 | 1.008 | 1.8 | 0.592 | 3898.943 | 200 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 779788.6 | [9 empty] | IS4
8 | 2023-01-18 13:42 | AB-230117-02 Cal Std 7 1 | M | 95.665 | #N/A | 95.727 | 93.802 | #N/A | 95.968 | 96.468 | #N/A | 97.43 | 1 | 95.665 | #N/A | 95.727 | 93.802 | #N/A | 95.968 | 96.468 | #N/A | 97.43 | No TrueV! | [12 empty] | 27 | 2023-01-19 00:05 | 217282-d200 | U | 0 | 0 | 0 | 0 | 0 | 0.97 | 1.8 | 0.143 | 4309.18 | 200 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 861836 | [9 empty] | IS4
9 | 2023-01-18 14:09 | AB-230117-02 Cal Std 8 1 | M | 163.759 | #N/A | 197.427 | 191.961 | #N/A | 187.35 | 189.259 | #N/A | 193.114 | 1 | 163.759 | #N/A | 197.427 | 191.961 | #N/A | 187.35 | 189.259 | #N/A | 193.114 | No TrueV! | [12 empty] | 28 | 2023-01-19 00:32 | 217283-d200 | U | 0 | 0 | 0 | 0 | 0 | 1.051 | 1.666 | 0.27 | 3931.95 | 200 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 786390 | [9 empty] | IS4
10 | 2023-01-18 14:36 | AB-230117-02 Cal Std 9 1 | M | 328.948 | #N/A | 402.773 | 384.687 | #N/A | 378.816 | 386.336 | #N/A | 385.495 | 1 | 328.948 | #N/A | 402.773 | 384.687 | #N/A | 378.816 | 386.336 | #N/A | 385.495 | No TrueV! | [12 empty] | 29 | 2023-01-19 00:59 | 217284-d200 | U | 0 | 0 | 0 | 0 | 0 | 0.95 | 1.794 | 0.079 | 4432.787 | 200 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 886557.4 | [9 empty] | IS4
11 | 2023-01-18 15:03 | AB-230117-02 Cal Std 10 1 | M | 951.413 | #N/A | 1029.861 | 1001.157 | #N/A | 980.955 | 982.021 | #N/A | 972.285 | 1 | 951.413 | #N/A | 1029.861 | 1001.157 | #N/A | 980.955 | 982.021 | #N/A | 972.285 | No TrueV! | [7 empty] | IS3 | IS3 |  |  |  | 32 | 2023-01-19 02:20 | 217285-d200 | U | 0 | 0 | 0.293 | 0 | 0 | 0.861 | 1.711 | 0.143 | 4146.818 | 200 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 829363.6 | [9 empty] | IS4
12 | 2023-01-18 15:30 | AB-230117-02 Cal Std 11 1 | M | 1822.989 | #N/A | 2030.625 | 1991.195 | #N/A | 1905.749 | 1944.508 | #N/A | 1949.845 | 1 | 1822.989 | #N/A | 2030.625 | 1991.195 | #N/A | 1905.749 | 1944.508 | #N/A | 1949.845 | No TrueV! |  |  |  |  | IS3 | IS3 | IS3 | IS3 | IS3 |  |  |  | 33 | 2023-01-19 02:47 | 217286-d200 | U | 0 | 0 | 0 | 0 | 0 | 0.96 | 1.709 | 0.272 | 4137.805 | 200 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 827561 | [9 empty] | IS4
13 | 2023-01-18 15:57 | AB-230117-02 Cal Std 12 1 | M | 4173.586 | #N/A | 4196.925 | 4184.365 | #N/A | 4022.699 | 3980.463 | #N/A | 4002.829 | 1 | 4173.586 | #N/A | 4196.925 | 4184.365 | #N/A | 4022.699 | 3980.463 | #N/A | 4002.829 | No TrueV! |  |  |  | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  |  |  | 34 | 2023-01-19 03:14 | 217287-d200 | U | 0 | 0 | 0.791 | 0 | 0 | 0.961 | 1.759 | 0.159 | 4255.588 | 200 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 851117.6 | [9 empty] | IS4
14 | 2023-01-18 16:24 | AB-230117-02 Cal Std 13 1 | M | 10164.091 | #N/A | 9751.461 | 9860.345 | #N/A | 10134.089 | 10124.793 | #N/A | 10103.42 | 1 | 10164.091 | #N/A | 9751.461 | 9860.345 | #N/A | 10134.089 | 10124.793 | #N/A | 10103.42 | No TrueV! |  |  | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  |  |  | 35 | 2023-01-19 03:41 | 217288-d4 | U | 0 | 4778.929 | 0 | 0 | 0 | 0.529 | 0.504 | 0.732 | 193.554 | 4 | <MRL | 19115.716 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 774.216 |  | IS3 | IS4 | [6 empty] | C | R1
15 | 2023-01-18 17:46 | AB-230117-02 LRB 1 | K | 0 | #N/A | 1.125 | 0 | #N/A | 3.537 | 4.565 | #N/A | 4.651 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [5 empty] | b | b | b | b | b |  |  |  | 36 | 2023-01-19 04:08 | 217289-d4 | U | 0 | 4665.408 | 0 | 0.817 | 0.969 | 0.502 | 0.488 | 0.966 | 7.75 | 4 | <MRL | 18661.632 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 31 |  | IS3 | IS4 | IS3 | IS4 |  |  |  |  | C | R1
16 | 2023-01-18 18:13 | AB-230117-02 LFB-H 1 | Z | 183.485 | #N/A | 178.384 | 194.206 | #N/A | 198.792 | 204.981 | #N/A | 191.802 | 1 | 183.485 | #N/A | 178.384 | 194.206 | #N/A | 198.792 | 204.981 | #N/A | 191.802 | No TrueV! | [12 empty] | 37 | 2023-01-19 04:35 | 217290-d4 | U | 0 | 4754.423 | 0.678 | 0.767 | 0.955 | 0.465 | 0.505 | 0.884 | 4.083 | 4 | <MRL | 19017.692 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4
17 | 2023-01-18 18:40 | AB-230117-02 QCS 1 | Q | 215.739 | #N/A | 219.599 | 212.607 | #N/A | 211.384 | 210.554 | #N/A | 213.968 | 1 | 215.739 | #N/A | 219.599 | 212.607 | #N/A | 211.384 | 210.554 | #N/A | 213.968 | [13 empty] | 38 | 2023-01-19 05:02 | 217291-d4 | U | 0 | 4685.686 | 0.851 | 0.834 | 0 | 0.579 | 0.469 | 0.981 | 2.358 | 4 | <MRL | 18742.744 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
18 | 2023-01-18 19:07 | AB-230117-02 IBC 1 | Q | 169.148 | #N/A | 189.323 | 190.496 | #N/A | 191.07 | 196.495 | #N/A | 195.489 | 1 | 169.148 | #N/A | 189.323 | 190.496 | #N/A | 191.07 | 196.495 | #N/A | 195.489 | [13 empty] | 39 | 2023-01-19 05:29 | 217292-d4 | U | 0 | 4759.817 | 0.314 | 0.925 | 1.021 | 0.69 | 0.508 | 0.883 | 1.905 | 4 | <MRL | 19039.268 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
19 | 2023-01-18 19:34 | AB-230117-02 CCC-H 1 | Q | 331.373 | #N/A | 392.93 | 384.696 | #N/A | 381.982 | 379.732 | #N/A | 378.211 | 1 | 331.373 | #N/A | 392.93 | 384.696 | #N/A | 381.982 | 379.732 | #N/A | 378.211 | [13 empty] | 40 | 2023-01-19 05:57 | 217293-d4 | U | 0 | 4724.366 | 0.589 | 0 | 0 | 0.681 | 0.569 | 0.885 | 1.605 | 4 | <MRL | 18897.464 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
20 | 2023-01-18 20:55 | WO-23003-01-A-1 @10 | U | 429.802 | #N/A | 0 | 1.148 | #N/A | 1.127 | 1.291 | #N/A | 1.464 | 1 | 429.802 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS3 | IS4 | [8 empty] | 41 | 2023-01-19 06:24 | 217294-d1000 | U | 0 | 0 | 0.342 | 0 | 1.865 | 0.524 | 0.531 | 1470.859 | 1.392 | 1000 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 1470859 | <MRL |  |  |  |  | IS4
21 | 2023-01-18 21:22 | WO-23003-02-A-1 | U | 458.755 | #N/A | 2.86 | 0 | #N/A | 0.97 | 1.012 | #N/A | 1.402 | 1 | 458.755 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS3 | IS4 | [8 empty] | 47 | 2023-01-19 09:33 | 217295-d1000 | U | 0 | 0 | 1.109 | 1.055 | 2.776 | 1.898 | 2.538 | 1416.32 | 3.141 | 1000 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 1416320 | <MRL
22 | 2023-01-18 21:49 | WO-23003-03-A-1 | U | 463.337 | #N/A | 0 | 0.902 | #N/A | 0.843 | 0.908 | #N/A | 1.12 | 1 | 463.337 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | [11 empty] | 48 | 2023-01-19 10:00 | 217296-d1000 | U | 0 | 0 | 0 | 0 | 1.996 | 0.665 | 0.601 | 1386.347 | 0.812 | 1000 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 1386347 | <MRL
23 | 2023-01-18 22:16 | WO-23003-04-A-1 | U | 434.867 | #N/A | 0 | 0 | #N/A | 0.757 | 0.772 | #N/A | 0.822 | 1 | 434.867 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 49 | 2023-01-19 10:27 | 217297-d1000 | U | 0 | 0 | 0 | 0.597 | 1.924 | 0.544 | 0.484 | 1540.035 | 0.718 | 1000 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 1540035 | <MRL
24 | 2023-01-18 22:43 | WO-23003-05-A-1 | U | 511.067 | #N/A | 0 | 0 | #N/A | 0.675 | 0.744 | #N/A | 1.093 | 1 | 511.067 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 50 | 2023-01-19 10:54 | 217298-d1000 | U | 0 | 0 | 0.512 | 0.905 | 2.015 | 0.863 | 0.727 | 1397.807 | 0.913 | 1000 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 1397807 | <MRL
25 | 2023-01-18 23:10 | WO-23003-06-A-1 | U | 462.818 | #N/A | 0 | 0 | #N/A | 0.657 | 0.771 | #N/A | 0.769 | 1 | 462.818 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 51 | 2023-01-19 11:21 | 217299-d1000 | U | 0 | 0 | 0.818 | 0.512 | 1.981 | 0.515 | 0.436 | 1542.714 | 0.578 | 1000 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 1542714 | <MRL
26 | 2023-01-18 23:37 | WO-23003-07-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 1.008 | 1.8 | #N/A | 3898.943 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 3898.943 | [9 empty] | IS4 |  |  |  | 52 | 2023-01-19 11:49 | 217300-d2 | U | 0 | 0 | 1.004 | 0 | 0 | 0.349 | 0.392 | 63.814 | 0.489 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 127.628 | <MRL |  | IS3 | IS4 | IS4 | IS4 |  |  |  | C
27 | 2023-01-19 00:05 | WO-23003-08-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.97 | 1.8 | #N/A | 4309.18 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4309.18 | [9 empty] | IS4 |  |  |  | 53 | 2023-01-19 12:16 | 217301-d2 | U | 0 | 4.255 | 0 | 0 | 0 | 0.424 | 0.363 | 2.775 | 0.461 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 |  | IS4 | IS4
28 | 2023-01-19 00:32 | WO-23003-09-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 1.051 | 1.666 | #N/A | 3931.95 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 3931.95 | [9 empty] | IS4 |  |  |  | 54 | 2023-01-19 12:43 | 217302-d2 | U | 0 | 0 | 0 | 0 | 0 | 0.354 | 0.397 | 1.4 | 0.444 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS4 | IS4
29 | 2023-01-19 00:59 | WO-23003-10-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.95 | 1.794 | #N/A | 4432.787 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4432.787 | [9 empty] | IS4 |  |  |  | 55 | 2023-01-19 13:10 | 217303-d2 | U | 0 | 2.843 | 0 | 0 | 0 | 0.28 | 0.324 | 0.915 | 0.406 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS4 | IS4
30 | 2023-01-19 01:26 | LRB2-Inj1 | K | 0 | #N/A | 0 | 0 | #N/A | 0.46 | 0.414 | #N/A | 204.132 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 204.132 | [9 empty] | b |  |  |  | 56 | 2023-01-19 13:37 | 217304-d2 | U | 0 | 0 | 0 | 0 | 0 | 0.295 | 0.37 | 0.911 | 0.484 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS4 | IS4
31 | 2023-01-19 01:53 | AB-230117-02 LRB 2 | K | 0 | #N/A | 0 | 0.364 | #N/A | 0.466 | 0.349 | #N/A | 7.368 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 7.368 | [9 empty] | b |  | Only detects within 10-fold are carryover and flagged |  | 59 | 2023-01-19 14:58 | 217305-d2 | U | 0 | 0 | 0 | 0 | 0 | 0.702 | 0.548 | 0.994 | 0.955 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4
32 | 2023-01-19 02:20 | WO-23003-11-A-1 | U | 0 | #N/A | 0.293 | 0 | #N/A | 0.861 | 1.711 | #N/A | 4146.818 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4146.818 | [9 empty] | IS4 |  |  |  | 60 | 2023-01-19 15:25 | 217306-d2 | U | 0 | 0 | 2.118 | 0 | 0 | 0.394 | 0.346 | 0.553 | 0.385 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4
33 | 2023-01-19 02:47 | WO-23003-12-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.96 | 1.709 | #N/A | 4137.805 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4137.805 | [9 empty] | IS4 |  |  |  | 61 | 2023-01-19 15:52 | 217307-d2 | U | 0 | 9.357 | 0 | 1.126 | 0 | 0.355 | 0.488 | 0.94 | 0.674 | 2 | <MRL | 18.714 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
34 | 2023-01-19 03:14 | WO-23003-13-A-1 | U | 0 | #N/A | 0.791 | 0 | #N/A | 0.961 | 1.759 | #N/A | 4255.588 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4255.588 | [9 empty] | IS4 |  |  |  | 62 | 2023-01-19 16:19 | 217308-d2 | U | 0 | 9.979 | 0 | 0.991 | 0.94 | 0.484 | 0.521 | 0.906 | 0.663 | 2 | <MRL | 19.958 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
35 | 2023-01-19 03:41 | WO-23003-14-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.529 | 0.504 | #N/A | 193.554 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 193.554 |  | IS3 | IS4 | [6 empty] | C | R1 |  |  | 63 | 2023-01-19 16:46 | 217309-d2 | U | 0 | 8.866 | 0 | 1.091 | 0.908 | 0.491 | 0.514 | 1.126 | 0.71 | 2 | <MRL | 17.732 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
36 | 2023-01-19 04:08 | WO-23003-15-A-1 | U | 0 | #N/A | 0 | 0.817 | #N/A | 0.502 | 0.488 | #N/A | 7.75 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 7.75 |  | IS3 | IS4 | IS3 | IS4 |  |  |  |  | C | R1 |  |  | 64 | 2023-01-19 17:14 | 217310-d2 | U | 0 | 8.682 | 0.99 | 0.885 | 0 | 0.516 | 0.587 | 1.314 | 0.763 | 2 | <MRL | 17.364 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
37 | 2023-01-19 04:35 | WO-23003-16-A-1 | U | 0 | #N/A | 0.678 | 0.767 | #N/A | 0.465 | 0.505 | #N/A | 4.083 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | [10 empty] | 65 | 2023-01-19 17:41 | 217311-d2 | U | 0 | 8.757 | 0 | 0.965 | 0.932 | 0.344 | 0.473 | 1.009 | 0.577 | 2 | <MRL | 17.514 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
38 | 2023-01-19 05:02 | WO-23003-17-A-1 | U | 0 | #N/A | 0.851 | 0.834 | #N/A | 0.579 | 0.469 | #N/A | 2.358 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 66 | 2023-01-19 18:08 | 217312-d2 | U | 0 | 10.034 | 0 | 0.903 | 0 | 0.433 | 0.435 | 1.074 | 0.639 | 2 | <MRL | 20.068 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 |  | IS3 | IS4
39 | 2023-01-19 05:29 | WO-23003-18-A-1 | U | 0 | #N/A | 0.314 | 0.925 | #N/A | 0.69 | 0.508 | #N/A | 1.905 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 67 | 2023-01-19 18:35 | 217313-d2 | U | 126.713 | 0 | 0.513 | 0 | 0 | 0.402 | 0.302 | 0.391 | 0.309 | 2 | 253.426 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4
40 | 2023-01-19 05:57 | WO-23003-19-A-1 | U | 0 | #N/A | 0.589 | 0 | #N/A | 0.681 | 0.569 | #N/A | 1.605 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 68 | 2023-01-19 19:02 | 217314-d2 | U | 134.042 | 0 | 1.336 | 0 | 0.769 | 0.33 | 0.443 | 0.454 | 0.518 | 2 | 268.084 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
41 | 2023-01-19 06:24 | WO-23003-20-A-1 | U | 0 | #N/A | 0.342 | 0 | #N/A | 0.524 | 0.531 | #N/A | 1.392 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  |  |  |  | IS4 | [8 empty] | 76 | 2023-01-19 22:38 | 217315-d2 | U | 150.848 | 2.138 | 0.582 | 0 | 1.08 | 0.665 | 0.764 | 10.144 | 0.677 | 2 | 301.696 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 20.288 | <MRL |  | IS3 | IS3 | IS4 | IS4 |  |  |  | C |  | R1 | Rerun due to carryover
42 | 2023-01-19 06:51 | AB-230117-02 LRB 3 | K | 0 | #N/A | 0 | 0 | #N/A | 0.452 | 0.37 | #N/A | 0.983 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [8 empty] | b |  |  |  |  | 77 | 2023-01-19 23:06 | 217316-d2 | U | 136.634 | 1.32 | 0 | 0 | 0.822 | 0.454 | 0.466 | 3.824 | 0.442 | 2 | 273.268 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
43 | 2023-01-19 07:18 | AB-230117-02 LD 1 | D | 574.418 | #N/A | 0 | 1.066 | #N/A | 0.424 | 0.632 | #N/A | 1.343 | 1 | 574.418 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | [11 empty] | 78 | 2023-01-19 23:33 | 217317-d2 | U | 111.294 | 0 | 0 | 0 | 0 | 0.354 | 0.416 | 2.172 | 0.4 | 2 | 222.588 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
44 | 2023-01-19 07:45 | AB-230117-02 LFB-H 2 | Z | 168.178 | #N/A | 196.167 | 190.739 | #N/A | 188.953 | 194.708 | #N/A | 191.306 | 1 | 168.178 | #N/A | 196.167 | 190.739 | #N/A | 188.953 | 194.708 | #N/A | 191.306 | No TrueV! | [10 empty] | Normal, just highlighted due to LRB carrryover |  | 79 | 2023-01-20 | 217318-d2 | U | 109.959 | 1.786 | 1.537 | 0 | 0.82 | 0.355 | 0.413 | 1.544 | 0.359 | 2 | 219.918 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
45 | 2023-01-19 08:12 | AB-230117-02 CCC-H 2 | Q | 342.723 | #N/A | 403.865 | 379.785 | #N/A | 381.16 | 387.516 | #N/A | 380.815 | 1 | 342.723 | #N/A | 403.865 | 379.785 | #N/A | 381.16 | 387.516 | #N/A | 380.815 | [13 empty] | 80 | 2023-01-20 00:27 | 217319-d2 | U | 126.129 | 1.622 | 0 | 0 | 0.782 | 0.401 | 0.366 | 1.121 | 0.368 | 2 | 252.258 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
46 | 2023-01-19 08:39 | AB-230117-02 LFSM-H 1 @5 | S | 570.426 | #N/A | 197.796 | 197.561 | #N/A | 202.268 | 209.365 | #N/A | 212.68 | 1 | 570.426 | #N/A | 197.796 | 197.561 | #N/A | 202.268 | 209.365 | #N/A | 212.68 | No TrueV! | IS3 | IS3 | IS3 | IS3 | [8 empty] | 81 | 2023-01-20 00:54 | 217320-d2 | U | 117.077 | 1.285 | 0 | 0 | 0.914 | 0.328 | 0.427 | 1.038 | 0.45 | 2 | 234.154 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4
47 | 2023-01-19 09:33 | WO-23003-21-A-1 | U | 0 | #N/A | 1.109 | 1.055 | #N/A | 1.898 | 2.538 | #N/A | 3.141 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 82 | 2023-01-20 01:21 | 217321-d2 | U | 268.234 | 0 | 1.131 | 0 | 0 | 0.429 | 0.355 | 1.061 | 0.318 | 2 | 536.468 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4 | IS4
48 | 2023-01-19 10:00 | WO-23003-22-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.665 | 0.601 | #N/A | 0.812 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 83 | 2023-01-20 01:48 | 217322-d2 | U | 253.554 | 0 | 0.421 | 0 | 0 | 0.491 | 0.363 | 0.744 | 0.334 | 2 | 507.108 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
49 | 2023-01-19 10:27 | WO-23003-23-A-1 | U | 0 | #N/A | 0 | 0.597 | #N/A | 0.544 | 0.484 | #N/A | 0.718 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 84 | 2023-01-20 02:15 | 217323-d2 | U | 250.815 | 1.26 | 0 | 0 | 0 | 0.807 | 0.381 | 0.719 | 0.323 | 2 | 501.63 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
50 | 2023-01-19 10:54 | WO-23003-24-A-1 | U | 0 | #N/A | 0.512 | 0.905 | #N/A | 0.863 | 0.727 | #N/A | 0.913 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 85 | 2023-01-20 02:42 | 217324-d2 | U | 214.775 | 2.099 | 2.007 | 0 | 0 | 0 | 0.388 | 0.744 | 0.264 | 2 | 429.55 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4 | IS4
51 | 2023-01-19 11:21 | WO-23003-25-A-1 | U | 0 | #N/A | 0.818 | 0.512 | #N/A | 0.515 | 0.436 | #N/A | 0.578 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 88 | 2023-01-20 04:04 | 217350-d2 | U | 498.597 | 2.96 | 0 | 0 | 0.784 | 0.375 | 0.324 | 0.737 | 0.295 | 2 | 997.194 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4
52 | 2023-01-19 11:49 | WO-23003-26-A-1 | U | 0 | #N/A | 1.004 | 0 | #N/A | 0.349 | 0.392 | #N/A | 0.489 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 |  |  |  | C |  |  |  |  | 89 | 2023-01-20 04:31 | 217351-d2 | U | 413.609 | 1.383 | 1.442 | 0 | 0 | 0.295 | 0.342 | 0.657 | 0.235 | 2 | 827.218 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
53 | 2023-01-19 12:16 | WO-23003-27-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.424 | 0.363 | #N/A | 0.461 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS4 | IS4 | [8 empty] | 90 | 2023-01-20 04:58 | 217352-d2 | U | 419.335 | 2.725 | 0 | 0 | 0.805 | 0.355 | 0.362 | 0.556 | 0.428 | 2 | 838.67 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
54 | 2023-01-19 12:43 | WO-23003-28-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.354 | 0.397 | #N/A | 0.444 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 91 | 2023-01-20 05:25 | 217353-d2 | U | 446.172 | 9.358 | 0 | 0 | 0.988 | 0.466 | 0.507 | 1.149 | 0.617 | 2 | 892.344 | 18.716 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS3 | IS4
55 | 2023-01-19 13:10 | WO-23003-29-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.28 | 0.324 | #N/A | 0.406 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 92 | 2023-01-20 05:52 | 217354-d2 | U | 456.386 | 7.447 | 0 | 0.584 | 0.901 | 0.523 | 0.541 | 1.348 | 1.323 | 2 | 912.772 | 14.894 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS3 | IS4
56 | 2023-01-19 13:37 | WO-23003-30-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.295 | 0.37 | #N/A | 0.484 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 93 | 2023-01-20 06:19 | 217355-d2 | U | 433.565 | 7.568 | 0 | 0 | 0.968 | 0.523 | 0.563 | 1.232 | 0.576 | 2 | 867.13 | 15.136 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS3 | IS4 | IS4
57 | 2023-01-19 14:04 | LRB4-inj1 | K | 0 | #N/A | 0 | 0.299 | #N/A | 0.368 | 0.282 | #N/A | 0.334 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 94 | 2023-01-20 06:46 | 217356-d2 | U | 0 | 1.166 | 0 | 0 | 0 | 0.337 | 0.365 | 0.509 | 4.897 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS4 | IS4
58 | 2023-01-19 14:31 | AB-230117-02 LRB 4 | K | 0 | #N/A | 0.341 | 0 | #N/A | 0.32 | 0.302 | #N/A | 0.359 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 95 | 2023-01-20 10:01 | 217300-d2-inj2 | U | 0 | 5.584 | 0 | 0 | 0 | 0.288 | 0.417 | 0.606 | 0.466 | 2 | <MRL | 11.168 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 |  |  |  |  | R1 | Analyzed again below to verify if this was 'real', repeated pattern
59 | 2023-01-19 14:58 | WO-23003-31-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.702 | 0.548 | #N/A | 0.955 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 96 | 2023-01-20 10:28 | 217357-d2-inj2 | U | 0 | 1.919 | 0 | 0 | 0.895 | 0.492 | 1.257 | 0.816 | 61.125 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 122.25 |  | IS3 | IS4 | IS4 | IS4 | [5 empty] | R1 | Analyzed again below to verify if this was 'real', repeated pattern
60 | 2023-01-19 15:25 | WO-23003-32-A-1 | U | 0 | #N/A | 2.118 | 0 | #N/A | 0.394 | 0.346 | #N/A | 0.385 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 97 | 2023-01-20 10:55 | 217358-d2-inj2 | U | 0 | 1.119 | 0.234 | 0 | 0.834 | 0.616 | 0.839 | 0.539 | 632.89 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 1265.78 |  | IS3 | IS4 | IS4 | IS4 | [5 empty] | R1 | Analyzed again below to verify if this was 'real', repeated pattern
61 | 2023-01-19 15:52 | WO-23003-33-A-1 | U | 0 | #N/A | 0 | 1.126 | #N/A | 0.355 | 0.488 | #N/A | 0.674 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 98 | 2023-01-20 11:23 | 217359-d2-inj2 | U | 0 | 1.537 | 0 | 0 | 0 | 1.587 | 2.809 | 0.336 | 4183.607 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 8367.214 |  | IS3 | IS4 | IS4 | IS4 | IS4 |  |  |  | IS4 | R1 | Analyzed again below to verify if this was 'real', repeated pattern
62 | 2023-01-19 16:19 | WO-23003-34-A-1 | U | 0 | #N/A | 0 | 0.991 | #N/A | 0.484 | 0.521 | #N/A | 0.663 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 101 | 2023-01-20 12:44 | 217288-d4-inj2 | U | 0 | 4236.334 | 0 | 0.614 | 0 | 0.47 | 0.383 | 0.898 | 1.981 | 4 | <MRL | 16945.336 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | [5 empty] | R1 | Reprepared and run below.
63 | 2023-01-19 16:46 | WO-23003-35-A-1 | U | 0 | #N/A | 0 | 1.091 | #N/A | 0.491 | 0.514 | #N/A | 0.71 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 103 | 2023-01-20 14:05 | 217357-d2 | U | 0 | 0 | 0 | 0 | 0.922 | 0.547 | 1.322 | 0.735 | 60.711 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 121.422 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b |  | Subsequently reprepared entirely, but showed no improvement. Probably contamination of samples
64 | 2023-01-19 17:14 | WO-23003-36-A-1 | U | 0 | #N/A | 0.99 | 0.885 | #N/A | 0.516 | 0.587 | #N/A | 0.763 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 104 | 2023-01-20 14:32 | 217358-d2 | U | 0 | 0 | 0 | 0 | 0.848 | 0.656 | 0.846 | 0.383 | 660.136 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 1320.272 |  | IS3 | IS4 | IS4 | IS4 | [6 empty] | Subsequently reprepared entirely, but showed no improvement. Probably contamination of samples
65 | 2023-01-19 17:41 | WO-23003-37-A-1 | U | 0 | #N/A | 0 | 0.965 | #N/A | 0.344 | 0.473 | #N/A | 0.577 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 105 | 2023-01-20 15:00 | 217359-d2 | U | 0 | 2.425 | 0 | 0 | 0.818 | 1.794 | 2.954 | 0.342 | 4256.541 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 8513.082 |  | IS3 | IS4 | IS4 | IS4 | IS4 |  |  |  | IS4 |  | Subsequently reprepared entirely, but showed no improvement. Probably contamination of samples
66 | 2023-01-19 18:08 | WO-23003-38-A-1 | U | 0 | #N/A | 0 | 0.903 | #N/A | 0.433 | 0.435 | #N/A | 0.639 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS3 | IS4 | [8 empty] | 112 | 2023-01-20 18:36 | 217360-d2 | U | 0 | 1.448 | 0 | 0 | 0 | 0.576 | 0.672 | 0.735 | 4.3 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS4 | IS4
67 | 2023-01-19 18:35 | WO-23003-39-A-1 | U | 126.713 | #N/A | 0.513 | 0 | #N/A | 0.402 | 0.302 | #N/A | 0.309 | 1 | 126.713 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 113 | 2023-01-20 19:03 | 217361-d2 | U | 0 | 2.118 | 0 | 0 | 0.924 | 0.494 | 0.45 | 0.508 | 5.986 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 11.972 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b
68 | 2023-01-19 19:02 | WO-23003-40-A-1 | U | 134.042 | #N/A | 1.336 | 0 | #N/A | 0.33 | 0.443 | #N/A | 0.518 | 1 | 134.042 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 114 | 2023-01-20 19:30 | 217362-d2 | U | 0 | 1.341 | 0.45 | 0 | 0.833 | 0.39 | 0.432 | 0.443 | 9.272 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 18.544 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b
69 | 2023-01-19 19:29 | LRB5-inj1 | K | 0 | #N/A | 0 | 0 | #N/A | 0.429 | 0.372 | #N/A | 0.473 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 115 | 2023-01-20 19:57 | 217363-d2 | U | 0 | 1.537 | 0 | 0 | 0.772 | 0.38 | 0.475 | 0.402 | 12.026 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 24.052 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b
70 | 2023-01-19 19:56 | AB-230117-02 LRB 5 | K | 0 | #N/A | 0 | 0 | #N/A | 0.37 | 0.288 | #N/A | 0.334 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 116 | 2023-01-20 20:24 | 217364-d2 | U | 0 | 1.691 | 0 | 0.25 | 0 | 0.317 | 0.34 | 0.359 | 8.885 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 17.77 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b
71 | 2023-01-19 20:23 | AB-230117-02 LD 2 | D | 0 | #N/A | 0 | 0 | #N/A | 0.553 | 0 | #N/A | 0.39 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [7 empty] | IS4 | IS4 |  |  |  |  | 117 | 2023-01-20 20:51 | 217365-d2 | U | 0 | 1.191 | 0 | 0 | 0 | 0.338 | 0.344 | 0.383 | 5.647 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 11.294 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b
72 | 2023-01-19 20:50 | AB-230117-02 LFB-H 3 | Z | 169.371 | #N/A | 198.181 | 191.991 | #N/A | 190.049 | 189.526 | #N/A | 186.355 | 1 | 169.371 | #N/A | 198.181 | 191.991 | #N/A | 190.049 | 189.526 | #N/A | 186.355 | No TrueV! | [7 empty] | C |  |  |  |  | 118 | 2023-01-20 21:19 | 217366-d2 | U | 0 | 1.384 | 0 | 0 | 0.804 | 0.362 | 0.338 | 0.308 | 5.711 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 11.422 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b
73 | 2023-01-19 21:17 | AB-230117-02 CCC-H 3 | Q | 366.428 | #N/A | 400.863 | 377.097 | #N/A | 377.585 | 383.968 | #N/A | 393.376 | 1 | 366.428 | #N/A | 400.863 | 377.097 | #N/A | 377.585 | 383.968 | #N/A | 393.376 | [13 empty] | 119 | 2023-01-20 21:46 | 217367-d2 | U | 0 | 1.443 | 0 | 0 | 0.812 | 0.382 | 0.389 | 0.434 | 10.214 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 20.428 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b
74 | 2023-01-19 21:44 | AB-230117-02 LFSM-H 2 | S | 169.438 | #N/A | 198.319 | 190.263 | #N/A | 196.626 | 202.465 | #N/A | 206.37 | 1 | 169.438 | #N/A | 198.319 | 190.263 | #N/A | 196.626 | 202.465 | #N/A | 206.37 | No TrueV! | [6 empty] | IS3 | IS3 |  |  |  |  | 120 | 2023-01-20 22:13 | 217368-d2 | U | 0 | 1.175 | 0 | 0 | 0.763 | 0.524 | 0.312 | 0.297 | 9.563 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 19.126 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b
75 | 2023-01-19 22:11 | Rinse5 | U | 0 | #N/A | 0 | 39.805 | #N/A | 160.825 | 169.964 | #N/A | 175.37 | 1 | <MRL | #N/A | <MRL | 39.805 | #N/A | 160.825 | 169.964 | #N/A | 175.37 |  | IS3 | IS4 | IS4 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3 |  |  |  | 121 | 2023-01-20 22:40 | 217369-d2 | U | 0 | 2.141 | 0 | 0 | 0 | 0.365 | 0.408 | 0.38 | 14.244 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 28.488 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b
76 | 2023-01-19 22:38 | WO-23003-41-A-1 | U | 150.848 | #N/A | 0.582 | 0 | #N/A | 0.665 | 0.764 | #N/A | 0.677 | 1 | 150.848 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 |  |  |  | C |  | R1 | Rerun due to carryover |  | 124 | 2023-01-21 00:01 | 217370-d2 | U | 0 | 0 | 0.602 | 0 | 0 | 0.33 | 0.328 | 0.134 | 14.541 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 29.082 |  | IS3 | IS4 | IS4 | IS4
77 | 2023-01-19 23:06 | WO-23003-42-A-1 | U | 136.634 | #N/A | 0 | 0 | #N/A | 0.454 | 0.466 | #N/A | 0.442 | 1 | 136.634 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 125 | 2023-01-21 00:28 | 217371-d2 | U | 0 | 1.154 | 0 | 0 | 0.795 | 0.515 | 0.315 | 0.446 | 19.333 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 38.666 |  | IS3 | IS4 | IS4 | IS4
78 | 2023-01-19 23:33 | WO-23003-43-A-1 | U | 111.294 | #N/A | 0 | 0 | #N/A | 0.354 | 0.416 | #N/A | 0.4 | 1 | 111.294 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 126 | 2023-01-21 00:55 | 217372-d2 | U | 0 | 1.375 | 0 | 0 | 0 | 0.532 | 0.365 | 0.31 | 16.405 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 32.81 |  | IS3 | IS4 | IS4 | IS4
79 | 2023-01-20 | WO-23003-44-A-1 | U | 109.959 | #N/A | 1.537 | 0 | #N/A | 0.355 | 0.413 | #N/A | 0.359 | 1 | 109.959 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 127 | 2023-01-21 01:22 | 217373-d2 | U | 0 | 0 | 0 | 0 | 0.765 | 0.311 | 0.357 | 0.3 | 33.419 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 66.838 |  | IS3 | IS4 | IS4 | IS4
80 | 2023-01-20 00:27 | WO-23003-45-A-1 | U | 126.129 | #N/A | 0 | 0 | #N/A | 0.401 | 0.366 | #N/A | 0.368 | 1 | 126.129 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 128 | 2023-01-21 01:49 | 217374-d2 | U | 0 | 1.73 | 0 | 0 | 0.805 | 0.403 | 0.51 | 0.486 | 46.694 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 93.388 |  | IS3 | IS4 | IS4 | IS4
81 | 2023-01-20 00:54 | WO-23003-46-A-1 | U | 117.077 | #N/A | 0 | 0 | #N/A | 0.328 | 0.427 | #N/A | 0.45 | 1 | 117.077 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 129 | 2023-01-21 02:16 | 217325-d2 | U | 201.051 | 1.534 | 0 | 0 | 0.781 | 0.32 | 0.35 | 0.24 | 1.751 | 2 | 402.102 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
82 | 2023-01-20 01:21 | WO-23003-47-A-1 | U | 268.234 | #N/A | 1.131 | 0 | #N/A | 0.429 | 0.355 | #N/A | 0.318 | 1 | 268.234 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 130 | 2023-01-21 02:43 | 217326-d2 | U | 212.264 | 1.2 | 0.689 | 0 | 0.81 | 0.35 | 0.385 | 0.067 | 0.393 | 2 | 424.528 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
83 | 2023-01-20 01:48 | WO-23003-48-A-1 | U | 253.554 | #N/A | 0.421 | 0 | #N/A | 0.491 | 0.363 | #N/A | 0.334 | 1 | 253.554 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 131 | 2023-01-21 03:11 | 217327-d2 | U | 343.62 | 1.06 | 0 | 0 | 0 | 0.336 | 0.301 | 0.047 | 0.283 | 2 | 687.24 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
84 | 2023-01-20 02:15 | WO-23003-49-A-1 | U | 250.815 | #N/A | 0 | 0 | #N/A | 0.807 | 0.381 | #N/A | 0.323 | 1 | 250.815 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 132 | 2023-01-21 03:38 | 217328-d2 | U | 341.696 | 1.386 | 0 | 0 | 0 | 0.404 | 0.384 | 0.161 | 0.401 | 2 | 683.392 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
85 | 2023-01-20 02:42 | WO-23003-50-A-1 | U | 214.775 | #N/A | 2.007 | 0 | #N/A | 0 | 0.388 | #N/A | 0.264 | 1 | 214.775 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 133 | 2023-01-21 04:05 | 217329-d2 | U | 354.309 | 2.611 | 0 | 0 | 0.781 | 0.352 | 0.326 | 0.069 | 0.25 | 2 | 708.618 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
86 | 2023-01-20 03:09 | LRB6-inj1 | K | 0 | #N/A | 0 | 0.352 | #N/A | 0.362 | 0.221 | #N/A | 0.343 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 140 | 2023-01-21 07:41 | 217330-d2 | U | 319.278 | 1.055 | 0 | 0 | 0.984 | 0.541 | 0.62 | 0.375 | 0.574 | 2 | 638.556 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
87 | 2023-01-20 03:36 | AB-230117-02 LRB 6 | K | 0 | #N/A | 0 | 0 | #N/A | 0.308 | 0.24 | #N/A | 0.346 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 141 | 2023-01-21 08:08 | 217331-d2 | U | 602.933 | 2.483 | 2.353 | 0 | 0.97 | 0.828 | 0.705 | 0.518 | 0.63 | 2 | 1205.866 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | R2 | Low internal standard in this sample, looks like around half.
88 | 2023-01-20 04:04 | WO-23003-51-A-1 | U | 498.597 | #N/A | 0 | 0 | #N/A | 0.375 | 0.324 | #N/A | 0.295 | 1 | 498.597 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | IS4 | [7 empty] | 142 | 2023-01-21 08:36 | 217332-d2 | U | 311.52 | 1.468 | 0 | 0 | 0.804 | 0.375 | 0.424 | 0.184 | 0.398 | 2 | 623.04 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
89 | 2023-01-20 04:31 | WO-23003-52-A-1 | U | 413.609 | #N/A | 1.442 | 0 | #N/A | 0.295 | 0.342 | #N/A | 0.235 | 1 | 413.609 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 143 | 2023-01-21 09:03 | 217333-d2 | U | 405.654 | 0 | 0 | 0 | 0.782 | 0.327 | 0.405 | 0.214 | 0.229 | 2 | 811.308 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 |  | IS4 | IS4
90 | 2023-01-20 04:58 | WO-23003-53-A-1 | U | 419.335 | #N/A | 0 | 0 | #N/A | 0.355 | 0.362 | #N/A | 0.428 | 1 | 419.335 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 144 | 2023-01-21 09:30 | 217334-d2 | U | 419.227 | 0 | 0 | 0 | 0.74 | 0.31 | 0.395 | 0.158 | 0.326 | 2 | 838.454 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 |  | IS4 | IS4
91 | 2023-01-20 05:25 | WO-23003-54-A-1 | U | 446.172 | #N/A | 0 | 0 | #N/A | 0.466 | 0.507 | #N/A | 0.617 | 1 | 446.172 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS3 | IS4 | [8 empty] | 145 | 2023-01-21 09:57 | 217335-d2 | U | 433.77 | 2.233 | 0.871 | 0 | 0.75 | 0.35 | 0.431 | 0.071 | 0.246 | 2 | 867.54 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
92 | 2023-01-20 05:52 | WO-23003-55-A-1 | U | 456.386 | #N/A | 0 | 0.584 | #N/A | 0.523 | 0.541 | #N/A | 1.323 | 1 | 456.386 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS3 | IS4 | [8 empty] | 146 | 2023-01-21 10:24 | 217336-d2 | U | 406.54 | 1.572 | 0 | 0 | 0.749 | 0.359 | 0.368 | 0.201 | 0.219 | 2 | 813.08 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
93 | 2023-01-20 06:19 | WO-23003-56-A-1 | U | 433.565 | #N/A | 0 | 0 | #N/A | 0.523 | 0.563 | #N/A | 0.576 | 1 | 433.565 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS3 | IS4 | IS4 | [7 empty] | 147 | 2023-01-21 10:51 | 217337-d2 | U | 382.709 | 0 | 0 | 0 | 0 | 0.287 | 0.342 | 0.092 | 0.223 | 2 | 765.418 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 |  | IS4 | IS4
94 | 2023-01-20 06:46 | WO-23003-57-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.337 | 0.365 | #N/A | 4.897 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 148 | 2023-01-21 11:18 | 217338-d2 | U | 415.919 | 1.046 | 0 | 0 | 0.752 | 0.284 | 0.291 | 0.075 | 0.192 | 2 | 831.838 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
95 | 2023-01-20 10:01 | WO-23003-26-A-1B | U | 0 | #N/A | 0 | 0 | #N/A | 0.288 | 0.417 | #N/A | 0.466 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 |  |  |  |  | R1 | Analyzed again below to verify if this was 'real', repeated pattern |  | 149 | 2023-01-21 11:45 | 217339-d2 | U | 365.485 | 2.291 | 0 | 0 | 0.792 | 0.344 | 0.321 | 0.152 | 0.371 | 2 | 730.97 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
96 | 2023-01-20 10:28 | WO-23003-58-A-1B | U | 0 | #N/A | 0 | 0 | #N/A | 0.492 | 1.257 | #N/A | 61.125 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 61.125 |  | IS3 | IS4 | IS4 | IS4 | [5 empty] | R1 | Analyzed again below to verify if this was 'real', repeated pattern |  | 152 | 2023-01-21 13:06 | 217340-d2 | U | 385.432 | 0 | 0 | 0 | 0.749 | 0.38 | 0.317 | 0.114 | 0.408 | 2 | 770.864 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
97 | 2023-01-20 10:55 | WO-23003-59-A-1B | U | 0 | #N/A | 0.234 | 0 | #N/A | 0.616 | 0.839 | #N/A | 632.89 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 632.89 |  | IS3 | IS4 | IS4 | IS4 | [5 empty] | R1 | Analyzed again below to verify if this was 'real', repeated pattern |  | 153 | 2023-01-21 13:33 | 217341-d2 | U | 385.868 | 0 | 0 | 0 | 0.76 | 0.312 | 0.343 | 0.23 | 0.174 | 2 | 771.736 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3
98 | 2023-01-20 11:23 | WO-23003-60-A-1B | U | 0 | #N/A | 0 | 0 | #N/A | 1.587 | 2.809 | #N/A | 4183.607 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4183.607 |  | IS3 | IS4 | IS4 | IS4 | IS4 |  |  |  | IS4 | R1 | Analyzed again below to verify if this was 'real', repeated pattern |  | 154 | 2023-01-21 14:00 | 217342-d2 | U | 377.744 | 2.816 | 0 | 0 | 0 | 0.277 | 0.391 | 0.097 | 0.324 | 2 | 755.488 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
99 | 2023-01-20 11:50 | LRB-7-inj2 | K | 0 | #N/A | 0 | 0 | #N/A | 0.294 | 0.334 | #N/A | 153.845 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 153.845 | [9 empty] | b |  |  |  | 155 | 2023-01-21 14:27 | 217343-d2 | U | 396.942 | 0 | 0 | 0 | 0 | 0.27 | 0.352 | 0.093 | 0.307 | 2 | 793.884 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 |  | IS4 | IS4
100 | 2023-01-20 12:17 | LRB-7-inj3 | K | 157.537 | #N/A | 1.41 | 0 | #N/A | 0.295 | 0.324 | #N/A | 4.078 | 1 | 157.537 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [6 empty] | Original sample had carryover, not a duplicate failure |  | 156 | 2023-01-21 14:54 | 217344-d2 | U | 396.948 | 0 | 0 | 0 | 0 | 0.24 | 0.271 | 0.007 | 0.195 | 2 | 793.896 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 |  | IS4 | IS4
101 | 2023-01-20 12:44 | WO-23003-14-A-1B | U | 0 | #N/A | 0 | 0.614 | #N/A | 0.47 | 0.383 | #N/A | 1.981 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [5 empty] | R1 | Reprepared and run below. |  | 157 | 2023-01-21 15:22 | 217345-d2 | U | 396.134 | 0 | 0 | 0 | 0.736 | 0.191 | 0.291 | 0.043 | 0.168 | 2 | 792.268 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
102 | 2023-01-20 13:11 | LFB-H-3-Inj2 | Z | 174.08 | #N/A | 196.02 | 191.147 | #N/A | 193.469 | 195.058 | #N/A | 197.419 | 1 | 174.08 | #N/A | 196.02 | 191.147 | #N/A | 193.469 | 195.058 | #N/A | 197.419 | No TrueV! | [12 empty] | 158 | 2023-01-21 15:49 | 217346-d2 | U | 382.008 | 0 | 0.489 | 0 | 0.753 | 0.314 | 0.355 | 0.188 | 0.572 | 2 | 764.016 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 |  | IS4 | IS4
103 | 2023-01-20 14:05 | WO-23003-58-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.547 | 1.322 | #N/A | 60.711 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 60.711 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b |  | Subsequently reprepared entirely, but showed no improvement. Probably contamination of samples |  | 159 | 2023-01-21 16:16 | 217347-d2 | U | 403.126 | 0 | 0 | 0 | 0.769 | 0.294 | 0.335 | 0.116 | 0.211 | 2 | 806.252 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 |  | IS4 | IS4
104 | 2023-01-20 14:32 | WO-23003-59-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.656 | 0.846 | #N/A | 660.136 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 660.136 |  | IS3 | IS4 | IS4 | IS4 | [6 empty] | Subsequently reprepared entirely, but showed no improvement. Probably contamination of samples |  | 160 | 2023-01-21 16:43 | 217348-d2 | U | 447.416 | 0 | 0 | 0 | 0.792 | 0.262 | 0.309 | -0.013 | 0.186 | 2 | 894.832 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
105 | 2023-01-20 15:00 | WO-23003-60-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 1.794 | 2.954 | #N/A | 4256.541 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4256.541 |  | IS3 | IS4 | IS4 | IS4 | IS4 |  |  |  | IS4 |  | Subsequently reprepared entirely, but showed no improvement. Probably contamination of samples |  | 161 | 2023-01-21 17:10 | 217349-d2 | U | 410.536 | 1.922 | 0.321 | 0 | 0.717 | 0.224 | 0.31 | 0.128 | 0.224 | 2 | 821.072 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
106 | 2023-01-20 15:27 | LRB7-inj1 | K | 0 | #N/A | 0 | 0 | #N/A | 0.326 | 0.308 | #N/A | 27.187 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 27.187 | [9 empty] | b |  |  |  | 168 | 2023-01-21 20:47 | 217375-d2 | U | 0 | 0 | 0 | 0.912 | 0.928 | 0.618 | 0.649 | 0.365 | 28.256 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 56.512 |  | IS3 |  | IS4 | IS4
107 | 2023-01-20 15:54 | AB-230117-02 LRB 7 | K | 0 | #N/A | 0 | 0 | #N/A | 14.3 | 4.824 | #N/A | 77.188 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 14.3 | <MRL | #N/A | 77.188 |  | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | b |  |  |  | 169 | 2023-01-21 21:14 | 217376-d2 | U | 0 | 0 | 0.788 | 1.032 | 0 | 0.732 | 0.666 | 0.352 | 64.801 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 129.602 |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
108 | 2023-01-20 16:21 | AB-230117-02 LD 3 | D | 140.88 | #N/A | 0 | 0 | #N/A | 0.387 | 0.41 | #N/A | 1.737 | 1 | 140.88 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [6 empty] | Original sample had carryover, not a duplicate failure |  | 170 | 2023-01-21 21:41 | 217377-d2 | U | 0 | 0 | 0 | 0.975 | 0 | 0.702 | 0.525 | 0.047 | 66.78 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 133.56 |  | IS3 | IS4 | IS4 | IS4 | IS4
109 | 2023-01-20 16:48 | AB-230117-02 LFB-H 4 | Z | 160.162 | #N/A | 199.569 | 187.965 | #N/A | 186.073 | 195.326 | #N/A | 192.196 | 1 | 160.162 | #N/A | 199.569 | 187.965 | #N/A | 186.073 | 195.326 | #N/A | 192.196 | No TrueV! | [12 empty] | 171 | 2023-01-21 22:08 | 217378-d2 | U | 0 | 0 | 0 | 0 | 0.815 | 0.713 | 0.618 | 0.093 | 128.717 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 257.434 |  | IS3
110 | 2023-01-20 17:15 | AB-230117-02 CCC-H 4 | Q | 333.561 | #N/A | 394.235 | 375.413 | #N/A | 375.67 | 378.57 | #N/A | 378.791 | 1 | 333.561 | #N/A | 394.235 | 375.413 | #N/A | 375.67 | 378.57 | #N/A | 378.791 | [13 empty] | 172 | 2023-01-21 22:35 | 217379-d2 | U | 0 | 0 | 0.4 | 0 | 0.781 | 0.819 | 0.695 | 0.119 | 128.03 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 256.06 |  | IS3 | IS4 | IS4 | IS4
111 | 2023-01-20 17:42 | AB-230117-02 LFSM-H 3 | S | 334.468 | #N/A | 193.406 | 190.483 | #N/A | 191.067 | 198.421 | #N/A | 203.756 | 1 | 334.468 | #N/A | 193.406 | 190.483 | #N/A | 191.067 | 198.421 | #N/A | 203.756 | No TrueV! | IS3 | IS3 | IS3 | IS3 | [8 empty] | 173 | 2023-01-21 23:02 | 217380-d2 | U | 0 | 0 | 0.23 | 0 | 0.727 | 0.742 | 0.515 | 0.016 | 137.351 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 274.702 |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
112 | 2023-01-20 18:36 | WO-23003-61-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.576 | 0.672 | #N/A | 4.3 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 174 | 2023-01-21 23:29 | 217381-d2 | U | 0 | 0 | 0 | 0 | 0 | 0.668 | 0.464 | 0.048 | 141.587 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 283.174 |  | IS3 | IS4 | IS4 | IS4
113 | 2023-01-20 19:03 | WO-23003-62-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.494 | 0.45 | #N/A | 5.986 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 5.986 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b |  |  |  | 175 | 2023-01-21 23:56 | 217382-d2 | U | 0 | 0 | 0 | 0 | 0.794 | 0.504 | 0.484 | -0.006 | 117.876 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 235.752 |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
114 | 2023-01-20 19:30 | WO-23003-63-A-1 | U | 0 | #N/A | 0.45 | 0 | #N/A | 0.39 | 0.432 | #N/A | 9.272 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 9.272 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b |  |  |  | 176 | 2023-01-22 00:23 | 217383-d2 | U | 0 | 1.175 | 0 | 0.53 | 0 | 0.575 | 0.506 | 0.117 | 137.434 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 274.868 |  | IS3 |  | IS4 | IS4
115 | 2023-01-20 19:57 | WO-23003-64-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.38 | 0.475 | #N/A | 12.026 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 12.026 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b |  |  |  | 177 | 2023-01-22 00:50 | 217384-d2 | U | 0 | 4.918 | 0 | 1.761 | 0 | 3.639 | 2.107 | -0.005 | 666.017 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 1332.034 |  | IS3 | IS4 | IS4 | IS4
116 | 2023-01-20 20:24 | WO-23003-65-A-1 | U | 0 | #N/A | 0 | 0.25 | #N/A | 0.317 | 0.34 | #N/A | 8.885 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 8.885 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b |  |  |  | 180 | 2023-01-22 02:12 | 217385-d2 | U | 0 | 4.153 | 0 | 1.35 | 0 | 4.279 | 2.348 | 0.379 | 740.764 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 1481.528 |  | IS3 | IS4 | IS4 | IS4 | [5 empty] | R1 | Electively rerun below.
117 | 2023-01-20 20:51 | WO-23003-66-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.338 | 0.344 | #N/A | 5.647 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 5.647 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b |  |  |  | 181 | 2023-01-22 02:39 | 217386-d2 | U | 0 | 4.125 | 0 | 1.139 | 0 | 2.811 | 1.711 | 0.084 | 536.317 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 1072.634 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b
118 | 2023-01-20 21:19 | WO-23003-67-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.362 | 0.338 | #N/A | 5.711 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 5.711 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b |  |  |  | 182 | 2023-01-22 03:06 | 217387-d2 | U | 0 | 2.379 | 0 | 0.765 | 0.746 | 2.136 | 1.514 | -0.04 | 641.876 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 1283.752 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b
119 | 2023-01-20 21:46 | WO-23003-68-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.382 | 0.389 | #N/A | 10.214 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 10.214 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b |  |  |  | 183 | 2023-01-22 03:33 | 217388-d2 | U | 0 | 2.409 | 0 | 1.029 | 0.762 | 2.034 | 1.569 | -0.089 | 622.53 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 1245.06 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b
120 | 2023-01-20 22:13 | WO-23003-69-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.524 | 0.312 | #N/A | 9.563 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 9.563 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b |  |  |  | 184 | 2023-01-22 04:00 | 217389-d2 | U | 0 | 0 | 0.437 | 1.365 | 0.841 | 2.371 | 1.895 | 0.047 | 714.013 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 1428.026 |  | IS3 | IS4 | IS4 | IS4
121 | 2023-01-20 22:40 | WO-23003-70-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.365 | 0.408 | #N/A | 14.244 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 14.244 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b |  |  |  | 185 | 2023-01-22 04:27 | 217390-d4 | U | 0 | 3.453 | 0 | 1.71 | 0.756 | 11.677 | 8.8 | 0.044 | 2477.13 | 4 | <MRL | <MRL | <MRL | <MRL | <MRL | 46.708 | 35.2 | <MRL | 9908.52 |  | IS3 | [7 empty] | IS4
122 | 2023-01-20 23:07 | LRB-8-inj1 | K | 0 | #N/A | 0.346 | 0 | #N/A | 0.355 | 0.28 | #N/A | 0.783 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 186 | 2023-01-22 04:54 | 217391-d4 | U | 0 | 0 | 0.376 | 1.608 | 0.828 | 10.943 | 8.243 | -0.151 | 2511.685 | 4 | <MRL | <MRL | <MRL | <MRL | <MRL | 43.772 | 32.972 | <MRL | 10046.74 |  | IS3 | [7 empty] | IS4
123 | 2023-01-20 23:34 | AB-230117-02 LRB 8 | K | 0 | #N/A | 0 | 0 | #N/A | 0.409 | 0.277 | #N/A | 0.257 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 187 | 2023-01-22 05:21 | 217392-d4 | U | 0 | 3.543 | 0 | 1.751 | 0.761 | 12.301 | 9.413 | -0.032 | 2857.187 | 4 | <MRL | <MRL | <MRL | <MRL | <MRL | 49.204 | 37.652 | <MRL | 11428.748 |  | IS3 | [7 empty] | IS4
124 | 2023-01-21 00:01 | WO-23003-71-A-1 | U | 0 | #N/A | 0.602 | 0 | #N/A | 0.33 | 0.328 | #N/A | 14.541 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 14.541 |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 188 | 2023-01-22 05:48 | 217393-d4 | U | 0 | 2.443 | 0 | 1.049 | 0.8 | 3.789 | 2.851 | -0.061 | 1085.45 | 4 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 4341.8 |  | IS3 | [8 empty] | R1 | Electively rerun below
125 | 2023-01-21 00:28 | WO-23003-72-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.515 | 0.315 | #N/A | 19.333 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 19.333 |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 189 | 2023-01-22 06:15 | 217394-d4 | U | 0 | 0 | 0.615 | 1.294 | 0.761 | 4.447 | 3.724 | 0 | 1329.015 | 4 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 5316.06 |  | IS3 | [8 empty] | R1 | Electively rerun below
126 | 2023-01-21 00:55 | WO-23003-73-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.532 | 0.365 | #N/A | 16.405 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 16.405 |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 196 | 2023-01-22 09:52 | 217395-d4 | U | 0 | 0 | 0 | 1.724 | 0.926 | 8.728 | 8.418 | 0.211 | 2598.122 | 4 | <MRL | <MRL | <MRL | <MRL | <MRL | 34.912 | 33.672 | <MRL | 10392.488 |  | IS3 | [7 empty] | IS4 | R1 | Electively rerun below
127 | 2023-01-21 01:22 | WO-23003-74-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.311 | 0.357 | #N/A | 33.419 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 33.419 |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 197 | 2023-01-22 10:19 | 217396-d200 | U | 0 | 0 | 0 | 0.942 | 0.815 | 1.505 | 3.298 | 0.037 | 4052.461 | 200 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 810492.2 | [9 empty] | IS4
128 | 2023-01-21 01:49 | WO-23003-75-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.403 | 0.51 | #N/A | 46.694 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 46.694 |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 198 | 2023-01-22 10:46 | 217397-d200 | U | 0 | 0 | 0 | 0.825 | 0.797 | 1.448 | 3.799 | -0.053 | 4114.272 | 200 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 822854.4 | [9 empty] | IS4
129 | 2023-01-21 02:16 | WO-23003-76-A-1 | U | 201.051 | #N/A | 0 | 0 | #N/A | 0.32 | 0.35 | #N/A | 1.751 | 1 | 201.051 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 199 | 2023-01-22 11:13 | 217398-d200 | U | 0 | 0 | 0 | 0.721 | 0.821 | 1.325 | 2.549 | -0.14 | 4254.003 | 200 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 850800.6 | [9 empty] | IS4
130 | 2023-01-21 02:43 | WO-23003-77-A-1 | U | 212.264 | #N/A | 0.689 | 0 | #N/A | 0.35 | 0.385 | #N/A | 0.393 | 1 | 212.264 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 203 | 2023-01-23 11:29 | 217385-d2-inj2 | U | 0 | 4.439 | 0 | 1.749 | 0.805 | 4.534 | 2.374 | 0.434 | 738.645 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 1477.29 |  | IS3 | IS4 | IS4 | IS4
131 | 2023-01-21 03:11 | WO-23003-78-A-1 | U | 343.62 | #N/A | 0 | 0 | #N/A | 0.336 | 0.301 | #N/A | 0.283 | 1 | 343.62 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 204 | 2023-01-23 11:56 | 217395-d4-inj2 | U | 0 | 0 | 0.285 | 1.486 | 0 | 8.541 | 7.928 | -0.034 | 2538.308 | 4 | <MRL | <MRL | <MRL | <MRL | <MRL | 34.164 | 31.712 | <MRL | 10153.232 | [9 empty] | IS4 | R1 | Reprepared and rerun below
132 | 2023-01-21 03:38 | WO-23003-79-A-1 | U | 341.696 | #N/A | 0 | 0 | #N/A | 0.404 | 0.384 | #N/A | 0.401 | 1 | 341.696 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 205 | 2023-01-23 12:24 | 217399-d2 | U | 0 | 11.776 | 0 | 0 | 0 | 0.512 | 0.413 | 0.089 | 46.972 | 2 | <MRL | 23.552 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 93.944 |  | IS3 | IS4 | IS3 | IS4 |  |  |  |  | C | R1 | Rerun due to carryover concern
133 | 2023-01-21 04:05 | WO-23003-80-A-1 | U | 354.309 | #N/A | 0 | 0 | #N/A | 0.352 | 0.326 | #N/A | 0.25 | 1 | 354.309 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 206 | 2023-01-23 12:51 | 217400-d2 | U | 0 | 9.643 | 0 | 0 | 0.809 | 0 | 0.412 | 0.267 | 4.716 | 2 | <MRL | 19.286 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  | R1 | Rerun due to carryover concern
134 | 2023-01-21 04:32 | AB-230117-02 LRB 9 | K | 0 | #N/A | 0 | 0.794 | #N/A | 0.552 | 0.229 | #N/A | 0.312 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 207 | 2023-01-23 13:18 | 217401-d2 | U | 0 | 9.456 | 0 | 0 | 0.839 | 0.451 | 0.375 | 0.029 | 1.887 | 2 | <MRL | 18.912 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4
135 | 2023-01-21 04:59 | LD4-inj1 | D | 0 | #N/A | 1.695 | 0 | #N/A | 0.397 | 0.342 | #N/A | 3.644 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 208 | 2023-01-23 13:45 | 217402-d2 | U | 0 | 9.095 | 0.672 | 0 | 0.774 | 0.583 | 0.347 | 0.093 | 1.569 | 2 | <MRL | 18.19 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
136 | 2023-01-21 05:26 | AB-230117-02 LD 4 | D | 0 | #N/A | 1.281 | 0 | #N/A | 0.502 | 0.314 | #N/A | 3.87 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 209 | 2023-01-23 14:12 | 217403-d2 | U | 0 | 12.556 | 0 | 0 | 0.771 | 0.451 | 0.345 | 0.236 | 1.157 | 2 | <MRL | 25.112 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4
137 | 2023-01-21 05:53 | AB-230117-02 LFB-H 5 | Z | 158.258 | #N/A | 179.705 | 179.093 | #N/A | 179.201 | 181.423 | #N/A | 180.317 | 1 | 158.258 | #N/A | 179.705 | 179.093 | #N/A | 179.201 | 181.423 | #N/A | 180.317 | No TrueV! | [12 empty] | 210 | 2023-01-23 14:39 | 217404-d2 | U | 0 | 13.469 | 0.482 | 0 | 0 | 0.481 | 0.296 | -0.025 | 0.778 | 2 | <MRL | 26.938 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
138 | 2023-01-21 06:20 | AB-230117-02 CCC-H 5 | Q | 309.052 | #N/A | 375.935 | 355.016 | #N/A | 353.128 | 357.216 | #N/A | 355.469 | 1 | 309.052 | #N/A | 375.935 | 355.016 | #N/A | 353.128 | 357.216 | #N/A | 355.469 | [13 empty] | 213 | 2023-01-23 16:00 | 217405-d2 | U | 0 | 14.312 | 0.329 | 0 | 0 | 0.366 | 0.306 | -0.101 | 0.748 | 2 | <MRL | 28.624 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
139 | 2023-01-21 06:47 | AB-230117-02 LFSM-H 4 | S | 189.256 | #N/A | 200.054 | 184.951 | #N/A | 197.378 | 204.634 | #N/A | 213.294 | 1 | 189.256 | #N/A | 200.054 | 184.951 | #N/A | 197.378 | 204.634 | #N/A | 213.294 | No TrueV! | IS3 | IS3 | IS3 | IS3 | [8 empty] | 214 | 2023-01-23 16:27 | 217406-d2 | U | 0 | 16.223 | 0 | 0 | 0 | 1.113 | 1.47 | -0.029 | 0.77 | 2 | <MRL | 32.446 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
140 | 2023-01-21 07:41 | WO-23003-81-A-1 | U | 319.278 | #N/A | 0 | 0 | #N/A | 0.541 | 0.62 | #N/A | 0.574 | 1 | 319.278 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 215 | 2023-01-23 16:54 | 217407-d2 | U | 0 | 33.78 | 1.203 | 0 | 0 | 0.367 | 0.295 | 0.106 | 0.523 | 2 | <MRL | 67.56 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4
141 | 2023-01-21 08:08 | WO-23003-82-A-1 | U | 602.933 | #N/A | 2.353 | 0 | #N/A | 0.828 | 0.705 | #N/A | 0.63 | 1 | 602.933 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | R2 | Low internal standard in this sample, looks like around half. |  | 216 | 2023-01-23 17:21 | 217408-d2 | U | 0 | 38.926 | 0.198 | 0 | 0.762 | 0.371 | 0.261 | -0.051 | 0.414 | 2 | <MRL | 77.852 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
142 | 2023-01-21 08:36 | WO-23003-83-A-1 | U | 311.52 | #N/A | 0 | 0 | #N/A | 0.375 | 0.424 | #N/A | 0.398 | 1 | 311.52 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 217 | 2023-01-23 17:49 | 217409-d2 | U | 0 | 34.102 | 0 | 0 | 0.754 | 0.412 | 0.333 | 0.159 | 0.443 | 2 | <MRL | 68.204 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
143 | 2023-01-21 09:03 | WO-23003-84-A-1 | U | 405.654 | #N/A | 0 | 0 | #N/A | 0.327 | 0.405 | #N/A | 0.229 | 1 | 405.654 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS4 | IS4 | [8 empty] | 218 | 2023-01-23 18:16 | 217410-d2 | U | 0 | 36.314 | 1.066 | 0 | 0.689 | 0.251 | 0.251 | -0.035 | 0.437 | 2 | <MRL | 72.628 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4
144 | 2023-01-21 09:30 | WO-23003-85-A-1 | U | 419.227 | #N/A | 0 | 0 | #N/A | 0.31 | 0.395 | #N/A | 0.326 | 1 | 419.227 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS4 | IS4 | [8 empty] | 219 | 2023-01-23 18:43 | 217411-d2 | U | 0 | 30.824 | 0 | 0 | 0 | 0.319 | 0.258 | -0.072 | 0.386 | 2 | <MRL | 61.648 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3
145 | 2023-01-21 09:57 | WO-23003-86-A-1 | U | 433.77 | #N/A | 0.871 | 0 | #N/A | 0.35 | 0.431 | #N/A | 0.246 | 1 | 433.77 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 220 | 2023-01-23 19:10 | 217412-d2 | U | 0 | 41.842 | 0 | 0 | 0.961 | 0.209 | 0.284 | 0.02 | 0.368 | 2 | <MRL | 83.684 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 |  | IS3 | IS4
146 | 2023-01-21 10:24 | WO-23003-87-A-1 | U | 406.54 | #N/A | 0 | 0 | #N/A | 0.359 | 0.368 | #N/A | 0.219 | 1 | 406.54 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 221 | 2023-01-27 08:27 | 217413-d2 | U | 0 | 141.212 | 3.112 | 0 | 0 | 0.695 | 0.57 | 0.75 | 133.922 | 2 | <MRL | 282.424 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 267.844 |  | IS3 | IS4 | IS3 | IS4 |  |  |  |  | C | R1 | Analyzed after instrument leak, some carryover observed and sample rerun
147 | 2023-01-21 10:51 | WO-23003-88-A-1 | U | 382.709 | #N/A | 0 | 0 | #N/A | 0.287 | 0.342 | #N/A | 0.223 | 1 | 382.709 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS4 | IS4 | [8 empty] | 222 | 2023-01-27 08:54 | 217413-d2-inj2 | U | 0 | 146.837 | 2.591 | 0 | 0.848 | 0.378 | 0.345 | 0.878 | 6.248 | 2 | <MRL | 293.674 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 12.496 |  | IS3 | IS4 | IS3 | IS4 |  |  |  |  | C
148 | 2023-01-21 11:18 | WO-23003-89-A-1 | U | 415.919 | #N/A | 0 | 0 | #N/A | 0.284 | 0.291 | #N/A | 0.192 | 1 | 415.919 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 223 | 2023-01-27 09:21 | 217414-d2 | U | 0 | 152.509 | 0.914 | 0 | 0 | 0.357 | 0.338 | -0.048 | 2.01 | 2 | <MRL | 305.018 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
149 | 2023-01-21 11:45 | WO-23003-90-A-1 | U | 365.485 | #N/A | 0 | 0 | #N/A | 0.344 | 0.321 | #N/A | 0.371 | 1 | 365.485 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 230 | 2023-01-27 12:58 | 217415-d2 | U | 0 | 113.991 | 0 | 0 | 0 | 0.824 | 0.724 | 0.287 | 0.925 | 2 | <MRL | 227.982 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4
150 | 2023-01-21 12:12 | LRB-10-inj1 | K | 0 | #N/A | 0 | 0.79 | #N/A | 0.515 | 0.261 | #N/A | 0.283 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 231 | 2023-01-27 13:25 | 217416-d2 | U | 0 | 151.9 | 0 | 1.467 | 0 | 0.453 | 0.486 | 0.024 | 0.92 | 2 | <MRL | 303.8 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
151 | 2023-01-21 12:39 | AB-230117-02 LRB 10 | K | 0 | #N/A | 0.079 | 0.848 | #N/A | 0.571 | 0.264 | #N/A | 0.321 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 232 | 2023-01-27 13:52 | 217417-d2 | U | 0 | 102.324 | 1.526 | 0 | 0.803 | 0.43 | 0.363 | 0.121 | 0.671 | 2 | <MRL | 204.648 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
152 | 2023-01-21 13:06 | WO-23003-91-A-1 | U | 385.432 | #N/A | 0 | 0 | #N/A | 0.38 | 0.317 | #N/A | 0.408 | 1 | 385.432 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 233 | 2023-01-27 14:19 | 217418-d2 | U | 0 | 127.716 | 0 | 0 | 0 | 0.442 | 0.337 | 0.056 | 0.465 | 2 | <MRL | 255.432 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
153 | 2023-01-21 13:33 | WO-23003-92-A-1 | U | 385.868 | #N/A | 0 | 0 | #N/A | 0.312 | 0.343 | #N/A | 0.174 | 1 | 385.868 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | [11 empty] | 234 | 2023-01-27 14:46 | 217419-d2 | U | 0 | 638.175 | 0 | 0 | 0.871 | 0.406 | 0.332 | 0.014 | 0.419 | 2 | <MRL | 1276.35 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
154 | 2023-01-21 14:00 | WO-23003-93-A-1 | U | 377.744 | #N/A | 0 | 0 | #N/A | 0.277 | 0.391 | #N/A | 0.324 | 1 | 377.744 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 235 | 2023-01-27 15:13 | 217420-d2 | U | 0 | 732.447 | 0 | 0 | 0.792 | 0.535 | 0.324 | 0.036 | 0.408 | 2 | <MRL | 1464.894 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
155 | 2023-01-21 14:27 | WO-23003-94-A-1 | U | 396.942 | #N/A | 0 | 0 | #N/A | 0.27 | 0.352 | #N/A | 0.307 | 1 | 396.942 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS4 | IS4 | [8 empty] | 236 | 2023-01-27 15:40 | 217421-d2 | U | 0 | 572.526 | 0 | 0 | 0 | 0.467 | 0.329 | -0.042 | 0.366 | 2 | <MRL | 1145.052 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
156 | 2023-01-21 14:54 | WO-23003-95-A-1 | U | 396.948 | #N/A | 0 | 0 | #N/A | 0.24 | 0.271 | #N/A | 0.195 | 1 | 396.948 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS4 | IS4 | [8 empty] | 237 | 2023-01-27 16:07 | 217422-d2 | U | 0 | 577.003 | 1.373 | 0 | 0 | 0.388 | 0.358 | -0.065 | 0.359 | 2 | <MRL | 1154.006 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
157 | 2023-01-21 15:22 | WO-23003-96-A-1 | U | 396.134 | #N/A | 0 | 0 | #N/A | 0.191 | 0.291 | #N/A | 0.168 | 1 | 396.134 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 238 | 2023-01-27 16:34 | 217423-d2 | U | 0 | 519.796 | 1.108 | 0 | 0.724 | 0.237 | 0.346 | 0 | 0.365 | 2 | <MRL | 1039.592 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
158 | 2023-01-21 15:49 | WO-23003-97-A-1 | U | 382.008 | #N/A | 0.489 | 0 | #N/A | 0.314 | 0.355 | #N/A | 0.572 | 1 | 382.008 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS4 | IS4 | [8 empty] | 239 | 2023-01-27 17:01 | 217424-d2 | U | 0 | 579.319 | 0 | 0 | 0.907 | 0.429 | 0.324 | 0.178 | 0.351 | 2 | <MRL | 1158.638 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
159 | 2023-01-21 16:16 | WO-23003-98-A-1 | U | 403.126 | #N/A | 0 | 0 | #N/A | 0.294 | 0.335 | #N/A | 0.211 | 1 | 403.126 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS4 | IS4 | [8 empty] | 242 | 2023-01-27 18:23 | 217425-d2 | U | 0 | 477.873 | 0 | 0 | 0 | 0.572 | 0.323 | 0.103 | 1.037 | 2 | <MRL | 955.746 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
160 | 2023-01-21 16:43 | WO-23003-99-A-1 | U | 447.416 | #N/A | 0 | 0 | #N/A | 0.262 | 0.309 | #N/A | 0.186 | 1 | 447.416 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 243 | 2023-01-27 18:50 | 217426-d2 | U | 0 | 521.5 | 1.181 | 0 | 0 | 0.4 | 0.266 | -0.223 | 0.27 | 2 | <MRL | 1043 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4
161 | 2023-01-21 17:10 | WO-23003-100-A-1 | U | 410.536 | #N/A | 0.321 | 0 | #N/A | 0.224 | 0.31 | #N/A | 0.224 | 1 | 410.536 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4 | [8 empty] | 244 | 2023-01-27 19:17 | 217427-d2 | U | 0 | 1340.929 | 0 | 0 | 0 | 0.36 | 0.347 | -0.135 | 0.404 | 2 | <MRL | 2681.858 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
162 | 2023-01-21 17:37 | AB-230117-02 LRB 11 | K | 0 | #N/A | 0.345 | 0.756 | #N/A | 0.55 | 0.267 | #N/A | 0.264 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 245 | 2023-01-27 19:44 | 217428-d2 | U | 0 | 2145.145 | 1.097 | 0 | 0.884 | 0.404 | 1.635 | -0.177 | 0.293 | 2 | <MRL | 4290.29 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
163 | 2023-01-21 18:04 | LD-5-inj1 | D | 0 | #N/A | 0 | 0.802 | #N/A | 0.657 | 0.477 | #N/A | 82.9 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 82.9 |  | IS3 |  | IS4 | IS4 |  |  |  |  | D |  | Duplicate failure. Reprepared below and passed. |  | 246 | 2023-01-27 20:11 | 217429-d2 | U | 0 | 2461.423 | 1.192 | 0 | 0.79 | 0.326 | 0.374 | 0 | 0.263 | 2 | <MRL | 4922.846 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
164 | 2023-01-21 18:31 | AB-230117-02 LD 5 | D | 0 | #N/A | 0 | 0.77 | #N/A | 0.645 | 0.438 | #N/A | 88.093 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 88.093 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | D |  | Duplicate failure. Reprepared below and passed. |  | 247 | 2023-01-27 20:38 | 217430-d2 | U | 0 | 1276.879 | 0 | 0 | 0 | 0.357 | 0.26 | -0.051 | 0.25 | 2 | <MRL | 2553.758 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 |  | IS3 | IS4
165 | 2023-01-21 18:58 | AB-230117-02 LFB-H 6 | Z | 158.298 | #N/A | 179.368 | 175.354 | #N/A | 173.631 | 176.742 | #N/A | 176.054 | 1 | 158.298 | #N/A | 179.368 | 175.354 | #N/A | 173.631 | 176.742 | #N/A | 176.054 | No TrueV! | [12 empty] | 248 | 2023-01-27 21:05 | 217431-d2 | U | 0 | 1266.582 | 0.409 | 0 | 0 | 0.38 | 0.286 | 0.041 | 0.251 | 2 | <MRL | 2533.164 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4
166 | 2023-01-21 19:25 | AB-230117-02 CCC-H 6 | Q | 304.261 | #N/A | 352.766 | 339.3 | #N/A | 333.08 | 341.007 | #N/A | 341.563 | 1 | 304.261 | #N/A | 352.766 | 339.3 | #N/A | 333.08 | 341.007 | #N/A | 341.563 | [13 empty] | 249 | 2023-01-27 21:32 | 217432-d2 | U | 0 | 1326.789 | 0.863 | 0 | 0.748 | 0.308 | 0.267 | -0.232 | 1.766 | 2 | <MRL | 2653.578 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 |  | IS3 | IS4
167 | 2023-01-21 19:52 | AB-230117-02 LFSM-H 5 | S | 182.099 | #N/A | 178.008 | 176.535 | #N/A | 182.066 | 186.755 | #N/A | 223.297 | 1 | 182.099 | #N/A | 178.008 | 176.535 | #N/A | 182.066 | 186.755 | #N/A | 223.297 | No TrueV! | IS3 | [11 empty] | 250 | 2023-01-27 21:59 | 217433-d2 | U | 0 | 4651.337 | 0 | 0 | 0 | 0.293 | 0.325 | -0.125 | 0.343 | 2 | <MRL | 9302.674 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
168 | 2023-01-21 20:47 | WO-23003-101-A-1 | U | 0 | #N/A | 0 | 0.912 | #N/A | 0.618 | 0.649 | #N/A | 28.256 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 28.256 |  | IS3 |  | IS4 | IS4 | [8 empty] | 251 | 2023-01-27 22:26 | 217434-d2 | U | 0 | 4609.936 | 0 | 1.797 | 0.834 | 0.361 | 0.313 | -0.103 | 0.487 | 2 | <MRL | 9219.872 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
169 | 2023-01-21 21:14 | WO-23003-102-A-1 | U | 0 | #N/A | 0.788 | 1.032 | #N/A | 0.732 | 0.666 | #N/A | 64.801 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 64.801 |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 258 | 2023-01-28 02:03 | 217435-d2 | U | 0 | 4700.929 | 0.794 | 0 | 0.954 | 0.552 | 0.561 | 0.212 | 0.674 | 2 | <MRL | 9401.858 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
170 | 2023-01-21 21:41 | WO-23003-103-A-1 | U | 0 | #N/A | 0 | 0.975 | #N/A | 0.702 | 0.525 | #N/A | 66.78 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 66.78 |  | IS3 | IS4 | IS4 | IS4 | IS4 | [7 empty] | 259 | 2023-01-28 02:30 | 217436-d2 | U | 0 | 2959.884 | 0.801 | 0 | 0.834 | 0.454 | 0.462 | 0.068 | 0.632 | 2 | <MRL | 5919.768 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
171 | 2023-01-21 22:08 | WO-23003-104-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.713 | 0.618 | #N/A | 128.717 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 128.717 |  | IS3 | [11 empty] | 260 | 2023-01-28 02:57 | 217437-d2 | U | 0 | 2697.215 | 1.814 | 0 | 0.846 | 0.449 | 0.346 | -0.04 | 0.298 | 2 | <MRL | 5394.43 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4
172 | 2023-01-21 22:35 | WO-23003-105-A-1 | U | 0 | #N/A | 0.4 | 0 | #N/A | 0.819 | 0.695 | #N/A | 128.03 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 128.03 |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 261 | 2023-01-28 03:24 | 217438-d2 | U | 0 | 2884.658 | 0.989 | 0 | 0.862 | 0.485 | 0.339 | -0.012 | 0.303 | 2 | <MRL | 5769.316 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
173 | 2023-01-21 23:02 | WO-23003-106-A-1 | U | 0 | #N/A | 0.23 | 0 | #N/A | 0.742 | 0.515 | #N/A | 137.351 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 137.351 |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 262 | 2023-01-28 03:51 | 217439-d2 | U | 0 | 8091.743 | 1.082 | 0.789 | 0.882 | 0.724 | 0.482 | 0.574 | 0.497 | 2 | <MRL | 16183.486 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
174 | 2023-01-21 23:29 | WO-23003-107-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.668 | 0.464 | #N/A | 141.587 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 141.587 |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 263 | 2023-01-28 04:18 | 217440-d2 | U | 0 | 8146.394 | 0 | 1.19 | 0.995 | 0.392 | 0.504 | 0.507 | 0.628 | 2 | <MRL | 16292.788 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4
175 | 2023-01-21 23:56 | WO-23003-108-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.504 | 0.484 | #N/A | 117.876 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 117.876 |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | [6 empty] | 264 | 2023-01-28 04:45 | 217441-d2 | U | 0 | 8310.781 | 0 | 1.668 | 0.912 | 0.488 | 0.566 | 0.567 | 0.71 | 2 | <MRL | 16621.562 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
176 | 2023-01-22 00:23 | WO-23003-109-A-1 | U | 0 | #N/A | 0 | 0.53 | #N/A | 0.575 | 0.506 | #N/A | 137.434 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 137.434 |  | IS3 |  | IS4 | IS4 | [8 empty] | 265 | 2023-01-28 05:12 | 217442-d2 | U | 0 | 7974.627 | 0.745 | 0.889 | 0 | 0.503 | 0.536 | 0.527 | 0.645 | 2 | <MRL | 15949.254 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4
177 | 2023-01-22 00:50 | WO-23003-110-A-1 | U | 0 | #N/A | 0 | 1.761 | #N/A | 3.639 | 2.107 | #N/A | 666.017 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 666.017 |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 266 | 2023-01-28 05:40 | 217443-d2 | U | 0 | 0 | 0.498 | 0 | 0.962 | 0.365 | 0.271 | 4.807 | 0.257 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
178 | 2023-01-22 01:17 | LRB-12-inj1 | K | 0 | #N/A | 0 | 0.799 | #N/A | 0.648 | 0.302 | #N/A | 24.998 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 24.998 | [9 empty] | b |  |  |  | 267 | 2023-01-28 06:07 | 217444-d2 | U | 0 | 0 | 0.735 | 0 | 0.689 | 0.311 | 0.245 | 3.409 | 0.221 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS4 | IS4
179 | 2023-01-22 01:44 | AB-230117-02 LRB 12 | K | 0 | #N/A | 0 | 0.758 | #N/A | 0.534 | 0.278 | #N/A | 1.003 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 270 | 2023-01-28 07:28 | 217445-d2 | U | 0 | 0 | 0 | 0 | 0 | 0.386 | 0.273 | 20.739 | 0.273 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 41.478 | <MRL |  | IS3 |  | IS4 | IS4
180 | 2023-01-22 02:12 | WO-23003-111-A-1 | U | 0 | #N/A | 0 | 1.35 | #N/A | 4.279 | 2.348 | #N/A | 740.764 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 740.764 |  | IS3 | IS4 | IS4 | IS4 | [5 empty] | R1 | Electively rerun below. |  | 271 | 2023-01-28 07:55 | 217446-d2 | U | 0 | 0 | 0.751 | 0 | 0.774 | 0.17 | 0.29 | 20.708 | 0.282 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 41.416 | <MRL |  | IS3 |  | IS4 | IS4
181 | 2023-01-22 02:39 | WO-23003-112-A-1 | U | 0 | #N/A | 0 | 1.139 | #N/A | 2.811 | 1.711 | #N/A | 536.317 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 536.317 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b |  |  |  | 272 | 2023-01-28 08:22 | 217447-d2 | U | 0 | 0 | 0.325 | 0 | 0.887 | 0.354 | 0.305 | 5.228 | 0.267 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 10.456 | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
182 | 2023-01-22 03:06 | WO-23003-113-A-1 | U | 0 | #N/A | 0 | 0.765 | #N/A | 2.136 | 1.514 | #N/A | 641.876 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 641.876 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b |  |  |  | 273 | 2023-01-28 08:49 | 217448-d2 | U | 0 | 0 | 0.845 | 0 | 0.855 | 0.338 | 0.271 | 4.496 | 0.229 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS4 | IS4
183 | 2023-01-22 03:33 | WO-23003-114-A-1 | U | 0 | #N/A | 0 | 1.029 | #N/A | 2.034 | 1.569 | #N/A | 622.53 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 622.53 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | b |  |  |  | 274 | 2023-01-28 09:16 | 217449-d2 | U | 0 | 0 | 0.341 | 0 | 0.81 | 0.339 | 0.286 | 5.044 | 0.19 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 10.088 | <MRL |  | IS3 | IS4 | IS4 | IS4
184 | 2023-01-22 04:00 | WO-23003-115-A-1 | U | 0 | #N/A | 0.437 | 1.365 | #N/A | 2.371 | 1.895 | #N/A | 714.013 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 714.013 |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 275 | 2023-01-28 09:43 | 217450-d2 | U | 0 | 0 | 1.229 | 0 | 0 | 0.374 | 0.249 | 5.099 | 0.183 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 10.198 | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4
185 | 2023-01-22 04:27 | WO-23003-116-A-1 | U | 0 | #N/A | 0 | 1.71 | #N/A | 11.677 | 8.8 | #N/A | 2477.13 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 11.677 | 8.8 | #N/A | 2477.13 |  | IS3 | [7 empty] | IS4 |  |  |  | 276 | 2023-01-28 10:10 | 217451-d2 | U | 0 | 0 | 0 | 0 | 0.877 | 0.378 | 0.285 | 10.965 | 0.246 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 21.93 | <MRL |  | IS3 | IS4 | IS4 | IS4
186 | 2023-01-22 04:54 | WO-23003-117-A-1 | U | 0 | #N/A | 0.376 | 1.608 | #N/A | 10.943 | 8.243 | #N/A | 2511.685 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 10.943 | 8.243 | #N/A | 2511.685 |  | IS3 | [7 empty] | IS4 |  |  |  | 277 | 2023-01-28 10:38 | 217452-d2 | U | 0 | 0 | 0.198 | 0 | 0.914 | 0.166 | 0.258 | 8.837 | 0.206 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 17.674 | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
187 | 2023-01-22 05:21 | WO-23003-118-A-1 | U | 0 | #N/A | 0 | 1.751 | #N/A | 12.301 | 9.413 | #N/A | 2857.187 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 12.301 | 9.413 | #N/A | 2857.187 |  | IS3 | [7 empty] | IS4 |  |  |  | 278 | 2023-01-28 11:05 | 217453-d2 | U | 0 | 0 | 0.544 | 0 | 0 | 0.376 | 0 | 7.783 | 0.249 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 15.566 | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
188 | 2023-01-22 05:48 | WO-23003-119-A-1 | U | 0 | #N/A | 0 | 1.049 | #N/A | 3.789 | 2.851 | #N/A | 1085.45 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 1085.45 |  | IS3 | [8 empty] | R1 | Electively rerun below |  | 279 | 2023-01-28 11:32 | 217454-d2 | U | 0 | 0 | 0 | 0 | 0 | 0.341 | 0.282 | 15.624 | 0.228 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 31.248 | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
189 | 2023-01-22 06:15 | WO-23003-120-A-1 | U | 0 | #N/A | 0.615 | 1.294 | #N/A | 4.447 | 3.724 | #N/A | 1329.015 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 1329.015 |  | IS3 | [8 empty] | R1 | Electively rerun below |  | 286 | 2023-01-28 15:08 | 217455-d2 | U | 0 | 0 | 0.867 | 0 | 1.055 | 0.611 | 0.524 | 15.408 | 0.526 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 30.816 | <MRL |  | IS3 |  | IS4 | IS4
190 | 2023-01-22 06:42 | AB-230117-02 LRB 13 | K | 0 | #N/A | 0.263 | 0.748 | #N/A | 0.62 | 0.374 | #N/A | 69.087 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 69.087 | [9 empty] | b |  |  |  | 287 | 2023-01-28 15:35 | 217456-d2 | U | 0 | 0 | 0.34 | 0 | 0.971 | 0.464 | 0.391 | 14.04 | 0.316 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 28.08 | <MRL |  | IS3 |  | IS4 | IS4
191 | 2023-01-22 07:10 | LD6-inj1 | D | 0 | #N/A | 0 | 0.883 | #N/A | 1.362 | 2.893 | #N/A | 3922.707 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 3922.707 | [9 empty] | IS4 |  |  |  | 288 | 2023-01-28 16:03 | 217457-d2 | U | 0 | 15.276 | 0.59 | 0 | 1.663 | 0.39 | 0.295 | 43.91 | 0.325 | 2 | <MRL | 30.552 | <MRL | <MRL | <MRL | <MRL | <MRL | 87.82 | <MRL |  | IS3 |  | IS3 | IS4
192 | 2023-01-22 07:37 | AB-230117-02 LD 6 | D | 0 | #N/A | 0 | 0.865 | #N/A | 1.374 | 3.079 | #N/A | 4142.057 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4142.057 | [9 empty] | IS4 |  |  |  | 289 | 2023-01-28 16:30 | 217458-d2 | U | 0 | 13.196 | 0.63 | 0 | 1.546 | 0.413 | 0.302 | 41.182 | 0.249 | 2 | <MRL | 26.392 | <MRL | <MRL | <MRL | <MRL | <MRL | 82.364 | <MRL |  | IS3 |  | IS3 | IS4
193 | 2023-01-22 08:04 | AB-230117-02 LFB-H 7 | Z | 157.151 | #N/A | 181.009 | 182.079 | #N/A | 177.55 | 177.827 | #N/A | 303.883 | 1 | 157.151 | #N/A | 181.009 | 182.079 | #N/A | 177.55 | 177.827 | #N/A | 303.883 | No TrueV! | [12 empty] | 290 | 2023-01-28 16:57 | 217459-d2 | U | 0 | 12.554 | 0.51 | 0.431 | 1.402 | 0.377 | 0.368 | 36.731 | 0.241 | 2 | <MRL | 25.108 | <MRL | <MRL | <MRL | <MRL | <MRL | 73.462 | <MRL |  | IS3 |  | IS3 | IS4
194 | 2023-01-22 08:31 | AB-230117-02 CCC-H 7 | Q | 307.839 | #N/A | 372.196 | 355.298 | #N/A | 349.262 | 350.888 | #N/A | 358.688 | 1 | 307.839 | #N/A | 372.196 | 355.298 | #N/A | 349.262 | 350.888 | #N/A | 358.688 | [13 empty] | 291 | 2023-01-28 17:24 | 217460-d2 | U | 0 | 5.009 | 0.239 | 0 | 1.147 | 0.37 | 0.258 | 53.928 | 0.224 | 2 | <MRL | 10.018 | <MRL | <MRL | <MRL | <MRL | <MRL | 107.856 | <MRL |  | IS3 | IS4 | IS3 | IS4
195 | 2023-01-22 08:58 | AB-230117-02 LFSM-H 6 | S | 160.732 | #N/A | 171.837 | 176.045 | #N/A | 184.547 | 191.689 | #N/A | 2843.453 | 1 | 160.732 | #N/A | 171.837 | 176.045 | #N/A | 184.547 | 191.689 | #N/A | 2843.453 | No TrueV! | IS3 | [7 empty] | IS3 |  |  |  | 292 | 2023-01-28 17:51 | 217461-d2 | U | 0 | 4.362 | 0 | 0 | 1.404 | 0.27 | 0.235 | 52.99 | 0.23 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 105.98 | <MRL |  | IS3 | IS4 | IS4 | IS4
196 | 2023-01-22 09:52 | WO-23003-121-A-1 | U | 0 | #N/A | 0 | 1.724 | #N/A | 8.728 | 8.418 | #N/A | 2598.122 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 8.728 | 8.418 | #N/A | 2598.122 |  | IS3 | [7 empty] | IS4 | R1 | Electively rerun below |  | 293 | 2023-01-28 18:18 | 217462-d2 | U | 0 | 6.698 | 0 | 0 | 1.423 | 0.342 | 0.293 | 61.169 | 0.203 | 2 | <MRL | 13.396 | <MRL | <MRL | <MRL | <MRL | <MRL | 122.338 | <MRL |  | IS3 | IS4 | IS3 | IS4
197 | 2023-01-22 10:19 | WO-23003-122-A-1 | U | 0 | #N/A | 0 | 0.942 | #N/A | 1.505 | 3.298 | #N/A | 4052.461 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4052.461 | [9 empty] | IS4 |  |  |  | 294 | 2023-01-28 18:45 | 217463-d2 | U | 0 | 91.382 | 0.778 | 1.662 | 16.124 | 0.995 | 0.474 | 457.157 | 0.38 | 2 | <MRL | 182.764 | <MRL | <MRL | 32.248 | <MRL | <MRL | 914.314 | <MRL |  | IS3 |  | IS3 | IS4
198 | 2023-01-22 10:46 | WO-23003-123-A-1 | U | 0 | #N/A | 0 | 0.825 | #N/A | 1.448 | 3.799 | #N/A | 4114.272 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4114.272 | [9 empty] | IS4 |  |  |  | 295 | 2023-01-28 19:12 | 217464-d2 | U | 0 | 86.45 | 0 | 1.55 | 13.315 | 1.102 | 0.329 | 377.598 | 0.385 | 2 | <MRL | 172.9 | <MRL | <MRL | 26.63 | <MRL | <MRL | 755.196 | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
199 | 2023-01-22 11:13 | WO-23003-124-A-1 | U | 0 | #N/A | 0 | 0.721 | #N/A | 1.325 | 2.549 | #N/A | 4254.003 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4254.003 | [9 empty] | IS4 |  |  |  | 298 | 2023-01-28 20:33 | 217465-d2 | U | 0 | 77.507 | 1.379 | 1.531 | 12.512 | 0.847 | 0.292 | 345.638 | 0.168 | 2 | <MRL | 155.014 | <MRL | <MRL | 25.024 | <MRL | <MRL | 691.276 | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS3
200 | 2023-01-23 10:08 | Rinse12 | U | 0 | #N/A | 40.266 | 0 | #N/A | 11.109 | 6.109 | #N/A | 3907 | 1 | <MRL | #N/A | 40.266 | <MRL | #N/A | 11.109 | 6.109 | #N/A | 3907 |  | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 |  | These are conditioning rinses, not true QC samples |  | 299 | 2023-01-28 21:01 | 217466-d2 | U | 0 | 81.655 | 0 | 1.124 | 11.749 | 0.889 | 0.327 | 359.261 | 0.175 | 2 | <MRL | 163.31 | <MRL | <MRL | 23.498 | <MRL | <MRL | 718.522 | <MRL |  | IS3 | IS4 | IS3 | IS4
201 | 2023-01-23 10:35 | AB-230117-02 LRB 12 20230123102918 | K | 0 | #N/A | 0 | 0 | #N/A | 49.489 | 26.569 | #N/A | 943.114 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 49.489 | 26.569 | #N/A | 943.114 |  | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 |  | These are conditioning rinses, not true QC samples |  | 300 | 2023-01-28 21:28 | 217467-d2 | U | 0 | 36.046 | 0.206 | 0.838 | 7.354 | 0.633 | 0.258 | 373.256 | 0.163 | 2 | <MRL | 72.092 | <MRL | <MRL | 14.708 | <MRL | <MRL | 746.512 | <MRL |  | IS3 | IS4 | IS3 | IS4
202 | 2023-01-23 11:02 | AB-230117-02 LRB 13 20230123105625 | K | 0 | #N/A | 0.169 | 0.805 | #N/A | 0.61 | 0.281 | #N/A | 3.33 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [9 empty] | b |  |  |  | 301 | 2023-01-28 21:55 | 217468-d2 | U | 0 | 48.031 | 1.246 | 0.838 | 9.528 | 0.905 | 0.277 | 504.393 | 0.209 | 2 | <MRL | 96.062 | <MRL | <MRL | 19.056 | <MRL | <MRL | 1008.786 | <MRL |  | IS3 |  | IS3 | IS4
203 | 2023-01-23 11:29 | WO-23003-111-A-1B | U | 0 | #N/A | 0 | 1.749 | #N/A | 4.534 | 2.374 | #N/A | 738.645 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 738.645 |  | IS3 | IS4 | IS4 | IS4 | [8 empty] | 302 | 2023-01-28 22:22 | 217469-d2 | U | 0 | 40.614 | 0.301 | 1.117 | 9.416 | 0.858 | 0 | 473.503 | 0.201 | 2 | <MRL | 81.228 | <MRL | <MRL | 18.832 | <MRL | <MRL | 947.006 | <MRL |  | IS3 | IS4 | IS3 | IS4
204 | 2023-01-23 11:56 | WO-23003-121-A-1B | U | 0 | #N/A | 0.285 | 1.486 | #N/A | 8.541 | 7.928 | #N/A | 2538.308 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 8.541 | 7.928 | #N/A | 2538.308 | [9 empty] | IS4 | R1 | Reprepared and rerun below |  | 303 | 2023-01-28 22:49 | 217470-d2 | U | 0 | 43.583 | 1.277 | 1.056 | 9.73 | 0.806 | 0.256 | 484.625 | 0.188 | 2 | <MRL | 87.166 | <MRL | <MRL | 19.46 | <MRL | <MRL | 969.25 | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
205 | 2023-01-23 12:24 | WO-23003-125-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.512 | 0.413 | #N/A | 46.972 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 46.972 |  | IS3 | IS4 | IS3 | IS4 |  |  |  |  | C | R1 | Rerun due to carryover concern |  | 304 | 2023-01-28 23:16 | 217471-d2 | U | 0 | 176.471 | 1.222 | 3.823 | 200.815 | 13.769 | 0 | 6886.659 | 0.24 | 2 | <MRL | 352.942 | <MRL | <MRL | 401.63 | 27.538 | <MRL | 13773.318 | <MRL |  | IS3 | IS4 | IS3 | IS4 |  | IS3 | IS3 | IS4
206 | 2023-01-23 12:51 | WO-23003-126-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0 | 0.412 | #N/A | 4.716 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  | R1 | Rerun due to carryover concern |  | 305 | 2023-01-28 23:43 | 217472-d2 | U | 0 | 178.451 | 0.694 | 4.452 | 193.513 | 14.717 | 0.435 | 6757.926 | 0.238 | 2 | <MRL | 356.902 | <MRL | <MRL | 387.026 | 29.434 | <MRL | 13515.852 | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS3 | IS3 | IS4
207 | 2023-01-23 13:18 | WO-23003-127-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.451 | 0.375 | #N/A | 1.887 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | [6 empty] | 306 | 2023-01-29 00:10 | 217473-d2 | U | 0 | 162.514 | 0.607 | 3.086 | 94.187 | 6.684 | 0.239 | 2990.294 | 0.248 | 2 | <MRL | 325.028 | <MRL | <MRL | 188.374 | 13.368 | <MRL | 5980.588 | <MRL |  | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | R2 | This one looks like it has twice the internal standard amount, cutting concentrations in half. I'll rerun it this weekend.
208 | 2023-01-23 13:45 | WO-23003-128-A-1 | U | 0 | #N/A | 0.672 | 0 | #N/A | 0.583 | 0.347 | #N/A | 1.569 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | [5 empty] | 307 | 2023-01-29 00:37 | 217474-d2 | U | 0 | 143.538 | 0 | 4.582 | 118.094 | 13.529 | 0 | 6059.278 | 0.206 | 2 | <MRL | 287.076 | <MRL | <MRL | 236.188 | 27.058 | <MRL | 12118.556 | <MRL |  | IS3 | IS4 | IS3 | IS4 |  |  | IS3 | IS4
209 | 2023-01-23 14:12 | WO-23003-129-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.451 | 0.345 | #N/A | 1.157 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | [6 empty] | 314 | 2023-01-29 04:14 | 217475-d2 | U | 0 | 152.553 | 1.429 | 3.72 | 123.987 | 11.225 | 0.809 | 6600.341 | 0.604 | 2 | <MRL | 305.106 | <MRL | <MRL | 247.974 | 22.45 | <MRL | 13200.682 | <MRL |  | IS3 | IS4 | IS3 | IS4 |  | IS3 | IS3 | IS4
210 | 2023-01-23 14:39 | WO-23003-130-A-1 | U | 0 | #N/A | 0.482 | 0 | #N/A | 0.481 | 0.296 | #N/A | 0.778 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 315 | 2023-01-29 04:41 | 217476-d2 | U | 0 | 150.15 | 0 | 4.087 | 127.028 | 11.115 | 0 | 6345.164 | 0.397 | 2 | <MRL | 300.3 | <MRL | <MRL | 254.056 | 22.23 | <MRL | 12690.328 | <MRL |  | IS3 | IS4 | IS3 | IS4 |  | IS3 | IS3 | IS4
211 | 2023-01-23 15:06 | LRB-14-inj1 | K | 0 | #N/A | 0 | 0.799 | #N/A | 0.486 | 0.247 | #N/A | 0.68 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 316 | 2023-01-29 05:08 | 217477-d4 | U | 0 | 206.316 | 1.134 | 6.695 | 359.206 | 39.631 | 0 | 20466.741 | 0.358 | 4 | <MRL | 825.264 | <MRL | 26.78 | 1436.824 | 158.524 | <MRL | 81866.964 | <MRL | UL! | IS3 | IS4 | IS3 | IS4 | IS3 | IS3 | IS3 | UL |  | R1
212 | 2023-01-23 15:33 | AB-230117-02 LRB 14 | K | 0 | #N/A | 0 | 0.844 | #N/A | 0.577 | 0.274 | #N/A | 0.664 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 317 | 2023-01-29 05:35 | 217478-d4 | U | 0 | 208.811 | 0.521 | 5.247 | 369.784 | 29.505 | 0 | 18254.898 | 0.31 | 4 | <MRL | 835.244 | <MRL | 20.988 | 1479.136 | 118.02 | <MRL | 73019.592 | <MRL | UL! | IS3 | IS4 | IS3 | IS4 | IS3 | IS3 | IS3 | UL |  | R1
213 | 2023-01-23 16:00 | WO-23003-131-A-1 | U | 0 | #N/A | 0.329 | 0 | #N/A | 0.366 | 0.306 | #N/A | 0.748 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 318 | 2023-01-29 06:02 | 217479-d4 | U | 0 | 95.979 | 0.479 | 5.258 | 192.84 | 35.231 | 0 | 19784.952 | 0.308 | 4 | <MRL | 383.916 | <MRL | 21.032 | 771.36 | 140.924 | <MRL | 79139.808 | <MRL | UL! | IS3 |  |  |  |  | IS3 | IS3 | UL |  | R1
214 | 2023-01-23 16:27 | WO-23003-132-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 1.113 | 1.47 | #N/A | 0.77 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | [7 empty] | 319 | 2023-01-29 06:29 | 217480-d4 | U | 0 | 96.687 | 0 | 4.559 | 165.874 | 26.636 | 0 | 13668.116 | 0.258 | 4 | <MRL | 386.748 | <MRL | <MRL | 663.496 | 106.544 | <MRL | 54672.464 | <MRL | UL! | IS3 |  |  |  |  | IS3 | IS3 | UL |  | R1
215 | 2023-01-23 16:54 | WO-23003-133-A-1 | U | 0 | #N/A | 1.203 | 0 | #N/A | 0.367 | 0.295 | #N/A | 0.523 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | [6 empty] | 320 | 2023-01-29 06:56 | 217481-d4 | U | 0 | 88.49 | 0.392 | 4.923 | 163.904 | 24.868 | 0 | 13164.028 | 0.26 | 4 | <MRL | 353.96 | <MRL | <MRL | 655.616 | 99.472 | <MRL | 52656.112 | <MRL | UL! | IS3 |  |  |  |  | IS3 | IS3 | UL |  | R1
216 | 2023-01-23 17:21 | WO-23003-134-A-1 | U | 0 | #N/A | 0.198 | 0 | #N/A | 0.371 | 0.261 | #N/A | 0.414 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 321 | 2023-01-29 07:23 | 217482-d4 | U | 0 | 95.224 | 0.264 | 4.472 | 166.268 | 25.391 | 0 | 14233.119 | 0.31 | 4 | <MRL | 380.896 | <MRL | <MRL | 665.072 | 101.564 | <MRL | 56932.476 | <MRL | UL! | [5 empty] | IS3 | IS3 | UL |  | R1
217 | 2023-01-23 17:49 | WO-23003-135-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.412 | 0.333 | #N/A | 0.443 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | [7 empty] | 322 | 2023-01-29 07:50 | 217483-d200 | U | 0 | 3.382 | 0.152 | 0.762 | 8.704 | 1.497 | 0 | 7113.022 | 0.289 | 200 | <MRL | <MRL | <MRL | <MRL | 1740.8 | <MRL | <MRL | 1422604.4 | <MRL | [7 empty] | IS4 | IS4
218 | 2023-01-23 18:16 | WO-23003-136-A-1 | U | 0 | #N/A | 1.066 | 0 | #N/A | 0.251 | 0.251 | #N/A | 0.437 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | [6 empty] | 323 | 2023-01-29 08:18 | 217484-d200 | U | 0 | 3.07 | 0 | 0.916 | 5.436 | 1.074 | 0 | 6424.609 | 0.265 | 200 | <MRL | <MRL | <MRL | <MRL | 1087.2 | <MRL | <MRL | 1284921.8 | <MRL | [7 empty] | IS4 | IS4
219 | 2023-01-23 18:43 | WO-23003-137-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.319 | 0.258 | #N/A | 0.386 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | [11 empty] | 325 | 2023-01-29 09:12 | 217485-d200 | U | 0 | 2.35 | 0.224 | 0.803 | 5.409 | 1.148 | 0 | 6474.482 | 0.284 | 200 | <MRL | <MRL | <MRL | <MRL | 1081.8 | <MRL | <MRL | 1294896.4 | <MRL | [7 empty] | IS4 | IS4
220 | 2023-01-23 19:10 | WO-23003-138-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.209 | 0.284 | #N/A | 0.368 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS3 | IS4 | [8 empty] | 330 | 2023-01-30 12:09 | 217479-d20 | U | 0 | 10.448 | 0.231 | 0.868 | 17.597 | 4.491 | 0.326 | 1809.619 | 0.378 | 20 | <MRL | 208.96 | <MRL | <MRL | 351.94 | <MRL | <MRL | 36192.38 | <MRL | [8 empty] | IS4
221 | 2023-01-27 08:27 | WO-23003-139-A-1 | U | 0 | #N/A | 3.112 | 0 | #N/A | 0.695 | 0.57 | #N/A | 133.922 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 133.922 |  | IS3 | IS4 | IS3 | IS4 |  |  |  |  | C | R1 | Analyzed after instrument leak, some carryover observed and sample rerun |  | 331 | 2023-01-30 12:36 | 217478-d10 | U | 0 | 44.764 | 0.313 | 1.758 | 73.237 | 9.339 | 0 | 5298.356 | 0.315 | 10 | <MRL | 447.64 | <MRL | <MRL | 732.37 | 93.39 | <MRL | 52983.56 | <MRL | [7 empty] | IS3 | IS4
222 | 2023-01-27 08:54 | WO-23003-139-A-1B | U | 0 | #N/A | 2.591 | 0 | #N/A | 0.378 | 0.345 | #N/A | 6.248 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 6.248 |  | IS3 | IS4 | IS3 | IS4 |  |  |  |  | C |  |  |  | 332 | 2023-01-30 13:03 | 217477-d10 | U | 0 | 44.219 | 0.35 | 1.411 | 74.136 | 7.698 | 0 | 4923.452 | 0.343 | 10 | <MRL | 442.19 | <MRL | <MRL | 741.36 | 76.98 | <MRL | 49234.52 | <MRL | [7 empty] | IS3 | IS4
223 | 2023-01-27 09:21 | WO-23003-140-A-1 | U | 0 | #N/A | 0.914 | 0 | #N/A | 0.357 | 0.338 | #N/A | 2.01 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 333 | 2023-01-30 13:30 | 217480-d20 | U | 0 | 11.594 | 0.356 | 0.794 | 20.732 | 4.89 | 0.362 | 2969.033 | 0.347 | 20 | <MRL | 231.88 | <MRL | <MRL | 414.64 | <MRL | <MRL | 59380.66 | <MRL | [7 empty] | IS4 | IS4
224 | 2023-01-27 09:48 | LRB-15-inj1 | K | 0 | #N/A | 0.374 | 0.91 | #N/A | 0.603 | 0.31 | #N/A | 1.417 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 334 | 2023-01-30 13:58 | 217481-d20 | U | 0 | 11.126 | 0 | 0.815 | 18.417 | 3.98 | 0.366 | 1989.268 | 0.34 | 20 | <MRL | 222.52 | <MRL | <MRL | 368.34 | <MRL | <MRL | 39785.36 | <MRL | [8 empty] | IS4
225 | 2023-01-27 10:15 | AB-230117-02 LRB 15 | K | 0 | #N/A | 0.658 | 0.9 | #N/A | 0.525 | 0.302 | #N/A | 1.286 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [13 empty] | 335 | 2023-01-30 14:25 | 217482-d20 | U | 0 | 10.574 | 0.718 | 0.753 | 17.222 | 3.878 | 0.294 | 1821.511 | 0.316 | 20 | <MRL | 211.48 | <MRL | <MRL | 344.44 | <MRL | <MRL | 36430.22 | <MRL | [8 empty] | IS4
226 | 2023-01-27 10:42 | AB-230117-02 LD 7 | D | 0 | #N/A | 0 | 1.76 | #N/A | 0.417 | 0.336 | #N/A | 1.197 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | [7 empty] | 336 | 2023-01-30 15:56 | 217288-d4-reprep | U | 0 | 4342.461 | 0.593 | 0.768 | 0.858 | 0.739 | 0.407 | 4.749 | 0.577 | 4 | <MRL | 17369.844 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3
227 | 2023-01-27 11:09 | AB-230117-02 LFB-H 8 | Z | 149.916 | #N/A | 178.569 | 174.977 | #N/A | 174.263 | 174.528 | #N/A | 172.643 | 1 | 149.916 | #N/A | 178.569 | 174.977 | #N/A | 174.263 | 174.528 | #N/A | 172.643 | No TrueV! | [12 empty] | 337 | 2023-01-30 16:23 | 217289-d4-reprep | U | 0 | 4355.518 | 0 | 0.642 | 0.848 | 0.655 | 0.343 | 2.443 | 0.368 | 4 | <MRL | 17422.072 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3
228 | 2023-01-27 11:36 | AB-230117-02 CCC-H 8 | Q | 306.068 | #N/A | 374.743 | 347.778 | #N/A | 345.655 | 353.314 | #N/A | 346.394 | 1 | 306.068 | #N/A | 374.743 | 347.778 | #N/A | 345.655 | 353.314 | #N/A | 346.394 | [13 empty] | 338 | 2023-01-30 16:50 | 217315-d2-reprep | U | 137.147 | 0 | 0 | 0 | 0.802 | 0.537 | 0.321 | 1.534 | 0.228 | 2 | 274.294 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS3 | IS4 | IS4
229 | 2023-01-27 12:04 | AB-230117-02 LFSM-H 7 | S | 188.826 | #N/A | 175.013 | 173.568 | #N/A | 182.658 | 188.523 | #N/A | 188.842 | 1 | 188.826 | #N/A | 175.013 | 173.568 | #N/A | 182.658 | 188.523 | #N/A | 188.842 | No TrueV! | IS3 | IS3 | IS3 | IS3 | IS3 | [5 empty] | PFDA okay vs second injection, as is PFBS |  | 339 | 2023-01-30 17:17 | 217356-d2-reprep | U | 0 | 0 | 0.299 | 0.477 | 0.792 | 0.546 | 0.326 | 1.099 | 4.268 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS4 | IS4
230 | 2023-01-27 12:58 | WO-23003-141-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.824 | 0.724 | #N/A | 0.925 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | [6 empty] | 340 | 2023-01-30 17:44 | 217357-d2-reprep | U | 0 | 0 | 0.835 | 0 | 0.82 | 0.607 | 0.929 | 1.143 | 48.327 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 96.654 |  | IS3 | IS4 | IS4 | IS4
231 | 2023-01-27 13:25 | WO-23003-142-A-1 | U | 0 | #N/A | 0 | 1.467 | #N/A | 0.453 | 0.486 | #N/A | 0.92 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 341 | 2023-01-30 18:12 | 217358-d2-reprep | U | 0 | 0 | 0 | 0 | 0.917 | 0.733 | 1.195 | 0.749 | 2575.277 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 5150.554 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | IS4
232 | 2023-01-27 13:52 | WO-23003-143-A-1 | U | 0 | #N/A | 1.526 | 0 | #N/A | 0.43 | 0.363 | #N/A | 0.671 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 342 | 2023-01-30 18:39 | 217359-d2-reprep | U | 0 | 0 | 0 | 0 | 0.815 | 1.217 | 3.011 | 0.671 | 9903.264 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 19806.528 |  | IS3 |  | IS4 | IS4 |  |  |  |  | IS4
233 | 2023-01-27 14:19 | WO-23003-144-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.442 | 0.337 | #N/A | 0.465 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 343 | 2023-01-30 19:06 | 217360-d2-reprep | U | 0 | 0 | 0 | 0 | 0.8 | 0.642 | 0.345 | 0.63 | 378.052 | 2 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 756.104 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | C | R1 | Electively rerun, but this has carrover, original sample did not. Retain original results
234 | 2023-01-27 14:46 | WO-23003-145-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.406 | 0.332 | #N/A | 0.419 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 344 | 2023-01-30 19:33 | 217393-d4-reprep | U | 0 | 0 | 0 | 0.775 | 0.771 | 3.747 | 2.949 | 0.47 | 1061.939 | 4 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 4247.756 |  | IS3
235 | 2023-01-27 15:13 | WO-23003-146-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.535 | 0.324 | #N/A | 0.408 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 345 | 2023-01-30 20:00 | 217394-d4-reprep | U | 0 | 0 | 0.511 | 1.218 | 0.752 | 4.579 | 4.034 | 0.294 | 1414.578 | 4 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | 5658.312 |  | IS3
236 | 2023-01-27 15:40 | WO-23003-147-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.467 | 0.329 | #N/A | 0.366 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 346 | 2023-01-30 21:21 | 217395-d4-reprep | U | 0 | 0 | 0.132 | 1.534 | 0.747 | 9.032 | 8.354 | 0.241 | 2714.735 | 4 | <MRL | <MRL | <MRL | <MRL | <MRL | 36.128 | 33.416 | <MRL | 10858.94 |  | IS3 | [7 empty] | IS4
237 | 2023-01-27 16:07 | WO-23003-148-A-1 | U | 0 | #N/A | 1.373 | 0 | #N/A | 0.388 | 0.358 | #N/A | 0.359 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 347 | 2023-01-30 22:15 | 217399-d2-reprep | U | 0 | 11.758 | 0 | 0 | 0 | 0.679 | 0.312 | 0.353 | 4.311 | 2 | <MRL | 23.516 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 | IS4 | IS3 | IS4
238 | 2023-01-27 16:34 | WO-23003-149-A-1 | U | 0 | #N/A | 1.108 | 0 | #N/A | 0.237 | 0.346 | #N/A | 0.365 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | [8 empty] | 348 | 2023-01-30 22:42 | 217400-d2-reprep | U | 0 | 11.796 | 0 | 0 | 0 | 0.707 | 0.321 | 0.213 | 1.945 | 2 | <MRL | 23.592 | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL | <MRL |  | IS3 |  | IS3 | IS4
239 | 2023-01-27 17:01 | WO-23003-150-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.429 | 0.324 | #N/A | 0.351 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
240 | 2023-01-27 17:28 | LRB-16-inj1 | K | 0 | #N/A | 0.811 | 0.741 | #N/A | 0.516 | 0.273 | #N/A | 0.375 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL
241 | 2023-01-27 17:56 | AB-230117-02 LRB 16 | K | 0 | #N/A | 0 | 0.749 | #N/A | 0.617 | 0.266 | #N/A | 0.431 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL
242 | 2023-01-27 18:23 | WO-23003-151-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.572 | 0.323 | #N/A | 1.037 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
243 | 2023-01-27 18:50 | WO-23003-152-A-1 | U | 0 | #N/A | 1.181 | 0 | #N/A | 0.4 | 0.266 | #N/A | 0.27 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4
244 | 2023-01-27 19:17 | WO-23003-153-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.36 | 0.347 | #N/A | 0.404 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4
245 | 2023-01-27 19:44 | WO-23003-154-A-1 | U | 0 | #N/A | 1.097 | 0 | #N/A | 0.404 | 1.635 | #N/A | 0.293 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
246 | 2023-01-27 20:11 | WO-23003-155-A-1 | U | 0 | #N/A | 1.192 | 0 | #N/A | 0.326 | 0.374 | #N/A | 0.263 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4
247 | 2023-01-27 20:38 | WO-23003-156-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.357 | 0.26 | #N/A | 0.25 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS3 | IS4
248 | 2023-01-27 21:05 | WO-23003-157-A-1 | U | 0 | #N/A | 0.409 | 0 | #N/A | 0.38 | 0.286 | #N/A | 0.251 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4
249 | 2023-01-27 21:32 | WO-23003-158-A-1 | U | 0 | #N/A | 0.863 | 0 | #N/A | 0.308 | 0.267 | #N/A | 1.766 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS3 | IS4
250 | 2023-01-27 21:59 | WO-23003-159-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.293 | 0.325 | #N/A | 0.343 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
251 | 2023-01-27 22:26 | WO-23003-160-A-1 | U | 0 | #N/A | 0 | 1.797 | #N/A | 0.361 | 0.313 | #N/A | 0.487 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4
252 | 2023-01-27 22:53 | LRB-17-inj1 | K | 0 | #N/A | 0 | 0.786 | #N/A | 0.588 | 0.256 | #N/A | 0.322 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL
253 | 2023-01-27 23:20 | AB-230117-02 LRB 17 | K | 0 | #N/A | 0 | 0.889 | #N/A | 0.626 | 0.23 | #N/A | 0.331 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL
254 | 2023-01-27 23:48 | AB-230117-02 LD 8 | D | 0 | #N/A | 0 | 0 | #N/A | 0.414 | 0.38 | #N/A | 0.49 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4
255 | 2023-01-28 00:15 | AB-230117-02 LFB-H 9 | Z | 160.935 | #N/A | 183.084 | 181.434 | #N/A | 178.355 | 179.694 | #N/A | 177.434 | 1 | 160.935 | #N/A | 183.084 | 181.434 | #N/A | 178.355 | 179.694 | #N/A | 177.434 | No TrueV! |  | IS1
256 | 2023-01-28 00:42 | AB-230117-02 CCC-H 9 | Q | 302.616 | #N/A | 351.796 | 350.18 | #N/A | 337.277 | 341.66 | #N/A | 345.726 | 1 | 302.616 | #N/A | 351.796 | 350.18 | #N/A | 337.277 | 341.66 | #N/A | 345.726
257 | 2023-01-28 01:09 | AB-230117-02 LFSM-H 8 | S | 177.244 | #N/A | 183.28 | 170.633 | #N/A | 178.895 | 185.896 | #N/A | 191.129 | 1 | 177.244 | #N/A | 183.28 | 170.633 | #N/A | 178.895 | 185.896 | #N/A | 191.129 | No TrueV! | IS3 | IS3 | IS3 | IS3
258 | 2023-01-28 02:03 | WO-23003-161-A-1 | U | 0 | #N/A | 0.794 | 0 | #N/A | 0.552 | 0.561 | #N/A | 0.674 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4
259 | 2023-01-28 02:30 | WO-23003-162-A-1 | U | 0 | #N/A | 0.801 | 0 | #N/A | 0.454 | 0.462 | #N/A | 0.632 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4
260 | 2023-01-28 02:57 | WO-23003-163-A-1 | U | 0 | #N/A | 1.814 | 0 | #N/A | 0.449 | 0.346 | #N/A | 0.298 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS4
261 | 2023-01-28 03:24 | WO-23003-164-A-1 | U | 0 | #N/A | 0.989 | 0 | #N/A | 0.485 | 0.339 | #N/A | 0.303 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
262 | 2023-01-28 03:51 | WO-23003-165-A-1 | U | 0 | #N/A | 1.082 | 0.789 | #N/A | 0.724 | 0.482 | #N/A | 0.497 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4
263 | 2023-01-28 04:18 | WO-23003-166-A-1 | U | 0 | #N/A | 0 | 1.19 | #N/A | 0.392 | 0.504 | #N/A | 0.628 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4
264 | 2023-01-28 04:45 | WO-23003-167-A-1 | U | 0 | #N/A | 0 | 1.668 | #N/A | 0.488 | 0.566 | #N/A | 0.71 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4
265 | 2023-01-28 05:12 | WO-23003-168-A-1 | U | 0 | #N/A | 0.745 | 0.889 | #N/A | 0.503 | 0.536 | #N/A | 0.645 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4
266 | 2023-01-28 05:40 | WO-23003-169-A-1 | U | 0 | #N/A | 0.498 | 0 | #N/A | 0.365 | 0.271 | #N/A | 0.257 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
267 | 2023-01-28 06:07 | WO-23003-170-A-1 | U | 0 | #N/A | 0.735 | 0 | #N/A | 0.311 | 0.245 | #N/A | 0.221 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4
268 | 2023-01-28 06:34 | LRB-18-inj1 | K | 0 | #N/A | 0 | 0.716 | #N/A | 0.512 | 0.231 | #N/A | 0.287 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL
269 | 2023-01-28 07:01 | AB-230117-02 LRB 18 | K | 0 | #N/A | 0 | 0.712 | #N/A | 0.506 | 0.239 | #N/A | 0.309 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL
270 | 2023-01-28 07:28 | WO-23003-171-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.386 | 0.273 | #N/A | 0.273 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS4 | IS4
271 | 2023-01-28 07:55 | WO-23003-172-A-1 | U | 0 | #N/A | 0.751 | 0 | #N/A | 0.17 | 0.29 | #N/A | 0.282 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS4 | IS4
272 | 2023-01-28 08:22 | WO-23003-173-A-1 | U | 0 | #N/A | 0.325 | 0 | #N/A | 0.354 | 0.305 | #N/A | 0.267 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
273 | 2023-01-28 08:49 | WO-23003-174-A-1 | U | 0 | #N/A | 0.845 | 0 | #N/A | 0.338 | 0.271 | #N/A | 0.229 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4
274 | 2023-01-28 09:16 | WO-23003-175-A-1 | U | 0 | #N/A | 0.341 | 0 | #N/A | 0.339 | 0.286 | #N/A | 0.19 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4
275 | 2023-01-28 09:43 | WO-23003-176-A-1 | U | 0 | #N/A | 1.229 | 0 | #N/A | 0.374 | 0.249 | #N/A | 0.183 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4
276 | 2023-01-28 10:10 | WO-23003-177-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.378 | 0.285 | #N/A | 0.246 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4
277 | 2023-01-28 10:38 | WO-23003-178-A-1 | U | 0 | #N/A | 0.198 | 0 | #N/A | 0.166 | 0.258 | #N/A | 0.206 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
278 | 2023-01-28 11:05 | WO-23003-179-A-1 | U | 0 | #N/A | 0.544 | 0 | #N/A | 0.376 | 0 | #N/A | 0.249 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4
279 | 2023-01-28 11:32 | WO-23003-180-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.341 | 0.282 | #N/A | 0.228 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4 | IS4 | IS4
280 | 2023-01-28 11:59 | LRB-19-inj1 | K | 0 | #N/A | 0.179 | 0.906 | #N/A | 0.519 | 0.249 | #N/A | 0.281 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL
281 | 2023-01-28 12:26 | AB-230117-02 LRB 19 | K | 0 | #N/A | 0 | 0.827 | #N/A | 0.512 | 0.277 | #N/A | 0.287 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL
282 | 2023-01-28 12:53 | AB-230117-02 LD 9 | D | 0 | #N/A | 0.632 | 0 | #N/A | 0.32 | 0.269 | #N/A | 0.236 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4
283 | 2023-01-28 13:20 | AB-230117-02 LFB-H 10 | Z | 151.291 | #N/A | 180.078 | 170.015 | #N/A | 171.506 | 172.512 | #N/A | 173.244 | 1 | 151.291 | #N/A | 180.078 | 170.015 | #N/A | 171.506 | 172.512 | #N/A | 173.244 | No TrueV!
284 | 2023-01-28 13:47 | AB-230117-02 CCC-H 10 | Q | 301.419 | #N/A | 356.398 | 350.601 | #N/A | 341.141 | 348.019 | #N/A | 342.023 | 1 | 301.419 | #N/A | 356.398 | 350.601 | #N/A | 341.141 | 348.019 | #N/A | 342.023
285 | 2023-01-28 14:14 | AB-230117-02 LFSM-H 9 | S | 168.758 | #N/A | 184.196 | 179.695 | #N/A | 182.596 | 190.975 | #N/A | 189.486 | 1 | 168.758 | #N/A | 184.196 | 179.695 | #N/A | 182.596 | 190.975 | #N/A | 189.486 | No TrueV! | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
286 | 2023-01-28 15:08 | WO-23003-181-A-1 | U | 0 | #N/A | 0.867 | 0 | #N/A | 0.611 | 0.524 | #N/A | 0.526 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS4 | IS4
287 | 2023-01-28 15:35 | WO-23003-182-A-1 | U | 0 | #N/A | 0.34 | 0 | #N/A | 0.464 | 0.391 | #N/A | 0.316 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS4 | IS4
288 | 2023-01-28 16:03 | WO-23003-183-A-1 | U | 0 | #N/A | 0.59 | 0 | #N/A | 0.39 | 0.295 | #N/A | 0.325 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS3 | IS4
289 | 2023-01-28 16:30 | WO-23003-184-A-1 | U | 0 | #N/A | 0.63 | 0 | #N/A | 0.413 | 0.302 | #N/A | 0.249 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS3 | IS4
290 | 2023-01-28 16:57 | WO-23003-185-A-1 | U | 0 | #N/A | 0.51 | 0.431 | #N/A | 0.377 | 0.368 | #N/A | 0.241 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS3 | IS4
291 | 2023-01-28 17:24 | WO-23003-186-A-1 | U | 0 | #N/A | 0.239 | 0 | #N/A | 0.37 | 0.258 | #N/A | 0.224 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4
292 | 2023-01-28 17:51 | WO-23003-187-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.27 | 0.235 | #N/A | 0.23 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4
293 | 2023-01-28 18:18 | WO-23003-188-A-1 | U | 0 | #N/A | 0 | 0 | #N/A | 0.342 | 0.293 | #N/A | 0.203 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4
294 | 2023-01-28 18:45 | WO-23003-189-A-1 | U | 0 | #N/A | 0.778 | 1.662 | #N/A | 0.995 | 0.474 | #N/A | 0.38 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS3 | IS4
295 | 2023-01-28 19:12 | WO-23003-190-A-1 | U | 0 | #N/A | 0 | 1.55 | #N/A | 1.102 | 0.329 | #N/A | 0.385 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
296 | 2023-01-28 19:39 | LRB-20-inj1 | K | 0 | #N/A | 0 | 0.763 | #N/A | 0.521 | 0.244 | #N/A | 0.307 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [8 empty] | b
297 | 2023-01-28 20:06 | AB-230117-02 LRB 20 | K | 0 | #N/A | 0.168 | 0.841 | #N/A | 0.497 | 0.242 | #N/A | 0.282 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL
298 | 2023-01-28 20:33 | WO-23003-191-A-1 | U | 0 | #N/A | 1.379 | 1.531 | #N/A | 0.847 | 0.292 | #N/A | 0.168 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS3
299 | 2023-01-28 21:01 | WO-23003-192-A-1 | U | 0 | #N/A | 0 | 1.124 | #N/A | 0.889 | 0.327 | #N/A | 0.175 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4
300 | 2023-01-28 21:28 | WO-23003-193-A-1 | U | 0 | #N/A | 0.206 | 0.838 | #N/A | 0.633 | 0.258 | #N/A | 0.163 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4
301 | 2023-01-28 21:55 | WO-23003-194-A-1 | U | 0 | #N/A | 1.246 | 0.838 | #N/A | 0.905 | 0.277 | #N/A | 0.209 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS3 | IS4
302 | 2023-01-28 22:22 | WO-23003-195-A-1 | U | 0 | #N/A | 0.301 | 1.117 | #N/A | 0.858 | 0 | #N/A | 0.201 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4
303 | 2023-01-28 22:49 | WO-23003-196-A-1 | U | 0 | #N/A | 1.277 | 1.056 | #N/A | 0.806 | 0.256 | #N/A | 0.188 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4
304 | 2023-01-28 23:16 | WO-23003-197-A-1 | U | 0 | #N/A | 1.222 | 3.823 | #N/A | 13.769 | 0 | #N/A | 0.24 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 13.769 | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 |  | IS3 | IS3 | IS4
305 | 2023-01-28 23:43 | WO-23003-198-A-1 | U | 0 | #N/A | 0.694 | 4.452 | #N/A | 14.717 | 0.435 | #N/A | 0.238 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 14.717 | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS3 | IS3 | IS4
306 | 2023-01-29 00:10 | WO-23003-199-A-1 | U | 0 | #N/A | 0.607 | 3.086 | #N/A | 6.684 | 0.239 | #N/A | 0.248 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 6.684 | <MRL | #N/A | <MRL |  | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | R2 | This one looks like it has twice the internal standard amount, cutting concentrations in half. I'll rerun it this weekend.
307 | 2023-01-29 00:37 | WO-23003-200-A-1 | U | 0 | #N/A | 0 | 4.582 | #N/A | 13.529 | 0 | #N/A | 0.206 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 13.529 | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 |  |  | IS3 | IS4
308 | 2023-01-29 01:04 | LRB-21-inj1 | K | 0 | #N/A | 0 | 0.929 | #N/A | 0.653 | 0.233 | #N/A | 0.291 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [8 empty] | b
309 | 2023-01-29 01:31 | AB-230117-02 LRB 21 | K | 0 | #N/A | 0.064 | 0.722 | #N/A | 0.609 | 0.237 | #N/A | 0.325 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [8 empty] | b
310 | 2023-01-29 01:58 | AB-230117-02 LD 10 | D | 0 | #N/A | 1.057 | 3.438 | #N/A | 10.794 | 0 | #N/A | 0.273 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 10.794 | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS4 | IS3 | IS3 | IS4
311 | 2023-01-29 02:26 | AB-230117-02 LFB-H 11 | Z | 163.487 | #N/A | 195.443 | 186.635 | #N/A | 183.897 | 185.609 | #N/A | 186.414 | 1 | 163.487 | #N/A | 195.443 | 186.635 | #N/A | 183.897 | 185.609 | #N/A | 186.414 | No TrueV!
312 | 2023-01-29 02:53 | AB-230117-02 CCC-H 11 | Q | 309.141 | #N/A | 366.619 | 355.509 | #N/A | 355.982 | 358.493 | #N/A | 360.917 | 1 | 309.141 | #N/A | 366.619 | 355.509 | #N/A | 355.982 | 358.493 | #N/A | 360.917
313 | 2023-01-29 03:20 | AB-230117-02 LFSM-H 10 | S | 194.12 | #N/A | 183.533 | 188.607 | #N/A | 190.703 | 192.309 | #N/A | 194.699 | 1 | 194.12 | #N/A | 183.533 | 188.607 | #N/A | 190.703 | 192.309 | #N/A | 194.699 | No TrueV! | IS3 |  | IS3 | IS3 |  | IS3 | IS3 | IS3
314 | 2023-01-29 04:14 | WO-23003-201-A-1 | U | 0 | #N/A | 1.429 | 3.72 | #N/A | 11.225 | 0.809 | #N/A | 0.604 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 11.225 | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 |  | IS3 | IS3 | IS4
315 | 2023-01-29 04:41 | WO-23003-202-A-1 | U | 0 | #N/A | 0 | 4.087 | #N/A | 11.115 | 0 | #N/A | 0.397 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 11.115 | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 |  | IS3 | IS3 | IS4
316 | 2023-01-29 05:08 | WO-23003-203-A-1 | U | 0 | #N/A | 1.134 | 6.695 | #N/A | 39.631 | 0 | #N/A | 0.358 | 1 | <MRL | #N/A | <MRL | 6.695 | #N/A | 39.631 | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS3 | IS3 | IS3 | UL |  | R1
317 | 2023-01-29 05:35 | WO-23003-204-A-1 | U | 0 | #N/A | 0.521 | 5.247 | #N/A | 29.505 | 0 | #N/A | 0.31 | 1 | <MRL | #N/A | <MRL | 5.247 | #N/A | 29.505 | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4 | IS3 | IS3 | IS3 | UL |  | R1
318 | 2023-01-29 06:02 | WO-23003-206-A-1 | U | 0 | #N/A | 0.479 | 5.258 | #N/A | 35.231 | 0 | #N/A | 0.308 | 1 | <MRL | #N/A | <MRL | 5.258 | #N/A | 35.231 | <MRL | #N/A | <MRL |  | IS3 |  |  |  |  | IS3 | IS3 | UL |  | R1
319 | 2023-01-29 06:29 | WO-23003-207-A-1 | U | 0 | #N/A | 0 | 4.559 | #N/A | 26.636 | 0 | #N/A | 0.258 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 26.636 | <MRL | #N/A | <MRL |  | IS3 |  |  |  |  | IS3 | IS3 | UL |  | R1
320 | 2023-01-29 06:56 | WO-23003-208-A-1 | U | 0 | #N/A | 0.392 | 4.923 | #N/A | 24.868 | 0 | #N/A | 0.26 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 24.868 | <MRL | #N/A | <MRL |  | IS3 |  |  |  |  | IS3 | IS3 | UL |  | R1
321 | 2023-01-29 07:23 | WO-23003-209-A-1 | U | 0 | #N/A | 0.264 | 4.472 | #N/A | 25.391 | 0 | #N/A | 0.31 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 25.391 | <MRL | #N/A | <MRL | [6 empty] | IS3 | IS3 | UL |  | R1
322 | 2023-01-29 07:50 | WO-23003-210-A-1 | U | 0 | #N/A | 0.152 | 0.762 | #N/A | 1.497 | 0 | #N/A | 0.289 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [7 empty] | IS4 | IS4
323 | 2023-01-29 08:18 | WO-23003-211-A-1 | U | 0 | #N/A | 0 | 0.916 | #N/A | 1.074 | 0 | #N/A | 0.265 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [7 empty] | IS4 | IS4
324 | 2023-01-29 08:45 | AB-230117-02 LRB 22 | K | 0 | #N/A | 0.161 | 0.756 | #N/A | 0.574 | 0.186 | #N/A | 0.267 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [8 empty] | b
325 | 2023-01-29 09:12 | WO-23003-212-A-1 | U | 0 | #N/A | 0.224 | 0.803 | #N/A | 1.148 | 0 | #N/A | 0.284 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [7 empty] | IS4 | IS4
326 | 2023-01-29 09:39 | AB-230117-02 CCC-H 12 | Q | 343.096 | #N/A | 391.777 | 383.581 | #N/A | 375.367 | 373.921 | #N/A | 372.448 | 1 | 343.096 | #N/A | 391.777 | 383.581 | #N/A | 375.367 | 373.921 | #N/A | 372.448 | [8 empty] | C |  |  | Carryover in CCC, not a cal failure. QCS Okay
327 | 2023-01-29 10:06 | AB-230117-02 QCS 2 | Q | 194.121 | #N/A | 224.431 | 212.761 | #N/A | 218.131 | 232.69 | #N/A | 220.165 | 1 | 194.121 | #N/A | 224.431 | 212.761 | #N/A | 218.131 | 232.69 | #N/A | 220.165
328 | 2023-01-30 11:15 | LRB-22-reprep | K | 0 | #N/A | 0 | 0.435 | #N/A | 2.334 | 0.439 | #N/A | 0.417 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS1 | [6 empty] | b
329 | 2023-01-30 11:42 | LRB-22-reprep-inj2 | K | 0 | #N/A | 0.212 | 0.494 | #N/A | 2.24 | 0.435 | #N/A | 0.404 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS1 | [6 empty] | b
330 | 2023-01-30 12:09 | WO-23003-203-A-1B | U | 0 | #N/A | 0.231 | 0.868 | #N/A | 4.491 | 0.326 | #N/A | 0.378 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [8 empty] | IS4
331 | 2023-01-30 12:36 | WO-23003-204-A-1B | U | 0 | #N/A | 0.313 | 1.758 | #N/A | 9.339 | 0 | #N/A | 0.315 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 9.339 | <MRL | #N/A | <MRL | [7 empty] | IS3 | IS4
332 | 2023-01-30 13:03 | WO-23003-205-A-1B | U | 0 | #N/A | 0.35 | 1.411 | #N/A | 7.698 | 0 | #N/A | 0.343 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | 7.698 | <MRL | #N/A | <MRL | [7 empty] | IS3 | IS4
333 | 2023-01-30 13:30 | WO-23003-206-A-1B | U | 0 | #N/A | 0.356 | 0.794 | #N/A | 4.89 | 0.362 | #N/A | 0.347 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [7 empty] | IS4 | IS4
334 | 2023-01-30 13:58 | WO-23003-207-A-1B | U | 0 | #N/A | 0 | 0.815 | #N/A | 3.98 | 0.366 | #N/A | 0.34 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [8 empty] | IS4
335 | 2023-01-30 14:25 | WO-23003-208-A-1B | U | 0 | #N/A | 0.718 | 0.753 | #N/A | 3.878 | 0.294 | #N/A | 0.316 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | [8 empty] | IS4
336 | 2023-01-30 15:56 | 217288-d4-reprep | U | 0 | #N/A | 0.593 | 0.768 | #N/A | 0.739 | 0.407 | #N/A | 0.577 | 4 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3
337 | 2023-01-30 16:23 | 217289-d4-reprep | U | 0 | #N/A | 0 | 0.642 | #N/A | 0.655 | 0.343 | #N/A | 0.368 | 4 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3
338 | 2023-01-30 16:50 | 217315-d2-reprep | U | 137.147 | #N/A | 0 | 0 | #N/A | 0.537 | 0.321 | #N/A | 0.228 | 2 | 274.294 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4
339 | 2023-01-30 17:17 | 217356-d2-reprep | U | 0 | #N/A | 0.299 | 0.477 | #N/A | 0.546 | 0.326 | #N/A | 4.268 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS4 | IS4
340 | 2023-01-30 17:44 | 217357-d2-reprep | U | 0 | #N/A | 0.835 | 0 | #N/A | 0.607 | 0.929 | #N/A | 48.327 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 96.654 |  | IS3 | IS4 | IS4 | IS4
341 | 2023-01-30 18:12 | 217358-d2-reprep | U | 0 | #N/A | 0 | 0 | #N/A | 0.733 | 1.195 | #N/A | 2575.277 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 5150.554 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | IS4
342 | 2023-01-30 18:39 | 217359-d2-reprep | U | 0 | #N/A | 0 | 0 | #N/A | 1.217 | 3.011 | #N/A | 9903.264 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 19806.528 |  | IS3 |  | IS4 | IS4 |  |  |  |  | IS4
343 | 2023-01-30 19:06 | 217360-d2-reprep | U | 0 | #N/A | 0 | 0 | #N/A | 0.642 | 0.345 | #N/A | 378.052 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 756.104 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | C | R1 | Electively rerun, but this has carrover, original sample did not. Retain original results
344 | 2023-01-30 19:33 | 217393-d4-reprep | U | 0 | #N/A | 0 | 0.775 | #N/A | 3.747 | 2.949 | #N/A | 1061.939 | 4 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 4247.756 |  | IS3
345 | 2023-01-30 20:00 | 217394-d4-reprep | U | 0 | #N/A | 0.511 | 1.218 | #N/A | 4.579 | 4.034 | #N/A | 1414.578 | 4 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 5658.312 |  | IS3
346 | 2023-01-30 21:21 | 217395-d4-reprep | U | 0 | #N/A | 0.132 | 1.534 | #N/A | 9.032 | 8.354 | #N/A | 2714.735 | 4 | <MRL | #N/A | <MRL | <MRL | #N/A | 36.128 | 33.416 | #N/A | 10858.94 |  | IS3 | [7 empty] | IS4
347 | 2023-01-30 22:15 | 217399-d2-reprep | U | 0 | #N/A | 0 | 0 | #N/A | 0.679 | 0.312 | #N/A | 4.311 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS4 | IS3 | IS4
348 | 2023-01-30 22:42 | 217400-d2-reprep | U | 0 | #N/A | 0 | 0 | #N/A | 0.707 | 0.321 | #N/A | 1.945 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 |  | IS3 | IS4
349 | 2023-01-30 23:09 | 217295-d1000-LD-reprep | D | 0 | #N/A | 0.177 | 0.308 | #N/A | 0.299 | 0.28 | #N/A | 1.087 | 1000 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS1 | IS2
350 | 2023-01-30 23:36 | 217335-d2-LD-reprep | D | 430.994 | #N/A | 0 | 0 | #N/A | 0.574 | 0.328 | #N/A | 1.008 | 2 | 861.988 | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS3 | IS3 | IS4 | IS4
351 | 2023-01-31 00:04 | 217375-d2-LD-reprep | D | 0 | #N/A | 0 | 0.931 | #N/A | 0.848 | 0.367 | #N/A | 28.506 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 57.012 |  | IS3 | IS4 | IS4 | IS4 | IS4
352 | 2023-01-31 00:31 | 217395-d4-LD-reprep | D | 0 | #N/A | 0 | 1.523 | #N/A | 8.776 | 8.406 | #N/A | 2607.572 | 4 | <MRL | #N/A | <MRL | <MRL | #N/A | 35.104 | 33.624 | #N/A | 10430.288 |  | IS3 | [7 empty] | IS4
353 | 2023-01-31 01:25 | 217359-d2-LD | D | 0 | #N/A | 0 | 0 | #N/A | 0.968 | 2.774 | #N/A | 8748.882 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 17497.764 |  | IS3 | IS4 | IS4 | IS4 |  |  |  |  | IS4
354 | 2023-01-31 02:19 | 217415-d2-LD-reprep | D | 0 | #N/A | 0.57 | 0 | #N/A | 0.48 | 0.339 | #N/A | 10.557 | 2 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | 21.114 |  | IS3 | IS4 | IS3 | IS4
355 | 2023-01-31 02:46 | 217296-d1000-LFSM-200-reprep | S | 166.141 | #N/A | 199.356 | 189.823 | #N/A | 180.753 | 184.164 | #N/A | 190.796 | 1000 | 166141 | #N/A | 199356 | 189823 | #N/A | 180753 | 184164 | #N/A | 190796 |  | IS1 | IS1 | [5 empty] | IS3
356 | 2023-01-31 03:13 | 217336-d2-LFSM-200-reprep | S | 606.641 | #N/A | 185.585 | 184.866 | #N/A | 175.24 | 187.674 | #N/A | 194.566 | 2 | 1213.282 | #N/A | 371.17 | 369.732 | #N/A | 350.48 | 375.348 | #N/A | 389.132 |  | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
357 | 2023-01-31 03:40 | 217376-d2-LFSM-200-reprep | S | 196.844 | #N/A | 173.032 | 173.811 | #N/A | 166.325 | 177.327 | #N/A | 266.746 | 2 | 393.688 | #N/A | 346.064 | 347.622 | #N/A | 332.65 | 354.654 | #N/A | 533.492 |  | IS3 |  | IS3 | IS3
358 | 2023-01-31 04:07 | 217396-d200-LFSM-200 | S | 153.37 | #N/A | 175.395 | 175.651 | #N/A | 168.361 | 178.6 | #N/A | 4040.399 | 200 | 30674 | #N/A | 35079 | 35130.2 | #N/A | 33672.2 | 35720 | #N/A | 808079.8 |  | IS1 | IS1 | [6 empty] | IS3
359 | 2023-01-31 04:34 | 217360-d2-LFSM-200 | S | 182.595 | #N/A | 187.722 | 186.158 | #N/A | 185.148 | 185.543 | #N/A | 384.137 | 2 | 365.19 | #N/A | 375.444 | 372.316 | #N/A | 370.296 | 371.086 | #N/A | 768.274 |  | IS3 | IS3 | IS3 | IS3 | IS3 |  |  |  | C
360 | 2023-01-31 05:01 | 217416-d2-LFSM-200-reprep | S | 199.788 | #N/A | 183.534 | 184.672 | #N/A | 186.591 | 187.492 | #N/A | 195.064 | 2 | 399.576 | #N/A | 367.068 | 369.344 | #N/A | 373.182 | 374.984 | #N/A | 390.128 |  | IS3 | IS3 | IS3 | IS3 |  | IS3 | IS3
361 | 2023-01-30 11:15 | LRB-22-reprep | K | 0 | #N/A | 0 | 0.435 | #N/A | 2.334 | 0.439 | #N/A | 0.417 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS1 | IS2 |  |  | b | b | b | b | b
362 | 2023-01-30 11:42 | LRB-22-reprep-inj2 | K | 0 | #N/A | 0.212 | 0.494 | #N/A | 2.24 | 0.435 | #N/A | 0.404 | 1 | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL | <MRL | #N/A | <MRL |  | IS1 | IS2
363 | 2023-01-31 06:23 | QCS-2-Reinject | Q | 174.903 | #N/A | 211.41 | 200.785 | #N/A | 189.964 | 202.109 | #N/A | 205.073 | 1 | 174.903 | #N/A | 211.41 | 200.785 | #N/A | 189.964 | 202.109 | #N/A | 205.073 |  | IS1
364 | 2023-01-31 06:50 | LFB-H-11-reinject | Z | 194.643 | #N/A | 225.808 | 219.041 | #N/A | 205.408 | 224.483 | #N/A | 213.202 | 1 | 194.643 | #N/A | 225.808 | 219.041 | #N/A | 205.408 | 224.483 | #N/A | 213.202 | No TrueV! | IS1
365 | 2023-01-31 07:17 | CCC-H-12-reinject | Q | 305.952 | #N/A | 364.623 | 349.782 | #N/A | 351.819 | 354.686 | #N/A | 353.248 | 1 | 305.952 | #N/A | 364.623 | 349.782 | #N/A | 351.819 | 354.686 | #N/A | 353.248 | No TrueV!
0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## Sheet1_1
Batch Folder Path: | D:\TraceFinderData\Projects\PFAS Online SPE\2023-01-18\2023-01-18
Batch Name: | 2023-01-18
Last Acquisition Date: | 2023-01-31 07:17
User Name: | Quantiva
Sample Count: | 365
Sample Amount - N/F
Data File | 11-Cl-PF3OUDS | 4-2:FTS | 6-2:FTS | 8-2:FTS | 9-Cl-PF3ONS | FBSA | FHxA | HFPO-DA | L-PFBS | L-PFHpS | L-PFHxS | L-PFOS | NaDONA | N-Et-FOSA | N-Me-FOSA | PFBA | PFDA | PFDoA | PFDS | PFHpA | PFHxA | PFHxDA | PFNA | PFNS | PFOA | PFODA | PFOSA | PFPeA | PFPeS | PFTeDA | PFTrDA | PFUdA | All Flags | Sample Type
AB-230117-02 Cal Blank 1 | N/F | 35.433 | 30.203 | 0.143 | 2.522 | N/F | 3560.748 | N/F | N/F | 2.128 | 1.927 | 3.298 | N/F | N/F | 0.734 | N/F | 2.511 | N/F | N/F | 2.616 | 0.566 | N/A | 2.445 | 0.906 | 2.166 | N/A | 1.03 | N/F | -4.716 | N/A | N/F | N/A |  | Unknown
AB-230117-02 Cal Std 1 1 | N/F | 6.433 | 3.598 | 2.526 | 1.239 | N/F | 4.842 | 1.078 | N/F | 0.897 | 1.451 | 1.593 | N/F | N/F | 1.215 | N/F | 1.325 | N/F | N/F | 2.116 | 1.415 | N/A | 1.368 | 1.526 | 1.437 | N/A | 1.481 | N/F | -3.231 | N/A | N/F | N/A |  | Cal Std
AB-230117-02 Cal Std 2 1 | N/F | -5.318 | -0.486 | 1.171 | 1.909 | N/F | -369.984 | 2.003 | 1.995 | 1.816 | 2.155 | 2.003 | N/F | N/A | 2.201 | N/F | 2.016 | N/F | N/A | 1.404 | 2.095 | N/A | 2.05 | 1.966 | 2.071 | N/A | 2.324 | N/F | -2.568 | N/A | N/F | N/A |  | Cal Std
AB-230117-02 Cal Std 3 1 | N/F | -2.501 | 0.199 | 1.353 | 3.936 | N/A | 534.951 | 4.412 | 3.707 | 3.299 | 3.709 | 3.798 | N/F | N/F | 3.778 | 5.762 | 4.043 | N/F | N/F | 3.081 | 3.86 | N/A | 3.843 | 4.03 | 3.96 | N/A | 4.016 | N/F | 0.921 | 47.308 | N/F | N/A |  | Cal Std
AB-230117-02 Cal Std 4 1 | N/F | 7.181 | 6.426 | 8.048 | 9.297 | N/F | 352.458 | 9.51 | 7.789 | 9.624 | 7.745 | 7.845 | N/F | N/A | 9.26 | 11.646 | 9.394 | N/F | N/F | 8.566 | 9.126 | N/A | 9.265 | 8.612 | 9.194 | N/A | 9.157 | N/F | 16.627 | N/A | N/F | N/A |  | Cal Std
AB-230117-02 Cal Std 5 1 | N/F | 23.955 | 16.593 | 17.373 | 19.189 | N/F | 185.587 | 18.606 | 18.452 | 16.778 | 15.779 | 15.192 | N/F | N/A | 18.987 | 18.831 | 18.415 | N/F | N/F | 17.177 | 17.853 | N/A | 18.63 | 16.924 | 18.158 | N/A | 17.349 | N/F | 57.566 | N/A | N/F | N/A |  | Cal Std
AB-230117-02 Cal Std 6 1 | N/F | 41.816 | 35.955 | 35.512 | 38.804 | N/F | 650.022 | 36.592 | 32.581 | 38.875 | 30.702 | 31.551 | N/F | N/A | 36.41 | 37.308 | 37.389 | N/F | N/F | 36.068 | 35.14 | N/A | 37.996 | 33.194 | 36.553 | N/A | 35.688 | N/F | 88.672 | N/A | N/F | N/A |  | Cal Std
AB-230117-02 Cal Std 7 1 | N/F | 104.514 | 100.259 | 101.538 | 92.867 | N/F | -293.284 | 95.727 | 87.282 | 104.133 | 79.193 | 79.635 | N/F | N/F | 95.217 | 95.665 | 97.43 | N/F | N/F | 94.573 | 93.802 | N/A | 96.468 | 90.032 | 95.968 | N/A | 90.01 | N/F | 194.53 | N/A | N/F | N/A |  | Cal Std
AB-230117-02 Cal Std 8 1 | N/F | 183.607 | 195.409 | 208.267 | 186.785 | N/F | -217.327 | 197.427 | 174.364 | 200.784 | 164.11 | 161.542 | N/F | N/F | 192.477 | 163.759 | 193.114 | N/F | N/F | 189.282 | 191.961 | N/A | 189.259 | 174.816 | 187.35 | N/A | 188.906 | N/F | 379.144 | N/A | N/F | N/A |  | Cal Std
AB-230117-02 Cal Std 9 1 | N/F | 352.895 | 427.559 | 368.173 | 379.575 | N/F | -502.067 | 402.773 | 352.84 | 396.639 | 336.272 | 345.567 | N/F | N/A | 391.958 | 328.948 | 385.495 | N/F | N/F | 377.454 | 384.687 | N/A | 386.336 | 339.641 | 378.816 | N/A | 359.431 | N/F | 790.225 | N/A | N/F | N/A |  | Cal Std
AB-230117-02 Cal Std 10 1 | N/F | 986.227 | 1112.502 | 835.797 | 987.021 | N/F | 632.176 | 1029.861 | 949.998 | 975.058 | 867.74 | 896.777 | 1310.263 | N/F | 1022.645 | 951.413 | 972.285 | N/F | N/A | 968.916 | 1001.157 | N/A | 982.021 | 858.007 | 980.955 | N/A | 914.041 | N/F | 1693.731 | N/A | N/F | N/A |  | Cal Std
AB-230117-02 Cal Std 11 1 | N/F | 1993.322 | 2033.349 | 2298.172 | 2306.061 | N/F | 2444.598 | 2030.625 | 1689.093 | 2154.985 | 1807.699 | 1876.965 | 1632.429 | N/F | 2041.136 | 1822.989 | 1949.845 | N/F | N/F | 1943.349 | 1991.195 | N/A | 1944.508 | 2064.871 | 1905.749 | N/A | 1988.411 | N/F | 2397.48 | 17779.806 | N/F | N/A |  | Cal Std
AB-230117-02 Cal Std 12 1 | N/F | 4354.331 | 4412.932 | 4246.579 | 4244.259 | N/F | 4140.162 | 4196.925 | 3581.739 | 3951.007 | 3852.301 | 3839.827 | 3118.615 | N/F | 4135.203 | 4173.586 | 4002.829 | N/F | N/F | 4032.427 | 4184.365 | N/A | 3980.463 | 3919.784 | 4022.699 | N/A | 4062.479 | N/F | 4013.728 | N/A | N/F | N/A |  | Cal Std
AB-230117-02 Cal Std 13 1 | N/F | 9730.538 | 9432.704 | 9652.493 | 9506.059 | N/F | 10214.865 | 9751.461 | 8867.332 | 9087.584 | 9077.736 | 9233.802 | 10938.694 | N/F | 9826.514 | 10164.091 | 10103.42 | N/F | N/F | 10102.586 | 9860.345 | N/A | 10124.793 | 9583.875 | 10134.089 | N/A | 10103.708 | N/F | 8150.176 | N/A | N/F | 17707.693 |  | Cal Std
AB-230117-02 CCC-H 1 | N/F | 385.362 | 398.734 | 404.887 | 418.575 | N/F | 292.39 | 392.93 | 328.415 | 389.833 | 331.89 | 338.658 | N/F | N/F | 389.955 | 331.373 | 378.211 | N/F | N/F | 372.53 | 384.696 | N/A | 379.732 | 344.117 | 381.982 | N/A | 349.518 | N/F | 792.398 | N/A | N/F | N/A |  | Unknown
AB-230117-02 CCC-H 2 | N/F | 345.949 | 405.237 | 356.5 | 448.375 | N/F | 350.502 | 403.865 | 354.539 | 395.824 | 338.298 | 349.146 | 1220.854 | N/F | 403.12 | 342.723 | 380.815 | N/A | N/F | 384.974 | 379.785 | N/A | 387.516 | 309.753 | 381.16 | N/A | 362.633 | N/F | 783.932 | N/A | N/F | N/A |  | Unknown
AB-230117-02 CCC-H 3 | N/F | 337.834 | 419.419 | 342.705 | 455.979 | N/F | 3615.811 | 400.863 | 332.997 | 408.594 | 344.125 | 354.775 | N/F | N/A | 405.278 | 366.428 | 393.376 | N/F | N/A | 378.782 | 377.097 | N/A | 383.968 | 320.066 | 377.585 | N/A | 357.124 | N/F | 561.845 | N/A | N/F | N/A |  | Unknown
AB-230117-02 IBC 1 | N/F | 190.793 | 208.097 | 192.396 | 197.738 | N/F | -132.401 | 189.323 | 167.778 | 199.011 | 161.272 | 160.013 | N/F | N/F | 188.721 | 169.148 | 195.489 | N/F | N/F | 187.723 | 190.496 | N/A | 196.495 | 155.883 | 191.07 | N/A | 167.237 | N/F | 467.225 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LD 1 | N/F | 19.279 | 4.977 | -1.672 | 0.391 | N/F | -2546.733 | N/F | 10.532 | 0.309 | 0.993 | 2.418 | N/F | N/A | 0.48 | 574.418 | 1.343 | N/F | N/F | 0.777 | 1.066 | N/A | 0.632 | 0.701 | 0.424 | N/A | 0.761 | N/F | -5.707 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LD 2 | N/F | -13.426 | N/F | -0.182 | 0.556 | N/F | -267.63 | N/F | 2.738 | 161.815 | 5.696 | 7256.85 | N/F | N/A | 0.528 | N/F | 0.39 | N/F | N/A | 0.236 | N/F | N/A | N/F | 9.276 | 0.553 | N/A | 0.718 | N/F | -5.837 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LFB-H 1 | N/F | 3.695 | -0.164 | 1.695 | 5.355 | N/F | -70.393 | 178.384 | 160.623 | 206.101 | 185.074 | 180.479 | N/F | N/F | 3.972 | 183.485 | 191.802 | N/F | N/F | 207.335 | 194.206 | N/A | 204.981 | 162.745 | 198.792 | N/A | 3.616 | N/F | 556.028 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LFB-H 2 | N/F | 175.463 | 205.745 | 202.468 | 203.342 | N/F | 177.92 | 196.167 | 168.796 | 194.454 | 164.256 | 163.131 | N/F | N/F | 197.893 | 168.178 | 191.306 | N/F | N/F | 187.116 | 190.739 | N/A | 194.708 | 156.284 | 188.953 | N/A | 178.779 | N/F | 365.417 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LFB-H 3 | N/F | 203.096 | 200.347 | 185.527 | 235.568 | N/F | 132.029 | 198.181 | 174.498 | 200.319 | 162.174 | 578.353 | 1133.244 | N/A | 194.01 | 169.371 | 186.355 | N/F | N/A | 189.889 | 191.991 | N/A | 189.526 | 148.021 | 190.049 | N/A | 171.972 | N/F | 399.14 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LFSM-H 1 | N/F | 311.403 | 222.11 | 233.218 | 219.826 | N/F | -2571.449 | 197.796 | 195.907 | 271.782 | 160.987 | 175.543 | N/F | N/A | 220.519 | 570.426 | 212.68 | N/F | N/F | 198.678 | 197.561 | N/A | 209.365 | 234.207 | 202.268 | N/A | 200.144 | N/F | 473.874 | 4509.996 | N/F | N/A |  | Unknown
AB-230117-02 LFSM-H 2 | N/F | 136.812 | 190.809 | 191.947 | 511.677 | N/F | 464.347 | 198.319 | 178.852 | 348.673 | 174.353 | 6985.084 | N/F | N/F | 210.273 | 169.438 | 206.37 | N/F | N/A | 186.077 | 190.263 | N/A | 202.465 | 174.947 | 196.626 | N/A | 176.42 | N/F | 399.43 | N/A | N/F | N/A |  | Unknown
LRB2-Inj1 | N/F | -5.457 | -4.446 | 0.14 | 0.461 | N/F | 777.915 | N/F | N/F | N/F | N/F | 0.532 | N/F | N/F | 0.553 | N/F | 204.132 | N/F | N/F | -0.213 | N/F | N/A | 0.414 | 0.745 | 0.46 | N/A | 0.828 | N/F | -5.669 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LRB 2 | N/F | -4.465 | -3.552 | 0.977 | 0.428 | N/F | -789.43 | N/F | N/F | N/F | N/F | 0.413 | N/F | N/F | 0.669 | N/F | 7.368 | N/F | N/F | 0.033 | 0.364 | N/A | 0.349 | N/F | 0.466 | N/A | 0.849 | N/F | -5.087 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LRB 3 | N/F | -5.609 | -4.808 | -1.395 | 0.446 | N/F | -119.75 | N/F | N/F | 1.341 | N/F | 67.797 | N/F | N/F | 0.413 | N/F | 0.983 | N/F | N/F | 0.028 | N/F | N/A | 0.37 | 0.784 | 0.452 | N/A | 0.866 | N/F | -5.484 | 1.367 | N/F | N/A |  | Unknown
LRB4-inj1 | N/F | 1.514 | -2.845 | -0.729 | 0.361 | N/F | 44.313 | N/F | N/F | N/F | N/F | 0.907 | N/F | N/A | 0.371 | N/F | 0.334 | N/F | N/A | -0.395 | 0.299 | N/A | 0.282 | N/F | 0.368 | N/A | 0.748 | N/F | -5.851 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LRB 4 | N/F | 29.293 | -3.76 | -0.358 | 0.365 | N/A | 527.884 | 0.341 | N/F | N/F | N/F | 0.692 | N/F | N/A | 0.409 | N/F | 0.359 | N/F | N/F | N/F | N/F | N/A | 0.302 | N/F | 0.32 | N/A | 0.762 | N/F | -5.869 | N/A | N/F | N/A |  | Unknown
LRB5-inj1 | N/F | -6.821 | -3.242 | -0.026 | 0.429 | N/F | -105.886 | N/F | N/F | 0.055 | 0.817 | 0.588 | N/F | N/A | 0.557 | N/F | 0.473 | N/A | N/A | -0.225 | N/F | N/A | 0.372 | 0.655 | 0.429 | N/A | 0.736 | N/F | -5.91 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LRB 5 | N/F | 6.597 | -2.882 | -0.354 | 0.355 | N/F | -325.863 | N/F | 1.083 | N/F | 0.774 | 0.513 | N/F | N/F | 0.46 | N/F | 0.334 | N/F | N/A | -0.234 | N/F | N/A | 0.288 | 0.651 | 0.37 | N/A | 0.759 | N/F | -5.285 | N/A | N/F | N/A |  | Unknown
Rinse5 | N/F | 721.957 | 331.464 | 92.14 | 219.16 | N/A | 4024.193 | N/F | N/F | 427.283 | 125.204 | 6694.884 | N/F | N/A | 190.002 | N/F | 175.37 | N/F | N/A | 139.192 | 39.805 | N/A | 169.964 | 174.028 | 160.825 | N/A | 144.437 | N/F | 277.142 | N/A | N/F | N/A |  | Unknown
LRB6-inj1 | N/F | -8.173 | -3.104 | -0.428 | N/F | N/F | -912.107 | N/F | 1.051 | N/F | N/F | 0.554 | N/F | N/A | 0.488 | N/F | 0.343 | N/F | N/A | -0.344 | 0.352 | N/A | 0.221 | 0.66 | 0.362 | N/A | 0.725 | N/F | -4.753 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LRB 6 | N/F | -3.834 | -4.704 | -0.241 | 0.354 | N/F | -156.963 | N/F | 1.452 | -0.028 | N/F | 0.657 | N/F | N/A | 0.447 | N/F | 0.346 | N/A | N/A | 0.04 | N/F | N/A | 0.24 | 0.756 | 0.308 | N/A | 0.729 | N/F | -5.088 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LRB 1 | N/F | -4.387 | -1.13 | 0.339 | 4.722 | N/F | 103.833 | 1.125 | N/F | 4.052 | 3.296 | 3.964 | N/F | N/F | 4.164 | N/F | 4.651 | N/F | N/F | 1.439 | N/F | N/A | 4.565 | 4.532 | 3.537 | N/A | 5.969 | N/F | -2.944 | N/A | N/F | N/A |  | Unknown
AB-230117-02 QCS 1 | N/F | 222.006 | 231.135 | 228.66 | 234.748 | N/F | 137.107 | 219.599 | 197.426 | 228.332 | 186.245 | 186.768 | N/F | N/F | 6.392 | 215.739 | 213.968 | N/F | N/F | 214.666 | 212.607 | N/A | 210.554 | 15.561 | 211.384 | N/A | 6.033 | N/F | 497.605 | N/A | N/F | N/A |  | Unknown
WO-23003-01-A-1 | N/F | 102.349 | 17.153 | -1.827 | 1.059 | N/A | -1781.897 | N/F | 11.936 | 1.204 | 1.272 | 1.714 | N/F | N/F | 0.859 | 429.802 | 1.464 | N/F | N/F | 0.641 | 1.148 | N/A | 1.291 | 1.371 | 1.127 | N/A | 1.506 | N/F | -4.44 | N/A | N/F | N/A |  | Unknown
WO-23003-02-A-1 | N/F | 87.645 | 9.739 | -2.227 | 0.779 | N/F | -2518.655 | 2.86 | 12.332 | 0.911 | 1.232 | 1.579 | N/F | N/F | 0.801 | 458.755 | 1.402 | N/F | N/F | 3.701 | N/F | N/A | 1.012 | 1.069 | 0.97 | N/A | 1.314 | N/F | -4.854 | N/A | N/F | N/A |  | Unknown
WO-23003-03-A-1 | N/F | 77.152 | 8.005 | -1.763 | 0.694 | N/A | -1606.626 | N/F | 11.056 | 0.525 | 1.223 | 1.271 | N/F | N/F | 0.728 | 463.337 | 1.12 | N/F | N/F | 0.072 | 0.902 | N/A | 0.908 | 1.036 | 0.843 | N/A | 1.185 | N/F | -5.105 | N/A | N/F | N/A |  | Unknown
WO-23003-04-A-1 | N/F | 169.578 | 17.44 | 0.46 | 0.63 | N/A | -3225.553 | N/F | N/F | 0.631 | N/F | 1.588 | N/F | N/F | 0.576 | 434.867 | 0.822 | N/F | N/A | 0.944 | N/F | N/A | 0.772 | 0.908 | 0.757 | N/A | 1.054 | N/F | N/F | N/A | N/F | N/A |  | Unknown
WO-23003-05-A-1 | N/F | 179.896 | 23.684 | -1.516 | 0.611 | N/A | -2833.139 | N/F | N/F | 0.873 | N/F | 1.348 | N/F | N/F | 0.497 | 511.067 | 1.093 | N/F | N/F | N/F | N/F | N/A | 0.744 | 0.892 | 0.675 | N/A | 0.955 | N/F | N/F | N/A | N/F | N/A |  | Unknown
WO-23003-06-A-1 | N/F | 194.87 | 21.382 | -1.044 | 0.584 | N/A | -3075.464 | N/F | N/F | N/F | N/F | 8.176 | N/F | N/F | 0.691 | 462.818 | 0.769 | N/F | N/F | 5.339 | N/F | N/A | 0.771 | 0.837 | 0.657 | N/A | 0.989 | N/F | N/F | N/A | N/F | N/A |  | Unknown
WO-23003-07-A-1 | N/F | -0.162 | -3.679 | -1.026 | 0.48 | N/F | 571.97 | N/F | N/F | 0.156 | 0.934 | 0.592 | N/F | N/F | 0.906 | N/F | 3898.943 | N/F | N/A | 0.318 | N/F | N/A | 1.8 | N/F | 1.008 | N/A | 0.944 | N/F | -5.957 | N/A | N/F | N/A |  | Unknown
WO-23003-08-A-1 | N/F | -7.845 | -3.222 | -0.974 | 0.484 | N/A | 327.435 | N/F | N/F | 0.062 | N/F | 0.143 | N/F | N/F | 0.938 | N/F | 4309.18 | N/F | N/F | -0.098 | N/F | N/A | 1.8 | N/F | 0.97 | N/A | 0.904 | N/F | -5.795 | N/A | N/F | N/A |  | Unknown
WO-23003-09-A-1 | N/F | 33.342 | -3.591 | -1.23 | 0.445 | N/F | 1224.806 | N/F | N/F | N/F | N/F | 0.27 | N/F | N/F | 0.988 | N/F | 3931.95 | N/F | N/F | N/F | N/F | N/A | 1.666 | N/F | 1.051 | N/A | 0.929 | N/F | -5.484 | N/A | N/F | N/A |  | Unknown
WO-23003-10-A-1 | N/F | -8.664 | -3.038 | 2.19 | 0.425 | N/F | -359.729 | N/F | N/F | 0.056 | N/F | 0.079 | N/F | N/A | 0.985 | N/F | 4432.787 | N/F | N/F | -0.097 | N/F | N/A | 1.794 | N/F | 0.95 | N/A | 0.854 | N/F | -5.715 | N/A | N/F | N/A |  | Unknown
WO-23003-11-A-1 | N/F | -16.634 | -3.581 | -0.534 | 0.389 | N/F | -438.895 | 0.293 | N/F | N/F | N/F | 0.143 | N/F | N/A | 1.04 | N/F | 4146.818 | N/F | N/F | 0.016 | N/F | N/A | 1.711 | N/F | 0.861 | N/A | 0.844 | N/F | -5.506 | N/A | N/F | N/A |  | Unknown
WO-23003-12-A-1 | N/F | -12.417 | -4.349 | -0.145 | 0.397 | N/F | -57.933 | N/F | N/F | N/F | N/F | 0.272 | N/F | N/F | 1.078 | N/F | 4137.805 | N/F | N/F | -0.04 | N/F | N/A | 1.709 | N/F | 0.96 | N/A | 0.753 | N/F | -5.271 | N/A | N/F | N/A |  | Unknown
WO-23003-13-A-1 | N/F | -5.993 | -3.172 | 1.969 | 0.414 | N/F | 174.572 | 0.791 | N/F | N/F | N/F | 0.159 | N/F | N/A | 1.442 | N/F | 4255.588 | N/F | N/F | -0.069 | N/F | N/A | 1.759 | N/F | 0.961 | N/A | 0.777 | N/F | -5.492 | N/A | N/A | N/A |  | Unknown
WO-23003-14-A-1 | N/F | 31.051 | 5.44 | -2.184 | 0.425 | N/F | -767.572 | N/F | 4778.929 | 0.187 | N/F | 0.732 | N/F | N/A | 0.348 | N/F | 193.554 | N/F | N/F | 4.04 | N/F | N/A | 0.504 | N/F | 0.529 | N/A | 0.867 | N/F | -5.131 | N/A | N/A | N/A |  | Unknown
WO-23003-15-A-1 | N/F | 122.795 | 6.429 | -1.146 | 0.445 | N/A | 1885.316 | N/F | 4665.408 | N/F | 0.969 | 0.966 | N/F | N/A | 0.18 | N/F | 7.75 | N/F | N/A | 3.858 | 0.817 | N/A | 0.488 | 0.754 | 0.502 | N/A | 0.936 | N/F | -4.624 | N/A | N/F | N/A |  | Unknown
WO-23003-16-A-1 | N/F | 32.463 | 4.202 | -1.835 | 0.421 | N/F | -903.118 | 0.678 | 4754.423 | N/F | 0.955 | 0.884 | N/F | N/A | 0.218 | N/F | 4.083 | N/F | N/F | 3.737 | 0.767 | N/A | 0.505 | N/F | 0.465 | N/A | 0.798 | N/F | -4.805 | N/A | N/F | N/A |  | Unknown
WO-23003-17-A-1 | N/F | 34.024 | 5.429 | -1.916 | 0.437 | N/F | -861.573 | 0.851 | 4685.686 | 0.165 | N/F | 0.981 | N/F | N/A | 0.58 | N/F | 2.358 | N/F | N/F | 6.124 | 0.834 | N/A | 0.469 | 0.695 | 0.579 | N/A | 0.802 | N/F | -4.72 | 2.366 | N/F | N/A |  | Unknown
WO-23003-18-A-1 | N/F | 41.081 | 4.209 | -1.447 | 0.484 | N/A | -900.536 | 0.314 | 4759.817 | 0.324 | 1.021 | 0.883 | N/F | N/F | 0.72 | N/F | 1.905 | N/F | N/A | 3.8 | 0.925 | N/A | 0.508 | 0.793 | 0.69 | N/A | 0.831 | N/F | -4.788 | N/A | N/F | N/A |  | Unknown
WO-23003-19-A-1 | N/F | 56.352 | 2.334 | -0.769 | 0.5 | N/A | 408.006 | 0.589 | 4724.366 | 0.205 | N/F | 0.885 | N/F | N/A | 0.68 | N/F | 1.605 | N/F | N/F | 3.673 | N/F | N/A | 0.569 | 0.725 | 0.681 | N/A | 0.888 | N/F | -5.025 | N/A | N/F | N/A |  | Unknown
WO-23003-20-A-1 | N/F | 28.894 | N/F | -0.452 | 0.613 | N/F | 3274.15 | 0.342 | N/F | 37.335 | 1.865 | 1470.859 | N/F | N/A | 0.505 | N/F | 1.392 | N/F | N/A | 0.209 | N/F | N/A | 0.531 | 2.73 | 0.524 | N/A | 0.872 | N/F | -5.293 | N/A | N/F | N/A |  | Unknown
WO-23003-21-A-1 | N/F | -9.125 | -4.909 | 5.342 | 5.371 | N/A | -500 | 1.109 | N/F | 34.83 | 2.776 | 1416.32 | N/F | N/A | 3.39 | N/F | 3.141 | N/F | N/F | 0.998 | 1.055 | N/A | 2.538 | 5.021 | 1.898 | N/A | 3.321 | N/F | -4.713 | N/A | N/F | N/A |  | Unknown
WO-23003-22-A-1 | N/F | -7.308 | N/F | -1.81 | 1.01 | N/F | 60.363 | N/F | N/F | 31.662 | 1.996 | 1386.347 | N/F | N/A | 0.828 | N/F | 0.812 | N/F | N/A | 1.349 | N/F | N/A | 0.601 | 2.62 | 0.665 | N/A | 1.119 | N/F | -5.541 | N/A | N/F | N/A |  | Unknown
WO-23003-23-A-1 | N/F | -7.589 | N/F | -1.758 | 0.729 | N/F | -272.342 | N/F | N/F | 35.451 | 1.924 | 1540.035 | N/F | N/F | 0.797 | N/F | 0.718 | N/F | N/F | 0.172 | 0.597 | N/A | 0.484 | 2.82 | 0.544 | N/A | 1.008 | N/F | -5.776 | N/A | N/F | N/A |  | Unknown
WO-23003-24-A-1 | N/F | 1.161 | N/F | 0.435 | 0.938 | N/F | 511.516 | 0.512 | N/F | 34.306 | 2.015 | 1397.807 | N/F | N/F | 0.887 | N/F | 0.913 | N/F | N/F | 0.732 | 0.905 | N/A | 0.727 | 2.772 | 0.863 | N/A | 0.845 | N/F | -5.033 | N/A | N/F | N/A |  | Unknown
WO-23003-25-A-1 | N/F | 0.707 | N/F | -1.022 | 0.592 | N/F | 199.062 | 0.818 | N/F | 37.372 | 1.981 | 1542.714 | N/F | N/A | 0.592 | N/F | 0.578 | N/F | N/A | 0.239 | 0.512 | N/A | 0.436 | 2.743 | 0.515 | N/A | 0.941 | N/F | -5.402 | N/A | N/F | N/A |  | Unknown
WO-23003-26-A-1 | N/F | 57.364 | 11.807 | -0.273 | 0.415 | N/A | -1640.296 | 1.004 | N/F | 1.855 | N/F | 63.814 | N/F | N/A | 0.53 | N/F | 0.489 | N/F | N/A | 0.119 | N/F | N/A | 0.392 | 0.891 | 0.349 | N/A | 0.787 | N/F | -4.638 | N/A | N/F | N/A |  | Unknown
WO-23003-27-A-1 | N/F | 43.187 | 10.628 | -2.156 | 0.383 | N/F | -1668.934 | N/F | 4.255 | 0.252 | N/F | 2.775 | N/F | N/A | 0.65 | N/F | 0.461 | N/F | N/F | -0.054 | N/F | N/A | 0.363 | N/F | 0.424 | N/A | 0.798 | N/F | -5.376 | N/A | N/F | N/A |  | Unknown
WO-23003-28-A-1 | N/F | 58.511 | 11.078 | -1.202 | 0.393 | N/F | -1322.278 | N/F | N/F | N/F | N/F | 1.4 | N/F | N/A | 0.44 | N/F | 0.444 | N/F | N/F | N/F | N/F | N/A | 0.397 | N/F | 0.354 | N/A | 0.802 | N/F | -5.076 | N/A | N/F | N/A |  | Unknown
WO-23003-29-A-1 | N/F | 65.149 | 12.443 | -1.327 | 0.385 | N/F | -1673.799 | N/F | 2.843 | 0.155 | N/F | 0.915 | N/F | N/A | 0.515 | N/F | 0.406 | N/F | N/A | N/F | N/F | N/A | 0.324 | N/F | 0.28 | N/A | 0.841 | N/F | N/F | N/A | N/F | N/A |  | Unknown
WO-23003-30-A-1 | N/F | 51.374 | 14.607 | -1.681 | 0.37 | N/F | -1872.615 | N/F | N/F | 0.21 | N/F | 0.911 | N/F | N/A | 0.607 | N/F | 0.484 | N/F | N/A | 0.163 | N/F | N/A | 0.37 | N/F | 0.295 | N/A | 0.776 | N/F | -4.781 | N/A | N/F | N/A |  | Unknown
WO-23003-31-A-1 | N/F | 129.1 | 12.705 | -0.898 | 0.389 | N/F | -746.072 | N/F | N/F | 0.265 | N/F | 0.994 | N/F | N/F | 0.337 | N/F | 0.955 | N/F | N/A | N/F | N/F | N/A | 0.548 | N/F | 0.702 | N/A | 0.749 | N/F | N/F | N/A | N/F | N/A |  | Unknown
WO-23003-32-A-1 | N/F | 167.609 | 11.98 | -0.548 | N/F | N/F | -1448.898 | 2.118 | N/F | N/F | N/F | 0.553 | N/F | N/A | 0.354 | N/F | 0.385 | N/F | N/A | N/F | N/F | N/A | 0.346 | N/F | 0.394 | N/A | 0.757 | N/F | -5.256 | N/A | N/F | N/A |  | Unknown
WO-23003-33-A-1 | N/F | 66.955 | 11.216 | -1.442 | 0.371 | N/F | -1359.876 | N/F | 9.357 | N/F | N/F | 0.94 | N/F | N/A | 0.553 | N/F | 0.674 | N/F | N/A | N/F | 1.126 | N/A | 0.488 | 0.64 | 0.355 | N/A | 0.717 | N/F | -4.694 | N/A | N/F | N/A |  | Unknown
WO-23003-34-A-1 | N/F | 75.414 | 12.203 | -2.275 | 0.377 | N/F | -1629.974 | N/F | 9.979 | N/F | 0.94 | 0.906 | N/F | N/A | 0.509 | N/F | 0.663 | N/F | N/A | N/F | 0.991 | N/A | 0.521 | 0.721 | 0.484 | N/A | 0.766 | N/F | -5.233 | N/A | N/A | N/A |  | Unknown
WO-23003-35-A-1 | N/F | 84.423 | 8.691 | -1.053 | 0.41 | N/F | -1568.524 | N/F | 8.866 | 0.316 | 0.908 | 1.126 | N/F | N/A | 0.458 | N/F | 0.71 | N/F | N/A | N/F | 1.091 | N/A | 0.514 | 0.635 | 0.491 | N/A | 0.766 | N/F | -5.555 | N/A | N/F | N/A |  | Unknown
WO-23003-36-A-1 | N/F | 72.702 | 12.557 | -0.235 | 0.455 | N/F | -1442.188 | 0.99 | 8.682 | 0.215 | N/F | 1.314 | N/F | N/A | 0.639 | N/F | 0.763 | N/F | N/A | N/F | 0.885 | N/A | 0.587 | 0.726 | 0.516 | N/A | 0.77 | N/F | -5.188 | N/A | N/F | N/A |  | Unknown
WO-23003-37-A-1 | N/F | 94.073 | 11.154 | -0.679 | 0.364 | N/F | -1722.622 | N/F | 8.757 | 0.064 | 0.932 | 1.009 | N/F | N/A | 0.277 | N/F | 0.577 | N/F | N/A | N/F | 0.965 | N/A | 0.473 | N/F | 0.344 | N/A | 0.762 | N/F | -5.242 | N/A | N/F | N/A |  | Unknown
WO-23003-38-A-1 | N/F | 71.857 | 11.321 | -0.351 | N/F | N/F | -1985.937 | N/F | 10.034 | N/F | N/F | 1.074 | N/F | N/A | 0.438 | N/F | 0.639 | N/F | N/A | N/F | 0.903 | N/A | 0.435 | 0.629 | 0.433 | N/A | 0.729 | N/F | -4.867 | N/A | N/F | N/A |  | Unknown
WO-23003-39-A-1 | N/F | 96.773 | 14.017 | -1.129 | 0.343 | N/F | -2379.923 | 0.513 | N/F | N/F | N/F | 0.391 | N/F | N/A | 0.301 | 126.713 | 0.309 | N/F | N/A | 0.517 | N/F | N/A | 0.302 | N/F | 0.402 | N/A | 0.729 | N/F | -3.654 | N/A | N/A | N/A |  | Unknown
WO-23003-40-A-1 | N/F | 53.886 | 12.627 | -1.151 | 0.35 | N/F | -1682.319 | 1.336 | N/F | 0.141 | 0.769 | 0.454 | N/F | N/F | 0.211 | 134.042 | 0.518 | N/F | N/A | 0.224 | N/F | N/A | 0.443 | 0.698 | 0.33 | N/A | 0.75 | N/F | -4.24 | N/A | N/F | N/A |  | Unknown
WO-23003-41-A-1 | N/F | 69.688 | 10.95 | -1.808 | 0.845 | N/A | -1986.736 | 0.582 | 2.138 | 0.843 | 1.08 | 10.144 | N/F | N/A | 0.882 | 150.848 | 0.677 | N/F | N/A | N/F | N/F | N/A | 0.764 | 1.108 | 0.665 | N/A | 1.315 | N/F | -5.144 | N/A | N/F | N/A |  | Unknown
WO-23003-42-A-1 | N/F | 87.99 | 12.286 | -1.332 | 0.534 | N/F | -2104.218 | N/F | 1.32 | 0.63 | 0.822 | 3.824 | N/F | N/A | 0.589 | 136.634 | 0.442 | N/F | N/A | 0.452 | N/F | N/A | 0.466 | 0.808 | 0.454 | N/A | 0.887 | N/F | -4.282 | N/A | N/F | N/A |  | Unknown
WO-23003-43-A-1 | N/F | 71.186 | 12.53 | -0.713 | 0.444 | N/F | -1882.501 | N/F | N/F | 0.197 | N/F | 2.172 | N/F | N/A | 0.683 | 111.294 | 0.4 | N/F | N/A | 0.326 | N/F | N/A | 0.416 | 0.78 | 0.354 | N/A | 0.934 | N/F | -5.579 | N/A | N/F | N/A |  | Unknown
WO-23003-44-A-1 | N/F | 132.635 | 11.786 | -1.048 | 0.408 | N/F | -303.851 | 1.537 | 1.786 | 0.765 | 0.82 | 1.544 | N/F | N/A | 0.622 | 109.959 | 0.359 | N/F | N/A | 0.56 | N/F | N/A | 0.413 | 0.957 | 0.355 | N/A | 0.842 | N/F | -4.166 | N/A | N/F | N/A |  | Unknown
WO-23003-45-A-1 | N/F | 134.043 | 10.332 | -1.235 | 0.393 | N/F | -77.53 | N/F | 1.622 | 0.19 | 0.782 | 1.121 | N/F | N/A | 0.43 | 126.129 | 0.368 | N/F | N/A | 0.39 | N/F | N/A | 0.366 | 0.709 | 0.401 | N/A | 0.83 | N/F | -3.679 | N/A | N/F | N/A |  | Unknown
WO-23003-46-A-1 | N/F | 70.431 | 14.225 | -1.2 | 0.377 | N/F | -2293.717 | N/F | 1.285 | 1.032 | 0.914 | 1.038 | N/F | N/A | 0.441 | 117.077 | 0.45 | N/F | N/A | 2.305 | N/F | N/A | 0.427 | 0.832 | 0.328 | N/A | 0.786 | N/F | -4.634 | 0.215 | N/F | N/A |  | Unknown
WO-23003-47-A-1 | N/F | 140.708 | 19.502 | -1.238 | 0.37 | N/F | -3246.202 | 1.131 | N/F | 0.901 | N/F | 1.061 | N/F | N/A | 0.296 | 268.234 | 0.318 | N/F | N/A | N/F | N/F | N/A | 0.355 | 1.047 | 0.429 | N/A | 0.778 | N/F | -2.226 | N/A | N/F | N/A |  | Unknown
WO-23003-48-A-1 | N/F | 53.173 | 12.796 | 0.517 | 0.362 | N/F | -1807.114 | 0.421 | N/F | 0.268 | N/F | 0.744 | N/F | N/A | 0.481 | 253.554 | 0.334 | N/F | N/A | 0.424 | N/F | N/A | 0.363 | 0.69 | 0.491 | N/A | 0.779 | N/F | -4.234 | N/A | N/F | N/A |  | Unknown
WO-23003-49-A-1 | N/F | 75.231 | 12.672 | -0.308 | 0.353 | N/F | -1555.867 | N/F | 1.26 | 0.27 | N/F | 0.719 | N/F | N/A | 0.719 | 250.815 | 0.323 | N/F | N/A | 0.418 | N/F | N/A | 0.381 | 0.851 | 0.807 | N/A | 0.74 | N/F | -3.44 | N/A | N/F | N/A |  | Unknown
WO-23003-50-A-1 | N/F | 189.432 | 19.964 | -0.698 | 0.352 | N/F | -3016.487 | 2.007 | 2.099 | 0.963 | N/F | 0.744 | N/F | N/A | 0.432 | 214.775 | 0.264 | N/F | N/A | 6.252 | N/F | N/A | 0.388 | 0.718 | N/F | N/A | 0.76 | N/F | -2.066 | N/A | N/F | N/A |  | Unknown
WO-23003-51-A-1 | N/F | 148.297 | 14.144 | -1.405 | 0.349 | N/F | -1019.414 | N/F | 2.96 | 0.287 | 0.784 | 0.737 | N/F | N/A | 0.394 | 498.597 | 0.295 | N/F | N/A | N/F | N/F | N/A | 0.324 | 0.73 | 0.375 | N/A | 0.736 | N/F | -3.027 | N/A | N/A | N/A |  | Unknown
WO-23003-52-A-1 | N/F | 121.947 | 9.354 | -1.494 | N/F | N/F | -497.088 | 1.442 | 1.383 | 0.249 | N/F | 0.657 | N/F | N/A | 0.441 | 413.609 | 0.235 | N/F | N/A | 0.949 | N/F | N/A | 0.342 | 0.788 | 0.295 | N/A | 0.748 | N/F | -4.26 | N/A | N/F | N/A |  | Unknown
WO-23003-53-A-1 | N/F | 63.869 | 8.815 | -1.373 | 0.353 | N/F | -1726.66 | N/F | 2.725 | 0.101 | 0.805 | 0.556 | N/F | N/A | 0.521 | 419.335 | 0.428 | N/A | N/A | 0.413 | N/F | N/A | 0.362 | 0.763 | 0.355 | N/A | 0.714 | N/F | -4.253 | N/A | N/F | N/A |  | Unknown
WO-23003-54-A-1 | N/F | 99.103 | 15.857 | -1.1 | 0.364 | N/F | -2169.25 | N/F | 9.358 | 0.777 | 0.988 | 1.149 | N/F | N/A | 0.502 | 446.172 | 0.617 | N/F | N/A | 0.612 | N/F | N/A | 0.507 | 0.848 | 0.466 | N/A | 0.737 | N/F | -2.451 | N/A | N/A | N/A |  | Unknown
WO-23003-55-A-1 | N/F | 102.639 | 8.308 | -1.367 | N/F | N/F | -2473.552 | N/F | 7.447 | 0.436 | 0.901 | 1.348 | N/F | N/A | 0.269 | 456.386 | 1.323 | N/F | N/A | N/F | 0.584 | N/A | 0.541 | 0.933 | 0.523 | N/A | 0.75 | N/F | -2.361 | N/A | N/F | N/A |  | Unknown
WO-23003-56-A-1 | N/F | 109.563 | 13.594 | -1.21 | 0.339 | N/F | -1865.556 | N/F | 7.568 | 0.717 | 0.968 | 1.232 | N/F | N/A | 0.34 | 433.565 | 0.576 | N/A | N/A | 3.78 | N/F | N/A | 0.563 | 0.759 | 0.523 | N/A | 0.723 | N/F | -4.251 | N/A | N/F | N/A |  | Unknown
LFB-H-3-Inj2 | N/F | 140.66 | 208.14 | 213.786 | 245.586 | N/F | -435.592 | 196.02 | 183.118 | 164.246 | 164.593 | 168.317 | N/F | N/A | 193.778 | 174.08 | 197.419 | N/A | N/A | 190.077 | 191.147 | N/A | 195.058 | 157.581 | 193.469 | N/A | 180.139 | N/F | 670.231 | N/A | N/F | N/A |  | Unknown
WO-23003-14-A-1B | N/F | 53.523 | 7.992 | -2.221 | N/F | N/F | -1440.617 | N/F | 4236.334 | 0.441 | N/F | 0.898 | N/F | N/A | 0.334 | N/F | 1.981 | N/F | N/A | 3.569 | 0.614 | N/A | 0.383 | 0.749 | 0.47 | N/A | 0.712 | N/F | -1.737 | N/A | N/F | N/A |  | Unknown
LRB-7-inj3 | N/F | 64.564 | 12.519 | -1.97 | N/F | N/F | -2183.995 | 1.41 | 1.355 | 0.01 | 0.761 | 0.42 | N/F | N/A | 0.416 | 157.537 | 4.078 | N/F | N/A | 0.667 | N/F | N/A | 0.324 | 0.651 | 0.295 | N/A | 0.728 | N/F | -3.946 | N/A | N/F | N/A |  | Unknown
LRB-7-inj2 | N/F | -3.489 | -5.103 | -0.39 | N/F | N/F | -370.969 | N/F | N/F | -0.026 | 0.766 | 0.387 | N/F | N/A | 0.317 | N/F | 153.845 | N/F | N/A | 0.088 | N/F | N/A | 0.334 | 0.769 | 0.294 | N/A | 0.719 | N/F | -4.659 | N/A | N/A | N/A |  | Unknown
WO-23003-60-A-1B | N/F | 143.791 | 16.675 | 2.795 | 0.347 | N/F | -3252.894 | N/F | 1.537 | 0.543 | N/F | 0.336 | N/F | N/A | 0.867 | N/F | 4183.607 | N/F | N/A | 2.882 | N/F | N/A | 2.809 | 1.324 | 1.587 | N/A | 0.78 | N/F | -1.869 | N/A | N/F | N/A |  | Unknown
WO-23003-59-A-1B | N/F | 80.547 | 12.996 | -0.138 | 0.383 | N/F | -1872.514 | 0.234 | 1.119 | 0.678 | 0.834 | 0.539 | N/F | N/A | 0.564 | N/F | 632.89 | N/F | N/A | 2.687 | N/F | N/A | 0.839 | 0.754 | 0.616 | N/A | 0.738 | N/F | -4.476 | 4.404 | N/F | N/A |  | Unknown
WO-23003-58-A-1B | N/F | 110.815 | 13.451 | 0.336 | 0.391 | N/F | -846.279 | N/F | 1.919 | 0.554 | 0.895 | 0.816 | N/F | N/A | 0.367 | N/F | 61.125 | N/F | N/A | 4.18 | N/F | N/A | 1.257 | 0.923 | 0.492 | N/A | 0.739 | N/F | -4.533 | N/A | N/F | N/A |  | Unknown
WO-23003-26-A-1B | N/F | 140.856 | 17.933 | -0.461 | 0.357 | N/F | -2195.753 | N/F | 5.584 | 0.376 | N/F | 0.606 | N/F | N/A | 0.501 | N/F | 0.466 | N/F | N/A | -0.059 | N/F | N/A | 0.417 | 0.78 | 0.288 | N/A | 0.78 | N/F | -3.887 | N/A | N/A | N/A |  | Unknown
WO-23003-124-A-1 | N/F | -23.237 | -5.273 | -0.824 | 0.382 | N/F | 81.578 | N/F | N/F | 0.002 | 0.821 | -0.14 | 1124.729 | N/A | 1.035 | N/F | 4254.003 | N/F | N/A | -0.116 | 0.721 | N/A | 2.549 | 0.788 | 1.325 | N/A | 0.787 | N/F | -5.628 | N/A | N/F | N/A |  | Unknown
WO-23003-123-A-1 | N/F | -24.267 | -4.586 | 0.034 | 0.398 | N/F | -115.917 | N/F | N/F | 0.182 | 0.797 | -0.053 | N/F | N/A | 1.14 | N/F | 4114.272 | N/F | N/A | -0.091 | 0.825 | N/A | 3.799 | 1.074 | 1.448 | N/A | 0.803 | N/F | -5.716 | N/A | N/A | N/A |  | Unknown
WO-23003-122-A-1 | N/F | -21.73 | -4.194 | 0.848 | 0.459 | N/A | -261.406 | N/F | N/F | 0.174 | 0.815 | 0.037 | N/F | N/A | 1.131 | N/F | 4052.461 | N/F | N/A | 0.023 | 0.942 | N/A | 3.298 | 1.329 | 1.505 | N/A | 0.895 | N/F | -5.715 | N/A | N/F | N/A |  | Unknown
WO-23003-121-A-1 | N/F | 12.878 | 0.073 | 1.206 | 0.658 | N/A | -592.908 | N/F | N/F | 0.742 | 0.926 | 0.211 | N/F | N/A | 1.054 | N/F | 2598.122 | N/F | N/A | 2.996 | 1.724 | N/A | 8.418 | 1.699 | 8.728 | N/A | 1.117 | N/F | -5.129 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LFSM-H 6 | N/F | 194.659 | 197.206 | 204.181 | 183.464 | N/F | -751.578 | 171.837 | 163.133 | 222.285 | 150.338 | 158.364 | N/F | N/A | 185.672 | 160.732 | 2843.453 | N/F | N/A | 176.546 | 176.045 | N/A | 191.689 | 199.234 | 184.547 | N/A | 179.992 | N/F | 362.027 | N/A | N/A | 3945.889 |  | Unknown
AB-230117-02 CCC-H 7 | N/F | 313.969 | 378.56 | 351.172 | 422.357 | N/F | 386.541 | 372.196 | 313.775 | 352.805 | 306.664 | 311.988 | N/F | N/A | 363.775 | 307.839 | 358.688 | N/F | N/A | 351.037 | 355.298 | N/A | 350.888 | 274.256 | 349.262 | N/A | 331.124 | N/F | 676.091 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LFB-H 7 | N/F | 171.999 | 188.773 | 197.499 | 200.108 | N/F | 247.442 | 181.009 | 155.689 | 179.63 | 152.841 | 149.508 | 1139.203 | N/A | 180.109 | 157.151 | 303.883 | N/F | N/F | 177.143 | 182.079 | N/A | 177.827 | 136.419 | 177.55 | N/A | 160.163 | N/F | 328.706 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LD 6 | N/F | -23.21 | -3.529 | 4.204 | 0.334 | N/F | -99.91 | N/F | N/F | -0.011 | N/F | -0.105 | N/F | N/A | 1.04 | N/F | 4142.057 | N/A | N/A | -0.128 | 0.865 | N/A | 3.079 | 0.865 | 1.374 | N/A | 0.695 | N/F | -5.699 | N/A | N/A | N/A |  | Unknown
LD6-inj1 | N/A | -20.918 | -4.217 | 3.755 | 0.344 | N/F | 37.367 | N/F | N/F | 0.014 | 0.795 | -0.052 | N/F | N/A | 1.43 | N/F | 3922.707 | N/F | N/F | -0.179 | 0.883 | N/A | 2.893 | 0.723 | 1.362 | N/A | 0.739 | N/F | -5.529 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LRB 13 | N/F | -17.496 | -3.378 | -1.698 | 0.327 | N/F | -172.351 | 0.263 | 0.997 | -0.03 | 0.752 | 0.061 | N/F | N/A | 0.362 | N/F | 69.087 | N/F | N/A | N/F | 0.748 | N/A | 0.374 | 0.643 | 0.62 | N/A | 0.702 | N/F | -5.731 | N/A | N/A | N/A |  | Unknown
WO-23003-120-A-1 | N/A | -0.741 | 5.699 | -1.784 | 0.331 | N/F | -1072.482 | 0.615 | N/F | 0.08 | 0.761 | N/F | N/F | N/F | 0.718 | N/F | 1329.015 | N/F | N/A | 1.273 | 1.294 | N/A | 3.724 | 0.736 | 4.447 | N/A | 0.716 | N/F | -4.723 | N/A | N/A | N/A |  | Unknown
WO-23003-119-A-1 | N/F | 7.47 | 8.163 | -1.42 | N/F | N/F | -835.051 | N/F | 2.443 | 0.196 | 0.8 | -0.061 | N/F | N/A | 0.446 | N/F | 1085.45 | N/F | N/A | 1.297 | 1.049 | N/A | 2.851 | 0.725 | 3.789 | N/A | 0.726 | N/F | -5.284 | N/A | N/F | N/A |  | Unknown
WO-23003-118-A-1 | N/F | 5.418 | 24.206 | 0.391 | N/F | N/F | -725.587 | N/F | 3.543 | 0.019 | 0.761 | -0.032 | N/F | N/A | 0.865 | N/F | 2857.187 | N/F | N/A | 3.116 | 1.751 | N/A | 9.413 | 0.989 | 12.301 | N/A | 0.702 | N/F | -4.834 | N/A | N/F | N/A |  | Unknown
WO-23003-117-A-1 | N/F | 6.359 | 40.521 | -0.501 | 0.33 | N/F | -952.521 | 0.376 | N/F | 0.092 | 0.828 | -0.151 | N/F | N/A | 0.778 | N/F | 2511.685 | N/F | N/A | 2.496 | 1.608 | N/A | 8.243 | 0.785 | 10.943 | N/A | 0.716 | N/F | -5.124 | N/A | N/A | N/A |  | Unknown
WO-23003-116-A-1 | N/F | 4.468 | 7.755 | -0.774 | N/F | N/F | -838.898 | N/F | 3.453 | 0.103 | 0.756 | 0.044 | N/F | N/A | 0.705 | N/F | 2477.13 | N/F | N/A | 2.716 | 1.71 | N/A | 8.8 | 0.795 | 11.677 | N/A | 0.701 | N/F | -5.132 | 3.168 | N/F | N/A |  | Unknown
WO-23003-115-A-1 | N/F | 42.735 | 153.375 | -0.792 | N/F | N/F | -1485.127 | 0.437 | N/F | 0.13 | 0.841 | 0.047 | N/F | N/A | 0.393 | N/F | 714.013 | N/F | N/A | 2.204 | 1.365 | N/A | 1.895 | 0.686 | 2.371 | N/A | 0.718 | N/F | -4.09 | N/A | N/F | N/A |  | Unknown
WO-23003-114-A-1 | N/F | 30.584 | 45.908 | 0.349 | N/F | N/F | -1797.346 | N/F | 2.409 | 0.151 | 0.762 | -0.089 | N/F | N/A | 0.418 | N/F | 622.53 | N/A | N/A | 2.214 | 1.029 | N/A | 1.569 | 0.801 | 2.034 | N/A | 0.725 | N/F | -4.916 | N/A | N/F | N/A |  | Unknown
WO-23003-113-A-1 | N/F | 29.409 | 29.561 | -0.972 | N/F | N/F | -1928.643 | N/F | 2.379 | 0.462 | 0.746 | -0.04 | N/F | N/A | 0.476 | N/F | 641.876 | N/A | N/A | 0.381 | 0.765 | N/A | 1.514 | 0.726 | 2.136 | N/A | 0.741 | N/F | -5.183 | N/A | N/A | N/A |  | Unknown
WO-23003-112-A-1 | N/F | 103.462 | 69.81 | 0.392 | 0.335 | N/F | -1720.371 | N/F | 4.125 | 0.349 | N/F | 0.084 | N/F | N/A | 0.277 | N/F | 536.317 | N/F | N/A | 2.721 | 1.139 | N/A | 1.711 | 0.793 | 2.811 | N/A | 0.692 | N/F | -4.792 | N/A | N/A | N/A |  | Unknown
WO-23003-111-A-1 | N/F | 38.064 | 63.909 | -0.484 | N/F | N/F | -1837.754 | N/F | 4.153 | 0.435 | N/F | 0.379 | N/F | N/A | 0.384 | N/F | 740.764 | N/A | N/A | 3.217 | 1.35 | N/A | 2.348 | 0.756 | 4.279 | N/A | 0.685 | N/F | -4.5 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LRB 12 | N/F | -15.737 | -3.219 | -0.499 | 0.342 | N/F | -64.152 | N/F | 1.042 | 0.073 | N/F | 0.02 | N/F | N/A | 0.438 | N/F | 1.003 | N/F | N/A | -0.25 | 0.758 | N/A | 0.278 | 0.675 | 0.534 | N/A | 0.762 | N/F | -5.279 | N/A | N/A | N/A |  | Unknown
LRB-12-inj1 | N/F | -21.937 | -0.563 | -0.016 | 0.337 | N/F | 40.195 | N/F | N/F | 0.016 | 0.791 | -0.026 | N/F | N/A | 0.274 | N/F | 24.998 | N/A | N/A | -0.156 | 0.799 | N/A | 0.302 | 0.713 | 0.648 | N/A | 0.744 | N/F | -5.699 | N/A | N/A | N/A |  | Unknown
WO-23003-110-A-1 | N/F | 44.332 | 222.256 | -1.25 | 0.341 | N/F | -1591.716 | N/F | 4.918 | 0.234 | N/F | -0.005 | N/F | N/A | 0.419 | N/F | 666.017 | N/F | N/A | 1.827 | 1.761 | N/A | 2.107 | 0.737 | 3.639 | N/A | 0.725 | N/F | -4.831 | N/A | N/F | N/A |  | Unknown
WO-23003-109-A-1 | N/F | 30.793 | 64.865 | -1.51 | 0.34 | N/F | -1755.039 | N/F | 1.175 | 0.294 | N/F | 0.117 | N/F | N/A | 0.497 | N/F | 137.434 | N/F | N/A | -0.263 | 0.53 | N/A | 0.506 | 0.719 | 0.575 | N/A | 0.712 | N/F | -5.335 | N/A | N/A | N/A |  | Unknown
WO-23003-108-A-1 | N/A | 66.078 | 142.55 | -1.305 | 0.357 | N/F | -2030.64 | N/F | N/F | 0.258 | 0.794 | -0.006 | N/F | N/A | 0.426 | N/F | 117.876 | N/A | N/A | 0.083 | N/F | N/A | 0.484 | 0.637 | 0.504 | N/A | 0.741 | N/F | -4.723 | N/A | N/A | N/A |  | Unknown
WO-23003-107-A-1 | N/F | 30.047 | 38.99 | -1.589 | 0.35 | N/F | -1440.237 | N/F | N/F | 0.52 | N/F | 0.048 | N/F | N/A | 0.271 | N/F | 141.587 | N/F | N/A | N/F | N/F | N/A | 0.464 | 0.805 | 0.668 | N/A | 0.742 | N/F | -5.388 | N/A | N/A | N/A |  | Unknown
WO-23003-106-A-1 | N/F | 43.438 | 42.492 | -1.36 | 0.367 | N/F | -1668.765 | 0.23 | N/F | 0.602 | 0.727 | 0.016 | N/F | N/A | 0.632 | N/F | 137.351 | N/A | N/A | N/F | N/F | N/A | 0.515 | 0.744 | 0.742 | N/A | 0.745 | N/F | -5.593 | N/A | N/F | N/A |  | Unknown
WO-23003-105-A-1 | N/F | 79.408 | 118.301 | -1.063 | 0.373 | N/F | -2176.952 | 0.4 | N/F | 0.267 | 0.781 | 0.119 | N/F | N/A | 0.427 | N/F | 128.03 | N/F | N/A | 1.54 | N/F | N/A | 0.695 | 0.707 | 0.819 | N/A | 0.776 | N/F | -5.072 | N/A | N/F | N/A |  | Unknown
WO-23003-104-A-1 | N/F | 28.431 | 42.646 | -1.607 | 0.387 | N/A | -1298.381 | N/F | N/F | 0.218 | 0.815 | 0.093 | N/F | N/A | 0.413 | N/F | 128.717 | N/A | N/A | 0.148 | N/F | N/A | 0.618 | 0.777 | 0.713 | N/A | 0.804 | N/F | -5.253 | N/A | N/F | N/A |  | Unknown
WO-23003-103-A-1 | N/F | 48.723 | 139.401 | -1.964 | 0.424 | N/F | -1609.503 | N/F | N/F | 0.241 | N/F | 0.047 | N/F | N/A | 0.5 | N/F | 66.78 | N/F | N/A | 1.829 | 0.975 | N/A | 0.525 | 0.761 | 0.702 | N/A | 0.833 | N/F | -6.185 | N/A | N/F | N/A |  | Unknown
WO-23003-102-A-1 | N/F | 69.052 | 172.024 | -0.869 | 0.541 | N/F | -1540.208 | 0.788 | N/F | 1.153 | N/F | 0.352 | N/F | N/A | 0.481 | N/F | 64.801 | N/F | N/A | 0.173 | 1.032 | N/A | 0.666 | 0.881 | 0.732 | N/A | 0.985 | N/F | -5.777 | N/A | N/F | N/A |  | Unknown
WO-23003-101-A-1 | N/F | 61.53 | 156.344 | -1.137 | 0.691 | N/F | -1834.692 | N/F | N/F | 0.578 | 0.928 | 0.365 | N/F | N/A | 0.599 | N/F | 28.256 | N/A | N/A | 0.093 | 0.912 | N/A | 0.649 | 0.911 | 0.618 | N/A | 1.162 | N/F | -5.518 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LFSM-H 5 | N/F | 107.228 | 345.773 | 191.07 | 211.268 | N/F | -2016.365 | 178.008 | 163.798 | 202.628 | 151.084 | 165.247 | N/F | N/A | 193.47 | 182.099 | 223.297 | N/F | N/A | 175.928 | 176.535 | N/A | 186.755 | 189.923 | 182.066 | N/A | 183.253 | N/F | 466.75 | N/A | N/F | N/A |  | Unknown
AB-230117-02 CCC-H 6 | N/F | 289.127 | 371.455 | 391.326 | 421.057 | N/F | 443.96 | 352.766 | 298.434 | 329.233 | 297.736 | 300.62 | N/F | N/A | 349.783 | 304.261 | 341.563 | N/F | N/A | 337.342 | 339.3 | N/A | 341.007 | 285.514 | 333.08 | N/A | 315.445 | N/F | 615.557 | N/A | N/F | 5650.988 |  | Unknown
AB-230117-02 LFB-H 6 | N/F | 155.833 | 177.057 | 189.557 | 210.176 | N/F | 137.426 | 179.368 | 149.25 | 176.216 | 148.382 | 151.608 | 1199.01 | N/A | 177.686 | 158.298 | 176.054 | N/F | N/A | 172.683 | 175.354 | N/A | 176.742 | 143.941 | 173.631 | N/A | 159.314 | N/F | 331.989 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LD 5 | N/F | 37.363 | 134.103 | -1.164 | N/F | N/F | -1974.742 | N/F | 1.186 | 0.249 | 0.756 | 0.096 | N/F | N/A | 0.293 | N/F | 88.093 | N/F | N/A | 0.066 | 0.77 | N/A | 0.438 | 0.74 | 0.645 | N/A | 0.725 | N/F | -4.489 | N/A | N/F | N/A |  | Unknown
LD-5-inj1 | N/F | 46.612 | 136.51 | -0.986 | 0.336 | N/F | -1800.431 | N/F | N/F | 0.062 | 0.744 | 0.089 | N/F | N/A | 0.216 | N/F | 82.9 | N/F | N/A | -0.175 | 0.802 | N/A | 0.477 | 0.657 | 0.657 | N/A | 0.713 | N/F | -4.724 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LRB 11 | N/F | -17.567 | -3.443 | -1.285 | N/F | N/A | 7.149 | 0.345 | N/F | -0.049 | 0.727 | 0.017 | N/F | N/A | 0.401 | N/F | 0.264 | N/F | N/A | N/F | 0.756 | N/A | 0.267 | 0.643 | 0.55 | N/A | 0.701 | N/F | -5.676 | N/A | N/A | N/A |  | Unknown
WO-23003-100-A-1 | N/F | 27.513 | 11.115 | -2.307 | N/F | N/F | -1773.975 | 0.321 | 1.922 | 0.295 | 0.717 | 0.128 | N/F | N/A | 0.451 | 410.536 | 0.224 | N/F | N/A | 0.007 | N/F | N/A | 0.31 | 0.629 | 0.224 | N/A | 0.696 | N/F | -4.966 | N/A | N/F | N/A |  | Unknown
WO-23003-99-A-1 | N/F | 45.806 | 12.489 | -1.836 | N/F | N/F | -2297.103 | N/F | N/F | 0.255 | 0.792 | -0.013 | N/F | N/A | 0.277 | 447.416 | 0.186 | N/A | N/A | N/F | N/F | N/A | 0.309 | 0.614 | 0.262 | N/A | 0.699 | N/F | -4.181 | N/A | N/F | N/A |  | Unknown
WO-23003-98-A-1 | N/A | 24.011 | 12.584 | -1.408 | N/F | N/F | -1446.733 | N/F | N/F | 0.341 | 0.769 | 0.116 | N/F | N/A | 0.3 | 403.126 | 0.211 | N/F | N/A | 1.427 | N/F | N/A | 0.335 | 0.728 | 0.294 | N/A | 0.725 | N/F | -5.191 | N/A | N/F | N/A |  | Unknown
WO-23003-97-A-1 | N/F | 46.302 | 12.895 | -1.532 | N/F | N/F | -1567.275 | 0.489 | N/F | 0.432 | 0.753 | 0.188 | N/F | N/A | 0.34 | 382.008 | 0.572 | N/F | N/A | 0.216 | N/F | N/A | 0.355 | 0.646 | 0.314 | N/A | 0.694 | N/F | -4.759 | N/A | N/F | N/A |  | Unknown
WO-23003-96-A-1 | N/F | 26.128 | 12.667 | -1.004 | N/F | N/F | -1481.869 | N/F | N/F | 0.283 | 0.736 | 0.043 | N/F | N/A | 0.249 | 396.134 | 0.168 | N/F | N/A | N/F | N/F | N/A | 0.291 | 0.681 | 0.191 | N/A | 0.715 | N/F | -4.851 | N/A | N/F | N/A |  | Unknown
WO-23003-95-A-1 | N/F | 29.325 | 9.807 | -1.682 | 0.333 | N/F | -1637.525 | N/F | N/F | 0.307 | N/F | 0.007 | N/F | N/A | 0.263 | 396.948 | 0.195 | N/F | N/A | N/F | N/F | N/A | 0.271 | 0.678 | 0.24 | N/A | 0.734 | N/F | -5.363 | N/A | N/A | N/A |  | Unknown
WO-23003-94-A-1 | N/F | 31.35 | 11.508 | -0.332 | N/F | N/F | -1478.751 | N/F | N/F | 0.304 | N/F | 0.093 | N/F | N/A | 0.419 | 396.942 | 0.307 | N/F | N/A | N/F | N/F | N/A | 0.352 | 0.713 | 0.27 | N/A | 0.716 | N/F | -4.846 | N/A | N/F | N/A |  | Unknown
WO-23003-93-A-1 | N/F | 49.528 | 11.324 | -1.572 | N/F | N/F | -2004.192 | N/F | 2.816 | 0.127 | N/F | 0.097 | N/F | N/A | 0.4 | 377.744 | 0.324 | N/F | N/A | -0.196 | N/F | N/A | 0.391 | 0.67 | 0.277 | N/A | 0.705 | N/F | -4.939 | N/A | N/F | N/A |  | Unknown
WO-23003-92-A-1 | N/F | 18.741 | 6.111 | -2.068 | N/F | N/F | -1393.945 | N/F | N/F | 0.215 | 0.76 | 0.23 | N/F | N/A | 0.493 | 385.868 | 0.174 | N/F | N/A | N/F | N/F | N/A | 0.343 | 0.787 | 0.312 | N/A | 0.704 | N/F | -4.863 | N/A | N/F | N/A |  | Unknown
WO-23003-91-A-1 | N/F | 64.998 | 15.662 | -2.262 | 0.337 | N/F | -2158.633 | N/F | N/F | 0.461 | 0.749 | 0.114 | N/F | N/A | 0.446 | 385.432 | 0.408 | N/F | N/A | N/F | N/F | N/A | 0.317 | 0.656 | 0.38 | N/A | 0.708 | N/F | -4.358 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LRB 10 | N/F | -16.448 | -2.021 | -1.941 | 0.34 | N/F | 161.946 | 0.079 | 1.395 | 0.014 | 0.766 | 0.228 | N/F | N/A | 0.357 | N/F | 0.321 | N/F | N/A | -0.318 | 0.848 | N/A | 0.264 | 0.653 | 0.571 | N/A | 0.757 | N/F | -5.396 | N/A | N/F | N/A |  | Unknown
LRB-10-inj1 | N/F | -16.798 | -2.485 | -2.314 | 0.34 | N/F | 96.671 | N/F | 1.09 | 0.021 | 0.754 | 0.056 | N/F | N/A | 0.419 | N/F | 0.283 | N/A | N/A | -0.337 | 0.79 | N/A | 0.261 | 0.669 | 0.515 | N/A | 0.733 | N/F | -5.676 | N/A | N/F | N/A |  | Unknown
WO-23003-90-A-1 | N/F | 66.885 | 13.107 | -1.876 | 0.341 | N/F | -1728.648 | N/F | 2.291 | 0.445 | 0.792 | 0.152 | N/F | N/A | 0.455 | 365.485 | 0.371 | N/F | N/A | -0.091 | N/F | N/A | 0.321 | 0.665 | 0.344 | N/A | 0.724 | N/F | -4.599 | N/A | N/F | N/A |  | Unknown
WO-23003-89-A-1 | N/F | 37.172 | 11.98 | -1.478 | 0.344 | N/F | -1555.855 | N/F | 1.046 | 0.322 | 0.752 | 0.075 | N/F | N/A | 0.244 | 415.919 | 0.192 | N/F | N/A | N/F | N/F | N/A | 0.291 | 0.721 | 0.284 | N/A | 0.744 | N/F | -4.289 | N/A | N/A | N/A |  | Unknown
WO-23003-88-A-1 | N/F | 28.868 | 14.826 | -1.622 | 0.346 | N/F | -1353.534 | N/F | N/F | 0.429 | N/F | 0.092 | N/F | N/A | 0.284 | 382.709 | 0.223 | N/F | N/A | -0.142 | N/F | N/A | 0.342 | 0.683 | 0.287 | N/A | 0.736 | N/F | -5.104 | N/A | N/F | N/A |  | Unknown
WO-23003-87-A-1 | N/F | 90.659 | 18.952 | -1.724 | 0.357 | N/F | -2512.26 | N/F | 1.572 | 0.25 | 0.749 | 0.201 | N/F | N/A | 0.414 | 406.54 | 0.219 | N/F | N/A | 0.359 | N/F | N/A | 0.368 | 0.735 | 0.359 | N/A | 0.705 | N/F | -4.907 | N/A | N/F | N/A |  | Unknown
WO-23003-86-A-1 | N/F | 52.311 | 17.014 | -1.452 | 0.369 | N/F | -2232.3 | 0.871 | 2.233 | 0.465 | 0.75 | 0.071 | N/F | N/A | 0.367 | 433.77 | 0.246 | N/F | N/A | 0.022 | N/F | N/A | 0.431 | 0.719 | 0.35 | N/A | 0.731 | N/F | -4.784 | N/A | N/F | N/A |  | Unknown
WO-23003-85-A-1 | N/F | 22.091 | 12.5 | -1.934 | 0.38 | N/F | -1405.98 | N/F | N/F | 0.353 | 0.74 | 0.158 | N/F | N/A | 0.456 | 419.227 | 0.326 | N/F | N/A | N/F | N/F | N/A | 0.395 | 0.739 | 0.31 | N/A | 0.75 | N/F | -4.307 | N/A | N/F | N/A |  | Unknown
WO-23003-84-A-1 | N/F | 31.744 | 10.608 | -1.128 | 0.392 | N/A | -1266.276 | N/F | N/F | 0.478 | 0.782 | 0.214 | N/F | N/A | 0.741 | 405.654 | 0.229 | N/A | N/A | 0.733 | N/F | N/A | 0.405 | 0.731 | 0.327 | N/A | 0.803 | N/F | -5.085 | N/A | N/F | N/A |  | Unknown
WO-23003-83-A-1 | N/F | 46.21 | 12.053 | -1.422 | 0.434 | N/F | -1406.442 | N/F | 1.468 | 0.54 | 0.804 | 0.184 | N/F | N/A | 0.45 | 311.52 | 0.398 | N/F | N/A | N/F | N/F | N/A | 0.424 | 0.927 | 0.375 | N/A | 0.845 | N/F | -5.075 | N/A | N/F | N/A |  | Unknown
WO-23003-82-A-1 | N/F | 93.384 | 33.548 | -0.169 | 0.642 | N/F | -1171.651 | 2.353 | 2.483 | 0.593 | 0.97 | 0.518 | N/F | N/A | 0.654 | 602.933 | 0.63 | N/F | N/A | N/F | N/F | N/A | 0.705 | 0.958 | 0.828 | N/A | 1.221 | N/F | -2.783 | N/A | N/F | N/A |  | Unknown
WO-23003-81-A-1 | N/F | 30.769 | 10.576 | -1.403 | 0.746 | N/F | -1414.582 | N/F | 1.055 | 1.117 | 0.984 | 0.375 | N/F | N/A | 0.575 | 319.278 | 0.574 | N/F | N/A | N/F | N/F | N/A | 0.62 | 1.164 | 0.541 | N/A | 1.121 | N/F | -4.163 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LFSM-H 4 | N/F | 141.547 | 230.253 | 235.884 | 231.311 | N/F | -1461.066 | 200.054 | 174.965 | 262.399 | 164.339 | 176.253 | N/F | N/A | 208.366 | 189.256 | 213.294 | N/F | N/A | 191.747 | 184.951 | N/A | 204.634 | 216.789 | 197.378 | N/A | 199.573 | N/F | 389.584 | N/A | N/F | N/A |  | Unknown
AB-230117-02 CCC-H 5 | N/F | 311.228 | 363.733 | 402.453 | 435.261 | N/F | 894.413 | 375.935 | 301.325 | 364.192 | 313.327 | 323.051 | 1260.747 | N/A | 357.871 | 309.052 | 355.469 | N/F | N/A | 349.56 | 355.016 | N/A | 357.216 | 299.892 | 353.128 | N/A | 332.741 | N/F | 629.614 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LFB-H 5 | N/F | 163.23 | 198.722 | 203.039 | 213.609 | N/F | 432.648 | 179.705 | 157 | 179.847 | 152.771 | 156.013 | 1216.497 | N/A | 182.108 | 158.258 | 180.317 | N/A | N/A | 177.279 | 179.093 | N/A | 181.423 | 157.635 | 179.201 | N/A | 168.236 | N/F | 345.389 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LD 4 | N/F | 118.358 | 18.244 | -0.735 | 0.334 | N/F | -2842.877 | 1.281 | 1.608 | 0.633 | 0.794 | 0.374 | N/F | N/A | 0.482 | N/F | 3.87 | N/F | N/A | 0.465 | N/F | N/A | 0.314 | 0.795 | 0.502 | N/A | 0.712 | N/F | -1.788 | N/A | N/F | N/A |  | Unknown
LD4-inj1 | N/F | 70.143 | 15.537 | -0.94 | N/F | N/F | -2317.693 | 1.695 | N/F | 0.488 | 0.807 | 0.243 | N/F | N/A | 0.8 | N/F | 3.644 | N/F | N/A | 0.552 | N/F | N/A | 0.342 | 0.68 | 0.397 | N/A | 0.714 | N/F | -4.845 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LRB 9 | N/F | -14.4 | -3.528 | -0.887 | N/F | N/F | 164.97 | N/F | 1.019 | 0.015 | 0.693 | 0.153 | N/F | N/A | 0.39 | N/F | 0.312 | N/F | N/A | -0.198 | 0.794 | N/A | 0.229 | 0.656 | 0.552 | N/A | 0.707 | N/F | -5.512 | N/A | N/F | N/A |  | Unknown
WO-23003-80-A-1 | N/F | 71.46 | 17.773 | -1.63 | N/F | N/F | -2164.04 | N/F | 2.611 | 0.199 | 0.781 | 0.069 | N/F | N/A | 0.358 | 354.309 | 0.25 | N/F | N/A | 1.633 | N/F | N/A | 0.326 | 0.642 | 0.352 | N/A | 0.685 | N/F | -4.458 | N/A | N/F | N/A |  | Unknown
WO-23003-79-A-1 | N/F | 50.524 | 12.127 | -1.481 | N/F | N/F | -1620.783 | N/F | 1.386 | 0.255 | N/F | 0.161 | N/F | N/A | 0.373 | 341.696 | 0.401 | N/A | N/A | 1.21 | N/F | N/A | 0.384 | 0.791 | 0.404 | N/A | 0.722 | N/F | -4.709 | N/A | N/F | N/A |  | Unknown
WO-23003-78-A-1 | N/F | 30.148 | 13.933 | -1.639 | N/F | N/F | -1435.327 | N/F | 1.06 | 0.288 | N/F | 0.047 | N/F | N/A | 0.318 | 343.62 | 0.283 | N/F | N/A | N/F | N/F | N/A | 0.301 | 0.741 | 0.336 | N/A | 0.717 | N/F | -4.589 | N/A | N/A | N/A |  | Unknown
WO-23003-77-A-1 | N/A | 34.735 | 9.843 | -1.371 | N/F | N/F | -1294.126 | 0.689 | 1.2 | 0.116 | 0.81 | 0.067 | N/F | N/A | 0.317 | 212.264 | 0.393 | N/F | N/A | 0.046 | N/F | N/A | 0.385 | 0.612 | 0.35 | N/A | 0.736 | N/F | -4.597 | N/A | N/F | N/A |  | Unknown
WO-23003-76-A-1 | N/F | 31.722 | 11.695 | -1.806 | N/F | N/F | -1538.305 | N/F | 1.534 | 0.457 | 0.781 | 0.24 | N/F | N/A | 0.587 | 201.051 | 1.751 | N/F | N/A | N/F | N/F | N/A | 0.35 | 0.724 | 0.32 | N/A | 0.7 | N/F | -4.458 | N/A | N/A | N/A |  | Unknown
WO-23003-75-A-1 | N/F | 65.003 | 10.994 | -1.445 | N/F | N/F | -1689.705 | N/F | 1.73 | 0.374 | 0.805 | 0.486 | N/F | N/A | 0.528 | N/F | 46.694 | N/F | N/A | 3.079 | N/F | N/A | 0.51 | 0.806 | 0.403 | N/A | 0.699 | N/F | -4.766 | N/A | N/F | N/A |  | Unknown
WO-23003-74-A-1 | N/F | 30.233 | 10.449 | -0.248 | N/F | N/F | -1296.292 | N/F | N/F | 0.448 | 0.765 | 0.3 | N/F | N/A | 0.382 | N/F | 33.419 | N/F | N/A | N/F | N/F | N/A | 0.357 | 0.722 | 0.311 | N/A | 0.713 | N/F | -5.195 | N/A | N/F | N/A |  | Unknown
WO-23003-73-A-1 | N/F | 50.056 | 14.758 | -0.628 | 0.339 | N/F | -1754.151 | N/F | 1.375 | 0.54 | N/F | 0.31 | N/F | N/A | 0.413 | N/F | 16.405 | N/F | N/A | 0.523 | N/F | N/A | 0.365 | 0.848 | 0.532 | N/A | 0.709 | N/F | -4.489 | N/A | N/F | N/A |  | Unknown
WO-23003-72-A-1 | N/F | 53.994 | 11.49 | -1.927 | 0.344 | N/F | -1730.007 | N/F | 1.154 | 0.445 | 0.795 | 0.446 | N/F | N/A | 0.468 | N/F | 19.333 | N/F | N/A | N/F | N/F | N/A | 0.315 | 0.869 | 0.515 | N/A | 0.715 | N/F | -4.144 | N/A | N/F | N/A |  | Unknown
WO-23003-71-A-1 | N/F | 40.276 | 11.754 | -1.05 | 0.339 | N/F | -1370.883 | 0.602 | N/F | 0.549 | N/F | 0.134 | N/F | N/A | 0.474 | N/F | 14.541 | N/F | N/A | N/F | N/F | N/A | 0.328 | 0.721 | 0.33 | N/A | 0.776 | N/F | -5.312 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LRB 8 | N/F | -6.171 | -1.453 | -1.693 | 0.344 | N/F | 85.351 | N/F | 0.941 | 0.211 | 0.791 | 0.287 | N/F | N/A | 0.334 | N/F | 0.257 | N/A | N/A | -0.341 | N/F | N/A | 0.277 | 0.805 | 0.409 | N/A | 0.74 | N/F | -4.928 | N/A | N/F | N/A |  | Unknown
LRB-8-inj1 | N/F | -10.631 | -3.094 | -1.403 | 0.344 | N/F | 115.431 | 0.346 | 1.025 | 0.204 | 0.693 | 0.344 | N/F | N/F | 0.49 | N/F | 0.783 | N/F | N/A | N/F | N/F | N/A | 0.28 | 0.796 | 0.355 | N/A | 0.752 | N/F | -5.019 | N/A | N/F | N/A |  | Unknown
WO-23003-70-A-1 | N/F | 59.943 | 12.318 | -0.121 | 0.344 | N/F | -1951.131 | N/F | 2.141 | 0.198 | N/F | 0.38 | N/F | N/A | 0.503 | N/F | 14.244 | N/F | N/A | 0.081 | N/F | N/A | 0.408 | 0.767 | 0.365 | N/A | 0.755 | N/F | -4.06 | N/A | N/F | N/A |  | Unknown
WO-23003-69-A-1 | N/F | 36.985 | 12.159 | -0.985 | 0.348 | N/F | -1775.526 | N/F | 1.175 | 0.533 | 0.763 | 0.297 | N/F | N/A | 0.532 | N/F | 9.563 | N/F | N/A | 0.446 | N/F | N/A | 0.312 | 0.873 | 0.524 | N/A | 0.751 | N/F | -4.755 | N/A | N/F | N/A |  | Unknown
WO-23003-68-A-1 | N/F | 38.538 | 11.795 | -0.159 | 0.352 | N/F | -1777.291 | N/F | 1.443 | 0.515 | 0.812 | 0.434 | N/F | N/A | 0.51 | N/F | 10.214 | N/F | N/A | 2.645 | N/F | N/A | 0.389 | 0.793 | 0.382 | N/A | 0.758 | N/F | -4.565 | N/A | N/F | N/A |  | Unknown
WO-23003-67-A-1 | N/F | 36.888 | 10.917 | -1.54 | 0.361 | N/F | -1821.798 | N/F | 1.384 | 0.543 | 0.804 | 0.308 | N/F | N/A | 0.484 | N/F | 5.711 | N/F | N/A | 0.304 | N/F | N/A | 0.338 | 0.721 | 0.362 | N/A | 0.788 | N/F | -4.29 | N/A | N/A | N/A |  | Unknown
WO-23003-66-A-1 | N/F | 35.802 | 10.592 | -1.263 | 0.365 | N/F | -1633.299 | N/F | 1.191 | 0.079 | N/F | 0.383 | N/F | N/A | 0.378 | N/F | 5.647 | N/F | N/A | -0.001 | N/F | N/A | 0.344 | 0.851 | 0.338 | N/A | 0.818 | N/F | -3.613 | N/A | N/F | N/A |  | Unknown
WO-23003-65-A-1 | N/A | 76.978 | 12.387 | -0.688 | 0.38 | N/F | -2378.334 | N/F | 1.691 | 0.572 | N/F | 0.359 | N/F | N/A | 0.693 | N/F | 8.885 | N/F | N/A | 0.579 | 0.25 | N/A | 0.34 | 0.96 | 0.317 | N/A | 0.787 | N/F | -3.606 | N/A | N/F | N/A |  | Unknown
WO-23003-64-A-1 | N/F | 82.787 | 13.315 | -0.483 | 0.4 | N/F | -2433.644 | N/F | 1.537 | 0.857 | 0.772 | 0.402 | N/F | N/A | 0.669 | N/F | 12.026 | N/F | N/A | 0.586 | N/F | N/A | 0.475 | 0.959 | 0.38 | N/A | 0.878 | N/F | -3.782 | N/A | N/F | N/A |  | Unknown
WO-23003-63-A-1 | N/F | 47.787 | 10.717 | -0.877 | 0.43 | N/F | -2266.741 | 0.45 | 1.341 | 0.673 | 0.833 | 0.443 | N/F | N/A | 0.672 | N/F | 9.272 | N/F | N/A | 0.739 | N/F | N/A | 0.432 | 0.993 | 0.39 | N/A | 0.873 | N/F | -4.976 | N/A | N/F | N/A |  | Unknown
WO-23003-62-A-1 | N/F | 35.68 | 11.608 | -0.986 | 0.489 | N/F | -2156.762 | N/F | 2.118 | 0.63 | 0.924 | 0.508 | N/F | N/A | 0.591 | N/F | 5.986 | N/F | N/A | 0.075 | N/F | N/A | 0.45 | 0.757 | 0.494 | N/A | 0.9 | N/F | -3.534 | N/A | N/F | N/A |  | Unknown
WO-23003-61-A-1 | N/F | 53.245 | 14.969 | 0.409 | 0.744 | N/F | -2454.472 | N/F | 1.448 | 1.301 | N/F | 0.735 | N/F | N/A | 1.275 | N/F | 4.3 | N/F | N/A | 1.168 | N/F | N/A | 0.672 | 1.339 | 0.576 | N/A | 1.23 | N/F | -2.834 | 10.286 | N/F | N/A |  | Unknown
AB-230117-02 LFSM-H 3 | N/F | 131.12 | 222.169 | 223.168 | 217.1 | N/F | -2625.739 | 193.406 | 157.155 | 288.344 | 163.49 | 169.947 | N/F | N/A | 205.752 | 334.468 | 203.756 | N/F | N/A | 187.492 | 190.483 | N/A | 198.421 | 216.121 | 191.067 | N/A | 195.334 | N/F | 578.701 | N/A | N/F | N/A |  | Unknown
AB-230117-02 CCC-H 4 | N/F | 394.354 | 374.225 | 425.492 | 444.887 | N/F | -782.4 | 394.235 | 330.463 | 406.982 | 340.248 | 335.865 | N/F | N/A | 391.26 | 333.561 | 378.791 | N/F | N/A | 372.36 | 375.413 | N/A | 378.57 | 356.247 | 375.67 | N/A | 346.062 | N/F | 1138.264 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LFB-H 4 | N/F | 128.037 | 202.802 | 228.56 | 218.871 | N/F | -919.956 | 199.569 | 176.623 | 191.501 | 161.298 | 165.666 | 1149.754 | N/A | 197.291 | 160.162 | 192.196 | N/F | N/A | 190.159 | 187.965 | N/A | 195.326 | 175.359 | 186.073 | N/A | 173.955 | N/F | 643.923 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LD 3 | N/F | 68.015 | 14.791 | -1.662 | 0.342 | N/F | -2643.539 | N/F | 1.333 | 0.803 | 0.799 | 0.394 | N/F | N/A | 0.269 | 140.88 | 1.737 | N/F | N/A | 0.617 | N/F | N/A | 0.41 | 0.646 | 0.387 | N/A | 0.724 | N/F | -1.95 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LRB 7 | N/A | 552.849 | 367.278 | 6.887 | 0.936 | N/F | 1001.123 | N/F | 13.867 | 8.257 | N/F | 16.487 | N/F | N/A | 2.82 | N/F | 77.188 | N/F | N/A | N/F | N/F | N/A | 4.824 | 2.663 | 14.3 | N/A | 2.626 | N/F | 212.947 | N/A | N/A | N/A |  | Unknown
LRB7-inj1 | N/A | 4.902 | -2.659 | -1.134 | 0.353 | N/F | -898.494 | N/F | 0.998 | 0.379 | 0.794 | 0.502 | N/F | N/A | 0.433 | N/F | 27.187 | N/F | N/A | -0.106 | N/F | N/A | 0.308 | 0.795 | 0.326 | N/A | 0.768 | N/F | -4.049 | N/A | N/F | N/A |  | Unknown
WO-23003-60-A-1 | N/F | 226.438 | 12.429 | 3.175 | 0.383 | N/F | -3516.853 | N/F | 2.425 | 0.686 | 0.818 | 0.342 | N/F | N/A | 0.794 | N/F | 4256.541 | N/F | N/A | 3.233 | N/F | N/A | 2.954 | 1.342 | 1.794 | N/A | 0.755 | N/F | -2.292 | N/A | N/A | N/A |  | Unknown
WO-23003-59-A-1 | N/F | 92.981 | 11.907 | -1.497 | 0.421 | N/A | -2262.582 | N/F | N/F | 0.208 | 0.848 | 0.383 | N/F | N/A | 0.69 | N/F | 660.136 | N/F | N/A | 0.057 | N/F | N/A | 0.846 | 0.977 | 0.656 | N/A | 0.951 | N/F | -3.04 | N/A | N/F | N/A |  | Unknown
WO-23003-58-A-1 | N/F | 108.162 | 14.295 | 1.014 | 0.5 | N/A | -2303.618 | N/F | N/F | 1.006 | 0.922 | 0.735 | N/F | N/A | 0.697 | N/F | 60.711 | N/F | N/A | 0.903 | N/F | N/A | 1.322 | 1.186 | 0.547 | N/A | 0.931 | N/F | -3.172 | N/A | N/F | N/A |  | Unknown
WO-23003-138-A-1 | N/F | -12.818 | 2.295 | -2.043 | N/F | N/F | -1442.374 | N/F | 41.842 | 0.238 | 0.961 | 0.02 | N/F | N/A | 0.288 | N/F | 0.368 | N/F | N/A | 0.174 | N/F | N/A | 0.284 | 0.651 | 0.209 | N/A | 0.735 | N/F | -4.874 | N/A | N/F | N/A |  | Unknown
WO-23003-137-A-1 | N/F | -9.688 | 3.611 | -1.534 | N/F | N/F | -824.309 | N/F | 30.824 | 0.201 | N/F | -0.072 | N/F | N/A | 0.283 | N/F | 0.386 | N/F | N/A | N/F | N/F | N/A | 0.258 | 0.628 | 0.319 | N/A | 0.714 | N/F | -5.245 | 1.639 | N/F | N/A |  | Unknown
WO-23003-136-A-1 | N/F | 52.844 | 13.284 | -1.427 | N/F | N/F | -3093.314 | 1.066 | 36.314 | 0.427 | 0.689 | -0.035 | N/F | N/A | 0.293 | N/F | 0.437 | N/F | N/A | N/F | N/F | N/A | 0.251 | 0.656 | 0.251 | N/A | 0.697 | N/F | -4.117 | N/A | N/F | N/A |  | Unknown
WO-23003-135-A-1 | N/F | -1.856 | 9.115 | -1.979 | N/F | N/F | -854.404 | N/F | 34.102 | 0.181 | 0.754 | 0.159 | N/F | N/A | 0.201 | N/F | 0.443 | N/F | N/A | 0.786 | N/F | N/A | 0.333 | 0.642 | 0.412 | N/A | 0.707 | N/F | -5.041 | 3.528 | N/A | N/A |  | Unknown
WO-23003-134-A-1 | N/F | -1.97 | 9.035 | -2.13 | N/F | N/F | -1117.89 | 0.198 | 38.926 | 0.282 | 0.762 | -0.051 | N/F | N/A | 0.318 | N/F | 0.414 | N/A | N/A | -0.151 | N/F | N/A | 0.261 | 0.73 | 0.371 | N/A | 0.703 | N/F | -4.286 | N/A | N/F | N/A |  | Unknown
WO-23003-133-A-1 | N/F | -55.247 | 14.845 | -1.049 | N/F | N/F | -1525.586 | 1.203 | 33.78 | 0.427 | N/F | 0.106 | N/F | N/A | 0.218 | N/F | 0.523 | N/A | N/A | N/F | N/F | N/A | 0.295 | 0.631 | 0.367 | N/A | 0.723 | N/F | -6.271 | N/A | N/A | N/A |  | Unknown
WO-23003-132-A-1 | N/F | 29.492 | 12.305 | -1.795 | N/F | N/F | -2502.402 | N/F | 16.223 | 0.256 | N/F | -0.029 | N/F | N/A | 0.233 | N/F | 0.77 | N/F | N/A | N/F | N/F | N/A | 1.47 | 0.715 | 1.113 | N/A | 0.71 | N/F | -4.262 | N/A | N/F | N/A |  | Unknown
WO-23003-131-A-1 | N/F | N/F | 5.382 | -1.716 | N/F | N/F | -1073.406 | 0.329 | 14.312 | 0.425 | N/F | -0.101 | N/F | N/A | 0.128 | N/F | 0.748 | N/F | N/A | -0.044 | N/F | N/A | 0.306 | 0.666 | 0.366 | N/A | 0.697 | N/F | -5.575 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LRB 14 | N/A | -38.979 | -2.48 | -2.26 | N/F | N/F | 847.72 | N/F | N/F | -0.009 | 0.756 | -0.011 | N/F | N/A | 0.294 | N/F | 0.664 | N/F | N/F | -0.311 | 0.844 | N/A | 0.274 | 0.639 | 0.577 | N/A | 0.713 | N/F | -5.985 | N/A | N/F | N/A |  | Unknown
LRB-14-inj1 | N/F | -40.573 | -3.646 | -1.938 | N/F | N/F | 1038.216 | N/F | 1.102 | 0.067 | 0.769 | -0.07 | N/F | N/A | 0.434 | N/F | 0.68 | N/F | N/A | -0.028 | 0.799 | N/A | 0.247 | 0.672 | 0.486 | N/A | 0.733 | N/F | -5.863 | 0.733 | N/A | N/A |  | Unknown
WO-23003-130-A-1 | N/F | 17.822 | 6.711 | -1.729 | N/F | N/F | -2282.252 | 0.482 | 13.469 | 0.118 | N/F | -0.025 | N/F | N/A | 0.414 | N/F | 0.778 | N/F | N/A | 0.613 | N/F | N/A | 0.296 | 0.69 | 0.481 | N/A | 0.696 | N/F | -3.997 | N/A | N/F | N/A |  | Unknown
WO-23003-129-A-1 | N/F | 10.44 | 16.63 | -1.245 | N/F | N/F | -1476.605 | N/F | 12.556 | 0.388 | 0.771 | 0.236 | N/F | N/A | 0.136 | N/F | 1.157 | N/A | N/A | 0.115 | N/F | N/A | 0.345 | 0.697 | 0.451 | N/A | 0.72 | N/F | -6.01 | N/A | N/A | N/A |  | Unknown
WO-23003-128-A-1 | N/F | 5.716 | 22.32 | -1.218 | N/F | N/F | -1838.63 | 0.672 | 9.095 | 0.61 | 0.774 | 0.093 | N/F | N/A | 0.346 | N/F | 1.569 | N/A | N/A | 0.138 | N/F | N/A | 0.347 | 0.664 | 0.583 | N/A | 0.734 | N/F | -6.272 | N/A | N/A | N/A |  | Unknown
WO-23003-127-A-1 | N/F | 1.443 | 19.169 | -1.18 | N/F | N/F | -1053.048 | N/F | 9.456 | 0.199 | 0.839 | 0.029 | N/F | N/A | 0.171 | N/F | 1.887 | N/F | N/A | N/F | N/F | N/A | 0.375 | 0.688 | 0.451 | N/A | 0.699 | N/F | -4.841 | N/A | N/F | N/A |  | Unknown
WO-23003-126-A-1 | N/F | 27.436 | 29.326 | -0.999 | N/F | N/F | -2192.888 | N/F | 9.643 | 0.757 | 0.809 | 0.267 | N/F | N/A | 0.331 | N/F | 4.716 | N/F | N/A | N/F | N/F | N/A | 0.412 | 0.73 | N/F | N/A | 0.778 | N/F | -6.01 | N/A | N/A | N/A |  | Unknown
WO-23003-125-A-1 | N/F | 14.694 | 7.22 | -1.089 | 0.334 | N/F | -2482.221 | N/F | 11.776 | 0.391 | N/F | 0.089 | N/F | N/A | 0.223 | N/F | 46.972 | N/F | N/A | N/F | N/F | N/A | 0.413 | 0.67 | 0.512 | N/A | 0.723 | N/F | -3.829 | N/A | N/F | N/A |  | Unknown
WO-23003-121-A-1B | N/F | -18.215 | 1.726 | 0.68 | 0.341 | N/F | -245.723 | 0.285 | N/F | 0.086 | N/F | -0.034 | N/F | N/A | 0.515 | N/F | 2538.308 | N/A | N/A | 2.383 | 1.486 | N/A | 7.928 | 0.987 | 8.541 | N/A | 0.72 | N/F | -4.768 | N/A | N/A | N/A |  | Unknown
WO-23003-111-A-1B | N/F | 9.45 | 65.578 | -0.942 | 0.346 | N/F | -2215.917 | N/F | 4.439 | 0.39 | 0.805 | 0.434 | N/F | N/A | 0.32 | N/F | 738.645 | N/F | N/A | 1.983 | 1.749 | N/A | 2.374 | 0.679 | 4.534 | N/A | 0.761 | N/F | -3.773 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LRB 13 20230123105625 | N/F | -66.085 | -2.845 | -1.953 | 0.358 | N/F | 991.182 | 0.169 | N/F | 0.139 | 0.82 | -0.001 | N/F | N/A | 0.342 | N/F | 3.33 | N/F | N/A | -0.178 | 0.805 | N/A | 0.281 | 0.728 | 0.61 | N/A | 0.713 | N/F | -5.692 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LRB 12 20230123102918 | N/F | 3916.074 | 1541.758 | 55.619 | 4.726 | N/F | 8701.812 | N/F | N/F | 12.919 | 10.094 | 58.764 | N/F | N/A | 4.153 | N/F | 943.114 | N/A | N/A | 37.615 | N/F | N/A | 26.569 | 22.213 | 49.489 | N/A | 7.11 | N/F | 145.848 | N/A | N/A | N/A | C | Unknown
AB-230117-02 QCS 2 | N/F | 194.759 | 241.344 | 253.37 | 299.949 | N/A | 955.974 | 224.431 | 202.348 | 231.625 | 192.851 | 217.642 | 1183.236 | N/A | 25.094 | 194.121 | 220.165 | N/F | N/F | 217.546 | 212.761 | N/A | 232.69 | 24.852 | 218.131 | N/A | 22.632 | N/A | 390.264 | N/A | N/F | 4510.311 |  | Unknown
AB-230117-02 CCC-H 12 | N/A | 344.411 | 399.094 | 436.054 | 553.34 | N/F | 1331.427 | 391.777 | 353.417 | 386.979 | 327.791 | 783.523 | N/F | N/A | 371.954 | 343.096 | 372.448 | N/F | N/A | 379.928 | 383.581 | N/A | 373.921 | 316.256 | 375.367 | N/A | 342.557 | N/A | 626.903 | N/A | N/A | N/A |  | Unknown
WO-23003-212-A-1 | N/F | -43.519 | N/F | -1.551 | 0.387 | N/F | 913.895 | 0.224 | 2.35 | 144.855 | 5.409 | 6474.482 | N/F | N/A | 0.304 | N/F | 0.284 | N/A | N/A | -0.046 | 0.803 | N/A | N/F | 9.022 | 1.148 | N/A | 0.702 | N/F | -5.8 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LRB 22 | N/F | -42.936 | -5.511 | -1.769 | 0.343 | N/F | 819.687 | 0.161 | N/F | 9.133 | 1.076 | 429.568 | N/F | N/A | 0.252 | N/F | 0.267 | N/F | N/F | -0.274 | 0.756 | N/A | 0.186 | 1.27 | 0.574 | N/A | 0.733 | N/F | -5.831 | N/A | N/A | N/A |  | Unknown
WO-23003-211-A-1 | N/F | -45.103 | N/F | -1.405 | 0.394 | N/F | 884.261 | N/F | 3.07 | 151.883 | 5.436 | 6424.609 | N/F | N/A | 0.337 | N/F | 0.265 | N/F | N/A | -0.043 | 0.916 | N/A | N/F | 9.216 | 1.074 | N/A | 0.764 | N/F | -6.139 | N/A | N/A | N/A |  | Unknown
WO-23003-210-A-1 | N/A | -41.937 | N/F | -1.724 | 0.407 | N/F | 584.76 | 0.152 | 3.382 | 204.552 | 8.704 | 7113.022 | N/F | N/A | 0.28 | N/F | 0.289 | N/F | N/F | 0.026 | 0.762 | N/A | N/F | 9.065 | 1.497 | N/A | 0.717 | N/F | -6.04 | N/A | N/A | N/A |  | Unknown
WO-23003-209-A-1 | N/F | -31.4 | 81.281 | -1.319 | 0.455 | N/F | -127.386 | 0.264 | 95.224 | 1176.751 | 166.268 | 14233.119 | N/F | N/A | 0.324 | N/F | 0.31 | N/F | N/A | 8.603 | 4.472 | N/A | N/F | 3.664 | 25.391 | N/A | 0.761 | N/A | -2.241 | N/A | N/A | N/A |  | Unknown
WO-23003-208-A-1 | N/F | -29.157 | 115.362 | -1.529 | 0.48 | N/F | -790.029 | 0.392 | 88.49 | 1223.778 | 163.904 | 13164.028 | N/F | N/A | 0.322 | N/F | 0.26 | N/F | N/A | 8.83 | 4.923 | N/A | N/F | 3.549 | 24.868 | N/A | 0.725 | N/F | -2.037 | N/A | N/A | N/A |  | Unknown
WO-23003-207-A-1 | N/F | -26.48 | 177.239 | -0.32 | 0.466 | N/F | -818.512 | N/F | 96.687 | 1200.616 | 165.874 | 13668.116 | N/F | N/A | 0.421 | N/F | 0.258 | N/F | N/A | 8.724 | 4.559 | N/A | N/F | 3.564 | 26.636 | N/A | 0.792 | N/F | -3.037 | N/A | N/A | N/A |  | Unknown
WO-23003-206-A-1 | N/F | -31.701 | 162.26 | -1.684 | 0.596 | N/F | -561.891 | 0.479 | 95.979 | 1412.768 | 192.84 | 19784.952 | N/F | N/A | 0.398 | N/F | 0.308 | N/F | N/A | 10.752 | 5.258 | N/A | N/F | 4.966 | 35.231 | N/A | 0.796 | N/F | -2.494 | 2.752 | N/A | N/A |  | Unknown
WO-23003-204-A-1 | N/F | 0.19 | 553.285 | -0.826 | 0.663 | N/F | -2157.675 | 0.521 | 208.811 | 2148.827 | 369.784 | 18254.898 | N/F | N/A | 0.196 | N/F | 0.31 | N/F | N/A | 10.348 | 5.247 | N/A | N/F | 4.485 | 29.505 | N/A | 0.768 | N/F | 4.524 | N/A | N/A | N/A |  | Unknown
WO-23003-203-A-1 | N/F | -0.832 | 200.704 | -1.725 | 0.876 | N/F | -2402.212 | 1.134 | 206.316 | 2263.57 | 359.206 | 20466.741 | N/F | N/A | 0.323 | N/F | 0.358 | N/A | N/A | 13.398 | 6.695 | N/A | N/F | 4.759 | 39.631 | N/A | 1.242 | N/F | 3.743 | N/A | N/F | N/A |  | Unknown
WO-23003-202-A-1 | N/F | -56.939 | 31.723 | -1.523 | 0.899 | N/A | -1431.887 | N/F | 150.15 | 990.295 | 127.028 | 6345.164 | N/F | N/A | 0.356 | N/F | 0.397 | N/F | N/A | 5.912 | 4.087 | N/A | N/F | 2.23 | 11.115 | N/A | 0.857 | N/F | -0.673 | N/A | N/F | N/A |  | Unknown
WO-23003-201-A-1 | N/F | -7.287 | 25.273 | -1.857 | 1.739 | N/F | -1705.164 | 1.429 | 152.553 | 1040.549 | 123.987 | 6600.341 | N/F | N/A | 0.544 | N/F | 0.604 | N/A | N/A | 6.13 | 3.72 | N/A | 0.809 | 2.858 | 11.225 | N/A | 1.249 | N/F | -1.174 | 7.232 | N/A | N/A |  | Unknown
AB-230117-02 LFSM-H 10 | N/F | 128.719 | 226.564 | 213.162 | 651.363 | N/F | -1322.557 | 183.533 | 306.812 | 1106.092 | 280.201 | 6432.403 | N/F | N/A | 190.036 | 194.12 | 194.699 | N/F | N/A | 188.5 | 188.607 | N/A | 192.309 | 203.781 | 190.703 | N/A | 216.072 | N/A | 328.134 | N/A | N/A | N/A |  | Unknown
AB-230117-02 CCC-H 11 | N/F | 300.023 | 395.726 | 396.305 | 487.898 | N/F | 1568.535 | 366.619 | 334.538 | 351.634 | 312.89 | 331.453 | N/F | N/A | 343.062 | 309.141 | 360.917 | N/A | N/A | 350.017 | 355.509 | N/A | 358.493 | 284.364 | 355.982 | N/A | 329.659 | N/A | 530.876 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LFB-H 11 | N/F | 152.333 | 204.71 | 204.191 | 257.431 | N/F | 1204.746 | 195.443 | 176.561 | 227.636 | 164.505 | 441.806 | N/F | N/A | 185.518 | 163.487 | 186.414 | N/F | N/A | 184.715 | 186.635 | N/A | 185.609 | 144.059 | 183.897 | N/A | 178.769 | N/A | 332.108 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LD 10 | N/F | 42.941 | 32.975 | 0.037 | N/F | N/F | -3038.081 | 1.057 | 133.374 | 1325.043 | 118.336 | 6584.537 | N/F | N/A | 0.25 | N/F | 0.273 | N/A | N/A | 6.23 | 3.438 | N/A | N/F | 2.272 | 10.794 | N/A | 0.839 | N/F | 1.278 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LRB 21 | N/A | -41.454 | -0.196 | -0.996 | N/F | N/F | 223.459 | 0.064 | 1.048 | 2.381 | 0.868 | 8.959 | 1147.569 | N/A | 0.233 | N/F | 0.325 | N/A | N/F | -0.166 | 0.722 | N/A | 0.237 | 0.645 | 0.609 | N/A | 0.703 | N/F | -5.54 | N/A | N/F | N/A |  | Unknown
LRB-21-inj1 | N/F | -39.945 | -4.307 | -0.301 | N/F | N/F | 680.04 | N/F | N/F | 28.612 | 2.409 | 238.564 | N/F | N/A | 0.154 | N/F | 0.291 | N/A | N/A | -0.279 | 0.929 | N/A | 0.233 | 0.701 | 0.653 | N/A | 0.695 | N/F | -5.831 | N/A | N/A | N/A |  | Unknown
WO-23003-200-A-1 | N/F | -21.673 | 11.614 | -1.999 | 0.354 | N/F | -1523.371 | N/F | 143.538 | 958.021 | 118.094 | 6059.278 | N/F | N/A | 0.212 | N/F | 0.206 | N/A | N/A | 7.358 | 4.582 | N/A | N/F | 1.891 | 13.529 | N/A | 0.68 | N/F | -0.853 | N/A | N/F | N/A |  | Unknown
WO-23003-199-A-1 | N/A | -18.023 | 3.719 | -1.723 | N/F | N/F | -1266.07 | 0.607 | 162.514 | 648.524 | 94.187 | 2990.294 | N/F | N/A | 0.102 | N/F | 0.248 | N/F | N/A | 4.615 | 3.086 | N/A | 0.239 | 1.29 | 6.684 | N/A | 0.714 | N/F | -1.323 | N/A | N/A | N/A |  | Unknown
WO-23003-198-A-1 | N/A | 12.293 | 28.004 | -1.55 | N/F | N/F | -2565.66 | 0.694 | 178.451 | 1475.424 | 193.513 | 6757.926 | N/F | N/A | 0.114 | N/F | 0.238 | N/A | N/A | 7.529 | 4.452 | N/A | 0.435 | 2.078 | 14.717 | N/A | 0.694 | N/F | 2.569 | N/A | N/A | 7.778 |  | Unknown
WO-23003-197-A-1 | N/A | -20.849 | 41.121 | -1.054 | N/F | N/F | -1826.768 | 1.222 | 176.471 | 1247.438 | 200.815 | 6886.659 | N/F | N/A | 0.229 | N/F | 0.24 | N/F | N/A | 6.997 | 3.823 | N/A | N/F | 1.888 | 13.769 | N/A | 0.683 | N/F | -0.26 | N/A | N/F | N/A |  | Unknown
WO-23003-196-A-1 | N/F | -1.925 | 12.065 | -1.276 | 0.337 | N/F | -2415.74 | 1.277 | 43.583 | 97.719 | 9.73 | 484.625 | N/F | N/A | 0.148 | N/F | 0.188 | N/F | N/A | 2.449 | 1.056 | N/A | 0.256 | 0.799 | 0.806 | N/A | 0.705 | N/F | -4.25 | N/A | N/A | N/A |  | Unknown
WO-23003-195-A-1 | N/F | -11.084 | 6.283 | -1.489 | N/F | N/F | -1778.528 | 0.301 | 40.614 | 91.362 | 9.416 | 473.503 | N/F | N/A | 0.194 | N/F | 0.201 | N/F | N/A | 0.492 | 1.117 | N/A | N/F | 0.76 | 0.858 | N/A | 0.722 | N/F | -5.232 | N/A | N/A | N/A |  | Unknown
WO-23003-194-A-1 | N/F | -12.964 | 4.344 | -1.359 | 0.335 | N/F | -1508.543 | 1.246 | 48.031 | 92.375 | 9.528 | 504.393 | N/F | N/A | 0.092 | N/F | 0.209 | N/A | N/A | 0.885 | 0.838 | N/A | 0.277 | 0.784 | 0.905 | N/A | 0.697 | N/F | -5.635 | N/A | N/A | N/A |  | Unknown
WO-23003-193-A-1 | N/F | -10.094 | 4.523 | -1.824 | N/F | N/F | -1802.627 | 0.206 | 36.046 | 68.422 | 7.354 | 373.256 | N/F | N/A | 0.227 | N/F | 0.163 | N/F | N/A | 0.623 | 0.838 | N/A | 0.258 | 0.791 | 0.633 | N/A | 0.708 | N/F | -5.368 | N/A | N/F | N/A |  | Unknown
WO-23003-192-A-1 | N/F | -13.081 | 5.086 | -1.145 | N/F | N/F | -1762.546 | N/F | 81.655 | 86.017 | 11.749 | 359.261 | N/F | N/A | 0.077 | N/F | 0.175 | N/F | N/A | 0.9 | 1.124 | N/A | 0.327 | 0.755 | 0.889 | N/A | 0.738 | N/F | -4.848 | N/A | N/F | N/A |  | Unknown
WO-23003-191-A-1 | N/F | 21.416 | 11.863 | -1.616 | N/F | N/F | -2832.986 | 1.379 | 77.507 | 103.961 | 12.512 | 345.638 | N/F | N/A | 0.146 | N/F | 0.168 | N/A | N/A | 2.901 | 1.531 | N/A | 0.292 | 0.702 | 0.847 | N/A | 0.706 | N/F | -4.537 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LRB 20 | N/F | -41.142 | -4.56 | -2.104 | 0.345 | N/F | 493.29 | 0.168 | N/F | 0.11 | 0.761 | 0.331 | N/F | N/A | 0.351 | N/F | 0.282 | N/F | N/A | -0.093 | 0.841 | N/A | 0.242 | 0.631 | 0.497 | N/A | 0.709 | N/F | -5.868 | N/A | N/F | N/A |  | Unknown
LRB-20-inj1 | N/A | -44.736 | -2.555 | -2.075 | 0.343 | N/F | 476.427 | N/F | N/F | 1.686 | 0.972 | 12.246 | N/F | N/A | 0.284 | N/F | 0.307 | N/A | N/A | -0.216 | 0.763 | N/A | 0.244 | 0.671 | 0.521 | N/A | 0.746 | N/F | -5.845 | N/A | N/A | N/A |  | Unknown
WO-23003-190-A-1 | N/F | 0.095 | 8.321 | -1.484 | 0.356 | N/F | -2279.318 | N/F | 86.45 | 102.394 | 13.315 | 377.598 | N/F | N/A | 0.212 | N/F | 0.385 | N/F | N/A | 0.705 | 1.55 | N/A | 0.329 | 0.8 | 1.102 | N/A | 0.733 | N/F | -4.737 | N/A | N/A | N/A |  | Unknown
WO-23003-189-A-1 | N/F | -13.482 | 5.643 | -2.237 | 0.356 | N/F | -1394.553 | 0.778 | 91.382 | 114.655 | 16.124 | 457.157 | N/F | N/A | 0.144 | N/F | 0.38 | N/F | N/A | 2.061 | 1.662 | N/A | 0.474 | 0.865 | 0.995 | N/A | 0.736 | N/F | -2.879 | N/A | N/A | N/A |  | Unknown
WO-23003-188-A-1 | N/F | -12.861 | 7.688 | -1.516 | 0.35 | N/F | -1705.624 | N/F | 6.698 | 8.309 | 1.423 | 61.169 | N/F | N/A | 0.211 | N/F | 0.203 | N/A | N/A | 1.083 | N/F | N/A | 0.293 | 0.645 | 0.342 | N/A | 0.737 | N/F | -6.02 | N/A | N/F | N/A |  | Unknown
WO-23003-187-A-1 | N/F | -20.249 | 6.438 | -1.089 | 0.35 | N/F | -1707.246 | N/F | 4.362 | 7.271 | 1.404 | 52.99 | N/F | N/A | 0.127 | N/F | 0.23 | N/F | N/A | -0.171 | N/F | N/A | 0.235 | 0.857 | 0.27 | N/A | 0.752 | N/F | -4.997 | N/A | N/F | N/A |  | Unknown
WO-23003-186-A-1 | N/F | -12.638 | 5.887 | -2.198 | 0.381 | N/F | -1683.477 | 0.239 | 5.009 | 7.175 | 1.147 | 53.928 | N/F | N/A | 0.193 | N/F | 0.224 | N/F | N/A | -0.237 | N/F | N/A | 0.258 | 0.724 | 0.37 | N/A | 0.755 | N/F | -4.55 | N/A | N/A | N/A |  | Unknown
WO-23003-185-A-1 | N/F | -13.235 | 5.964 | -1.515 | 0.377 | N/F | -1633.735 | 0.51 | 12.554 | 6.019 | 1.402 | 36.731 | N/F | N/A | 0.324 | N/F | 0.241 | N/F | N/A | N/F | 0.431 | N/A | 0.368 | 0.754 | 0.377 | N/A | 0.775 | N/F | -4.406 | N/A | N/A | N/A |  | Unknown
WO-23003-184-A-1 | N/F | -5.964 | 5.151 | -1.562 | 0.391 | N/F | -1776.215 | 0.63 | 13.196 | 6.464 | 1.546 | 41.182 | N/F | N/A | 0.277 | N/F | 0.249 | N/A | N/A | N/F | N/F | N/A | 0.302 | 0.692 | 0.413 | N/A | 0.84 | N/F | -4.461 | N/A | N/F | N/A |  | Unknown
WO-23003-183-A-1 | N/F | -14.563 | 7.857 | -1.479 | 0.429 | N/F | -1486.633 | 0.59 | 15.276 | 7.612 | 1.663 | 43.91 | N/F | N/A | 0.144 | N/F | 0.325 | N/F | N/A | N/F | N/F | N/A | 0.295 | 0.943 | 0.39 | N/A | 0.866 | N/F | -5.135 | 10.713 | N/F | N/A |  | Unknown
WO-23003-182-A-1 | N/F | -11.502 | 5.53 | -1.84 | 0.504 | N/F | -1690.644 | 0.34 | N/F | 2.138 | 0.971 | 14.04 | N/F | N/A | 0.386 | N/F | 0.316 | N/F | N/A | -0.123 | N/F | N/A | 0.391 | 0.763 | 0.464 | N/A | 0.899 | N/F | -5.7 | N/A | N/F | N/A |  | Unknown
WO-23003-181-A-1 | N/F | -10.962 | 5.199 | -2.194 | 0.758 | N/A | -1618.915 | 0.867 | N/F | 2.185 | 1.055 | 15.408 | N/F | N/A | 0.746 | N/F | 0.526 | N/A | N/A | N/F | N/F | N/A | 0.524 | 0.917 | 0.611 | N/A | 1.146 | N/F | -5.297 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LFSM-H 9 | N/F | 139.186 | 217.799 | 212.529 | 238.907 | N/F | -2901.434 | 184.196 | 171.966 | 299.168 | 152.805 | 184.114 | N/F | N/A | 196.369 | 168.758 | 189.486 | N/A | N/A | 181.188 | 179.695 | N/A | 190.975 | 189.6 | 182.596 | N/A | 184.04 | N/A | 438.485 | N/A | N/F | N/A |  | Unknown
AB-230117-02 CCC-H 10 | N/F | 295.609 | 365.898 | 365.56 | 456.743 | N/F | 1643.252 | 356.398 | 324.182 | 328.58 | 311.762 | 317.881 | N/F | N/A | 334.928 | 301.419 | 342.023 | N/A | N/F | 331.083 | 350.601 | N/A | 348.019 | 287.328 | 341.141 | N/A | 315.458 | N/A | 530.431 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LFB-H 10 | N/F | 133.231 | 182.03 | 184.535 | 232.188 | N/F | 1127.595 | 180.078 | 158.45 | 182.873 | 149.734 | 146.83 | N/F | N/A | 169.126 | 151.291 | 173.244 | N/A | N/A | 172.611 | 170.015 | N/A | 172.512 | 141.182 | 171.506 | N/A | 157.495 | N/A | 274.234 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LD 9 | N/F | -5.796 | 6.395 | -0.961 | N/F | N/F | -2158.372 | 0.632 | N/F | 1.929 | N/F | 14.245 | 1183.691 | N/A | 0.302 | N/F | 0.236 | N/A | N/A | N/F | N/F | N/A | 0.269 | 0.765 | 0.32 | N/A | 0.707 | N/F | -5.318 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LRB 19 | N/F | -43.385 | -3.285 | -1.295 | N/F | N/F | 684.629 | N/F | N/F | 0.011 | 0.753 | -0.2 | N/F | N/A | 0.192 | N/F | 0.287 | N/F | N/A | -0.299 | 0.827 | N/A | 0.277 | 0.674 | 0.512 | N/A | 0.7 | N/F | -5.845 | N/A | N/F | N/A |  | Unknown
LRB-19-inj1 | N/F | -42.612 | -3.18 | -1.619 | N/F | N/F | 776.597 | 0.179 | N/F | 0.022 | 0.723 | 0.66 | N/F | N/A | 0.22 | N/F | 0.281 | N/F | N/A | -0.201 | 0.906 | N/A | 0.249 | 0.637 | 0.519 | N/A | 0.72 | N/F | -5.705 | N/A | N/F | N/A |  | Unknown
WO-23003-180-A-1 | N/F | 6.748 | 18.366 | -1.707 | N/F | N/F | -2101.02 | N/F | N/F | 2.451 | N/F | 15.624 | N/F | N/A | 0.125 | N/F | 0.228 | N/F | N/A | N/F | N/F | N/A | 0.282 | 0.678 | 0.341 | N/A | 0.743 | N/F | -5.381 | N/A | N/A | N/A |  | Unknown
WO-23003-179-A-1 | N/F | -5.418 | 39.223 | -0.005 | N/F | N/F | -1779.836 | 0.544 | N/F | 1.11 | N/F | 7.783 | N/F | N/A | 0.377 | N/F | 0.249 | N/A | N/A | 2.121 | N/F | N/A | N/F | 0.719 | 0.376 | N/A | 0.706 | N/F | -6.269 | N/A | N/A | N/A |  | Unknown
WO-23003-178-A-1 | N/F | -8.932 | 19.196 | -0.41 | N/F | N/F | -1737.963 | 0.198 | N/F | 1.536 | 0.914 | 8.837 | N/F | N/A | 0.171 | N/F | 0.206 | N/A | N/A | 1.865 | N/F | N/A | 0.258 | 0.645 | 0.166 | N/A | 0.706 | N/F | -6.199 | 0.018 | N/A | 12.207 |  | Unknown
WO-23003-177-A-1 | N/F | -5.386 | 12.386 | -0.782 | N/F | N/F | -2103.018 | N/F | N/F | 1.644 | 0.877 | 10.965 | N/F | N/A | 0.279 | N/F | 0.246 | N/F | N/A | 0.36 | N/F | N/A | 0.285 | 0.681 | 0.378 | N/A | 0.694 | N/F | -5.397 | N/A | N/A | N/A |  | Unknown
WO-23003-176-A-1 | N/F | -13.509 | 9.567 | -2.164 | N/F | N/F | -1751.345 | 1.229 | N/F | 0.837 | N/F | 5.099 | N/F | N/A | 0.154 | N/F | 0.183 | N/F | N/A | N/F | N/F | N/A | 0.249 | 0.726 | 0.374 | N/A | 0.69 | N/F | -5.457 | N/A | N/A | N/A |  | Unknown
WO-23003-175-A-1 | N/F | -11.424 | 24.809 | -1.512 | N/F | N/A | -1674.736 | 0.341 | N/F | 0.766 | 0.81 | 5.044 | N/F | N/A | 0.243 | N/F | 0.19 | N/F | N/A | -0.345 | N/F | N/A | 0.286 | 0.716 | 0.339 | N/A | 0.708 | N/F | -5.193 | N/A | N/F | N/A |  | Unknown
WO-23003-174-A-1 | N/F | -1.624 | 63.983 | -1.911 | N/F | N/F | -1906.857 | 0.845 | N/F | 0.723 | 0.855 | 4.496 | N/F | N/A | 0.085 | N/F | 0.229 | N/F | N/A | -0.199 | N/F | N/A | 0.271 | 0.695 | 0.338 | N/A | 0.722 | N/F | -5.019 | N/A | N/F | N/A |  | Unknown
WO-23003-173-A-1 | N/F | 58.578 | 306.028 | -1.23 | N/F | N/F | -2842.957 | 0.325 | N/F | 0.877 | 0.887 | 5.228 | N/F | N/A | 0.183 | N/F | 0.267 | N/A | N/A | 1.921 | N/F | N/A | 0.305 | 0.725 | 0.354 | N/A | 0.693 | N/F | -4.188 | N/A | N/A | N/A |  | Unknown
WO-23003-172-A-1 | N/F | -10.455 | 31.79 | -1.081 | N/F | N/F | -1965.934 | 0.751 | N/F | 1.06 | 0.774 | 20.708 | N/F | N/A | 0.207 | N/F | 0.282 | N/A | N/A | 0.213 | N/F | N/A | 0.29 | 0.719 | 0.17 | N/A | 0.691 | N/F | -4.895 | N/A | N/A | N/A |  | Unknown
WO-23003-171-A-1 | N/F | -17.248 | 11.19 | -1.526 | N/F | N/F | -1354.166 | N/F | N/F | 1.033 | N/F | 20.739 | N/F | N/A | 0.07 | N/F | 0.273 | N/F | N/A | N/F | N/F | N/A | 0.273 | 0.679 | 0.386 | N/A | 0.697 | N/F | -4.584 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LRB 18 | N/F | -43.766 | -2.391 | -1.653 | 0.341 | N/F | 793.961 | N/F | N/F | -0.045 | 0.751 | -0.299 | N/F | N/A | 0.092 | N/F | 0.309 | N/F | N/A | -0.254 | 0.712 | N/A | 0.239 | 0.633 | 0.506 | N/A | 0.716 | N/F | -5.804 | 1.665 | N/F | N/A |  | Unknown
LRB-18-inj1 | N/F | -41.06 | -3.08 | -1.581 | 0.339 | N/F | 636.977 | N/F | N/F | -0.008 | 0.741 | -0.008 | N/F | N/A | 0.162 | N/F | 0.287 | N/A | N/A | N/F | 0.716 | N/A | 0.231 | 0.658 | 0.512 | N/A | 0.71 | N/F | -5.768 | N/A | N/A | N/A |  | Unknown
WO-23003-170-A-1 | N/F | -9.324 | 29.873 | -0.79 | 0.345 | N/F | -2215.853 | 0.735 | N/F | 0.538 | 0.689 | 3.409 | N/F | N/A | 0.255 | N/F | 0.221 | N/F | N/A | 0.165 | N/F | N/A | 0.245 | 0.746 | 0.311 | N/A | 0.723 | N/F | -4.796 | -2.137 | N/A | N/A |  | Unknown
WO-23003-169-A-1 | N/F | 59.074 | 40.365 | -2.062 | 0.346 | N/F | -2903.292 | 0.498 | N/F | 1.05 | 0.962 | 4.807 | N/F | N/A | 0.114 | N/F | 0.257 | N/A | N/A | -0.138 | N/F | N/A | 0.271 | 0.803 | 0.365 | N/A | 0.734 | N/F | -5.445 | N/A | N/A | N/A |  | Unknown
WO-23003-168-A-1 | N/F | 26.912 | 35.829 | -1.473 | 0.346 | N/F | -1362.474 | 0.745 | 7974.627 | 0.194 | N/F | 0.527 | N/F | N/A | 0.388 | N/F | 0.645 | N/F | N/A | 8.287 | 0.889 | N/A | 0.536 | 0.656 | 0.503 | N/A | 0.72 | N/F | -4.761 | N/A | N/A | N/A |  | Unknown
WO-23003-167-A-1 | N/F | 49.292 | 208.606 | -1.445 | 0.355 | N/F | -1865.867 | N/F | 8310.781 | 0.357 | 0.912 | 0.567 | N/F | N/A | 0.257 | N/F | 0.71 | N/F | N/A | 8.59 | 1.668 | N/A | 0.566 | 0.702 | 0.488 | N/A | 0.776 | N/F | -3.784 | N/A | N/A | N/A |  | Unknown
WO-23003-166-A-1 | N/F | 7.499 | 104.483 | -1.646 | 0.359 | N/F | -1020.625 | N/F | 8146.394 | 0.438 | 0.995 | 0.507 | N/F | N/A | 0.187 | N/F | 0.628 | N/F | N/A | 8.573 | 1.19 | N/A | 0.504 | 0.773 | 0.392 | N/A | 0.766 | N/F | -4.66 | N/A | N/F | N/A |  | Unknown
WO-23003-165-A-1 | N/F | 20.959 | 6.048 | -1.383 | 0.369 | N/F | -1797.05 | 1.082 | 8091.743 | 0.594 | 0.882 | 0.574 | N/F | N/A | 0.261 | N/F | 0.497 | N/F | N/A | 8.648 | 0.789 | N/A | 0.482 | 0.85 | 0.724 | N/A | 0.792 | N/F | -4.275 | N/A | N/F | N/A |  | Unknown
WO-23003-164-A-1 | N/F | 23.227 | 6.026 | -1.718 | 0.39 | N/F | -2519.815 | 0.989 | 2884.658 | 0.34 | 0.862 | -0.012 | N/F | N/A | 0.276 | N/F | 0.303 | N/A | N/A | 1.161 | N/F | N/A | 0.339 | 0.734 | 0.485 | N/A | 0.797 | N/F | -4.182 | N/A | N/F | N/A |  | Unknown
WO-23003-163-A-1 | N/F | 35.591 | 18.214 | -1.823 | 0.438 | N/A | -2891.911 | 1.814 | 2697.215 | 0.733 | 0.846 | -0.04 | N/F | N/A | 0.364 | N/F | 0.298 | N/A | N/A | N/F | N/F | N/A | 0.346 | 0.945 | 0.449 | N/A | 0.795 | N/F | -5.284 | N/A | N/F | N/A |  | Unknown
WO-23003-162-A-1 | N/F | 0.88 | 5.28 | -1.502 | 0.476 | N/F | -2149.442 | 0.801 | 2959.884 | 0.647 | 0.834 | 0.068 | N/F | N/A | 0.226 | N/F | 0.632 | N/A | N/A | 1.984 | N/F | N/A | 0.462 | 0.822 | 0.454 | N/A | 0.955 | N/F | -4.851 | N/A | N/F | N/A |  | Unknown
WO-23003-161-A-1 | N/F | -12.179 | 3.242 | -2.173 | 0.718 | N/A | -1834.274 | 0.794 | 4700.929 | 0.876 | 0.954 | 0.212 | N/F | N/A | 0.727 | N/F | 0.674 | N/F | N/A | 2.057 | N/F | N/A | 0.561 | 1.089 | 0.552 | N/A | 1.145 | N/F | -5.135 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LFSM-H 8 | N/F | 134.487 | 204.898 | 207.338 | 205.794 | N/F | -1865.78 | 183.28 | 4891.74 | 245.396 | 144.805 | 156.716 | N/F | N/A | 189.553 | 177.244 | 191.129 | N/F | N/A | 176.725 | 170.633 | N/A | 185.896 | 213.588 | 178.895 | N/A | 179.016 | N/A | 403.986 | N/A | N/F | N/A |  | Unknown
AB-230117-02 CCC-H 9 | N/F | 299.118 | 382.686 | 385.836 | 449.642 | N/F | 1459.244 | 351.796 | 322.214 | 343.428 | 302.345 | 306.496 | N/F | N/A | 342.172 | 302.616 | 345.726 | N/A | N/A | 338.625 | 350.18 | N/A | 341.66 | 286.071 | 337.277 | N/A | 309.308 | N/A | 528.598 | 3537.041 | N/F | N/A |  | Unknown
AB-230117-02 LFB-H 9 | N/F | 157.285 | 192.491 | 211.342 | 218.588 | N/F | 910.687 | 183.084 | 171.888 | 180.175 | 155.518 | 152.895 | N/F | N/A | 173.617 | 160.935 | 177.434 | N/F | N/A | 175.811 | 181.434 | N/A | 179.694 | 151.15 | 178.355 | N/A | 164.696 | N/A | 277.512 | N/A | N/A | N/A |  | Unknown
AB-230117-02 LD 8 | N/A | 11.19 | 7.824 | -1.357 | 0.351 | N/F | -2344.617 | N/F | 3160.606 | 0.238 | N/F | -0.179 | N/F | N/A | 0.243 | N/F | 0.49 | N/A | N/A | 1.032 | N/F | N/A | 0.38 | 0.674 | 0.414 | N/A | 0.732 | N/F | -5.182 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LRB 17 | N/F | -43.046 | -2.369 | -1.964 | N/F | N/F | 461.817 | N/F | N/F | 0.041 | 0.833 | -0.144 | N/F | N/A | 0.341 | N/F | 0.331 | N/F | N/A | 0.07 | 0.889 | N/A | 0.23 | 0.629 | 0.626 | N/A | 0.695 | N/F | -5.701 | N/A | N/A | N/A |  | Unknown
LRB-17-inj1 | N/F | -40.691 | -1.752 | -1.648 | N/F | N/F | 680.421 | N/F | N/F | 0.101 | 0.767 | -0.151 | N/F | N/A | 0.12 | N/F | 0.322 | N/F | N/A | -0.215 | 0.786 | N/A | 0.256 | 0.705 | 0.588 | N/A | 0.704 | N/F | -5.985 | N/A | N/A | N/A |  | Unknown
WO-23003-160-A-1 | N/F | -4.395 | 4.97 | -1.852 | N/F | N/F | -2016.166 | N/F | 4609.936 | 0.462 | 0.834 | -0.103 | N/F | N/A | 0.086 | N/F | 0.487 | N/A | N/A | 2.311 | 1.797 | N/A | 0.313 | 0.647 | 0.361 | N/A | 0.737 | N/F | -5.155 | N/A | N/A | N/A |  | Unknown
WO-23003-159-A-1 | N/F | -6.585 | 9.144 | -1.088 | N/F | N/F | -2384.286 | N/F | 4651.337 | 0.446 | N/F | -0.125 | N/F | N/A | 0.252 | N/F | 0.343 | N/A | N/A | 2.729 | N/F | N/A | 0.325 | 0.726 | 0.293 | N/A | 0.718 | N/F | -4.858 | N/A | N/F | N/A |  | Unknown
WO-23003-158-A-1 | N/F | -18.739 | 4.796 | -1.865 | N/F | N/F | -1384.915 | 0.863 | 1326.789 | 0.367 | 0.748 | -0.232 | 1226.073 | N/A | 0.213 | N/F | 1.766 | N/F | N/A | 1.412 | N/F | N/A | 0.267 | 0.655 | 0.308 | N/A | 0.7 | N/F | -5.172 | N/A | N/A | N/A |  | Unknown
WO-23003-157-A-1 | N/F | 22.948 | 11.395 | -2.167 | N/F | N/F | -2694.916 | 0.409 | 1266.582 | 0.396 | N/F | 0.041 | N/F | N/A | 0.144 | N/F | 0.251 | N/A | N/A | -0.207 | N/F | N/A | 0.286 | 0.661 | 0.38 | N/A | 0.724 | N/F | -5.04 | N/A | N/A | N/A |  | Unknown
WO-23003-156-A-1 | N/F | -46.945 | 3.245 | -1.922 | N/F | N/F | -1693.141 | N/F | 1276.879 | 0.464 | N/F | -0.051 | N/F | N/A | 0.206 | N/F | 0.25 | N/F | N/A | 1.092 | N/F | N/A | 0.26 | 0.647 | 0.357 | N/A | 0.702 | N/F | -4.76 | N/A | N/A | N/A |  | Unknown
WO-23003-155-A-1 | N/F | -10.24 | 8.673 | -1.567 | N/F | N/F | -1793.372 | 1.192 | 2461.423 | 0.367 | 0.79 | N/F | 1220.91 | N/A | 0.207 | N/F | 0.263 | N/F | N/A | 0.289 | N/F | N/A | 0.374 | 0.667 | 0.326 | N/A | 0.724 | N/F | -4.517 | N/A | N/A | N/A |  | Unknown
WO-23003-154-A-1 | N/F | -10.764 | 11.307 | -1.803 | N/F | N/F | -1821.495 | 1.097 | 2145.145 | 0.36 | 0.884 | -0.177 | N/F | N/A | 0.171 | N/F | 0.293 | N/F | N/A | N/F | N/F | N/A | 1.635 | 0.656 | 0.404 | N/A | 0.707 | N/F | -5.882 | N/A | N/A | N/A |  | Unknown
WO-23003-153-A-1 | N/F | -6.914 | 23.34 | -1.817 | 0.342 | N/F | -1881.413 | N/F | 1340.929 | 0.718 | N/F | -0.135 | N/F | N/A | 0.246 | N/F | 0.404 | N/A | N/A | N/F | N/F | N/A | 0.347 | 0.7 | 0.36 | N/A | 0.72 | N/F | -5.994 | N/A | N/A | N/A |  | Unknown
WO-23003-152-A-1 | N/F | -3.145 | 14.397 | -1.565 | N/F | N/F | -2322.605 | 1.181 | 521.5 | 0.483 | N/F | -0.223 | N/F | N/A | 0.161 | N/F | 0.27 | N/F | N/A | N/F | N/F | N/A | 0.266 | 0.803 | 0.4 | N/A | 0.726 | N/F | -4.929 | N/A | N/F | N/A |  | Unknown
WO-23003-151-A-1 | N/F | -6.063 | 12.117 | -1.227 | 0.393 | N/F | -1585.015 | N/F | 477.873 | 0.208 | N/F | 0.103 | N/F | N/A | 0.316 | N/F | 1.037 | N/F | N/A | N/F | N/F | N/A | 0.323 | 0.665 | 0.572 | N/A | 0.756 | N/F | -5.976 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LRB 16 | N/F | -33.958 | -2.887 | -1.569 | 0.341 | N/F | -302.333 | N/F | N/F | -0.016 | 0.781 | -0.003 | N/F | N/A | 0.315 | N/F | 0.431 | N/F | N/A | -0.316 | 0.749 | N/A | 0.266 | 0.745 | 0.617 | N/A | 0.739 | N/F | -5.814 | N/A | N/A | N/A |  | Unknown
LRB-16-inj1 | N/F | -23.075 | -3.753 | -1.774 | 0.338 | N/F | 259.3 | 0.811 | N/F | 0.011 | 0.71 | -0.069 | N/F | N/A | 0.306 | N/F | 0.375 | N/F | N/A | -0.235 | 0.741 | N/A | 0.273 | 0.654 | 0.516 | N/A | 0.741 | N/F | -5.258 | N/A | N/A | N/A |  | Unknown
WO-23003-150-A-1 | N/F | 24.516 | 7.99 | -1.791 | 0.34 | N/F | -2360.054 | N/F | 579.319 | 0.241 | 0.907 | 0.178 | N/F | N/A | 0.436 | N/F | 0.351 | N/F | N/A | N/F | N/F | N/A | 0.324 | 0.7 | 0.429 | N/A | 0.711 | N/F | -5.05 | N/A | N/F | N/A |  | Unknown
WO-23003-149-A-1 | N/F | 46.828 | 7.993 | -1.577 | N/F | N/F | -2330.049 | 1.108 | 519.796 | 0.552 | 0.724 | N/F | N/F | N/A | 0.262 | N/F | 0.365 | N/F | N/A | 0.032 | N/F | N/A | 0.346 | 0.811 | 0.237 | N/A | 0.749 | N/F | -4.16 | N/A | N/F | N/A |  | Unknown
WO-23003-148-A-1 | N/F | 39.824 | 6.918 | -2.099 | 0.343 | N/F | -2144.673 | 1.373 | 577.003 | 0.485 | N/F | -0.065 | 1173.655 | N/A | 0.24 | N/F | 0.359 | N/A | N/A | 0.106 | N/F | N/A | 0.358 | 0.672 | 0.388 | N/A | 0.725 | N/F | -4.126 | N/A | N/A | N/A |  | Unknown
WO-23003-147-A-1 | N/F | 24.359 | 4.832 | -2.056 | 0.349 | N/F | -1954.718 | N/F | 572.526 | 0.624 | N/F | -0.042 | N/F | N/A | 0.28 | N/F | 0.366 | N/F | N/A | N/F | N/F | N/A | 0.329 | 0.711 | 0.467 | N/A | 0.722 | N/F | -5.078 | N/A | N/A | N/A |  | Unknown
WO-23003-146-A-1 | N/F | 49.173 | 3.129 | -1.39 | 0.357 | N/F | -1857.046 | N/F | 732.447 | 0.457 | 0.792 | 0.036 | N/F | N/A | 0.29 | N/F | 0.408 | N/F | N/A | 1.461 | N/F | N/A | 0.324 | 0.761 | 0.535 | N/A | 0.721 | N/F | -5.148 | N/A | N/F | N/A |  | Unknown
WO-23003-145-A-1 | N/F | 21.135 | 6.036 | -0.113 | 0.369 | N/F | -1574.062 | N/F | 638.175 | 0.552 | 0.871 | 0.014 | N/F | N/A | 0.347 | N/F | 0.419 | N/F | N/A | N/F | N/F | N/A | 0.332 | 0.681 | 0.406 | N/A | 0.792 | N/F | -4.668 | N/A | N/F | N/A |  | Unknown
WO-23003-144-A-1 | N/F | 31.261 | 7.157 | -0.794 | 0.387 | N/A | -1853.622 | N/F | 127.716 | 0.492 | N/F | 0.056 | N/F | N/A | 0.415 | N/F | 0.465 | N/F | N/A | 0.237 | N/F | N/A | 0.337 | 0.759 | 0.442 | N/A | 0.861 | N/F | -4.129 | N/A | N/F | 17.888 |  | Unknown
WO-23003-143-A-1 | N/F | 25.016 | 6.142 | -2.023 | 0.423 | N/F | -1858.381 | 1.526 | 102.324 | 0.633 | 0.803 | 0.121 | N/F | N/A | 0.235 | N/F | 0.671 | N/F | N/A | -0.387 | N/F | N/A | 0.363 | 0.756 | 0.43 | N/A | 0.794 | N/F | -4.687 | N/A | N/F | N/A |  | Unknown
WO-23003-142-A-1 | N/F | 46.024 | 4.891 | -2.065 | 0.499 | N/F | -2160.174 | N/F | 151.9 | 0.636 | N/F | 0.024 | N/F | N/A | 0.527 | N/F | 0.92 | N/F | N/A | 0.445 | 1.467 | N/A | 0.486 | 1.038 | 0.453 | N/A | 0.896 | N/F | -4.971 | 4.655 | N/F | N/A |  | Unknown
WO-23003-141-A-1 | N/F | 112.953 | 11.896 | -1.589 | 0.896 | N/F | -2955.93 | N/F | 113.991 | 1.564 | N/F | 0.287 | N/F | N/A | 0.615 | N/F | 0.925 | N/F | N/A | 1.69 | N/F | N/A | 0.724 | 1.115 | 0.824 | N/A | 1.191 | N/F | -4.586 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LFSM-H 7 | N/F | 250.956 | 199.912 | 215.304 | 200.939 | N/F | -2487.807 | 175.013 | 277.866 | 282.112 | 145.268 | 159.093 | 1237.967 | N/A | 190.18 | 188.826 | 188.842 | N/F | N/A | 175.116 | 173.568 | N/A | 188.523 | 202.11 | 182.658 | N/A | 186.394 | N/F | 416.861 | N/A | N/F | N/A |  | Unknown
AB-230117-02 CCC-H 8 | N/F | 314.441 | 371.867 | 399.548 | 445.243 | N/F | 782.316 | 374.743 | 302.112 | 355.866 | 310.833 | 321.519 | 1183.254 | N/A | 346.466 | 306.068 | 346.394 | N/F | N/A | 345.124 | 347.778 | N/A | 353.314 | 248.6 | 345.655 | N/A | 322.674 | N/F | 616.443 | N/A | N/F | 17128.031 |  | Unknown
AB-230117-02 LFB-H 8 | N/F | 159.103 | 180.56 | 189.665 | 205.316 | N/F | 568.024 | 178.569 | 157.405 | 174.967 | 149.213 | 147.72 | N/F | N/A | 170.147 | 149.916 | 172.643 | N/A | N/A | 170.567 | 174.977 | N/A | 174.528 | 132.033 | 174.263 | N/A | 157.989 | N/F | 343.211 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LD 7 | N/F | 95.852 | 9.315 | -2.14 | 0.345 | N/F | -2682.455 | N/F | 151.363 | 0.531 | N/F | N/F | N/F | N/A | 0.262 | N/F | 1.197 | N/A | N/A | 0.083 | 1.76 | N/A | 0.336 | 0.745 | 0.417 | N/A | 0.725 | N/F | -5.26 | N/A | N/F | N/A |  | Unknown
AB-230117-02 LRB 15 | N/F | -21.294 | -0.741 | -1.558 | N/F | N/F | -156.018 | 0.658 | N/F | 0.037 | N/F | 0.002 | N/F | N/A | 0.374 | N/F | 1.286 | N/F | N/A | N/F | 0.9 | N/A | 0.302 | 0.686 | 0.525 | N/A | 0.736 | N/F | -5.853 | N/A | N/A | N/A |  | Unknown
LRB-15-inj1 | N/F | -25.7 | -4.702 | -1.943 | 0.344 | N/F | 233.873 | 0.374 | N/F | -0.071 | N/F | -0.083 | 1143.28 | N/A | 0.306 | N/F | 1.417 | N/F | N/A | -0.287 | 0.91 | N/A | 0.31 | 0.689 | 0.603 | N/A | 0.727 | N/F | -5.644 | N/A | N/F | N/A |  | Unknown
WO-23003-140-A-1 | N/F | 21.135 | 6.787 | -2.18 | N/F | N/F | -1913.783 | 0.914 | 152.509 | 0.531 | N/F | -0.048 | N/F | N/A | 0.192 | N/F | 2.01 | N/F | N/A | N/F | N/F | N/A | 0.338 | 0.7 | 0.357 | N/A | 0.736 | N/F | -5.123 | N/A | N/A | 10.995 |  | Unknown
WO-23003-139-A-1B | N/F | 34.286 | 6.982 | -1.819 | 0.36 | N/F | -2000.414 | 2.591 | 146.837 | 0.517 | 0.848 | 0.878 | N/F | N/A | 0.229 | N/F | 6.248 | N/F | N/A | N/F | N/F | N/A | 0.345 | 0.892 | 0.378 | N/A | 0.743 | N/F | -4.155 | N/A | N/F | N/A |  | Unknown
WO-23003-139-A-1 | N/F | 42.601 | 24.366 | -1.978 | 0.475 | N/F | -1388.169 | 3.112 | 141.212 | 0.236 | N/F | 0.75 | N/F | N/A | 0.428 | N/F | 133.922 | N/A | N/A | 0.217 | N/F | N/A | 0.57 | 0.947 | 0.695 | N/A | 0.941 | N/F | -0.443 | N/A | N/A | N/A |  | Unknown
Rinse12 | N/F | 4264.452 | 312.292 | 3.56 | 1.902 | N/F | 6368.875 | 40.266 | N/F | 4.619 | 4.956 | 5.493 | N/F | N/A | 2.482 | N/F | 3907 | N/F | N/F | 16.054 | N/F | N/A | 6.109 | 3.51 | 11.109 | N/A | 4.679 | N/F | 210.223 | -1.202 | N/A | N/A | C | Unknown
WO-23003-203-A-1B | N/A | -18.384 | 1.188 | 0.39 | 0.447 | N/F | 1183.271 | 0.231 | 10.448 | 215.936 | 17.597 | 1809.619 | N/F | N/A | 0.225 | N/F | 0.378 | N/F | N/F | 1.058 | 0.868 | N/A | 0.326 | 0.992 | 4.491 | N/A | 2.892 | N/F | -5.409 | N/A | N/A | N/A |  | Unknown
WO-23003-208-A-1B | N/F | -10.214 | 3.565 | -1.073 | 0.361 | N/F | 756.265 | 0.718 | 10.574 | 216.322 | 17.222 | 1821.511 | N/F | N/A | 0.219 | N/F | 0.316 | N/F | N/A | 0.964 | 0.753 | N/A | 0.294 | 0.999 | 3.878 | N/A | 2.513 | N/F | -5.806 | N/A | N/A | N/A |  | Unknown
WO-23003-207-A-1B | N/A | 5.897 | 11.164 | 0.53 | 0.395 | N/F | 913.324 | N/F | 11.126 | 232.734 | 18.417 | 1989.268 | N/F | N/A | 0.242 | N/F | 0.34 | N/F | N/A | 0.967 | 0.815 | N/A | 0.366 | 1.009 | 3.98 | N/A | 2.809 | N/F | -5.891 | N/A | N/A | 12.585 |  | Unknown
WO-23003-206-A-1B | N/F | -17.296 | 5.148 | -1.224 | 0.394 | N/F | 1064.734 | 0.356 | 11.594 | 296.595 | 20.732 | 2969.033 | N/F | N/A | 0.282 | N/F | 0.347 | N/A | N/A | 1.012 | 0.794 | N/A | 0.362 | 1.199 | 4.89 | N/A | 2.78 | N/F | -6.146 | 1.669 | N/A | N/A |  | Unknown
WO-23003-205-A-1B | N/F | 12.562 | 105.739 | 0.794 | 0.449 | N/F | 641.965 | 0.35 | 44.219 | 640.608 | 74.136 | 4923.452 | N/F | N/A | 0.233 | N/F | 0.343 | N/A | N/A | 2.017 | 1.411 | N/A | N/F | 1.626 | 7.698 | N/A | 3.51 | N/F | -4.735 | N/A | N/A | N/A |  | Unknown
WO-23003-204-A-1B | N/F | -11.839 | 27.602 | -0.253 | 0.485 | N/F | 614.722 | 0.313 | 44.764 | 643.996 | 73.237 | 5298.356 | N/F | N/A | 0.303 | N/F | 0.315 | N/A | N/A | 2.904 | 1.758 | N/A | N/F | 1.389 | 9.339 | N/A | 2.867 | N/F | -4.387 | N/A | N/A | N/A |  | Unknown
LRB-22-reprep-inj2 | N/F | -26.773 | 7.198 | 0.367 | 0.419 | N/F | 1845.449 | 0.212 | N/F | 0.3 | 0.799 | 3.082 | N/F | N/A | 0.411 | N/F | 0.404 | N/F | N/A | 0.041 | 0.494 | N/A | 0.435 | 0.678 | 2.24 | N/A | 2.98 | N/F | -5.507 | N/A | N/A | N/A |  | Unknown
LRB-22-reprep | N/F | -19.826 | 6.092 | 1.198 | 0.465 | N/F | 1516.417 | N/F | N/F | 0.273 | 0.85 | 3.365 | N/F | N/A | 0.378 | N/F | 0.417 | N/F | N/A | 0.042 | 0.435 | N/A | 0.439 | 0.68 | 2.334 | N/A | 2.759 | N/F | -5.898 | N/A | N/A | N/A |  | Unknown
217288-d4-reprep | N/F | -9.074 | 10.978 | -1.272 | 0.339 | N/F | -596.574 | 0.593 | 4342.461 | 0.77 | 0.858 | 4.749 | N/F | N/A | 0.029 | N/F | 0.577 | N/F | N/A | 3.487 | 0.768 | N/A | 0.407 | 0.641 | 0.739 | N/A | 1.391 | N/F | -4.475 | N/A | N/A | N/A |  | Unknown
217289-d4-reprep | N/A | -13.822 | 12.459 | -1.583 | 0.339 | N/F | -537.227 | N/F | 4355.518 | 0.624 | 0.848 | 2.443 | N/F | N/A | 0.202 | N/F | 0.368 | N/F | N/A | 3.221 | 0.642 | N/A | 0.343 | 0.723 | 0.655 | N/A | 1.237 | N/F | -4.853 | N/A | N/A | N/A |  | Unknown
217295-d1000-LD-reprep | N/A | -47.866 | -0.431 | -1.582 | N/F | N/F | 1063.903 | 0.177 | N/F | 26.189 | 1.598 | 1222.979 | N/F | N/A | 0.126 | N/F | 1.087 | N/F | N/A | -0.202 | 0.308 | N/A | 0.28 | 2.205 | 0.299 | N/A | 0.725 | N/F | -5.991 | N/A | N/F | N/A |  | Unknown
217296-d1000-LFSM-200-reprep | N/A | 145.29 | 189.732 | 202.065 | 404.612 | N/F | 1364.617 | 199.356 | 171.849 | 209.69 | 164.478 | 1611.097 | N/F | N/A | 177.964 | 166.141 | 190.796 | N/F | N/A | 182.061 | 189.823 | N/A | 184.164 | 144.268 | 180.753 | N/A | 177.374 | N/A | 265.776 | N/A | N/A | N/A |  | Unknown
217315-d2-reprep | N/F | 4.342 | 16.17 | -1.427 | 0.34 | N/F | -2029.531 | N/F | N/F | 0.515 | 0.802 | 1.534 | N/F | N/A | 0.108 | 137.147 | 0.228 | N/A | N/A | 1.512 | N/F | N/A | 0.321 | 0.741 | 0.537 | N/A | 1.215 | N/F | -4.692 | N/A | N/A | N/A |  | Unknown
217335-d2-LD-reprep | N/A | -0.844 | 14.757 | -0.376 | N/F | N/F | -2349.483 | N/F | N/F | 1.315 | 0.725 | 53.394 | N/F | N/A | 0.006 | 430.994 | 1.008 | N/A | N/A | -0.13 | N/F | N/A | 0.328 | 0.762 | 0.574 | N/A | 1.18 | N/F | -5.296 | N/A | N/A | N/A |  | Unknown
217336-d2-LFSM-200-reprep | N/F | 262.827 | 229.683 | 208.752 | 216.508 | N/F | -3057.123 | 185.585 | 175.682 | 310.298 | 149.657 | 211.498 | N/F | N/A | 185.41 | 606.641 | 194.566 | N/A | N/A | 180.656 | 184.866 | N/A | 187.674 | 202.297 | 175.24 | N/A | 239.752 | N/A | 475.374 | N/A | N/A | N/A |  | Unknown
217356-d2-reprep | N/F | -9.612 | 12.982 | 0.302 | 0.333 | N/A | -1843.624 | 0.299 | N/F | 0.256 | 0.792 | 1.099 | N/F | N/A | 0.195 | N/F | 4.268 | N/A | N/A | -0.193 | 0.477 | N/A | 0.326 | 0.753 | 0.546 | N/A | 1.376 | N/F | -5.28 | N/A | N/A | N/A |  | Unknown
217357-d2-reprep | N/F | 27.503 | 12.796 | 6.657 | 0.338 | N/F | -2048.512 | 0.835 | N/F | 0.222 | 0.82 | 1.143 | N/F | N/A | 0.173 | N/F | 48.327 | N/F | N/A | -0.174 | N/F | N/A | 0.929 | 0.68 | 0.607 | N/A | 1.21 | N/F | -4.977 | N/A | N/F | N/A |  | Unknown
217358-d2-reprep | N/F | 3.484 | 13.334 | 12.64 | N/F | N/F | -1817.146 | N/F | N/F | 0.686 | 0.917 | 0.749 | N/F | N/A | 0.188 | N/F | 2575.277 | N/A | N/A | N/F | N/F | N/A | 1.195 | 1.03 | 0.733 | N/A | 1.123 | N/F | -5.422 | N/A | N/A | N/A |  | Unknown
217359-d2-LD | N/A | -10.065 | 18.448 | 6.068 | N/F | N/F | -1977.266 | N/F | N/F | 0.176 | 0.759 | 0.25 | N/F | N/A | 0.479 | N/F | 8748.882 | N/F | N/A | -0.061 | N/F | N/A | 2.774 | 0.772 | 0.968 | N/A | 0.704 | N/F | -5.043 | N/A | N/A | N/A |  | Unknown
217359-d2-reprep | N/F | -11.12 | 14.511 | 11.553 | N/F | N/F | -1783.186 | N/F | N/F | 0.343 | 0.815 | 0.671 | N/F | N/A | 0.688 | N/F | 9903.264 | N/F | N/A | 0.399 | N/F | N/A | 3.011 | 1.398 | 1.217 | N/A | 1.23 | N/F | -5.261 | N/A | N/A | N/A |  | Unknown
217360-d2-LFSM-200 | N/F | 132.551 | 225.235 | 210.292 | 218.738 | N/F | -2191.951 | 187.722 | 168.197 | 267.17 | 144.38 | 160.884 | N/F | N/A | 178.944 | 182.595 | 384.137 | N/F | N/A | 182.089 | 186.158 | N/A | 185.543 | 185.019 | 185.148 | N/A | 182.929 | N/A | 336.899 | N/A | N/A | N/A |  | Unknown
217360-d2-reprep | N/A | 6.774 | 17.627 | -0.119 | N/F | N/A | -2344.388 | N/F | N/F | 0.313 | 0.8 | 0.63 | N/F | N/A | 0.221 | N/F | 378.052 | N/A | N/A | 0.348 | N/F | N/A | 0.345 | 0.666 | 0.642 | N/A | 1.125 | N/F | -4.772 | 2.774 | N/A | N/A |  | Unknown
217375-d2-LD-reprep | N/F | 30.848 | 190.929 | 3.309 | 0.335 | N/F | -2467.426 | N/F | N/F | 0.233 | 0.772 | 1.407 | N/F | N/A | 0.205 | N/F | 28.506 | N/A | N/A | -0.006 | 0.931 | N/A | 0.367 | 0.646 | 0.848 | N/A | 1.16 | N/F | -5.024 | N/A | N/F | N/A |  | Unknown
217376-d2-LFSM-200-reprep | N/F | 131.198 | 322.854 | 204.968 | 212.858 | N/F | -2446.857 | 173.032 | 158.711 | 211.603 | 146.373 | 153.34 | N/F | N/A | 174.657 | 196.844 | 266.746 | N/F | N/A | 170.416 | 173.811 | N/A | 177.327 | 175.577 | 166.325 | N/A | 230.508 | N/A | 384.556 | N/A | N/F | N/A |  | Unknown
217393-d4-reprep | N/F | -18.662 | 12.629 | 5.858 | N/F | N/F | -693.821 | N/F | N/F | 0.055 | 0.771 | 0.47 | N/F | N/A | 0.256 | N/F | 1061.939 | N/F | N/A | 1.123 | 0.775 | N/A | 2.949 | 0.697 | 3.747 | N/A | 1.235 | N/F | -5.328 | N/A | N/F | 13.117 |  | Unknown
217394-d4-reprep | N/F | -16.32 | 11.577 | 4.122 | N/F | N/F | -612.677 | 0.511 | N/F | 0.105 | 0.752 | 0.294 | N/F | N/A | 0.327 | N/F | 1414.578 | N/F | N/A | 1.642 | 1.218 | N/A | 4.034 | 0.749 | 4.579 | N/A | 1.236 | N/F | -5.208 | N/A | N/F | N/A |  | Unknown
217395-d4-LD-reprep | N/A | -18.905 | 9.055 | 1.839 | N/F | N/F | -729.492 | N/F | N/F | 0.302 | 0.717 | 0.313 | 1164.357 | N/A | 0.254 | N/F | 2607.572 | N/A | N/A | 2.363 | 1.523 | N/A | 8.406 | 0.779 | 8.776 | N/A | 1.381 | N/F | -5.074 | N/A | N/F | N/A |  | Unknown
217395-d4-reprep | N/F | -26.093 | 5.597 | 3.305 | N/F | N/F | -685.422 | 0.132 | N/F | 0.028 | 0.747 | 0.241 | N/F | N/A | 0.287 | N/F | 2714.735 | N/A | N/A | 2.849 | 1.534 | N/A | 8.354 | 0.741 | 9.032 | N/A | 1.16 | N/F | -5.257 | N/A | N/A | N/A |  | Unknown
217396-d200-LFSM-200 | N/A | 133.178 | 192.08 | 248.904 | 228.127 | N/F | 1291.449 | 175.395 | 162.554 | 165.612 | 144.704 | 154.142 | N/F | N/A | 160.84 | 153.37 | 4040.399 | N/F | N/A | 171.237 | 175.651 | N/A | 178.6 | 189.747 | 168.361 | N/A | 221.107 | N/A | 249.6 | N/A | N/A | N/A |  | Unknown
217399-d2-reprep | N/A | -3.205 | 5.743 | -0.074 | 0.332 | N/F | -1948.401 | N/F | 11.758 | 0.21 | N/F | 0.353 | N/F | N/A | 0.101 | N/F | 4.311 | N/F | N/A | 1.515 | N/F | N/A | 0.312 | 0.64 | 0.679 | N/A | 1.382 | N/F | -4.933 | -0.95 | N/F | N/A |  | Unknown
217400-d2-reprep | N/A | -14.742 | 2.805 | -0.786 | N/F | N/F | -2338.926 | N/F | 11.796 | 0.423 | N/F | 0.213 | N/F | N/A | 0.136 | N/F | 1.945 | N/F | N/A | 0.695 | N/F | N/A | 0.321 | 0.667 | 0.707 | N/A | 1.275 | N/F | -4.006 | N/A | N/F | N/A |  | Unknown
217415-d2-LD-reprep | N/A | 14.772 | 7.025 | -2.034 | N/F | N/F | -2375.359 | 0.57 | 121.83 | 0.572 | N/F | 0.179 | N/F | N/A | 0.021 | N/F | 10.557 | N/F | N/A | 2.022 | N/F | N/A | 0.339 | 0.658 | 0.48 | N/A | 0.698 | N/F | -4.645 | N/A | N/F | N/A |  | Unknown
217416-d2-LFSM-200-reprep | N/F | 230.751 | 209.798 | 211.635 | 218.234 | N/F | -2292.378 | 183.534 | 312.093 | 254.925 | 152.563 | 161.847 | N/F | N/A | 180.684 | 199.788 | 195.064 | N/F | N/A | 181.643 | 184.672 | N/A | 187.492 | 208.736 | 186.591 | N/A | 183.094 | N/A | 369.985 | N/A | N/F | N/A |  | Unknown
CCC-H-12-reinject | N/F | 319.537 | 380.522 | 398.231 | 516.567 | N/F | 827.953 | 364.623 | 333.02 | 348.29 | 324.299 | 326.958 | N/F | N/A | 337.682 | 305.952 | 353.248 | N/F | N/A | 350.683 | 349.782 | N/A | 354.686 | 246.665 | 351.819 | N/A | 325.504 | N/A | 536.228 | N/A | N/F | N/A |  | Unknown
LFB-H-11-reinject | N/F | 156.333 | 231.008 | 191.795 | 284.856 | N/A | 1952.581 | 225.808 | 199.836 | 220.341 | 195.122 | 194.608 | N/F | N/A | 17.623 | 194.643 | 213.202 | N/F | N/A | 209.699 | 219.041 | N/A | 224.483 | 16.591 | 205.408 | N/A | 27.003 | N/A | 305.009 | 530.887 | N/F | N/A |  | Unknown
QCS-2-Reinject | N/F | 160.3 | 228.079 | 220.864 | 255.969 | N/F | 1773.213 | 211.41 | 192.951 | 198.851 | 176.719 | 175.82 | N/F | N/A | 203.422 | 174.903 | 205.073 | N/A | N/A | 197.138 | 200.785 | N/A | 202.109 | 158.331 | 189.964 | N/A | 290.941 | N/A | 289.231 | N/A | N/A | N/A |  | Unknown
LRB-22-reprep-inj2 | N/F | -44.147 | 5.235 | -0.648 | 0.721 | N/F | 1478.318 | 0.24 | N/F | 0.395 | 0.899 | 0.384 | N/F | N/A | 0.653 | N/F | 1.636 | N/F | N/A | 0.203 | 0.315 | N/A | 0.537 | 1.116 | 0.764 | N/A | 1.633 | N/F | -6.141 | N/A | N/A | N/A |  | Unknown
LRB-22-reprep | N/F | -45.825 | 8.435 | 5.787 | 11.563 | N/A | 889.628 | N/F | N/F | 5.062 | 3.048 | 6.458 | N/F | N/A | 8.185 | N/F | 9.672 | N/F | N/A | 1.213 | 0.383 | N/A | 5.822 | 9.12 | 3.603 | N/A | 9.906 | N/F | -5.673 | N/A | N/A | N/A |  | Unknown
WO-23003-57-A-1 | N/F | 54.949 | 16.549 | -1.681 | N/F | N/F | -2289.125 | N/F | 1.166 | 0.59 | N/F | 0.509 | N/F | N/A | 0.42 | N/F | 4.897 | N/F | N/A | 0.591 | N/F | N/A | 0.365 | 0.678 | 0.337 | N/A | 0.698 | N/F | -3.46 | N/A | N/A | N/A |  | Unknown

## Sheet2
Batch Folder Path: | D:\TraceFinderData\Projects\PFAS Online SPE\2023-01-18\2023-01-18
Batch Name: | 2023-01-18
Last Acquisition Date: | 2023-01-31 07:17
User Name: | Quantiva
Sample Count: | 365
Sample Amount - Zero
Data File | 11Cl-PF3OUdS | 4:2 FTS | 6:2 FTS | 8:2 FTS | 9Cl-PF3ONS | FBSA | FHxA | HFPO-DA | PFBS | PFHpS | PFHxS | PFOS | NaDONA | N-EtFOSAA | N-MeFOSA | PFBA | PFDA | PFDoA | PFDS | PFHpA | PFHxA | PFHxDA | PFNA | PFNS | PFOA | PFODA | PFOSA | PFPeA | PFPeS | PFTeDA | PFTrDA | PFUdA | All Flags | Dilution Factor | Sample Type | Sample Acquisition Date
AB-230117-02 Cal Blank 1 | 0 | 35.433 | 30.203 | 0.143 | 2.522 | 0 | 3560.748 | 0 | 0 | 2.128 | 1.927 | 3.298 | 0 | 0 | 0.734 | 0 | 2.511 | 0 | 0 | 2.616 | 0.566 | 0 | 2.445 | 0.906 | 2.166 | 0 | 1.03 | 0 | -4.716 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-18 10:32
AB-230117-02 Cal Std 1 1 | 0 | 6.433 | 3.598 | 2.526 | 1.239 | 0 | 4.842 | 1.078 | 0 | 0.897 | 1.451 | 1.593 | 0 | 0 | 1.215 | 0 | 1.325 | 0 | 0 | 2.116 | 1.415 | 0 | 1.368 | 1.526 | 1.437 | 0 | 1.481 | 0 | -3.231 | 0 | 0 | 0 |  | 1 | Cal Std | 2023-01-18 11:00
AB-230117-02 Cal Std 2 1 | 0 | -5.318 | -0.486 | 1.171 | 1.909 | 0 | -369.984 | 2.003 | 1.995 | 1.816 | 2.155 | 2.003 | 0 | 0 | 2.201 | 0 | 2.016 | 0 | 0 | 1.404 | 2.095 | 0 | 2.05 | 1.966 | 2.071 | 0 | 2.324 | 0 | -2.568 | 0 | 0 | 0 |  | 1 | Cal Std | 2023-01-18 11:27
AB-230117-02 Cal Std 3 1 | 0 | -2.501 | 0.199 | 1.353 | 3.936 | 0 | 534.951 | 4.412 | 3.707 | 3.299 | 3.709 | 3.798 | 0 | 0 | 3.778 | 5.762 | 4.043 | 0 | 0 | 3.081 | 3.86 | 0 | 3.843 | 4.03 | 3.96 | 0 | 4.016 | 0 | 0.921 | 47.308 | 0 | 0 |  | 1 | Cal Std | 2023-01-18 11:54
AB-230117-02 Cal Std 4 1 | 0 | 7.181 | 6.426 | 8.048 | 9.297 | 0 | 352.458 | 9.51 | 7.789 | 9.624 | 7.745 | 7.845 | 0 | 0 | 9.26 | 11.646 | 9.394 | 0 | 0 | 8.566 | 9.126 | 0 | 9.265 | 8.612 | 9.194 | 0 | 9.157 | 0 | 16.627 | 0 | 0 | 0 |  | 1 | Cal Std | 2023-01-18 12:21
AB-230117-02 Cal Std 5 1 | 0 | 23.955 | 16.593 | 17.373 | 19.189 | 0 | 185.587 | 18.606 | 18.452 | 16.778 | 15.779 | 15.192 | 0 | 0 | 18.987 | 18.831 | 18.415 | 0 | 0 | 17.177 | 17.853 | 0 | 18.63 | 16.924 | 18.158 | 0 | 17.349 | 0 | 57.566 | 0 | 0 | 0 |  | 1 | Cal Std | 2023-01-18 12:48
AB-230117-02 Cal Std 6 1 | 0 | 41.816 | 35.955 | 35.512 | 38.804 | 0 | 650.022 | 36.592 | 32.581 | 38.875 | 30.702 | 31.551 | 0 | 0 | 36.41 | 37.308 | 37.389 | 0 | 0 | 36.068 | 35.14 | 0 | 37.996 | 33.194 | 36.553 | 0 | 35.688 | 0 | 88.672 | 0 | 0 | 0 |  | 1 | Cal Std | 2023-01-18 13:15
AB-230117-02 Cal Std 7 1 | 0 | 104.514 | 100.259 | 101.538 | 92.867 | 0 | -293.284 | 95.727 | 87.282 | 104.133 | 79.193 | 79.635 | 0 | 0 | 95.217 | 95.665 | 97.43 | 0 | 0 | 94.573 | 93.802 | 0 | 96.468 | 90.032 | 95.968 | 0 | 90.01 | 0 | 194.53 | 0 | 0 | 0 |  | 1 | Cal Std | 2023-01-18 13:42
AB-230117-02 Cal Std 8 1 | 0 | 183.607 | 195.409 | 208.267 | 186.785 | 0 | -217.327 | 197.427 | 174.364 | 200.784 | 164.11 | 161.542 | 0 | 0 | 192.477 | 163.759 | 193.114 | 0 | 0 | 189.282 | 191.961 | 0 | 189.259 | 174.816 | 187.35 | 0 | 188.906 | 0 | 379.144 | 0 | 0 | 0 |  | 1 | Cal Std | 2023-01-18 14:09
AB-230117-02 Cal Std 9 1 | 0 | 352.895 | 427.559 | 368.173 | 379.575 | 0 | -502.067 | 402.773 | 352.84 | 396.639 | 336.272 | 345.567 | 0 | 0 | 391.958 | 328.948 | 385.495 | 0 | 0 | 377.454 | 384.687 | 0 | 386.336 | 339.641 | 378.816 | 0 | 359.431 | 0 | 790.225 | 0 | 0 | 0 |  | 1 | Cal Std | 2023-01-18 14:36
AB-230117-02 Cal Std 10 1 | 0 | 986.227 | 1112.502 | 835.797 | 987.021 | 0 | 632.176 | 1029.861 | 949.998 | 975.058 | 867.74 | 896.777 | 1310.263 | 0 | 1022.645 | 951.413 | 972.285 | 0 | 0 | 968.916 | 1001.157 | 0 | 982.021 | 858.007 | 980.955 | 0 | 914.041 | 0 | 1693.731 | 0 | 0 | 0 |  | 1 | Cal Std | 2023-01-18 15:03
AB-230117-02 Cal Std 11 1 | 0 | 1993.322 | 2033.349 | 2298.172 | 2306.061 | 0 | 2444.598 | 2030.625 | 1689.093 | 2154.985 | 1807.699 | 1876.965 | 1632.429 | 0 | 2041.136 | 1822.989 | 1949.845 | 0 | 0 | 1943.349 | 1991.195 | 0 | 1944.508 | 2064.871 | 1905.749 | 0 | 1988.411 | 0 | 2397.48 | 17779.806 | 0 | 0 |  | 1 | Cal Std | 2023-01-18 15:30
AB-230117-02 Cal Std 12 1 | 0 | 4354.331 | 4412.932 | 4246.579 | 4244.259 | 0 | 4140.162 | 4196.925 | 3581.739 | 3951.007 | 3852.301 | 3839.827 | 3118.615 | 0 | 4135.203 | 4173.586 | 4002.829 | 0 | 0 | 4032.427 | 4184.365 | 0 | 3980.463 | 3919.784 | 4022.699 | 0 | 4062.479 | 0 | 4013.728 | 0 | 0 | 0 |  | 1 | Cal Std | 2023-01-18 15:57
AB-230117-02 Cal Std 13 1 | 0 | 9730.538 | 9432.704 | 9652.493 | 9506.059 | 0 | 10214.865 | 9751.461 | 8867.332 | 9087.584 | 9077.736 | 9233.802 | 10938.694 | 0 | 9826.514 | 10164.091 | 10103.42 | 0 | 0 | 10102.586 | 9860.345 | 0 | 10124.793 | 9583.875 | 10134.089 | 0 | 10103.708 | 0 | 8150.176 | 0 | 0 | 17707.693 |  | 1 | Cal Std | 2023-01-18 16:24
AB-230117-02 CCC-H 1 | 0 | 385.362 | 398.734 | 404.887 | 418.575 | 0 | 292.39 | 392.93 | 328.415 | 389.833 | 331.89 | 338.658 | 0 | 0 | 389.955 | 331.373 | 378.211 | 0 | 0 | 372.53 | 384.696 | 0 | 379.732 | 344.117 | 381.982 | 0 | 349.518 | 0 | 792.398 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-18 19:34
AB-230117-02 CCC-H 2 | 0 | 345.949 | 405.237 | 356.5 | 448.375 | 0 | 350.502 | 403.865 | 354.539 | 395.824 | 338.298 | 349.146 | 1220.854 | 0 | 403.12 | 342.723 | 380.815 | 0 | 0 | 384.974 | 379.785 | 0 | 387.516 | 309.753 | 381.16 | 0 | 362.633 | 0 | 783.932 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 08:12
AB-230117-02 CCC-H 3 | 0 | 337.834 | 419.419 | 342.705 | 455.979 | 0 | 3615.811 | 400.863 | 332.997 | 408.594 | 344.125 | 354.775 | 0 | 0 | 405.278 | 366.428 | 393.376 | 0 | 0 | 378.782 | 377.097 | 0 | 383.968 | 320.066 | 377.585 | 0 | 357.124 | 0 | 561.845 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 21:17
AB-230117-02 IBC 1 | 0 | 190.793 | 208.097 | 192.396 | 197.738 | 0 | -132.401 | 189.323 | 167.778 | 199.011 | 161.272 | 160.013 | 0 | 0 | 188.721 | 169.148 | 195.489 | 0 | 0 | 187.723 | 190.496 | 0 | 196.495 | 155.883 | 191.07 | 0 | 167.237 | 0 | 467.225 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-18 19:07
AB-230117-02 LD 1 | 0 | 19.279 | 4.977 | -1.672 | 0.391 | 0 | -2546.733 | 0 | 10.532 | 0.309 | 0.993 | 2.418 | 0 | 0 | 0.48 | 574.418 | 1.343 | 0 | 0 | 0.777 | 1.066 | 0 | 0.632 | 0.701 | 0.424 | 0 | 0.761 | 0 | -5.707 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-19 07:18
AB-230117-02 LD 2 | 0 | -13.426 | 0 | -0.182 | 0.556 | 0 | -267.63 | 0 | 2.738 | 161.815 | 5.696 | 7256.85 | 0 | 0 | 0.528 | 0 | 0.39 | 0 | 0 | 0.236 | 0 | 0 | 0 | 9.276 | 0.553 | 0 | 0.718 | 0 | -5.837 | 0 | 0 | 0 |  | 200 | Unknown | 2023-01-19 20:23
AB-230117-02 LFB-H 1 | 0 | 3.695 | -0.164 | 1.695 | 5.355 | 0 | -70.393 | 178.384 | 160.623 | 206.101 | 185.074 | 180.479 | 0 | 0 | 3.972 | 183.485 | 191.802 | 0 | 0 | 207.335 | 194.206 | 0 | 204.981 | 162.745 | 198.792 | 0 | 3.616 | 0 | 556.028 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-18 18:13
AB-230117-02 LFB-H 2 | 0 | 175.463 | 205.745 | 202.468 | 203.342 | 0 | 177.92 | 196.167 | 168.796 | 194.454 | 164.256 | 163.131 | 0 | 0 | 197.893 | 168.178 | 191.306 | 0 | 0 | 187.116 | 190.739 | 0 | 194.708 | 156.284 | 188.953 | 0 | 178.779 | 0 | 365.417 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 07:45
AB-230117-02 LFB-H 3 | 0 | 203.096 | 200.347 | 185.527 | 235.568 | 0 | 132.029 | 198.181 | 174.498 | 200.319 | 162.174 | 578.353 | 1133.244 | 0 | 194.01 | 169.371 | 186.355 | 0 | 0 | 189.889 | 191.991 | 0 | 189.526 | 148.021 | 190.049 | 0 | 171.972 | 0 | 399.14 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 20:50
AB-230117-02 LFSM-H 1 @5 | 0 | 311.403 | 222.11 | 233.218 | 219.826 | 0 | -2571.449 | 197.796 | 195.907 | 271.782 | 160.987 | 175.543 | 0 | 0 | 220.519 | 570.426 | 212.68 | 0 | 0 | 198.678 | 197.561 | 0 | 209.365 | 234.207 | 202.268 | 0 | 200.144 | 0 | 473.874 | 4509.996 | 0 | 0 |  | 2 | Unknown | 2023-01-19 08:39
AB-230117-02 LFSM-H 2 | 0 | 136.812 | 190.809 | 191.947 | 511.677 | 0 | 464.347 | 198.319 | 178.852 | 348.673 | 174.353 | 6985.084 | 0 | 0 | 210.273 | 169.438 | 206.37 | 0 | 0 | 186.077 | 190.263 | 0 | 202.465 | 174.947 | 196.626 | 0 | 176.42 | 0 | 399.43 | 0 | 0 | 0 |  | 200 | Unknown | 2023-01-19 21:44
LRB2-Inj1 | 0 | -5.457 | -4.446 | 0.14 | 0.461 | 0 | 777.915 | 0 | 0 | 0 | 0 | 0.532 | 0 | 0 | 0.553 | 0 | 204.132 | 0 | 0 | -0.213 | 0 | 0 | 0.414 | 0.745 | 0.46 | 0 | 0.828 | 0 | -5.669 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 01:26
AB-230117-02 LRB 2 | 0 | -4.465 | -3.552 | 0.977 | 0.428 | 0 | -789.43 | 0 | 0 | 0 | 0 | 0.413 | 0 | 0 | 0.669 | 0 | 7.368 | 0 | 0 | 0.033 | 0.364 | 0 | 0.349 | 0 | 0.466 | 0 | 0.849 | 0 | -5.087 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 01:53
AB-230117-02 LRB 3 | 0 | -5.609 | -4.808 | -1.395 | 0.446 | 0 | -119.75 | 0 | 0 | 1.341 | 0 | 67.797 | 0 | 0 | 0.413 | 0 | 0.983 | 0 | 0 | 0.028 | 0 | 0 | 0.37 | 0.784 | 0.452 | 0 | 0.866 | 0 | -5.484 | 1.367 | 0 | 0 |  | 1 | Unknown | 2023-01-19 06:51
LRB4-inj1 | 0 | 1.514 | -2.845 | -0.729 | 0.361 | 0 | 44.313 | 0 | 0 | 0 | 0 | 0.907 | 0 | 0 | 0.371 | 0 | 0.334 | 0 | 0 | -0.395 | 0.299 | 0 | 0.282 | 0 | 0.368 | 0 | 0.748 | 0 | -5.851 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 14:04
AB-230117-02 LRB 4 | 0 | 29.293 | -3.76 | -0.358 | 0.365 | 0 | 527.884 | 0.341 | 0 | 0 | 0 | 0.692 | 0 | 0 | 0.409 | 0 | 0.359 | 0 | 0 | 0 | 0 | 0 | 0.302 | 0 | 0.32 | 0 | 0.762 | 0 | -5.869 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 14:31
LRB5-inj1 | 0 | -6.821 | -3.242 | -0.026 | 0.429 | 0 | -105.886 | 0 | 0 | 0.055 | 0.817 | 0.588 | 0 | 0 | 0.557 | 0 | 0.473 | 0 | 0 | -0.225 | 0 | 0 | 0.372 | 0.655 | 0.429 | 0 | 0.736 | 0 | -5.91 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 19:29
AB-230117-02 LRB 5 | 0 | 6.597 | -2.882 | -0.354 | 0.355 | 0 | -325.863 | 0 | 1.083 | 0 | 0.774 | 0.513 | 0 | 0 | 0.46 | 0 | 0.334 | 0 | 0 | -0.234 | 0 | 0 | 0.288 | 0.651 | 0.37 | 0 | 0.759 | 0 | -5.285 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 19:56
Rinse5 | 0 | 721.957 | 331.464 | 92.14 | 219.16 | 0 | 4024.193 | 0 | 0 | 427.283 | 125.204 | 6694.884 | 0 | 0 | 190.002 | 0 | 175.37 | 0 | 0 | 139.192 | 39.805 | 0 | 169.964 | 174.028 | 160.825 | 0 | 144.437 | 0 | 277.142 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 22:11
LRB6-inj1 | 0 | -8.173 | -3.104 | -0.428 | 0 | 0 | -912.107 | 0 | 1.051 | 0 | 0 | 0.554 | 0 | 0 | 0.488 | 0 | 0.343 | 0 | 0 | -0.344 | 0.352 | 0 | 0.221 | 0.66 | 0.362 | 0 | 0.725 | 0 | -4.753 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 03:09
AB-230117-02 LRB 6 | 0 | -3.834 | -4.704 | -0.241 | 0.354 | 0 | -156.963 | 0 | 1.452 | -0.028 | 0 | 0.657 | 0 | 0 | 0.447 | 0 | 0.346 | 0 | 0 | 0.04 | 0 | 0 | 0.24 | 0.756 | 0.308 | 0 | 0.729 | 0 | -5.088 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 03:36
AB-230117-02 LRB 1 | 0 | -4.387 | -1.13 | 0.339 | 4.722 | 0 | 103.833 | 1.125 | 0 | 4.052 | 3.296 | 3.964 | 0 | 0 | 4.164 | 0 | 4.651 | 0 | 0 | 1.439 | 0 | 0 | 4.565 | 4.532 | 3.537 | 0 | 5.969 | 0 | -2.944 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-18 17:46
AB-230117-02 QCS 1 | 0 | 222.006 | 231.135 | 228.66 | 234.748 | 0 | 137.107 | 219.599 | 197.426 | 228.332 | 186.245 | 186.768 | 0 | 0 | 6.392 | 215.739 | 213.968 | 0 | 0 | 214.666 | 212.607 | 0 | 210.554 | 15.561 | 211.384 | 0 | 6.033 | 0 | 497.605 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-18 18:40
WO-23003-01-A-1 @10 | 0 | 102.349 | 17.153 | -1.827 | 1.059 | 0 | -1781.897 | 0 | 11.936 | 1.204 | 1.272 | 1.714 | 0 | 0 | 0.859 | 429.802 | 1.464 | 0 | 0 | 0.641 | 1.148 | 0 | 1.291 | 1.371 | 1.127 | 0 | 1.506 | 0 | -4.44 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-18 20:55
WO-23003-02-A-1 | 0 | 87.645 | 9.739 | -2.227 | 0.779 | 0 | -2518.655 | 2.86 | 12.332 | 0.911 | 1.232 | 1.579 | 0 | 0 | 0.801 | 458.755 | 1.402 | 0 | 0 | 3.701 | 0 | 0 | 1.012 | 1.069 | 0.97 | 0 | 1.314 | 0 | -4.854 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-18 21:22
WO-23003-03-A-1 | 0 | 77.152 | 8.005 | -1.763 | 0.694 | 0 | -1606.626 | 0 | 11.056 | 0.525 | 1.223 | 1.271 | 0 | 0 | 0.728 | 463.337 | 1.12 | 0 | 0 | 0.072 | 0.902 | 0 | 0.908 | 1.036 | 0.843 | 0 | 1.185 | 0 | -5.105 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-18 21:49
WO-23003-04-A-1 | 0 | 169.578 | 17.44 | 0.46 | 0.63 | 0 | -3225.553 | 0 | 0 | 0.631 | 0 | 1.588 | 0 | 0 | 0.576 | 434.867 | 0.822 | 0 | 0 | 0.944 | 0 | 0 | 0.772 | 0.908 | 0.757 | 0 | 1.054 | 0 | 0 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-18 22:16
WO-23003-05-A-1 | 0 | 179.896 | 23.684 | -1.516 | 0.611 | 0 | -2833.139 | 0 | 0 | 0.873 | 0 | 1.348 | 0 | 0 | 0.497 | 511.067 | 1.093 | 0 | 0 | 0 | 0 | 0 | 0.744 | 0.892 | 0.675 | 0 | 0.955 | 0 | 0 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-18 22:43
WO-23003-06-A-1 | 0 | 194.87 | 21.382 | -1.044 | 0.584 | 0 | -3075.464 | 0 | 0 | 0 | 0 | 8.176 | 0 | 0 | 0.691 | 462.818 | 0.769 | 0 | 0 | 5.339 | 0 | 0 | 0.771 | 0.837 | 0.657 | 0 | 0.989 | 0 | 0 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-18 23:10
WO-23003-07-A-1 | 0 | -0.162 | -3.679 | -1.026 | 0.48 | 0 | 571.97 | 0 | 0 | 0.156 | 0.934 | 0.592 | 0 | 0 | 0.906 | 0 | 3898.943 | 0 | 0 | 0.318 | 0 | 0 | 1.8 | 0 | 1.008 | 0 | 0.944 | 0 | -5.957 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-18 23:37
WO-23003-08-A-1 | 0 | -7.845 | -3.222 | -0.974 | 0.484 | 0 | 327.435 | 0 | 0 | 0.062 | 0 | 0.143 | 0 | 0 | 0.938 | 0 | 4309.18 | 0 | 0 | -0.098 | 0 | 0 | 1.8 | 0 | 0.97 | 0 | 0.904 | 0 | -5.795 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 00:05
WO-23003-09-A-1 | 0 | 33.342 | -3.591 | -1.23 | 0.445 | 0 | 1224.806 | 0 | 0 | 0 | 0 | 0.27 | 0 | 0 | 0.988 | 0 | 3931.95 | 0 | 0 | 0 | 0 | 0 | 1.666 | 0 | 1.051 | 0 | 0.929 | 0 | -5.484 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 00:32
WO-23003-10-A-1 | 0 | -8.664 | -3.038 | 2.19 | 0.425 | 0 | -359.729 | 0 | 0 | 0.056 | 0 | 0.079 | 0 | 0 | 0.985 | 0 | 4432.787 | 0 | 0 | -0.097 | 0 | 0 | 1.794 | 0 | 0.95 | 0 | 0.854 | 0 | -5.715 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 00:59
WO-23003-11-A-1 | 0 | -16.634 | -3.581 | -0.534 | 0.389 | 0 | -438.895 | 0.293 | 0 | 0 | 0 | 0.143 | 0 | 0 | 1.04 | 0 | 4146.818 | 0 | 0 | 0.016 | 0 | 0 | 1.711 | 0 | 0.861 | 0 | 0.844 | 0 | -5.506 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 02:20
WO-23003-12-A-1 | 0 | -12.417 | -4.349 | -0.145 | 0.397 | 0 | -57.933 | 0 | 0 | 0 | 0 | 0.272 | 0 | 0 | 1.078 | 0 | 4137.805 | 0 | 0 | -0.04 | 0 | 0 | 1.709 | 0 | 0.96 | 0 | 0.753 | 0 | -5.271 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 02:47
WO-23003-13-A-1 | 0 | -5.993 | -3.172 | 1.969 | 0.414 | 0 | 174.572 | 0.791 | 0 | 0 | 0 | 0.159 | 0 | 0 | 1.442 | 0 | 4255.588 | 0 | 0 | -0.069 | 0 | 0 | 1.759 | 0 | 0.961 | 0 | 0.777 | 0 | -5.492 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 03:14
WO-23003-14-A-1 | 0 | 31.051 | 5.44 | -2.184 | 0.425 | 0 | -767.572 | 0 | 4778.929 | 0.187 | 0 | 0.732 | 0 | 0 | 0.348 | 0 | 193.554 | 0 | 0 | 4.04 | 0 | 0 | 0.504 | 0 | 0.529 | 0 | 0.867 | 0 | -5.131 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 03:41
WO-23003-15-A-1 | 0 | 122.795 | 6.429 | -1.146 | 0.445 | 0 | 1885.316 | 0 | 4665.408 | 0 | 0.969 | 0.966 | 0 | 0 | 0.18 | 0 | 7.75 | 0 | 0 | 3.858 | 0.817 | 0 | 0.488 | 0.754 | 0.502 | 0 | 0.936 | 0 | -4.624 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 04:08
WO-23003-16-A-1 | 0 | 32.463 | 4.202 | -1.835 | 0.421 | 0 | -903.118 | 0.678 | 4754.423 | 0 | 0.955 | 0.884 | 0 | 0 | 0.218 | 0 | 4.083 | 0 | 0 | 3.737 | 0.767 | 0 | 0.505 | 0 | 0.465 | 0 | 0.798 | 0 | -4.805 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 04:35
WO-23003-17-A-1 | 0 | 34.024 | 5.429 | -1.916 | 0.437 | 0 | -861.573 | 0.851 | 4685.686 | 0.165 | 0 | 0.981 | 0 | 0 | 0.58 | 0 | 2.358 | 0 | 0 | 6.124 | 0.834 | 0 | 0.469 | 0.695 | 0.579 | 0 | 0.802 | 0 | -4.72 | 2.366 | 0 | 0 |  | 1 | Unknown | 2023-01-19 05:02
WO-23003-18-A-1 | 0 | 41.081 | 4.209 | -1.447 | 0.484 | 0 | -900.536 | 0.314 | 4759.817 | 0.324 | 1.021 | 0.883 | 0 | 0 | 0.72 | 0 | 1.905 | 0 | 0 | 3.8 | 0.925 | 0 | 0.508 | 0.793 | 0.69 | 0 | 0.831 | 0 | -4.788 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 05:29
WO-23003-19-A-1 | 0 | 56.352 | 2.334 | -0.769 | 0.5 | 0 | 408.006 | 0.589 | 4724.366 | 0.205 | 0 | 0.885 | 0 | 0 | 0.68 | 0 | 1.605 | 0 | 0 | 3.673 | 0 | 0 | 0.569 | 0.725 | 0.681 | 0 | 0.888 | 0 | -5.025 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 05:57
WO-23003-20-A-1 | 0 | 28.894 | 0 | -0.452 | 0.613 | 0 | 3274.15 | 0.342 | 0 | 37.335 | 1.865 | 1470.859 | 0 | 0 | 0.505 | 0 | 1.392 | 0 | 0 | 0.209 | 0 | 0 | 0.531 | 2.73 | 0.524 | 0 | 0.872 | 0 | -5.293 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 06:24
WO-23003-21-A-1 | 0 | -9.125 | -4.909 | 5.342 | 5.371 | 0 | -500 | 1.109 | 0 | 34.83 | 2.776 | 1416.32 | 0 | 0 | 3.39 | 0 | 3.141 | 0 | 0 | 0.998 | 1.055 | 0 | 2.538 | 5.021 | 1.898 | 0 | 3.321 | 0 | -4.713 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 09:33
WO-23003-22-A-1 | 0 | -7.308 | 0 | -1.81 | 1.01 | 0 | 60.363 | 0 | 0 | 31.662 | 1.996 | 1386.347 | 0 | 0 | 0.828 | 0 | 0.812 | 0 | 0 | 1.349 | 0 | 0 | 0.601 | 2.62 | 0.665 | 0 | 1.119 | 0 | -5.541 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 10:00
WO-23003-23-A-1 | 0 | -7.589 | 0 | -1.758 | 0.729 | 0 | -272.342 | 0 | 0 | 35.451 | 1.924 | 1540.035 | 0 | 0 | 0.797 | 0 | 0.718 | 0 | 0 | 0.172 | 0.597 | 0 | 0.484 | 2.82 | 0.544 | 0 | 1.008 | 0 | -5.776 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 10:27
WO-23003-24-A-1 | 0 | 1.161 | 0 | 0.435 | 0.938 | 0 | 511.516 | 0.512 | 0 | 34.306 | 2.015 | 1397.807 | 0 | 0 | 0.887 | 0 | 0.913 | 0 | 0 | 0.732 | 0.905 | 0 | 0.727 | 2.772 | 0.863 | 0 | 0.845 | 0 | -5.033 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 10:54
WO-23003-25-A-1 | 0 | 0.707 | 0 | -1.022 | 0.592 | 0 | 199.062 | 0.818 | 0 | 37.372 | 1.981 | 1542.714 | 0 | 0 | 0.592 | 0 | 0.578 | 0 | 0 | 0.239 | 0.512 | 0 | 0.436 | 2.743 | 0.515 | 0 | 0.941 | 0 | -5.402 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 11:21
WO-23003-26-A-1 | 0 | 57.364 | 11.807 | -0.273 | 0.415 | 0 | -1640.296 | 1.004 | 0 | 1.855 | 0 | 63.814 | 0 | 0 | 0.53 | 0 | 0.489 | 0 | 0 | 0.119 | 0 | 0 | 0.392 | 0.891 | 0.349 | 0 | 0.787 | 0 | -4.638 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 11:49
WO-23003-27-A-1 | 0 | 43.187 | 10.628 | -2.156 | 0.383 | 0 | -1668.934 | 0 | 4.255 | 0.252 | 0 | 2.775 | 0 | 0 | 0.65 | 0 | 0.461 | 0 | 0 | -0.054 | 0 | 0 | 0.363 | 0 | 0.424 | 0 | 0.798 | 0 | -5.376 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 12:16
WO-23003-28-A-1 | 0 | 58.511 | 11.078 | -1.202 | 0.393 | 0 | -1322.278 | 0 | 0 | 0 | 0 | 1.4 | 0 | 0 | 0.44 | 0 | 0.444 | 0 | 0 | 0 | 0 | 0 | 0.397 | 0 | 0.354 | 0 | 0.802 | 0 | -5.076 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 12:43
WO-23003-29-A-1 | 0 | 65.149 | 12.443 | -1.327 | 0.385 | 0 | -1673.799 | 0 | 2.843 | 0.155 | 0 | 0.915 | 0 | 0 | 0.515 | 0 | 0.406 | 0 | 0 | 0 | 0 | 0 | 0.324 | 0 | 0.28 | 0 | 0.841 | 0 | 0 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 13:10
WO-23003-30-A-1 | 0 | 51.374 | 14.607 | -1.681 | 0.37 | 0 | -1872.615 | 0 | 0 | 0.21 | 0 | 0.911 | 0 | 0 | 0.607 | 0 | 0.484 | 0 | 0 | 0.163 | 0 | 0 | 0.37 | 0 | 0.295 | 0 | 0.776 | 0 | -4.781 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 13:37
WO-23003-31-A-1 | 0 | 129.1 | 12.705 | -0.898 | 0.389 | 0 | -746.072 | 0 | 0 | 0.265 | 0 | 0.994 | 0 | 0 | 0.337 | 0 | 0.955 | 0 | 0 | 0 | 0 | 0 | 0.548 | 0 | 0.702 | 0 | 0.749 | 0 | 0 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 14:58
WO-23003-32-A-1 | 0 | 167.609 | 11.98 | -0.548 | 0 | 0 | -1448.898 | 2.118 | 0 | 0 | 0 | 0.553 | 0 | 0 | 0.354 | 0 | 0.385 | 0 | 0 | 0 | 0 | 0 | 0.346 | 0 | 0.394 | 0 | 0.757 | 0 | -5.256 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 15:25
WO-23003-33-A-1 | 0 | 66.955 | 11.216 | -1.442 | 0.371 | 0 | -1359.876 | 0 | 9.357 | 0 | 0 | 0.94 | 0 | 0 | 0.553 | 0 | 0.674 | 0 | 0 | 0 | 1.126 | 0 | 0.488 | 0.64 | 0.355 | 0 | 0.717 | 0 | -4.694 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 15:52
WO-23003-34-A-1 | 0 | 75.414 | 12.203 | -2.275 | 0.377 | 0 | -1629.974 | 0 | 9.979 | 0 | 0.94 | 0.906 | 0 | 0 | 0.509 | 0 | 0.663 | 0 | 0 | 0 | 0.991 | 0 | 0.521 | 0.721 | 0.484 | 0 | 0.766 | 0 | -5.233 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 16:19
WO-23003-35-A-1 | 0 | 84.423 | 8.691 | -1.053 | 0.41 | 0 | -1568.524 | 0 | 8.866 | 0.316 | 0.908 | 1.126 | 0 | 0 | 0.458 | 0 | 0.71 | 0 | 0 | 0 | 1.091 | 0 | 0.514 | 0.635 | 0.491 | 0 | 0.766 | 0 | -5.555 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 16:46
WO-23003-36-A-1 | 0 | 72.702 | 12.557 | -0.235 | 0.455 | 0 | -1442.188 | 0.99 | 8.682 | 0.215 | 0 | 1.314 | 0 | 0 | 0.639 | 0 | 0.763 | 0 | 0 | 0 | 0.885 | 0 | 0.587 | 0.726 | 0.516 | 0 | 0.77 | 0 | -5.188 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 17:14
WO-23003-37-A-1 | 0 | 94.073 | 11.154 | -0.679 | 0.364 | 0 | -1722.622 | 0 | 8.757 | 0.064 | 0.932 | 1.009 | 0 | 0 | 0.277 | 0 | 0.577 | 0 | 0 | 0 | 0.965 | 0 | 0.473 | 0 | 0.344 | 0 | 0.762 | 0 | -5.242 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 17:41
WO-23003-38-A-1 | 0 | 71.857 | 11.321 | -0.351 | 0 | 0 | -1985.937 | 0 | 10.034 | 0 | 0 | 1.074 | 0 | 0 | 0.438 | 0 | 0.639 | 0 | 0 | 0 | 0.903 | 0 | 0.435 | 0.629 | 0.433 | 0 | 0.729 | 0 | -4.867 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 18:08
WO-23003-39-A-1 | 0 | 96.773 | 14.017 | -1.129 | 0.343 | 0 | -2379.923 | 0.513 | 0 | 0 | 0 | 0.391 | 0 | 0 | 0.301 | 126.713 | 0.309 | 0 | 0 | 0.517 | 0 | 0 | 0.302 | 0 | 0.402 | 0 | 0.729 | 0 | -3.654 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 18:35
WO-23003-40-A-1 | 0 | 53.886 | 12.627 | -1.151 | 0.35 | 0 | -1682.319 | 1.336 | 0 | 0.141 | 0.769 | 0.454 | 0 | 0 | 0.211 | 134.042 | 0.518 | 0 | 0 | 0.224 | 0 | 0 | 0.443 | 0.698 | 0.33 | 0 | 0.75 | 0 | -4.24 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 19:02
WO-23003-41-A-1 | 0 | 69.688 | 10.95 | -1.808 | 0.845 | 0 | -1986.736 | 0.582 | 2.138 | 0.843 | 1.08 | 10.144 | 0 | 0 | 0.882 | 150.848 | 0.677 | 0 | 0 | 0 | 0 | 0 | 0.764 | 1.108 | 0.665 | 0 | 1.315 | 0 | -5.144 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 22:38
WO-23003-42-A-1 | 0 | 87.99 | 12.286 | -1.332 | 0.534 | 0 | -2104.218 | 0 | 1.32 | 0.63 | 0.822 | 3.824 | 0 | 0 | 0.589 | 136.634 | 0.442 | 0 | 0 | 0.452 | 0 | 0 | 0.466 | 0.808 | 0.454 | 0 | 0.887 | 0 | -4.282 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 23:06
WO-23003-43-A-1 | 0 | 71.186 | 12.53 | -0.713 | 0.444 | 0 | -1882.501 | 0 | 0 | 0.197 | 0 | 2.172 | 0 | 0 | 0.683 | 111.294 | 0.4 | 0 | 0 | 0.326 | 0 | 0 | 0.416 | 0.78 | 0.354 | 0 | 0.934 | 0 | -5.579 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-19 23:33
WO-23003-44-A-1 | 0 | 132.635 | 11.786 | -1.048 | 0.408 | 0 | -303.851 | 1.537 | 1.786 | 0.765 | 0.82 | 1.544 | 0 | 0 | 0.622 | 109.959 | 0.359 | 0 | 0 | 0.56 | 0 | 0 | 0.413 | 0.957 | 0.355 | 0 | 0.842 | 0 | -4.166 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20
WO-23003-45-A-1 | 0 | 134.043 | 10.332 | -1.235 | 0.393 | 0 | -77.53 | 0 | 1.622 | 0.19 | 0.782 | 1.121 | 0 | 0 | 0.43 | 126.129 | 0.368 | 0 | 0 | 0.39 | 0 | 0 | 0.366 | 0.709 | 0.401 | 0 | 0.83 | 0 | -3.679 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 00:27
WO-23003-46-A-1 | 0 | 70.431 | 14.225 | -1.2 | 0.377 | 0 | -2293.717 | 0 | 1.285 | 1.032 | 0.914 | 1.038 | 0 | 0 | 0.441 | 117.077 | 0.45 | 0 | 0 | 2.305 | 0 | 0 | 0.427 | 0.832 | 0.328 | 0 | 0.786 | 0 | -4.634 | 0.215 | 0 | 0 |  | 1 | Unknown | 2023-01-20 00:54
WO-23003-47-A-1 | 0 | 140.708 | 19.502 | -1.238 | 0.37 | 0 | -3246.202 | 1.131 | 0 | 0.901 | 0 | 1.061 | 0 | 0 | 0.296 | 268.234 | 0.318 | 0 | 0 | 0 | 0 | 0 | 0.355 | 1.047 | 0.429 | 0 | 0.778 | 0 | -2.226 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 01:21
WO-23003-48-A-1 | 0 | 53.173 | 12.796 | 0.517 | 0.362 | 0 | -1807.114 | 0.421 | 0 | 0.268 | 0 | 0.744 | 0 | 0 | 0.481 | 253.554 | 0.334 | 0 | 0 | 0.424 | 0 | 0 | 0.363 | 0.69 | 0.491 | 0 | 0.779 | 0 | -4.234 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 01:48
WO-23003-49-A-1 | 0 | 75.231 | 12.672 | -0.308 | 0.353 | 0 | -1555.867 | 0 | 1.26 | 0.27 | 0 | 0.719 | 0 | 0 | 0.719 | 250.815 | 0.323 | 0 | 0 | 0.418 | 0 | 0 | 0.381 | 0.851 | 0.807 | 0 | 0.74 | 0 | -3.44 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 02:15
WO-23003-50-A-1 | 0 | 189.432 | 19.964 | -0.698 | 0.352 | 0 | -3016.487 | 2.007 | 2.099 | 0.963 | 0 | 0.744 | 0 | 0 | 0.432 | 214.775 | 0.264 | 0 | 0 | 6.252 | 0 | 0 | 0.388 | 0.718 | 0 | 0 | 0.76 | 0 | -2.066 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 02:42
WO-23003-51-A-1 | 0 | 148.297 | 14.144 | -1.405 | 0.349 | 0 | -1019.414 | 0 | 2.96 | 0.287 | 0.784 | 0.737 | 0 | 0 | 0.394 | 498.597 | 0.295 | 0 | 0 | 0 | 0 | 0 | 0.324 | 0.73 | 0.375 | 0 | 0.736 | 0 | -3.027 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 04:04
WO-23003-52-A-1 | 0 | 121.947 | 9.354 | -1.494 | 0 | 0 | -497.088 | 1.442 | 1.383 | 0.249 | 0 | 0.657 | 0 | 0 | 0.441 | 413.609 | 0.235 | 0 | 0 | 0.949 | 0 | 0 | 0.342 | 0.788 | 0.295 | 0 | 0.748 | 0 | -4.26 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 04:31
WO-23003-53-A-1 | 0 | 63.869 | 8.815 | -1.373 | 0.353 | 0 | -1726.66 | 0 | 2.725 | 0.101 | 0.805 | 0.556 | 0 | 0 | 0.521 | 419.335 | 0.428 | 0 | 0 | 0.413 | 0 | 0 | 0.362 | 0.763 | 0.355 | 0 | 0.714 | 0 | -4.253 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 04:58
WO-23003-54-A-1 | 0 | 99.103 | 15.857 | -1.1 | 0.364 | 0 | -2169.25 | 0 | 9.358 | 0.777 | 0.988 | 1.149 | 0 | 0 | 0.502 | 446.172 | 0.617 | 0 | 0 | 0.612 | 0 | 0 | 0.507 | 0.848 | 0.466 | 0 | 0.737 | 0 | -2.451 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 05:25
WO-23003-55-A-1 | 0 | 102.639 | 8.308 | -1.367 | 0 | 0 | -2473.552 | 0 | 7.447 | 0.436 | 0.901 | 1.348 | 0 | 0 | 0.269 | 456.386 | 1.323 | 0 | 0 | 0 | 0.584 | 0 | 0.541 | 0.933 | 0.523 | 0 | 0.75 | 0 | -2.361 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 05:52
WO-23003-56-A-1 | 0 | 109.563 | 13.594 | -1.21 | 0.339 | 0 | -1865.556 | 0 | 7.568 | 0.717 | 0.968 | 1.232 | 0 | 0 | 0.34 | 433.565 | 0.576 | 0 | 0 | 3.78 | 0 | 0 | 0.563 | 0.759 | 0.523 | 0 | 0.723 | 0 | -4.251 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 06:19
LFB-H-3-Inj2 | 0 | 140.66 | 208.14 | 213.786 | 245.586 | 0 | -435.592 | 196.02 | 183.118 | 164.246 | 164.593 | 168.317 | 0 | 0 | 193.778 | 174.08 | 197.419 | 0 | 0 | 190.077 | 191.147 | 0 | 195.058 | 157.581 | 193.469 | 0 | 180.139 | 0 | 670.231 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 13:11
WO-23003-14-A-1B | 0 | 53.523 | 7.992 | -2.221 | 0 | 0 | -1440.617 | 0 | 4236.334 | 0.441 | 0 | 0.898 | 0 | 0 | 0.334 | 0 | 1.981 | 0 | 0 | 3.569 | 0.614 | 0 | 0.383 | 0.749 | 0.47 | 0 | 0.712 | 0 | -1.737 | 0 | 0 | 0 |  | 4 | Unknown | 2023-01-20 12:44
LRB-7-inj3 | 0 | 64.564 | 12.519 | -1.97 | 0 | 0 | -2183.995 | 1.41 | 1.355 | 0.01 | 0.761 | 0.42 | 0 | 0 | 0.416 | 157.537 | 4.078 | 0 | 0 | 0.667 | 0 | 0 | 0.324 | 0.651 | 0.295 | 0 | 0.728 | 0 | -3.946 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-20 12:17
LRB-7-inj2 | 0 | -3.489 | -5.103 | -0.39 | 0 | 0 | -370.969 | 0 | 0 | -0.026 | 0.766 | 0.387 | 0 | 0 | 0.317 | 0 | 153.845 | 0 | 0 | 0.088 | 0 | 0 | 0.334 | 0.769 | 0.294 | 0 | 0.719 | 0 | -4.659 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 11:50
WO-23003-60-A-1B | 0 | 143.791 | 16.675 | 2.795 | 0.347 | 0 | -3252.894 | 0 | 1.537 | 0.543 | 0 | 0.336 | 0 | 0 | 0.867 | 0 | 4183.607 | 0 | 0 | 2.882 | 0 | 0 | 2.809 | 1.324 | 1.587 | 0 | 0.78 | 0 | -1.869 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-20 11:23
WO-23003-59-A-1B | 0 | 80.547 | 12.996 | -0.138 | 0.383 | 0 | -1872.514 | 0.234 | 1.119 | 0.678 | 0.834 | 0.539 | 0 | 0 | 0.564 | 0 | 632.89 | 0 | 0 | 2.687 | 0 | 0 | 0.839 | 0.754 | 0.616 | 0 | 0.738 | 0 | -4.476 | 4.404 | 0 | 0 |  | 2 | Unknown | 2023-01-20 10:55
WO-23003-58-A-1B | 0 | 110.815 | 13.451 | 0.336 | 0.391 | 0 | -846.279 | 0 | 1.919 | 0.554 | 0.895 | 0.816 | 0 | 0 | 0.367 | 0 | 61.125 | 0 | 0 | 4.18 | 0 | 0 | 1.257 | 0.923 | 0.492 | 0 | 0.739 | 0 | -4.533 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-20 10:28
WO-23003-26-A-1B | 0 | 140.856 | 17.933 | -0.461 | 0.357 | 0 | -2195.753 | 0 | 5.584 | 0.376 | 0 | 0.606 | 0 | 0 | 0.501 | 0 | 0.466 | 0 | 0 | -0.059 | 0 | 0 | 0.417 | 0.78 | 0.288 | 0 | 0.78 | 0 | -3.887 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-20 10:01
WO-23003-124-A-1 | 0 | -23.237 | -5.273 | -0.824 | 0.382 | 0 | 81.578 | 0 | 0 | 0.002 | 0.821 | -0.14 | 1124.729 | 0 | 1.035 | 0 | 4254.003 | 0 | 0 | -0.116 | 0.721 | 0 | 2.549 | 0.788 | 1.325 | 0 | 0.787 | 0 | -5.628 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 11:13
WO-23003-123-A-1 | 0 | -24.267 | -4.586 | 0.034 | 0.398 | 0 | -115.917 | 0 | 0 | 0.182 | 0.797 | -0.053 | 0 | 0 | 1.14 | 0 | 4114.272 | 0 | 0 | -0.091 | 0.825 | 0 | 3.799 | 1.074 | 1.448 | 0 | 0.803 | 0 | -5.716 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 10:46
WO-23003-122-A-1 | 0 | -21.73 | -4.194 | 0.848 | 0.459 | 0 | -261.406 | 0 | 0 | 0.174 | 0.815 | 0.037 | 0 | 0 | 1.131 | 0 | 4052.461 | 0 | 0 | 0.023 | 0.942 | 0 | 3.298 | 1.329 | 1.505 | 0 | 0.895 | 0 | -5.715 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 10:19
WO-23003-121-A-1 | 0 | 12.878 | 0.073 | 1.206 | 0.658 | 0 | -592.908 | 0 | 0 | 0.742 | 0.926 | 0.211 | 0 | 0 | 1.054 | 0 | 2598.122 | 0 | 0 | 2.996 | 1.724 | 0 | 8.418 | 1.699 | 8.728 | 0 | 1.117 | 0 | -5.129 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 09:52
AB-230117-02 LFSM-H 6 | 0 | 194.659 | 197.206 | 204.181 | 183.464 | 0 | -751.578 | 171.837 | 163.133 | 222.285 | 150.338 | 158.364 | 0 | 0 | 185.672 | 160.732 | 2843.453 | 0 | 0 | 176.546 | 176.045 | 0 | 191.689 | 199.234 | 184.547 | 0 | 179.992 | 0 | 362.027 | 0 | 0 | 3945.889 |  | 4 | Unknown | 2023-01-22 08:58
AB-230117-02 CCC-H 7 | 0 | 313.969 | 378.56 | 351.172 | 422.357 | 0 | 386.541 | 372.196 | 313.775 | 352.805 | 306.664 | 311.988 | 0 | 0 | 363.775 | 307.839 | 358.688 | 0 | 0 | 351.037 | 355.298 | 0 | 350.888 | 274.256 | 349.262 | 0 | 331.124 | 0 | 676.091 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 08:31
AB-230117-02 LFB-H 7 | 0 | 171.999 | 188.773 | 197.499 | 200.108 | 0 | 247.442 | 181.009 | 155.689 | 179.63 | 152.841 | 149.508 | 1139.203 | 0 | 180.109 | 157.151 | 303.883 | 0 | 0 | 177.143 | 182.079 | 0 | 177.827 | 136.419 | 177.55 | 0 | 160.163 | 0 | 328.706 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 08:04
AB-230117-02 LD 6 | 0 | -23.21 | -3.529 | 4.204 | 0.334 | 0 | -99.91 | 0 | 0 | -0.011 | 0 | -0.105 | 0 | 0 | 1.04 | 0 | 4142.057 | 0 | 0 | -0.128 | 0.865 | 0 | 3.079 | 0.865 | 1.374 | 0 | 0.695 | 0 | -5.699 | 0 | 0 | 0 |  | 200 | Unknown | 2023-01-22 07:37
LD6-inj1 | 0 | -20.918 | -4.217 | 3.755 | 0.344 | 0 | 37.367 | 0 | 0 | 0.014 | 0.795 | -0.052 | 0 | 0 | 1.43 | 0 | 3922.707 | 0 | 0 | -0.179 | 0.883 | 0 | 2.893 | 0.723 | 1.362 | 0 | 0.739 | 0 | -5.529 | 0 | 0 | 0 |  | 200 | Unknown | 2023-01-22 07:10
AB-230117-02 LRB 13 | 0 | -17.496 | -3.378 | -1.698 | 0.327 | 0 | -172.351 | 0.263 | 0.997 | -0.03 | 0.752 | 0.061 | 0 | 0 | 0.362 | 0 | 69.087 | 0 | 0 | 0 | 0.748 | 0 | 0.374 | 0.643 | 0.62 | 0 | 0.702 | 0 | -5.731 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 06:42
WO-23003-120-A-1 | 0 | -0.741 | 5.699 | -1.784 | 0.331 | 0 | -1072.482 | 0.615 | 0 | 0.08 | 0.761 | 0 | 0 | 0 | 0.718 | 0 | 1329.015 | 0 | 0 | 1.273 | 1.294 | 0 | 3.724 | 0.736 | 4.447 | 0 | 0.716 | 0 | -4.723 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 06:15
WO-23003-119-A-1 | 0 | 7.47 | 8.163 | -1.42 | 0 | 0 | -835.051 | 0 | 2.443 | 0.196 | 0.8 | -0.061 | 0 | 0 | 0.446 | 0 | 1085.45 | 0 | 0 | 1.297 | 1.049 | 0 | 2.851 | 0.725 | 3.789 | 0 | 0.726 | 0 | -5.284 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 05:48
WO-23003-118-A-1 | 0 | 5.418 | 24.206 | 0.391 | 0 | 0 | -725.587 | 0 | 3.543 | 0.019 | 0.761 | -0.032 | 0 | 0 | 0.865 | 0 | 2857.187 | 0 | 0 | 3.116 | 1.751 | 0 | 9.413 | 0.989 | 12.301 | 0 | 0.702 | 0 | -4.834 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 05:21
WO-23003-117-A-1 | 0 | 6.359 | 40.521 | -0.501 | 0.33 | 0 | -952.521 | 0.376 | 0 | 0.092 | 0.828 | -0.151 | 0 | 0 | 0.778 | 0 | 2511.685 | 0 | 0 | 2.496 | 1.608 | 0 | 8.243 | 0.785 | 10.943 | 0 | 0.716 | 0 | -5.124 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 04:54
WO-23003-116-A-1 | 0 | 4.468 | 7.755 | -0.774 | 0 | 0 | -838.898 | 0 | 3.453 | 0.103 | 0.756 | 0.044 | 0 | 0 | 0.705 | 0 | 2477.13 | 0 | 0 | 2.716 | 1.71 | 0 | 8.8 | 0.795 | 11.677 | 0 | 0.701 | 0 | -5.132 | 3.168 | 0 | 0 |  | 1 | Unknown | 2023-01-22 04:27
WO-23003-115-A-1 | 0 | 42.735 | 153.375 | -0.792 | 0 | 0 | -1485.127 | 0.437 | 0 | 0.13 | 0.841 | 0.047 | 0 | 0 | 0.393 | 0 | 714.013 | 0 | 0 | 2.204 | 1.365 | 0 | 1.895 | 0.686 | 2.371 | 0 | 0.718 | 0 | -4.09 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 04:00
WO-23003-114-A-1 | 0 | 30.584 | 45.908 | 0.349 | 0 | 0 | -1797.346 | 0 | 2.409 | 0.151 | 0.762 | -0.089 | 0 | 0 | 0.418 | 0 | 622.53 | 0 | 0 | 2.214 | 1.029 | 0 | 1.569 | 0.801 | 2.034 | 0 | 0.725 | 0 | -4.916 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 03:33
WO-23003-113-A-1 | 0 | 29.409 | 29.561 | -0.972 | 0 | 0 | -1928.643 | 0 | 2.379 | 0.462 | 0.746 | -0.04 | 0 | 0 | 0.476 | 0 | 641.876 | 0 | 0 | 0.381 | 0.765 | 0 | 1.514 | 0.726 | 2.136 | 0 | 0.741 | 0 | -5.183 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 03:06
WO-23003-112-A-1 | 0 | 103.462 | 69.81 | 0.392 | 0.335 | 0 | -1720.371 | 0 | 4.125 | 0.349 | 0 | 0.084 | 0 | 0 | 0.277 | 0 | 536.317 | 0 | 0 | 2.721 | 1.139 | 0 | 1.711 | 0.793 | 2.811 | 0 | 0.692 | 0 | -4.792 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 02:39
WO-23003-111-A-1 | 0 | 38.064 | 63.909 | -0.484 | 0 | 0 | -1837.754 | 0 | 4.153 | 0.435 | 0 | 0.379 | 0 | 0 | 0.384 | 0 | 740.764 | 0 | 0 | 3.217 | 1.35 | 0 | 2.348 | 0.756 | 4.279 | 0 | 0.685 | 0 | -4.5 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 02:12
AB-230117-02 LRB 12 | 0 | -15.737 | -3.219 | -0.499 | 0.342 | 0 | -64.152 | 0 | 1.042 | 0.073 | 0 | 0.02 | 0 | 0 | 0.438 | 0 | 1.003 | 0 | 0 | -0.25 | 0.758 | 0 | 0.278 | 0.675 | 0.534 | 0 | 0.762 | 0 | -5.279 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 01:44
LRB-12-inj1 | 0 | -21.937 | -0.563 | -0.016 | 0.337 | 0 | 40.195 | 0 | 0 | 0.016 | 0.791 | -0.026 | 0 | 0 | 0.274 | 0 | 24.998 | 0 | 0 | -0.156 | 0.799 | 0 | 0.302 | 0.713 | 0.648 | 0 | 0.744 | 0 | -5.699 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 01:17
WO-23003-110-A-1 | 0 | 44.332 | 222.256 | -1.25 | 0.341 | 0 | -1591.716 | 0 | 4.918 | 0.234 | 0 | -0.005 | 0 | 0 | 0.419 | 0 | 666.017 | 0 | 0 | 1.827 | 1.761 | 0 | 2.107 | 0.737 | 3.639 | 0 | 0.725 | 0 | -4.831 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 00:50
WO-23003-109-A-1 | 0 | 30.793 | 64.865 | -1.51 | 0.34 | 0 | -1755.039 | 0 | 1.175 | 0.294 | 0 | 0.117 | 0 | 0 | 0.497 | 0 | 137.434 | 0 | 0 | -0.263 | 0.53 | 0 | 0.506 | 0.719 | 0.575 | 0 | 0.712 | 0 | -5.335 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-22 00:23
WO-23003-108-A-1 | 0 | 66.078 | 142.55 | -1.305 | 0.357 | 0 | -2030.64 | 0 | 0 | 0.258 | 0.794 | -0.006 | 0 | 0 | 0.426 | 0 | 117.876 | 0 | 0 | 0.083 | 0 | 0 | 0.484 | 0.637 | 0.504 | 0 | 0.741 | 0 | -4.723 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 23:56
WO-23003-107-A-1 | 0 | 30.047 | 38.99 | -1.589 | 0.35 | 0 | -1440.237 | 0 | 0 | 0.52 | 0 | 0.048 | 0 | 0 | 0.271 | 0 | 141.587 | 0 | 0 | 0 | 0 | 0 | 0.464 | 0.805 | 0.668 | 0 | 0.742 | 0 | -5.388 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 23:29
WO-23003-106-A-1 | 0 | 43.438 | 42.492 | -1.36 | 0.367 | 0 | -1668.765 | 0.23 | 0 | 0.602 | 0.727 | 0.016 | 0 | 0 | 0.632 | 0 | 137.351 | 0 | 0 | 0 | 0 | 0 | 0.515 | 0.744 | 0.742 | 0 | 0.745 | 0 | -5.593 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 23:02
WO-23003-105-A-1 | 0 | 79.408 | 118.301 | -1.063 | 0.373 | 0 | -2176.952 | 0.4 | 0 | 0.267 | 0.781 | 0.119 | 0 | 0 | 0.427 | 0 | 128.03 | 0 | 0 | 1.54 | 0 | 0 | 0.695 | 0.707 | 0.819 | 0 | 0.776 | 0 | -5.072 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 22:35
WO-23003-104-A-1 | 0 | 28.431 | 42.646 | -1.607 | 0.387 | 0 | -1298.381 | 0 | 0 | 0.218 | 0.815 | 0.093 | 0 | 0 | 0.413 | 0 | 128.717 | 0 | 0 | 0.148 | 0 | 0 | 0.618 | 0.777 | 0.713 | 0 | 0.804 | 0 | -5.253 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 22:08
WO-23003-103-A-1 | 0 | 48.723 | 139.401 | -1.964 | 0.424 | 0 | -1609.503 | 0 | 0 | 0.241 | 0 | 0.047 | 0 | 0 | 0.5 | 0 | 66.78 | 0 | 0 | 1.829 | 0.975 | 0 | 0.525 | 0.761 | 0.702 | 0 | 0.833 | 0 | -6.185 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 21:41
WO-23003-102-A-1 | 0 | 69.052 | 172.024 | -0.869 | 0.541 | 0 | -1540.208 | 0.788 | 0 | 1.153 | 0 | 0.352 | 0 | 0 | 0.481 | 0 | 64.801 | 0 | 0 | 0.173 | 1.032 | 0 | 0.666 | 0.881 | 0.732 | 0 | 0.985 | 0 | -5.777 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 21:14
WO-23003-101-A-1 | 0 | 61.53 | 156.344 | -1.137 | 0.691 | 0 | -1834.692 | 0 | 0 | 0.578 | 0.928 | 0.365 | 0 | 0 | 0.599 | 0 | 28.256 | 0 | 0 | 0.093 | 0.912 | 0 | 0.649 | 0.911 | 0.618 | 0 | 1.162 | 0 | -5.518 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 20:47
AB-230117-02 LFSM-H 5 | 0 | 107.228 | 345.773 | 191.07 | 211.268 | 0 | -2016.365 | 178.008 | 163.798 | 202.628 | 151.084 | 165.247 | 0 | 0 | 193.47 | 182.099 | 223.297 | 0 | 0 | 175.928 | 176.535 | 0 | 186.755 | 189.923 | 182.066 | 0 | 183.253 | 0 | 466.75 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-21 19:52
AB-230117-02 CCC-H 6 | 0 | 289.127 | 371.455 | 391.326 | 421.057 | 0 | 443.96 | 352.766 | 298.434 | 329.233 | 297.736 | 300.62 | 0 | 0 | 349.783 | 304.261 | 341.563 | 0 | 0 | 337.342 | 339.3 | 0 | 341.007 | 285.514 | 333.08 | 0 | 315.445 | 0 | 615.557 | 0 | 0 | 5650.988 |  | 1 | Unknown | 2023-01-21 19:25
AB-230117-02 LFB-H 6 | 0 | 155.833 | 177.057 | 189.557 | 210.176 | 0 | 137.426 | 179.368 | 149.25 | 176.216 | 148.382 | 151.608 | 1199.01 | 0 | 177.686 | 158.298 | 176.054 | 0 | 0 | 172.683 | 175.354 | 0 | 176.742 | 143.941 | 173.631 | 0 | 159.314 | 0 | 331.989 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 18:58
AB-230117-02 LD 5 | 0 | 37.363 | 134.103 | -1.164 | 0 | 0 | -1974.742 | 0 | 1.186 | 0.249 | 0.756 | 0.096 | 0 | 0 | 0.293 | 0 | 88.093 | 0 | 0 | 0.066 | 0.77 | 0 | 0.438 | 0.74 | 0.645 | 0 | 0.725 | 0 | -4.489 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-21 18:31
LD-5-inj1 | 0 | 46.612 | 136.51 | -0.986 | 0.336 | 0 | -1800.431 | 0 | 0 | 0.062 | 0.744 | 0.089 | 0 | 0 | 0.216 | 0 | 82.9 | 0 | 0 | -0.175 | 0.802 | 0 | 0.477 | 0.657 | 0.657 | 0 | 0.713 | 0 | -4.724 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-21 18:04
AB-230117-02 LRB 11 | 0 | -17.567 | -3.443 | -1.285 | 0 | 0 | 7.149 | 0.345 | 0 | -0.049 | 0.727 | 0.017 | 0 | 0 | 0.401 | 0 | 0.264 | 0 | 0 | 0 | 0.756 | 0 | 0.267 | 0.643 | 0.55 | 0 | 0.701 | 0 | -5.676 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 17:37
WO-23003-100-A-1 | 0 | 27.513 | 11.115 | -2.307 | 0 | 0 | -1773.975 | 0.321 | 1.922 | 0.295 | 0.717 | 0.128 | 0 | 0 | 0.451 | 410.536 | 0.224 | 0 | 0 | 0.007 | 0 | 0 | 0.31 | 0.629 | 0.224 | 0 | 0.696 | 0 | -4.966 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 17:10
WO-23003-99-A-1 | 0 | 45.806 | 12.489 | -1.836 | 0 | 0 | -2297.103 | 0 | 0 | 0.255 | 0.792 | -0.013 | 0 | 0 | 0.277 | 447.416 | 0.186 | 0 | 0 | 0 | 0 | 0 | 0.309 | 0.614 | 0.262 | 0 | 0.699 | 0 | -4.181 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 16:43
WO-23003-98-A-1 | 0 | 24.011 | 12.584 | -1.408 | 0 | 0 | -1446.733 | 0 | 0 | 0.341 | 0.769 | 0.116 | 0 | 0 | 0.3 | 403.126 | 0.211 | 0 | 0 | 1.427 | 0 | 0 | 0.335 | 0.728 | 0.294 | 0 | 0.725 | 0 | -5.191 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 16:16
WO-23003-97-A-1 | 0 | 46.302 | 12.895 | -1.532 | 0 | 0 | -1567.275 | 0.489 | 0 | 0.432 | 0.753 | 0.188 | 0 | 0 | 0.34 | 382.008 | 0.572 | 0 | 0 | 0.216 | 0 | 0 | 0.355 | 0.646 | 0.314 | 0 | 0.694 | 0 | -4.759 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 15:49
WO-23003-96-A-1 | 0 | 26.128 | 12.667 | -1.004 | 0 | 0 | -1481.869 | 0 | 0 | 0.283 | 0.736 | 0.043 | 0 | 0 | 0.249 | 396.134 | 0.168 | 0 | 0 | 0 | 0 | 0 | 0.291 | 0.681 | 0.191 | 0 | 0.715 | 0 | -4.851 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 15:22
WO-23003-95-A-1 | 0 | 29.325 | 9.807 | -1.682 | 0.333 | 0 | -1637.525 | 0 | 0 | 0.307 | 0 | 0.007 | 0 | 0 | 0.263 | 396.948 | 0.195 | 0 | 0 | 0 | 0 | 0 | 0.271 | 0.678 | 0.24 | 0 | 0.734 | 0 | -5.363 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 14:54
WO-23003-94-A-1 | 0 | 31.35 | 11.508 | -0.332 | 0 | 0 | -1478.751 | 0 | 0 | 0.304 | 0 | 0.093 | 0 | 0 | 0.419 | 396.942 | 0.307 | 0 | 0 | 0 | 0 | 0 | 0.352 | 0.713 | 0.27 | 0 | 0.716 | 0 | -4.846 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 14:27
WO-23003-93-A-1 | 0 | 49.528 | 11.324 | -1.572 | 0 | 0 | -2004.192 | 0 | 2.816 | 0.127 | 0 | 0.097 | 0 | 0 | 0.4 | 377.744 | 0.324 | 0 | 0 | -0.196 | 0 | 0 | 0.391 | 0.67 | 0.277 | 0 | 0.705 | 0 | -4.939 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 14:00
WO-23003-92-A-1 | 0 | 18.741 | 6.111 | -2.068 | 0 | 0 | -1393.945 | 0 | 0 | 0.215 | 0.76 | 0.23 | 0 | 0 | 0.493 | 385.868 | 0.174 | 0 | 0 | 0 | 0 | 0 | 0.343 | 0.787 | 0.312 | 0 | 0.704 | 0 | -4.863 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 13:33
WO-23003-91-A-1 | 0 | 64.998 | 15.662 | -2.262 | 0.337 | 0 | -2158.633 | 0 | 0 | 0.461 | 0.749 | 0.114 | 0 | 0 | 0.446 | 385.432 | 0.408 | 0 | 0 | 0 | 0 | 0 | 0.317 | 0.656 | 0.38 | 0 | 0.708 | 0 | -4.358 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 13:06
AB-230117-02 LRB 10 | 0 | -16.448 | -2.021 | -1.941 | 0.34 | 0 | 161.946 | 0.079 | 1.395 | 0.014 | 0.766 | 0.228 | 0 | 0 | 0.357 | 0 | 0.321 | 0 | 0 | -0.318 | 0.848 | 0 | 0.264 | 0.653 | 0.571 | 0 | 0.757 | 0 | -5.396 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 12:39
LRB-10-inj1 | 0 | -16.798 | -2.485 | -2.314 | 0.34 | 0 | 96.671 | 0 | 1.09 | 0.021 | 0.754 | 0.056 | 0 | 0 | 0.419 | 0 | 0.283 | 0 | 0 | -0.337 | 0.79 | 0 | 0.261 | 0.669 | 0.515 | 0 | 0.733 | 0 | -5.676 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 12:12
WO-23003-90-A-1 | 0 | 66.885 | 13.107 | -1.876 | 0.341 | 0 | -1728.648 | 0 | 2.291 | 0.445 | 0.792 | 0.152 | 0 | 0 | 0.455 | 365.485 | 0.371 | 0 | 0 | -0.091 | 0 | 0 | 0.321 | 0.665 | 0.344 | 0 | 0.724 | 0 | -4.599 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 11:45
WO-23003-89-A-1 | 0 | 37.172 | 11.98 | -1.478 | 0.344 | 0 | -1555.855 | 0 | 1.046 | 0.322 | 0.752 | 0.075 | 0 | 0 | 0.244 | 415.919 | 0.192 | 0 | 0 | 0 | 0 | 0 | 0.291 | 0.721 | 0.284 | 0 | 0.744 | 0 | -4.289 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 11:18
WO-23003-88-A-1 | 0 | 28.868 | 14.826 | -1.622 | 0.346 | 0 | -1353.534 | 0 | 0 | 0.429 | 0 | 0.092 | 0 | 0 | 0.284 | 382.709 | 0.223 | 0 | 0 | -0.142 | 0 | 0 | 0.342 | 0.683 | 0.287 | 0 | 0.736 | 0 | -5.104 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 10:51
WO-23003-87-A-1 | 0 | 90.659 | 18.952 | -1.724 | 0.357 | 0 | -2512.26 | 0 | 1.572 | 0.25 | 0.749 | 0.201 | 0 | 0 | 0.414 | 406.54 | 0.219 | 0 | 0 | 0.359 | 0 | 0 | 0.368 | 0.735 | 0.359 | 0 | 0.705 | 0 | -4.907 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 10:24
WO-23003-86-A-1 | 0 | 52.311 | 17.014 | -1.452 | 0.369 | 0 | -2232.3 | 0.871 | 2.233 | 0.465 | 0.75 | 0.071 | 0 | 0 | 0.367 | 433.77 | 0.246 | 0 | 0 | 0.022 | 0 | 0 | 0.431 | 0.719 | 0.35 | 0 | 0.731 | 0 | -4.784 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 09:57
WO-23003-85-A-1 | 0 | 22.091 | 12.5 | -1.934 | 0.38 | 0 | -1405.98 | 0 | 0 | 0.353 | 0.74 | 0.158 | 0 | 0 | 0.456 | 419.227 | 0.326 | 0 | 0 | 0 | 0 | 0 | 0.395 | 0.739 | 0.31 | 0 | 0.75 | 0 | -4.307 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 09:30
WO-23003-84-A-1 | 0 | 31.744 | 10.608 | -1.128 | 0.392 | 0 | -1266.276 | 0 | 0 | 0.478 | 0.782 | 0.214 | 0 | 0 | 0.741 | 405.654 | 0.229 | 0 | 0 | 0.733 | 0 | 0 | 0.405 | 0.731 | 0.327 | 0 | 0.803 | 0 | -5.085 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 09:03
WO-23003-83-A-1 | 0 | 46.21 | 12.053 | -1.422 | 0.434 | 0 | -1406.442 | 0 | 1.468 | 0.54 | 0.804 | 0.184 | 0 | 0 | 0.45 | 311.52 | 0.398 | 0 | 0 | 0 | 0 | 0 | 0.424 | 0.927 | 0.375 | 0 | 0.845 | 0 | -5.075 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 08:36
WO-23003-82-A-1 | 0 | 93.384 | 33.548 | -0.169 | 0.642 | 0 | -1171.651 | 2.353 | 2.483 | 0.593 | 0.97 | 0.518 | 0 | 0 | 0.654 | 602.933 | 0.63 | 0 | 0 | 0 | 0 | 0 | 0.705 | 0.958 | 0.828 | 0 | 1.221 | 0 | -2.783 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 08:08
WO-23003-81-A-1 | 0 | 30.769 | 10.576 | -1.403 | 0.746 | 0 | -1414.582 | 0 | 1.055 | 1.117 | 0.984 | 0.375 | 0 | 0 | 0.575 | 319.278 | 0.574 | 0 | 0 | 0 | 0 | 0 | 0.62 | 1.164 | 0.541 | 0 | 1.121 | 0 | -4.163 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 07:41
AB-230117-02 LFSM-H 4 | 0 | 141.547 | 230.253 | 235.884 | 231.311 | 0 | -1461.066 | 200.054 | 174.965 | 262.399 | 164.339 | 176.253 | 0 | 0 | 208.366 | 189.256 | 213.294 | 0 | 0 | 191.747 | 184.951 | 0 | 204.634 | 216.789 | 197.378 | 0 | 199.573 | 0 | 389.584 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-21 06:47
AB-230117-02 CCC-H 5 | 0 | 311.228 | 363.733 | 402.453 | 435.261 | 0 | 894.413 | 375.935 | 301.325 | 364.192 | 313.327 | 323.051 | 1260.747 | 0 | 357.871 | 309.052 | 355.469 | 0 | 0 | 349.56 | 355.016 | 0 | 357.216 | 299.892 | 353.128 | 0 | 332.741 | 0 | 629.614 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 06:20
AB-230117-02 LFB-H 5 | 0 | 163.23 | 198.722 | 203.039 | 213.609 | 0 | 432.648 | 179.705 | 157 | 179.847 | 152.771 | 156.013 | 1216.497 | 0 | 182.108 | 158.258 | 180.317 | 0 | 0 | 177.279 | 179.093 | 0 | 181.423 | 157.635 | 179.201 | 0 | 168.236 | 0 | 345.389 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 05:53
AB-230117-02 LD 4 | 0 | 118.358 | 18.244 | -0.735 | 0.334 | 0 | -2842.877 | 1.281 | 1.608 | 0.633 | 0.794 | 0.374 | 0 | 0 | 0.482 | 0 | 3.87 | 0 | 0 | 0.465 | 0 | 0 | 0.314 | 0.795 | 0.502 | 0 | 0.712 | 0 | -1.788 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-21 05:26
LD4-inj1 | 0 | 70.143 | 15.537 | -0.94 | 0 | 0 | -2317.693 | 1.695 | 0 | 0.488 | 0.807 | 0.243 | 0 | 0 | 0.8 | 0 | 3.644 | 0 | 0 | 0.552 | 0 | 0 | 0.342 | 0.68 | 0.397 | 0 | 0.714 | 0 | -4.845 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-21 04:59
AB-230117-02 LRB 9 | 0 | -14.4 | -3.528 | -0.887 | 0 | 0 | 164.97 | 0 | 1.019 | 0.015 | 0.693 | 0.153 | 0 | 0 | 0.39 | 0 | 0.312 | 0 | 0 | -0.198 | 0.794 | 0 | 0.229 | 0.656 | 0.552 | 0 | 0.707 | 0 | -5.512 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 04:32
WO-23003-80-A-1 | 0 | 71.46 | 17.773 | -1.63 | 0 | 0 | -2164.04 | 0 | 2.611 | 0.199 | 0.781 | 0.069 | 0 | 0 | 0.358 | 354.309 | 0.25 | 0 | 0 | 1.633 | 0 | 0 | 0.326 | 0.642 | 0.352 | 0 | 0.685 | 0 | -4.458 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 04:05
WO-23003-79-A-1 | 0 | 50.524 | 12.127 | -1.481 | 0 | 0 | -1620.783 | 0 | 1.386 | 0.255 | 0 | 0.161 | 0 | 0 | 0.373 | 341.696 | 0.401 | 0 | 0 | 1.21 | 0 | 0 | 0.384 | 0.791 | 0.404 | 0 | 0.722 | 0 | -4.709 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 03:38
WO-23003-78-A-1 | 0 | 30.148 | 13.933 | -1.639 | 0 | 0 | -1435.327 | 0 | 1.06 | 0.288 | 0 | 0.047 | 0 | 0 | 0.318 | 343.62 | 0.283 | 0 | 0 | 0 | 0 | 0 | 0.301 | 0.741 | 0.336 | 0 | 0.717 | 0 | -4.589 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 03:11
WO-23003-77-A-1 | 0 | 34.735 | 9.843 | -1.371 | 0 | 0 | -1294.126 | 0.689 | 1.2 | 0.116 | 0.81 | 0.067 | 0 | 0 | 0.317 | 212.264 | 0.393 | 0 | 0 | 0.046 | 0 | 0 | 0.385 | 0.612 | 0.35 | 0 | 0.736 | 0 | -4.597 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 02:43
WO-23003-76-A-1 | 0 | 31.722 | 11.695 | -1.806 | 0 | 0 | -1538.305 | 0 | 1.534 | 0.457 | 0.781 | 0.24 | 0 | 0 | 0.587 | 201.051 | 1.751 | 0 | 0 | 0 | 0 | 0 | 0.35 | 0.724 | 0.32 | 0 | 0.7 | 0 | -4.458 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 02:16
WO-23003-75-A-1 | 0 | 65.003 | 10.994 | -1.445 | 0 | 0 | -1689.705 | 0 | 1.73 | 0.374 | 0.805 | 0.486 | 0 | 0 | 0.528 | 0 | 46.694 | 0 | 0 | 3.079 | 0 | 0 | 0.51 | 0.806 | 0.403 | 0 | 0.699 | 0 | -4.766 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 01:49
WO-23003-74-A-1 | 0 | 30.233 | 10.449 | -0.248 | 0 | 0 | -1296.292 | 0 | 0 | 0.448 | 0.765 | 0.3 | 0 | 0 | 0.382 | 0 | 33.419 | 0 | 0 | 0 | 0 | 0 | 0.357 | 0.722 | 0.311 | 0 | 0.713 | 0 | -5.195 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 01:22
WO-23003-73-A-1 | 0 | 50.056 | 14.758 | -0.628 | 0.339 | 0 | -1754.151 | 0 | 1.375 | 0.54 | 0 | 0.31 | 0 | 0 | 0.413 | 0 | 16.405 | 0 | 0 | 0.523 | 0 | 0 | 0.365 | 0.848 | 0.532 | 0 | 0.709 | 0 | -4.489 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 00:55
WO-23003-72-A-1 | 0 | 53.994 | 11.49 | -1.927 | 0.344 | 0 | -1730.007 | 0 | 1.154 | 0.445 | 0.795 | 0.446 | 0 | 0 | 0.468 | 0 | 19.333 | 0 | 0 | 0 | 0 | 0 | 0.315 | 0.869 | 0.515 | 0 | 0.715 | 0 | -4.144 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 00:28
WO-23003-71-A-1 | 0 | 40.276 | 11.754 | -1.05 | 0.339 | 0 | -1370.883 | 0.602 | 0 | 0.549 | 0 | 0.134 | 0 | 0 | 0.474 | 0 | 14.541 | 0 | 0 | 0 | 0 | 0 | 0.328 | 0.721 | 0.33 | 0 | 0.776 | 0 | -5.312 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-21 00:01
AB-230117-02 LRB 8 | 0 | -6.171 | -1.453 | -1.693 | 0.344 | 0 | 85.351 | 0 | 0.941 | 0.211 | 0.791 | 0.287 | 0 | 0 | 0.334 | 0 | 0.257 | 0 | 0 | -0.341 | 0 | 0 | 0.277 | 0.805 | 0.409 | 0 | 0.74 | 0 | -4.928 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 23:34
LRB-8-inj1 | 0 | -10.631 | -3.094 | -1.403 | 0.344 | 0 | 115.431 | 0.346 | 1.025 | 0.204 | 0.693 | 0.344 | 0 | 0 | 0.49 | 0 | 0.783 | 0 | 0 | 0 | 0 | 0 | 0.28 | 0.796 | 0.355 | 0 | 0.752 | 0 | -5.019 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 23:07
WO-23003-70-A-1 | 0 | 59.943 | 12.318 | -0.121 | 0.344 | 0 | -1951.131 | 0 | 2.141 | 0.198 | 0 | 0.38 | 0 | 0 | 0.503 | 0 | 14.244 | 0 | 0 | 0.081 | 0 | 0 | 0.408 | 0.767 | 0.365 | 0 | 0.755 | 0 | -4.06 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 22:40
WO-23003-69-A-1 | 0 | 36.985 | 12.159 | -0.985 | 0.348 | 0 | -1775.526 | 0 | 1.175 | 0.533 | 0.763 | 0.297 | 0 | 0 | 0.532 | 0 | 9.563 | 0 | 0 | 0.446 | 0 | 0 | 0.312 | 0.873 | 0.524 | 0 | 0.751 | 0 | -4.755 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 22:13
WO-23003-68-A-1 | 0 | 38.538 | 11.795 | -0.159 | 0.352 | 0 | -1777.291 | 0 | 1.443 | 0.515 | 0.812 | 0.434 | 0 | 0 | 0.51 | 0 | 10.214 | 0 | 0 | 2.645 | 0 | 0 | 0.389 | 0.793 | 0.382 | 0 | 0.758 | 0 | -4.565 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 21:46
WO-23003-67-A-1 | 0 | 36.888 | 10.917 | -1.54 | 0.361 | 0 | -1821.798 | 0 | 1.384 | 0.543 | 0.804 | 0.308 | 0 | 0 | 0.484 | 0 | 5.711 | 0 | 0 | 0.304 | 0 | 0 | 0.338 | 0.721 | 0.362 | 0 | 0.788 | 0 | -4.29 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 21:19
WO-23003-66-A-1 | 0 | 35.802 | 10.592 | -1.263 | 0.365 | 0 | -1633.299 | 0 | 1.191 | 0.079 | 0 | 0.383 | 0 | 0 | 0.378 | 0 | 5.647 | 0 | 0 | -0.001 | 0 | 0 | 0.344 | 0.851 | 0.338 | 0 | 0.818 | 0 | -3.613 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 20:51
WO-23003-65-A-1 | 0 | 76.978 | 12.387 | -0.688 | 0.38 | 0 | -2378.334 | 0 | 1.691 | 0.572 | 0 | 0.359 | 0 | 0 | 0.693 | 0 | 8.885 | 0 | 0 | 0.579 | 0.25 | 0 | 0.34 | 0.96 | 0.317 | 0 | 0.787 | 0 | -3.606 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 20:24
WO-23003-64-A-1 | 0 | 82.787 | 13.315 | -0.483 | 0.4 | 0 | -2433.644 | 0 | 1.537 | 0.857 | 0.772 | 0.402 | 0 | 0 | 0.669 | 0 | 12.026 | 0 | 0 | 0.586 | 0 | 0 | 0.475 | 0.959 | 0.38 | 0 | 0.878 | 0 | -3.782 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 19:57
WO-23003-63-A-1 | 0 | 47.787 | 10.717 | -0.877 | 0.43 | 0 | -2266.741 | 0.45 | 1.341 | 0.673 | 0.833 | 0.443 | 0 | 0 | 0.672 | 0 | 9.272 | 0 | 0 | 0.739 | 0 | 0 | 0.432 | 0.993 | 0.39 | 0 | 0.873 | 0 | -4.976 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 19:30
WO-23003-62-A-1 | 0 | 35.68 | 11.608 | -0.986 | 0.489 | 0 | -2156.762 | 0 | 2.118 | 0.63 | 0.924 | 0.508 | 0 | 0 | 0.591 | 0 | 5.986 | 0 | 0 | 0.075 | 0 | 0 | 0.45 | 0.757 | 0.494 | 0 | 0.9 | 0 | -3.534 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 19:03
WO-23003-61-A-1 | 0 | 53.245 | 14.969 | 0.409 | 0.744 | 0 | -2454.472 | 0 | 1.448 | 1.301 | 0 | 0.735 | 0 | 0 | 1.275 | 0 | 4.3 | 0 | 0 | 1.168 | 0 | 0 | 0.672 | 1.339 | 0.576 | 0 | 1.23 | 0 | -2.834 | 10.286 | 0 | 0 |  | 1 | Unknown | 2023-01-20 18:36
AB-230117-02 LFSM-H 3 | 0 | 131.12 | 222.169 | 223.168 | 217.1 | 0 | -2625.739 | 193.406 | 157.155 | 288.344 | 163.49 | 169.947 | 0 | 0 | 205.752 | 334.468 | 203.756 | 0 | 0 | 187.492 | 190.483 | 0 | 198.421 | 216.121 | 191.067 | 0 | 195.334 | 0 | 578.701 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-20 17:42
AB-230117-02 CCC-H 4 | 0 | 394.354 | 374.225 | 425.492 | 444.887 | 0 | -782.4 | 394.235 | 330.463 | 406.982 | 340.248 | 335.865 | 0 | 0 | 391.26 | 333.561 | 378.791 | 0 | 0 | 372.36 | 375.413 | 0 | 378.57 | 356.247 | 375.67 | 0 | 346.062 | 0 | 1138.264 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 17:15
AB-230117-02 LFB-H 4 | 0 | 128.037 | 202.802 | 228.56 | 218.871 | 0 | -919.956 | 199.569 | 176.623 | 191.501 | 161.298 | 165.666 | 1149.754 | 0 | 197.291 | 160.162 | 192.196 | 0 | 0 | 190.159 | 187.965 | 0 | 195.326 | 175.359 | 186.073 | 0 | 173.955 | 0 | 643.923 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 16:48
AB-230117-02 LD 3 | 0 | 68.015 | 14.791 | -1.662 | 0.342 | 0 | -2643.539 | 0 | 1.333 | 0.803 | 0.799 | 0.394 | 0 | 0 | 0.269 | 140.88 | 1.737 | 0 | 0 | 0.617 | 0 | 0 | 0.41 | 0.646 | 0.387 | 0 | 0.724 | 0 | -1.95 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-20 16:21
AB-230117-02 LRB 7 | 0 | 552.849 | 367.278 | 6.887 | 0.936 | 0 | 1001.123 | 0 | 13.867 | 8.257 | 0 | 16.487 | 0 | 0 | 2.82 | 0 | 77.188 | 0 | 0 | 0 | 0 | 0 | 4.824 | 2.663 | 14.3 | 0 | 2.626 | 0 | 212.947 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 15:54
LRB7-inj1 | 0 | 4.902 | -2.659 | -1.134 | 0.353 | 0 | -898.494 | 0 | 0.998 | 0.379 | 0.794 | 0.502 | 0 | 0 | 0.433 | 0 | 27.187 | 0 | 0 | -0.106 | 0 | 0 | 0.308 | 0.795 | 0.326 | 0 | 0.768 | 0 | -4.049 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 15:27
WO-23003-60-A-1 | 0 | 226.438 | 12.429 | 3.175 | 0.383 | 0 | -3516.853 | 0 | 2.425 | 0.686 | 0.818 | 0.342 | 0 | 0 | 0.794 | 0 | 4256.541 | 0 | 0 | 3.233 | 0 | 0 | 2.954 | 1.342 | 1.794 | 0 | 0.755 | 0 | -2.292 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 15:00
WO-23003-59-A-1 | 0 | 92.981 | 11.907 | -1.497 | 0.421 | 0 | -2262.582 | 0 | 0 | 0.208 | 0.848 | 0.383 | 0 | 0 | 0.69 | 0 | 660.136 | 0 | 0 | 0.057 | 0 | 0 | 0.846 | 0.977 | 0.656 | 0 | 0.951 | 0 | -3.04 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 14:32
WO-23003-58-A-1 | 0 | 108.162 | 14.295 | 1.014 | 0.5 | 0 | -2303.618 | 0 | 0 | 1.006 | 0.922 | 0.735 | 0 | 0 | 0.697 | 0 | 60.711 | 0 | 0 | 0.903 | 0 | 0 | 1.322 | 1.186 | 0.547 | 0 | 0.931 | 0 | -3.172 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 14:05
WO-23003-138-A-1 | 0 | -12.818 | 2.295 | -2.043 | 0 | 0 | -1442.374 | 0 | 41.842 | 0.238 | 0.961 | 0.02 | 0 | 0 | 0.288 | 0 | 0.368 | 0 | 0 | 0.174 | 0 | 0 | 0.284 | 0.651 | 0.209 | 0 | 0.735 | 0 | -4.874 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-23 19:10
WO-23003-137-A-1 | 0 | -9.688 | 3.611 | -1.534 | 0 | 0 | -824.309 | 0 | 30.824 | 0.201 | 0 | -0.072 | 0 | 0 | 0.283 | 0 | 0.386 | 0 | 0 | 0 | 0 | 0 | 0.258 | 0.628 | 0.319 | 0 | 0.714 | 0 | -5.245 | 1.639 | 0 | 0 |  | 1 | Unknown | 2023-01-23 18:43
WO-23003-136-A-1 | 0 | 52.844 | 13.284 | -1.427 | 0 | 0 | -3093.314 | 1.066 | 36.314 | 0.427 | 0.689 | -0.035 | 0 | 0 | 0.293 | 0 | 0.437 | 0 | 0 | 0 | 0 | 0 | 0.251 | 0.656 | 0.251 | 0 | 0.697 | 0 | -4.117 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-23 18:16
WO-23003-135-A-1 | 0 | -1.856 | 9.115 | -1.979 | 0 | 0 | -854.404 | 0 | 34.102 | 0.181 | 0.754 | 0.159 | 0 | 0 | 0.201 | 0 | 0.443 | 0 | 0 | 0.786 | 0 | 0 | 0.333 | 0.642 | 0.412 | 0 | 0.707 | 0 | -5.041 | 3.528 | 0 | 0 |  | 1 | Unknown | 2023-01-23 17:49
WO-23003-134-A-1 | 0 | -1.97 | 9.035 | -2.13 | 0 | 0 | -1117.89 | 0.198 | 38.926 | 0.282 | 0.762 | -0.051 | 0 | 0 | 0.318 | 0 | 0.414 | 0 | 0 | -0.151 | 0 | 0 | 0.261 | 0.73 | 0.371 | 0 | 0.703 | 0 | -4.286 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-23 17:21
WO-23003-133-A-1 | 0 | -55.247 | 14.845 | -1.049 | 0 | 0 | -1525.586 | 1.203 | 33.78 | 0.427 | 0 | 0.106 | 0 | 0 | 0.218 | 0 | 0.523 | 0 | 0 | 0 | 0 | 0 | 0.295 | 0.631 | 0.367 | 0 | 0.723 | 0 | -6.271 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-23 16:54
WO-23003-132-A-1 | 0 | 29.492 | 12.305 | -1.795 | 0 | 0 | -2502.402 | 0 | 16.223 | 0.256 | 0 | -0.029 | 0 | 0 | 0.233 | 0 | 0.77 | 0 | 0 | 0 | 0 | 0 | 1.47 | 0.715 | 1.113 | 0 | 0.71 | 0 | -4.262 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-23 16:27
WO-23003-131-A-1 | 0 | 0 | 5.382 | -1.716 | 0 | 0 | -1073.406 | 0.329 | 14.312 | 0.425 | 0 | -0.101 | 0 | 0 | 0.128 | 0 | 0.748 | 0 | 0 | -0.044 | 0 | 0 | 0.306 | 0.666 | 0.366 | 0 | 0.697 | 0 | -5.575 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-23 16:00
AB-230117-02 LRB 14 | 0 | -38.979 | -2.48 | -2.26 | 0 | 0 | 847.72 | 0 | 0 | -0.009 | 0.756 | -0.011 | 0 | 0 | 0.294 | 0 | 0.664 | 0 | 0 | -0.311 | 0.844 | 0 | 0.274 | 0.639 | 0.577 | 0 | 0.713 | 0 | -5.985 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-23 15:33
LRB-14-inj1 | 0 | -40.573 | -3.646 | -1.938 | 0 | 0 | 1038.216 | 0 | 1.102 | 0.067 | 0.769 | -0.07 | 0 | 0 | 0.434 | 0 | 0.68 | 0 | 0 | -0.028 | 0.799 | 0 | 0.247 | 0.672 | 0.486 | 0 | 0.733 | 0 | -5.863 | 0.733 | 0 | 0 |  | 1 | Unknown | 2023-01-23 15:06
WO-23003-130-A-1 | 0 | 17.822 | 6.711 | -1.729 | 0 | 0 | -2282.252 | 0.482 | 13.469 | 0.118 | 0 | -0.025 | 0 | 0 | 0.414 | 0 | 0.778 | 0 | 0 | 0.613 | 0 | 0 | 0.296 | 0.69 | 0.481 | 0 | 0.696 | 0 | -3.997 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-23 14:39
WO-23003-129-A-1 | 0 | 10.44 | 16.63 | -1.245 | 0 | 0 | -1476.605 | 0 | 12.556 | 0.388 | 0.771 | 0.236 | 0 | 0 | 0.136 | 0 | 1.157 | 0 | 0 | 0.115 | 0 | 0 | 0.345 | 0.697 | 0.451 | 0 | 0.72 | 0 | -6.01 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-23 14:12
WO-23003-128-A-1 | 0 | 5.716 | 22.32 | -1.218 | 0 | 0 | -1838.63 | 0.672 | 9.095 | 0.61 | 0.774 | 0.093 | 0 | 0 | 0.346 | 0 | 1.569 | 0 | 0 | 0.138 | 0 | 0 | 0.347 | 0.664 | 0.583 | 0 | 0.734 | 0 | -6.272 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-23 13:45
WO-23003-127-A-1 | 0 | 1.443 | 19.169 | -1.18 | 0 | 0 | -1053.048 | 0 | 9.456 | 0.199 | 0.839 | 0.029 | 0 | 0 | 0.171 | 0 | 1.887 | 0 | 0 | 0 | 0 | 0 | 0.375 | 0.688 | 0.451 | 0 | 0.699 | 0 | -4.841 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-23 13:18
WO-23003-126-A-1 | 0 | 27.436 | 29.326 | -0.999 | 0 | 0 | -2192.888 | 0 | 9.643 | 0.757 | 0.809 | 0.267 | 0 | 0 | 0.331 | 0 | 4.716 | 0 | 0 | 0 | 0 | 0 | 0.412 | 0.73 | 0 | 0 | 0.778 | 0 | -6.01 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-23 12:51
WO-23003-125-A-1 | 0 | 14.694 | 7.22 | -1.089 | 0.334 | 0 | -2482.221 | 0 | 11.776 | 0.391 | 0 | 0.089 | 0 | 0 | 0.223 | 0 | 46.972 | 0 | 0 | 0 | 0 | 0 | 0.413 | 0.67 | 0.512 | 0 | 0.723 | 0 | -3.829 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-23 12:24
WO-23003-121-A-1B | 0 | -18.215 | 1.726 | 0.68 | 0.341 | 0 | -245.723 | 0.285 | 0 | 0.086 | 0 | -0.034 | 0 | 0 | 0.515 | 0 | 2538.308 | 0 | 0 | 2.383 | 1.486 | 0 | 7.928 | 0.987 | 8.541 | 0 | 0.72 | 0 | -4.768 | 0 | 0 | 0 |  | 4 | Unknown | 2023-01-23 11:56
WO-23003-111-A-1B | 0 | 9.45 | 65.578 | -0.942 | 0.346 | 0 | -2215.917 | 0 | 4.439 | 0.39 | 0.805 | 0.434 | 0 | 0 | 0.32 | 0 | 738.645 | 0 | 0 | 1.983 | 1.749 | 0 | 2.374 | 0.679 | 4.534 | 0 | 0.761 | 0 | -3.773 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-23 11:29
AB-230117-02 LRB 13 20230123105625 | 0 | -66.085 | -2.845 | -1.953 | 0.358 | 0 | 991.182 | 0.169 | 0 | 0.139 | 0.82 | -0.001 | 0 | 0 | 0.342 | 0 | 3.33 | 0 | 0 | -0.178 | 0.805 | 0 | 0.281 | 0.728 | 0.61 | 0 | 0.713 | 0 | -5.692 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-23 11:02
AB-230117-02 LRB 12 20230123102918 | 0 | 3916.074 | 1541.758 | 55.619 | 4.726 | 0 | 8701.812 | 0 | 0 | 12.919 | 10.094 | 58.764 | 0 | 0 | 4.153 | 0 | 943.114 | 0 | 0 | 37.615 | 0 | 0 | 26.569 | 22.213 | 49.489 | 0 | 7.11 | 0 | 145.848 | 0 | 0 | 0 | C | 1 | Unknown | 2023-01-23 10:35
AB-230117-02 QCS 2 | 0 | 194.759 | 241.344 | 253.37 | 299.949 | 0 | 955.974 | 224.431 | 202.348 | 231.625 | 192.851 | 217.642 | 1183.236 | 0 | 25.094 | 194.121 | 220.165 | 0 | 0 | 217.546 | 212.761 | 0 | 232.69 | 24.852 | 218.131 | 0 | 22.632 | 0 | 390.264 | 0 | 0 | 4510.311 |  | 1 | Unknown | 2023-01-29 10:06
AB-230117-02 CCC-H 12 | 0 | 344.411 | 399.094 | 436.054 | 553.34 | 0 | 1331.427 | 391.777 | 353.417 | 386.979 | 327.791 | 783.523 | 0 | 0 | 371.954 | 343.096 | 372.448 | 0 | 0 | 379.928 | 383.581 | 0 | 373.921 | 316.256 | 375.367 | 0 | 342.557 | 0 | 626.903 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 09:39
WO-23003-212-A-1 | 0 | -43.519 | 0 | -1.551 | 0.387 | 0 | 913.895 | 0.224 | 2.35 | 144.855 | 5.409 | 6474.482 | 0 | 0 | 0.304 | 0 | 0.284 | 0 | 0 | -0.046 | 0.803 | 0 | 0 | 9.022 | 1.148 | 0 | 0.702 | 0 | -5.8 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 09:12
AB-230117-02 LRB 22 | 0 | -42.936 | -5.511 | -1.769 | 0.343 | 0 | 819.687 | 0.161 | 0 | 9.133 | 1.076 | 429.568 | 0 | 0 | 0.252 | 0 | 0.267 | 0 | 0 | -0.274 | 0.756 | 0 | 0.186 | 1.27 | 0.574 | 0 | 0.733 | 0 | -5.831 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 08:45
WO-23003-211-A-1 | 0 | -45.103 | 0 | -1.405 | 0.394 | 0 | 884.261 | 0 | 3.07 | 151.883 | 5.436 | 6424.609 | 0 | 0 | 0.337 | 0 | 0.265 | 0 | 0 | -0.043 | 0.916 | 0 | 0 | 9.216 | 1.074 | 0 | 0.764 | 0 | -6.139 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 08:18
WO-23003-210-A-1 | 0 | -41.937 | 0 | -1.724 | 0.407 | 0 | 584.76 | 0.152 | 3.382 | 204.552 | 8.704 | 7113.022 | 0 | 0 | 0.28 | 0 | 0.289 | 0 | 0 | 0.026 | 0.762 | 0 | 0 | 9.065 | 1.497 | 0 | 0.717 | 0 | -6.04 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 07:50
WO-23003-209-A-1 | 0 | -31.4 | 81.281 | -1.319 | 0.455 | 0 | -127.386 | 0.264 | 95.224 | 1176.751 | 166.268 | 14233.119 | 0 | 0 | 0.324 | 0 | 0.31 | 0 | 0 | 8.603 | 4.472 | 0 | 0 | 3.664 | 25.391 | 0 | 0.761 | 0 | -2.241 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 07:23
WO-23003-208-A-1 | 0 | -29.157 | 115.362 | -1.529 | 0.48 | 0 | -790.029 | 0.392 | 88.49 | 1223.778 | 163.904 | 13164.028 | 0 | 0 | 0.322 | 0 | 0.26 | 0 | 0 | 8.83 | 4.923 | 0 | 0 | 3.549 | 24.868 | 0 | 0.725 | 0 | -2.037 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 06:56
WO-23003-207-A-1 | 0 | -26.48 | 177.239 | -0.32 | 0.466 | 0 | -818.512 | 0 | 96.687 | 1200.616 | 165.874 | 13668.116 | 0 | 0 | 0.421 | 0 | 0.258 | 0 | 0 | 8.724 | 4.559 | 0 | 0 | 3.564 | 26.636 | 0 | 0.792 | 0 | -3.037 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 06:29
WO-23003-206-A-1 | 0 | -31.701 | 162.26 | -1.684 | 0.596 | 0 | -561.891 | 0.479 | 95.979 | 1412.768 | 192.84 | 19784.952 | 0 | 0 | 0.398 | 0 | 0.308 | 0 | 0 | 10.752 | 5.258 | 0 | 0 | 4.966 | 35.231 | 0 | 0.796 | 0 | -2.494 | 2.752 | 0 | 0 |  | 1 | Unknown | 2023-01-29 06:02
WO-23003-204-A-1 | 0 | 0.19 | 553.285 | -0.826 | 0.663 | 0 | -2157.675 | 0.521 | 208.811 | 2148.827 | 369.784 | 18254.898 | 0 | 0 | 0.196 | 0 | 0.31 | 0 | 0 | 10.348 | 5.247 | 0 | 0 | 4.485 | 29.505 | 0 | 0.768 | 0 | 4.524 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 05:35
WO-23003-203-A-1 | 0 | -0.832 | 200.704 | -1.725 | 0.876 | 0 | -2402.212 | 1.134 | 206.316 | 2263.57 | 359.206 | 20466.741 | 0 | 0 | 0.323 | 0 | 0.358 | 0 | 0 | 13.398 | 6.695 | 0 | 0 | 4.759 | 39.631 | 0 | 1.242 | 0 | 3.743 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 05:08
WO-23003-202-A-1 | 0 | -56.939 | 31.723 | -1.523 | 0.899 | 0 | -1431.887 | 0 | 150.15 | 990.295 | 127.028 | 6345.164 | 0 | 0 | 0.356 | 0 | 0.397 | 0 | 0 | 5.912 | 4.087 | 0 | 0 | 2.23 | 11.115 | 0 | 0.857 | 0 | -0.673 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 04:41
WO-23003-201-A-1 | 0 | -7.287 | 25.273 | -1.857 | 1.739 | 0 | -1705.164 | 1.429 | 152.553 | 1040.549 | 123.987 | 6600.341 | 0 | 0 | 0.544 | 0 | 0.604 | 0 | 0 | 6.13 | 3.72 | 0 | 0.809 | 2.858 | 11.225 | 0 | 1.249 | 0 | -1.174 | 7.232 | 0 | 0 |  | 1 | Unknown | 2023-01-29 04:14
AB-230117-02 LFSM-H 10 | 0 | 128.719 | 226.564 | 213.162 | 651.363 | 0 | -1322.557 | 183.533 | 306.812 | 1106.092 | 280.201 | 6432.403 | 0 | 0 | 190.036 | 194.12 | 194.699 | 0 | 0 | 188.5 | 188.607 | 0 | 192.309 | 203.781 | 190.703 | 0 | 216.072 | 0 | 328.134 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-29 03:20
AB-230117-02 CCC-H 11 | 0 | 300.023 | 395.726 | 396.305 | 487.898 | 0 | 1568.535 | 366.619 | 334.538 | 351.634 | 312.89 | 331.453 | 0 | 0 | 343.062 | 309.141 | 360.917 | 0 | 0 | 350.017 | 355.509 | 0 | 358.493 | 284.364 | 355.982 | 0 | 329.659 | 0 | 530.876 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 02:53
AB-230117-02 LFB-H 11 | 0 | 152.333 | 204.71 | 204.191 | 257.431 | 0 | 1204.746 | 195.443 | 176.561 | 227.636 | 164.505 | 441.806 | 0 | 0 | 185.518 | 163.487 | 186.414 | 0 | 0 | 184.715 | 186.635 | 0 | 185.609 | 144.059 | 183.897 | 0 | 178.769 | 0 | 332.108 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 02:26
AB-230117-02 LD 10 | 0 | 42.941 | 32.975 | 0.037 | 0 | 0 | -3038.081 | 1.057 | 133.374 | 1325.043 | 118.336 | 6584.537 | 0 | 0 | 0.25 | 0 | 0.273 | 0 | 0 | 6.23 | 3.438 | 0 | 0 | 2.272 | 10.794 | 0 | 0.839 | 0 | 1.278 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-29 01:58
AB-230117-02 LRB 21 | 0 | -41.454 | -0.196 | -0.996 | 0 | 0 | 223.459 | 0.064 | 1.048 | 2.381 | 0.868 | 8.959 | 1147.569 | 0 | 0.233 | 0 | 0.325 | 0 | 0 | -0.166 | 0.722 | 0 | 0.237 | 0.645 | 0.609 | 0 | 0.703 | 0 | -5.54 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 01:31
LRB-21-inj1 | 0 | -39.945 | -4.307 | -0.301 | 0 | 0 | 680.04 | 0 | 0 | 28.612 | 2.409 | 238.564 | 0 | 0 | 0.154 | 0 | 0.291 | 0 | 0 | -0.279 | 0.929 | 0 | 0.233 | 0.701 | 0.653 | 0 | 0.695 | 0 | -5.831 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 01:04
WO-23003-200-A-1 | 0 | -21.673 | 11.614 | -1.999 | 0.354 | 0 | -1523.371 | 0 | 143.538 | 958.021 | 118.094 | 6059.278 | 0 | 0 | 0.212 | 0 | 0.206 | 0 | 0 | 7.358 | 4.582 | 0 | 0 | 1.891 | 13.529 | 0 | 0.68 | 0 | -0.853 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 00:37
WO-23003-199-A-1 | 0 | -18.023 | 3.719 | -1.723 | 0 | 0 | -1266.07 | 0.607 | 162.514 | 648.524 | 94.187 | 2990.294 | 0 | 0 | 0.102 | 0 | 0.248 | 0 | 0 | 4.615 | 3.086 | 0 | 0.239 | 1.29 | 6.684 | 0 | 0.714 | 0 | -1.323 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-29 00:10
WO-23003-198-A-1 | 0 | 12.293 | 28.004 | -1.55 | 0 | 0 | -2565.66 | 0.694 | 178.451 | 1475.424 | 193.513 | 6757.926 | 0 | 0 | 0.114 | 0 | 0.238 | 0 | 0 | 7.529 | 4.452 | 0 | 0.435 | 2.078 | 14.717 | 0 | 0.694 | 0 | 2.569 | 0 | 0 | 7.778 |  | 1 | Unknown | 2023-01-28 23:43
WO-23003-197-A-1 | 0 | -20.849 | 41.121 | -1.054 | 0 | 0 | -1826.768 | 1.222 | 176.471 | 1247.438 | 200.815 | 6886.659 | 0 | 0 | 0.229 | 0 | 0.24 | 0 | 0 | 6.997 | 3.823 | 0 | 0 | 1.888 | 13.769 | 0 | 0.683 | 0 | -0.26 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 23:16
WO-23003-196-A-1 | 0 | -1.925 | 12.065 | -1.276 | 0.337 | 0 | -2415.74 | 1.277 | 43.583 | 97.719 | 9.73 | 484.625 | 0 | 0 | 0.148 | 0 | 0.188 | 0 | 0 | 2.449 | 1.056 | 0 | 0.256 | 0.799 | 0.806 | 0 | 0.705 | 0 | -4.25 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 22:49
WO-23003-195-A-1 | 0 | -11.084 | 6.283 | -1.489 | 0 | 0 | -1778.528 | 0.301 | 40.614 | 91.362 | 9.416 | 473.503 | 0 | 0 | 0.194 | 0 | 0.201 | 0 | 0 | 0.492 | 1.117 | 0 | 0 | 0.76 | 0.858 | 0 | 0.722 | 0 | -5.232 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 22:22
WO-23003-194-A-1 | 0 | -12.964 | 4.344 | -1.359 | 0.335 | 0 | -1508.543 | 1.246 | 48.031 | 92.375 | 9.528 | 504.393 | 0 | 0 | 0.092 | 0 | 0.209 | 0 | 0 | 0.885 | 0.838 | 0 | 0.277 | 0.784 | 0.905 | 0 | 0.697 | 0 | -5.635 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 21:55
WO-23003-193-A-1 | 0 | -10.094 | 4.523 | -1.824 | 0 | 0 | -1802.627 | 0.206 | 36.046 | 68.422 | 7.354 | 373.256 | 0 | 0 | 0.227 | 0 | 0.163 | 0 | 0 | 0.623 | 0.838 | 0 | 0.258 | 0.791 | 0.633 | 0 | 0.708 | 0 | -5.368 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 21:28
WO-23003-192-A-1 | 0 | -13.081 | 5.086 | -1.145 | 0 | 0 | -1762.546 | 0 | 81.655 | 86.017 | 11.749 | 359.261 | 0 | 0 | 0.077 | 0 | 0.175 | 0 | 0 | 0.9 | 1.124 | 0 | 0.327 | 0.755 | 0.889 | 0 | 0.738 | 0 | -4.848 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 21:01
WO-23003-191-A-1 | 0 | 21.416 | 11.863 | -1.616 | 0 | 0 | -2832.986 | 1.379 | 77.507 | 103.961 | 12.512 | 345.638 | 0 | 0 | 0.146 | 0 | 0.168 | 0 | 0 | 2.901 | 1.531 | 0 | 0.292 | 0.702 | 0.847 | 0 | 0.706 | 0 | -4.537 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 20:33
AB-230117-02 LRB 20 | 0 | -41.142 | -4.56 | -2.104 | 0.345 | 0 | 493.29 | 0.168 | 0 | 0.11 | 0.761 | 0.331 | 0 | 0 | 0.351 | 0 | 0.282 | 0 | 0 | -0.093 | 0.841 | 0 | 0.242 | 0.631 | 0.497 | 0 | 0.709 | 0 | -5.868 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 20:06
LRB-20-inj1 | 0 | -44.736 | -2.555 | -2.075 | 0.343 | 0 | 476.427 | 0 | 0 | 1.686 | 0.972 | 12.246 | 0 | 0 | 0.284 | 0 | 0.307 | 0 | 0 | -0.216 | 0.763 | 0 | 0.244 | 0.671 | 0.521 | 0 | 0.746 | 0 | -5.845 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 19:39
WO-23003-190-A-1 | 0 | 0.095 | 8.321 | -1.484 | 0.356 | 0 | -2279.318 | 0 | 86.45 | 102.394 | 13.315 | 377.598 | 0 | 0 | 0.212 | 0 | 0.385 | 0 | 0 | 0.705 | 1.55 | 0 | 0.329 | 0.8 | 1.102 | 0 | 0.733 | 0 | -4.737 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 19:12
WO-23003-189-A-1 | 0 | -13.482 | 5.643 | -2.237 | 0.356 | 0 | -1394.553 | 0.778 | 91.382 | 114.655 | 16.124 | 457.157 | 0 | 0 | 0.144 | 0 | 0.38 | 0 | 0 | 2.061 | 1.662 | 0 | 0.474 | 0.865 | 0.995 | 0 | 0.736 | 0 | -2.879 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 18:45
WO-23003-188-A-1 | 0 | -12.861 | 7.688 | -1.516 | 0.35 | 0 | -1705.624 | 0 | 6.698 | 8.309 | 1.423 | 61.169 | 0 | 0 | 0.211 | 0 | 0.203 | 0 | 0 | 1.083 | 0 | 0 | 0.293 | 0.645 | 0.342 | 0 | 0.737 | 0 | -6.02 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 18:18
WO-23003-187-A-1 | 0 | -20.249 | 6.438 | -1.089 | 0.35 | 0 | -1707.246 | 0 | 4.362 | 7.271 | 1.404 | 52.99 | 0 | 0 | 0.127 | 0 | 0.23 | 0 | 0 | -0.171 | 0 | 0 | 0.235 | 0.857 | 0.27 | 0 | 0.752 | 0 | -4.997 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 17:51
WO-23003-186-A-1 | 0 | -12.638 | 5.887 | -2.198 | 0.381 | 0 | -1683.477 | 0.239 | 5.009 | 7.175 | 1.147 | 53.928 | 0 | 0 | 0.193 | 0 | 0.224 | 0 | 0 | -0.237 | 0 | 0 | 0.258 | 0.724 | 0.37 | 0 | 0.755 | 0 | -4.55 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 17:24
WO-23003-185-A-1 | 0 | -13.235 | 5.964 | -1.515 | 0.377 | 0 | -1633.735 | 0.51 | 12.554 | 6.019 | 1.402 | 36.731 | 0 | 0 | 0.324 | 0 | 0.241 | 0 | 0 | 0 | 0.431 | 0 | 0.368 | 0.754 | 0.377 | 0 | 0.775 | 0 | -4.406 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 16:57
WO-23003-184-A-1 | 0 | -5.964 | 5.151 | -1.562 | 0.391 | 0 | -1776.215 | 0.63 | 13.196 | 6.464 | 1.546 | 41.182 | 0 | 0 | 0.277 | 0 | 0.249 | 0 | 0 | 0 | 0 | 0 | 0.302 | 0.692 | 0.413 | 0 | 0.84 | 0 | -4.461 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 16:30
WO-23003-183-A-1 | 0 | -14.563 | 7.857 | -1.479 | 0.429 | 0 | -1486.633 | 0.59 | 15.276 | 7.612 | 1.663 | 43.91 | 0 | 0 | 0.144 | 0 | 0.325 | 0 | 0 | 0 | 0 | 0 | 0.295 | 0.943 | 0.39 | 0 | 0.866 | 0 | -5.135 | 10.713 | 0 | 0 |  | 1 | Unknown | 2023-01-28 16:03
WO-23003-182-A-1 | 0 | -11.502 | 5.53 | -1.84 | 0.504 | 0 | -1690.644 | 0.34 | 0 | 2.138 | 0.971 | 14.04 | 0 | 0 | 0.386 | 0 | 0.316 | 0 | 0 | -0.123 | 0 | 0 | 0.391 | 0.763 | 0.464 | 0 | 0.899 | 0 | -5.7 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 15:35
WO-23003-181-A-1 | 0 | -10.962 | 5.199 | -2.194 | 0.758 | 0 | -1618.915 | 0.867 | 0 | 2.185 | 1.055 | 15.408 | 0 | 0 | 0.746 | 0 | 0.526 | 0 | 0 | 0 | 0 | 0 | 0.524 | 0.917 | 0.611 | 0 | 1.146 | 0 | -5.297 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 15:08
AB-230117-02 LFSM-H 9 | 0 | 139.186 | 217.799 | 212.529 | 238.907 | 0 | -2901.434 | 184.196 | 171.966 | 299.168 | 152.805 | 184.114 | 0 | 0 | 196.369 | 168.758 | 189.486 | 0 | 0 | 181.188 | 179.695 | 0 | 190.975 | 189.6 | 182.596 | 0 | 184.04 | 0 | 438.485 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-28 14:14
AB-230117-02 CCC-H 10 | 0 | 295.609 | 365.898 | 365.56 | 456.743 | 0 | 1643.252 | 356.398 | 324.182 | 328.58 | 311.762 | 317.881 | 0 | 0 | 334.928 | 301.419 | 342.023 | 0 | 0 | 331.083 | 350.601 | 0 | 348.019 | 287.328 | 341.141 | 0 | 315.458 | 0 | 530.431 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 13:47
AB-230117-02 LFB-H 10 | 0 | 133.231 | 182.03 | 184.535 | 232.188 | 0 | 1127.595 | 180.078 | 158.45 | 182.873 | 149.734 | 146.83 | 0 | 0 | 169.126 | 151.291 | 173.244 | 0 | 0 | 172.611 | 170.015 | 0 | 172.512 | 141.182 | 171.506 | 0 | 157.495 | 0 | 274.234 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 13:20
AB-230117-02 LD 9 | 0 | -5.796 | 6.395 | -0.961 | 0 | 0 | -2158.372 | 0.632 | 0 | 1.929 | 0 | 14.245 | 1183.691 | 0 | 0.302 | 0 | 0.236 | 0 | 0 | 0 | 0 | 0 | 0.269 | 0.765 | 0.32 | 0 | 0.707 | 0 | -5.318 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-28 12:53
AB-230117-02 LRB 19 | 0 | -43.385 | -3.285 | -1.295 | 0 | 0 | 684.629 | 0 | 0 | 0.011 | 0.753 | -0.2 | 0 | 0 | 0.192 | 0 | 0.287 | 0 | 0 | -0.299 | 0.827 | 0 | 0.277 | 0.674 | 0.512 | 0 | 0.7 | 0 | -5.845 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 12:26
LRB-19-inj1 | 0 | -42.612 | -3.18 | -1.619 | 0 | 0 | 776.597 | 0.179 | 0 | 0.022 | 0.723 | 0.66 | 0 | 0 | 0.22 | 0 | 0.281 | 0 | 0 | -0.201 | 0.906 | 0 | 0.249 | 0.637 | 0.519 | 0 | 0.72 | 0 | -5.705 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 11:59
WO-23003-180-A-1 | 0 | 6.748 | 18.366 | -1.707 | 0 | 0 | -2101.02 | 0 | 0 | 2.451 | 0 | 15.624 | 0 | 0 | 0.125 | 0 | 0.228 | 0 | 0 | 0 | 0 | 0 | 0.282 | 0.678 | 0.341 | 0 | 0.743 | 0 | -5.381 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 11:32
WO-23003-179-A-1 | 0 | -5.418 | 39.223 | -0.005 | 0 | 0 | -1779.836 | 0.544 | 0 | 1.11 | 0 | 7.783 | 0 | 0 | 0.377 | 0 | 0.249 | 0 | 0 | 2.121 | 0 | 0 | 0 | 0.719 | 0.376 | 0 | 0.706 | 0 | -6.269 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 11:05
WO-23003-178-A-1 | 0 | -8.932 | 19.196 | -0.41 | 0 | 0 | -1737.963 | 0.198 | 0 | 1.536 | 0.914 | 8.837 | 0 | 0 | 0.171 | 0 | 0.206 | 0 | 0 | 1.865 | 0 | 0 | 0.258 | 0.645 | 0.166 | 0 | 0.706 | 0 | -6.199 | 0.018 | 0 | 12.207 |  | 1 | Unknown | 2023-01-28 10:38
WO-23003-177-A-1 | 0 | -5.386 | 12.386 | -0.782 | 0 | 0 | -2103.018 | 0 | 0 | 1.644 | 0.877 | 10.965 | 0 | 0 | 0.279 | 0 | 0.246 | 0 | 0 | 0.36 | 0 | 0 | 0.285 | 0.681 | 0.378 | 0 | 0.694 | 0 | -5.397 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 10:10
WO-23003-176-A-1 | 0 | -13.509 | 9.567 | -2.164 | 0 | 0 | -1751.345 | 1.229 | 0 | 0.837 | 0 | 5.099 | 0 | 0 | 0.154 | 0 | 0.183 | 0 | 0 | 0 | 0 | 0 | 0.249 | 0.726 | 0.374 | 0 | 0.69 | 0 | -5.457 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 09:43
WO-23003-175-A-1 | 0 | -11.424 | 24.809 | -1.512 | 0 | 0 | -1674.736 | 0.341 | 0 | 0.766 | 0.81 | 5.044 | 0 | 0 | 0.243 | 0 | 0.19 | 0 | 0 | -0.345 | 0 | 0 | 0.286 | 0.716 | 0.339 | 0 | 0.708 | 0 | -5.193 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 09:16
WO-23003-174-A-1 | 0 | -1.624 | 63.983 | -1.911 | 0 | 0 | -1906.857 | 0.845 | 0 | 0.723 | 0.855 | 4.496 | 0 | 0 | 0.085 | 0 | 0.229 | 0 | 0 | -0.199 | 0 | 0 | 0.271 | 0.695 | 0.338 | 0 | 0.722 | 0 | -5.019 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 08:49
WO-23003-173-A-1 | 0 | 58.578 | 306.028 | -1.23 | 0 | 0 | -2842.957 | 0.325 | 0 | 0.877 | 0.887 | 5.228 | 0 | 0 | 0.183 | 0 | 0.267 | 0 | 0 | 1.921 | 0 | 0 | 0.305 | 0.725 | 0.354 | 0 | 0.693 | 0 | -4.188 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 08:22
WO-23003-172-A-1 | 0 | -10.455 | 31.79 | -1.081 | 0 | 0 | -1965.934 | 0.751 | 0 | 1.06 | 0.774 | 20.708 | 0 | 0 | 0.207 | 0 | 0.282 | 0 | 0 | 0.213 | 0 | 0 | 0.29 | 0.719 | 0.17 | 0 | 0.691 | 0 | -4.895 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 07:55
WO-23003-171-A-1 | 0 | -17.248 | 11.19 | -1.526 | 0 | 0 | -1354.166 | 0 | 0 | 1.033 | 0 | 20.739 | 0 | 0 | 0.07 | 0 | 0.273 | 0 | 0 | 0 | 0 | 0 | 0.273 | 0.679 | 0.386 | 0 | 0.697 | 0 | -4.584 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 07:28
AB-230117-02 LRB 18 | 0 | -43.766 | -2.391 | -1.653 | 0.341 | 0 | 793.961 | 0 | 0 | -0.045 | 0.751 | -0.299 | 0 | 0 | 0.092 | 0 | 0.309 | 0 | 0 | -0.254 | 0.712 | 0 | 0.239 | 0.633 | 0.506 | 0 | 0.716 | 0 | -5.804 | 1.665 | 0 | 0 |  | 1 | Unknown | 2023-01-28 07:01
LRB-18-inj1 | 0 | -41.06 | -3.08 | -1.581 | 0.339 | 0 | 636.977 | 0 | 0 | -0.008 | 0.741 | -0.008 | 0 | 0 | 0.162 | 0 | 0.287 | 0 | 0 | 0 | 0.716 | 0 | 0.231 | 0.658 | 0.512 | 0 | 0.71 | 0 | -5.768 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 06:34
WO-23003-170-A-1 | 0 | -9.324 | 29.873 | -0.79 | 0.345 | 0 | -2215.853 | 0.735 | 0 | 0.538 | 0.689 | 3.409 | 0 | 0 | 0.255 | 0 | 0.221 | 0 | 0 | 0.165 | 0 | 0 | 0.245 | 0.746 | 0.311 | 0 | 0.723 | 0 | -4.796 | -2.137 | 0 | 0 |  | 1 | Unknown | 2023-01-28 06:07
WO-23003-169-A-1 | 0 | 59.074 | 40.365 | -2.062 | 0.346 | 0 | -2903.292 | 0.498 | 0 | 1.05 | 0.962 | 4.807 | 0 | 0 | 0.114 | 0 | 0.257 | 0 | 0 | -0.138 | 0 | 0 | 0.271 | 0.803 | 0.365 | 0 | 0.734 | 0 | -5.445 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 05:40
WO-23003-168-A-1 | 0 | 26.912 | 35.829 | -1.473 | 0.346 | 0 | -1362.474 | 0.745 | 7974.627 | 0.194 | 0 | 0.527 | 0 | 0 | 0.388 | 0 | 0.645 | 0 | 0 | 8.287 | 0.889 | 0 | 0.536 | 0.656 | 0.503 | 0 | 0.72 | 0 | -4.761 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 05:12
WO-23003-167-A-1 | 0 | 49.292 | 208.606 | -1.445 | 0.355 | 0 | -1865.867 | 0 | 8310.781 | 0.357 | 0.912 | 0.567 | 0 | 0 | 0.257 | 0 | 0.71 | 0 | 0 | 8.59 | 1.668 | 0 | 0.566 | 0.702 | 0.488 | 0 | 0.776 | 0 | -3.784 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 04:45
WO-23003-166-A-1 | 0 | 7.499 | 104.483 | -1.646 | 0.359 | 0 | -1020.625 | 0 | 8146.394 | 0.438 | 0.995 | 0.507 | 0 | 0 | 0.187 | 0 | 0.628 | 0 | 0 | 8.573 | 1.19 | 0 | 0.504 | 0.773 | 0.392 | 0 | 0.766 | 0 | -4.66 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 04:18
WO-23003-165-A-1 | 0 | 20.959 | 6.048 | -1.383 | 0.369 | 0 | -1797.05 | 1.082 | 8091.743 | 0.594 | 0.882 | 0.574 | 0 | 0 | 0.261 | 0 | 0.497 | 0 | 0 | 8.648 | 0.789 | 0 | 0.482 | 0.85 | 0.724 | 0 | 0.792 | 0 | -4.275 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 03:51
WO-23003-164-A-1 | 0 | 23.227 | 6.026 | -1.718 | 0.39 | 0 | -2519.815 | 0.989 | 2884.658 | 0.34 | 0.862 | -0.012 | 0 | 0 | 0.276 | 0 | 0.303 | 0 | 0 | 1.161 | 0 | 0 | 0.339 | 0.734 | 0.485 | 0 | 0.797 | 0 | -4.182 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 03:24
WO-23003-163-A-1 | 0 | 35.591 | 18.214 | -1.823 | 0.438 | 0 | -2891.911 | 1.814 | 2697.215 | 0.733 | 0.846 | -0.04 | 0 | 0 | 0.364 | 0 | 0.298 | 0 | 0 | 0 | 0 | 0 | 0.346 | 0.945 | 0.449 | 0 | 0.795 | 0 | -5.284 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 02:57
WO-23003-162-A-1 | 0 | 0.88 | 5.28 | -1.502 | 0.476 | 0 | -2149.442 | 0.801 | 2959.884 | 0.647 | 0.834 | 0.068 | 0 | 0 | 0.226 | 0 | 0.632 | 0 | 0 | 1.984 | 0 | 0 | 0.462 | 0.822 | 0.454 | 0 | 0.955 | 0 | -4.851 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 02:30
WO-23003-161-A-1 | 0 | -12.179 | 3.242 | -2.173 | 0.718 | 0 | -1834.274 | 0.794 | 4700.929 | 0.876 | 0.954 | 0.212 | 0 | 0 | 0.727 | 0 | 0.674 | 0 | 0 | 2.057 | 0 | 0 | 0.561 | 1.089 | 0.552 | 0 | 1.145 | 0 | -5.135 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 02:03
AB-230117-02 LFSM-H 8 | 0 | 134.487 | 204.898 | 207.338 | 205.794 | 0 | -1865.78 | 183.28 | 4891.74 | 245.396 | 144.805 | 156.716 | 0 | 0 | 189.553 | 177.244 | 191.129 | 0 | 0 | 176.725 | 170.633 | 0 | 185.896 | 213.588 | 178.895 | 0 | 179.016 | 0 | 403.986 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-28 01:09
AB-230117-02 CCC-H 9 | 0 | 299.118 | 382.686 | 385.836 | 449.642 | 0 | 1459.244 | 351.796 | 322.214 | 343.428 | 302.345 | 306.496 | 0 | 0 | 342.172 | 302.616 | 345.726 | 0 | 0 | 338.625 | 350.18 | 0 | 341.66 | 286.071 | 337.277 | 0 | 309.308 | 0 | 528.598 | 3537.041 | 0 | 0 |  | 1 | Unknown | 2023-01-28 00:42
AB-230117-02 LFB-H 9 | 0 | 157.285 | 192.491 | 211.342 | 218.588 | 0 | 910.687 | 183.084 | 171.888 | 180.175 | 155.518 | 152.895 | 0 | 0 | 173.617 | 160.935 | 177.434 | 0 | 0 | 175.811 | 181.434 | 0 | 179.694 | 151.15 | 178.355 | 0 | 164.696 | 0 | 277.512 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-28 00:15
AB-230117-02 LD 8 | 0 | 11.19 | 7.824 | -1.357 | 0.351 | 0 | -2344.617 | 0 | 3160.606 | 0.238 | 0 | -0.179 | 0 | 0 | 0.243 | 0 | 0.49 | 0 | 0 | 1.032 | 0 | 0 | 0.38 | 0.674 | 0.414 | 0 | 0.732 | 0 | -5.182 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-27 23:48
AB-230117-02 LRB 17 | 0 | -43.046 | -2.369 | -1.964 | 0 | 0 | 461.817 | 0 | 0 | 0.041 | 0.833 | -0.144 | 0 | 0 | 0.341 | 0 | 0.331 | 0 | 0 | 0.07 | 0.889 | 0 | 0.23 | 0.629 | 0.626 | 0 | 0.695 | 0 | -5.701 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 23:20
LRB-17-inj1 | 0 | -40.691 | -1.752 | -1.648 | 0 | 0 | 680.421 | 0 | 0 | 0.101 | 0.767 | -0.151 | 0 | 0 | 0.12 | 0 | 0.322 | 0 | 0 | -0.215 | 0.786 | 0 | 0.256 | 0.705 | 0.588 | 0 | 0.704 | 0 | -5.985 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 22:53
WO-23003-160-A-1 | 0 | -4.395 | 4.97 | -1.852 | 0 | 0 | -2016.166 | 0 | 4609.936 | 0.462 | 0.834 | -0.103 | 0 | 0 | 0.086 | 0 | 0.487 | 0 | 0 | 2.311 | 1.797 | 0 | 0.313 | 0.647 | 0.361 | 0 | 0.737 | 0 | -5.155 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 22:26
WO-23003-159-A-1 | 0 | -6.585 | 9.144 | -1.088 | 0 | 0 | -2384.286 | 0 | 4651.337 | 0.446 | 0 | -0.125 | 0 | 0 | 0.252 | 0 | 0.343 | 0 | 0 | 2.729 | 0 | 0 | 0.325 | 0.726 | 0.293 | 0 | 0.718 | 0 | -4.858 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 21:59
WO-23003-158-A-1 | 0 | -18.739 | 4.796 | -1.865 | 0 | 0 | -1384.915 | 0.863 | 1326.789 | 0.367 | 0.748 | -0.232 | 1226.073 | 0 | 0.213 | 0 | 1.766 | 0 | 0 | 1.412 | 0 | 0 | 0.267 | 0.655 | 0.308 | 0 | 0.7 | 0 | -5.172 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 21:32
WO-23003-157-A-1 | 0 | 22.948 | 11.395 | -2.167 | 0 | 0 | -2694.916 | 0.409 | 1266.582 | 0.396 | 0 | 0.041 | 0 | 0 | 0.144 | 0 | 0.251 | 0 | 0 | -0.207 | 0 | 0 | 0.286 | 0.661 | 0.38 | 0 | 0.724 | 0 | -5.04 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 21:05
WO-23003-156-A-1 | 0 | -46.945 | 3.245 | -1.922 | 0 | 0 | -1693.141 | 0 | 1276.879 | 0.464 | 0 | -0.051 | 0 | 0 | 0.206 | 0 | 0.25 | 0 | 0 | 1.092 | 0 | 0 | 0.26 | 0.647 | 0.357 | 0 | 0.702 | 0 | -4.76 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 20:38
WO-23003-155-A-1 | 0 | -10.24 | 8.673 | -1.567 | 0 | 0 | -1793.372 | 1.192 | 2461.423 | 0.367 | 0.79 | 0 | 1220.91 | 0 | 0.207 | 0 | 0.263 | 0 | 0 | 0.289 | 0 | 0 | 0.374 | 0.667 | 0.326 | 0 | 0.724 | 0 | -4.517 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 20:11
WO-23003-154-A-1 | 0 | -10.764 | 11.307 | -1.803 | 0 | 0 | -1821.495 | 1.097 | 2145.145 | 0.36 | 0.884 | -0.177 | 0 | 0 | 0.171 | 0 | 0.293 | 0 | 0 | 0 | 0 | 0 | 1.635 | 0.656 | 0.404 | 0 | 0.707 | 0 | -5.882 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 19:44
WO-23003-153-A-1 | 0 | -6.914 | 23.34 | -1.817 | 0.342 | 0 | -1881.413 | 0 | 1340.929 | 0.718 | 0 | -0.135 | 0 | 0 | 0.246 | 0 | 0.404 | 0 | 0 | 0 | 0 | 0 | 0.347 | 0.7 | 0.36 | 0 | 0.72 | 0 | -5.994 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 19:17
WO-23003-152-A-1 | 0 | -3.145 | 14.397 | -1.565 | 0 | 0 | -2322.605 | 1.181 | 521.5 | 0.483 | 0 | -0.223 | 0 | 0 | 0.161 | 0 | 0.27 | 0 | 0 | 0 | 0 | 0 | 0.266 | 0.803 | 0.4 | 0 | 0.726 | 0 | -4.929 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 18:50
WO-23003-151-A-1 | 0 | -6.063 | 12.117 | -1.227 | 0.393 | 0 | -1585.015 | 0 | 477.873 | 0.208 | 0 | 0.103 | 0 | 0 | 0.316 | 0 | 1.037 | 0 | 0 | 0 | 0 | 0 | 0.323 | 0.665 | 0.572 | 0 | 0.756 | 0 | -5.976 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 18:23
AB-230117-02 LRB 16 | 0 | -33.958 | -2.887 | -1.569 | 0.341 | 0 | -302.333 | 0 | 0 | -0.016 | 0.781 | -0.003 | 0 | 0 | 0.315 | 0 | 0.431 | 0 | 0 | -0.316 | 0.749 | 0 | 0.266 | 0.745 | 0.617 | 0 | 0.739 | 0 | -5.814 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 17:56
LRB-16-inj1 | 0 | -23.075 | -3.753 | -1.774 | 0.338 | 0 | 259.3 | 0.811 | 0 | 0.011 | 0.71 | -0.069 | 0 | 0 | 0.306 | 0 | 0.375 | 0 | 0 | -0.235 | 0.741 | 0 | 0.273 | 0.654 | 0.516 | 0 | 0.741 | 0 | -5.258 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 17:28
WO-23003-150-A-1 | 0 | 24.516 | 7.99 | -1.791 | 0.34 | 0 | -2360.054 | 0 | 579.319 | 0.241 | 0.907 | 0.178 | 0 | 0 | 0.436 | 0 | 0.351 | 0 | 0 | 0 | 0 | 0 | 0.324 | 0.7 | 0.429 | 0 | 0.711 | 0 | -5.05 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 17:01
WO-23003-149-A-1 | 0 | 46.828 | 7.993 | -1.577 | 0 | 0 | -2330.049 | 1.108 | 519.796 | 0.552 | 0.724 | 0 | 0 | 0 | 0.262 | 0 | 0.365 | 0 | 0 | 0.032 | 0 | 0 | 0.346 | 0.811 | 0.237 | 0 | 0.749 | 0 | -4.16 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 16:34
WO-23003-148-A-1 | 0 | 39.824 | 6.918 | -2.099 | 0.343 | 0 | -2144.673 | 1.373 | 577.003 | 0.485 | 0 | -0.065 | 1173.655 | 0 | 0.24 | 0 | 0.359 | 0 | 0 | 0.106 | 0 | 0 | 0.358 | 0.672 | 0.388 | 0 | 0.725 | 0 | -4.126 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 16:07
WO-23003-147-A-1 | 0 | 24.359 | 4.832 | -2.056 | 0.349 | 0 | -1954.718 | 0 | 572.526 | 0.624 | 0 | -0.042 | 0 | 0 | 0.28 | 0 | 0.366 | 0 | 0 | 0 | 0 | 0 | 0.329 | 0.711 | 0.467 | 0 | 0.722 | 0 | -5.078 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 15:40
WO-23003-146-A-1 | 0 | 49.173 | 3.129 | -1.39 | 0.357 | 0 | -1857.046 | 0 | 732.447 | 0.457 | 0.792 | 0.036 | 0 | 0 | 0.29 | 0 | 0.408 | 0 | 0 | 1.461 | 0 | 0 | 0.324 | 0.761 | 0.535 | 0 | 0.721 | 0 | -5.148 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 15:13
WO-23003-145-A-1 | 0 | 21.135 | 6.036 | -0.113 | 0.369 | 0 | -1574.062 | 0 | 638.175 | 0.552 | 0.871 | 0.014 | 0 | 0 | 0.347 | 0 | 0.419 | 0 | 0 | 0 | 0 | 0 | 0.332 | 0.681 | 0.406 | 0 | 0.792 | 0 | -4.668 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 14:46
WO-23003-144-A-1 | 0 | 31.261 | 7.157 | -0.794 | 0.387 | 0 | -1853.622 | 0 | 127.716 | 0.492 | 0 | 0.056 | 0 | 0 | 0.415 | 0 | 0.465 | 0 | 0 | 0.237 | 0 | 0 | 0.337 | 0.759 | 0.442 | 0 | 0.861 | 0 | -4.129 | 0 | 0 | 17.888 |  | 1 | Unknown | 2023-01-27 14:19
WO-23003-143-A-1 | 0 | 25.016 | 6.142 | -2.023 | 0.423 | 0 | -1858.381 | 1.526 | 102.324 | 0.633 | 0.803 | 0.121 | 0 | 0 | 0.235 | 0 | 0.671 | 0 | 0 | -0.387 | 0 | 0 | 0.363 | 0.756 | 0.43 | 0 | 0.794 | 0 | -4.687 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 13:52
WO-23003-142-A-1 | 0 | 46.024 | 4.891 | -2.065 | 0.499 | 0 | -2160.174 | 0 | 151.9 | 0.636 | 0 | 0.024 | 0 | 0 | 0.527 | 0 | 0.92 | 0 | 0 | 0.445 | 1.467 | 0 | 0.486 | 1.038 | 0.453 | 0 | 0.896 | 0 | -4.971 | 4.655 | 0 | 0 |  | 1 | Unknown | 2023-01-27 13:25
WO-23003-141-A-1 | 0 | 112.953 | 11.896 | -1.589 | 0.896 | 0 | -2955.93 | 0 | 113.991 | 1.564 | 0 | 0.287 | 0 | 0 | 0.615 | 0 | 0.925 | 0 | 0 | 1.69 | 0 | 0 | 0.724 | 1.115 | 0.824 | 0 | 1.191 | 0 | -4.586 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 12:58
AB-230117-02 LFSM-H 7 | 0 | 250.956 | 199.912 | 215.304 | 200.939 | 0 | -2487.807 | 175.013 | 277.866 | 282.112 | 145.268 | 159.093 | 1237.967 | 0 | 190.18 | 188.826 | 188.842 | 0 | 0 | 175.116 | 173.568 | 0 | 188.523 | 202.11 | 182.658 | 0 | 186.394 | 0 | 416.861 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-27 12:04
AB-230117-02 CCC-H 8 | 0 | 314.441 | 371.867 | 399.548 | 445.243 | 0 | 782.316 | 374.743 | 302.112 | 355.866 | 310.833 | 321.519 | 1183.254 | 0 | 346.466 | 306.068 | 346.394 | 0 | 0 | 345.124 | 347.778 | 0 | 353.314 | 248.6 | 345.655 | 0 | 322.674 | 0 | 616.443 | 0 | 0 | 17128.031 |  | 1 | Unknown | 2023-01-27 11:36
AB-230117-02 LFB-H 8 | 0 | 159.103 | 180.56 | 189.665 | 205.316 | 0 | 568.024 | 178.569 | 157.405 | 174.967 | 149.213 | 147.72 | 0 | 0 | 170.147 | 149.916 | 172.643 | 0 | 0 | 170.567 | 174.977 | 0 | 174.528 | 132.033 | 174.263 | 0 | 157.989 | 0 | 343.211 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 11:09
AB-230117-02 LD 7 | 0 | 95.852 | 9.315 | -2.14 | 0.345 | 0 | -2682.455 | 0 | 151.363 | 0.531 | 0 | 0 | 0 | 0 | 0.262 | 0 | 1.197 | 0 | 0 | 0.083 | 1.76 | 0 | 0.336 | 0.745 | 0.417 | 0 | 0.725 | 0 | -5.26 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-27 10:42
AB-230117-02 LRB 15 | 0 | -21.294 | -0.741 | -1.558 | 0 | 0 | -156.018 | 0.658 | 0 | 0.037 | 0 | 0.002 | 0 | 0 | 0.374 | 0 | 1.286 | 0 | 0 | 0 | 0.9 | 0 | 0.302 | 0.686 | 0.525 | 0 | 0.736 | 0 | -5.853 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 10:15
LRB-15-inj1 | 0 | -25.7 | -4.702 | -1.943 | 0.344 | 0 | 233.873 | 0.374 | 0 | -0.071 | 0 | -0.083 | 1143.28 | 0 | 0.306 | 0 | 1.417 | 0 | 0 | -0.287 | 0.91 | 0 | 0.31 | 0.689 | 0.603 | 0 | 0.727 | 0 | -5.644 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 09:48
WO-23003-140-A-1 | 0 | 21.135 | 6.787 | -2.18 | 0 | 0 | -1913.783 | 0.914 | 152.509 | 0.531 | 0 | -0.048 | 0 | 0 | 0.192 | 0 | 2.01 | 0 | 0 | 0 | 0 | 0 | 0.338 | 0.7 | 0.357 | 0 | 0.736 | 0 | -5.123 | 0 | 0 | 10.995 |  | 1 | Unknown | 2023-01-27 09:21
WO-23003-139-A-1B | 0 | 34.286 | 6.982 | -1.819 | 0.36 | 0 | -2000.414 | 2.591 | 146.837 | 0.517 | 0.848 | 0.878 | 0 | 0 | 0.229 | 0 | 6.248 | 0 | 0 | 0 | 0 | 0 | 0.345 | 0.892 | 0.378 | 0 | 0.743 | 0 | -4.155 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-27 08:54
WO-23003-139-A-1 | 0 | 42.601 | 24.366 | -1.978 | 0.475 | 0 | -1388.169 | 3.112 | 141.212 | 0.236 | 0 | 0.75 | 0 | 0 | 0.428 | 0 | 133.922 | 0 | 0 | 0.217 | 0 | 0 | 0.57 | 0.947 | 0.695 | 0 | 0.941 | 0 | -0.443 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-27 08:27
Rinse12 | 0 | 4264.452 | 312.292 | 3.56 | 1.902 | 0 | 6368.875 | 40.266 | 0 | 4.619 | 4.956 | 5.493 | 0 | 0 | 2.482 | 0 | 3907 | 0 | 0 | 16.054 | 0 | 0 | 6.109 | 3.51 | 11.109 | 0 | 4.679 | 0 | 210.223 | -1.202 | 0 | 0 | C | 1 | Unknown | 2023-01-23 10:08
WO-23003-203-A-1B | 0 | -18.384 | 1.188 | 0.39 | 0.447 | 0 | 1183.271 | 0.231 | 10.448 | 215.936 | 17.597 | 1809.619 | 0 | 0 | 0.225 | 0 | 0.378 | 0 | 0 | 1.058 | 0.868 | 0 | 0.326 | 0.992 | 4.491 | 0 | 2.892 | 0 | -5.409 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-30 12:09
WO-23003-208-A-1B | 0 | -10.214 | 3.565 | -1.073 | 0.361 | 0 | 756.265 | 0.718 | 10.574 | 216.322 | 17.222 | 1821.511 | 0 | 0 | 0.219 | 0 | 0.316 | 0 | 0 | 0.964 | 0.753 | 0 | 0.294 | 0.999 | 3.878 | 0 | 2.513 | 0 | -5.806 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-30 14:25
WO-23003-207-A-1B | 0 | 5.897 | 11.164 | 0.53 | 0.395 | 0 | 913.324 | 0 | 11.126 | 232.734 | 18.417 | 1989.268 | 0 | 0 | 0.242 | 0 | 0.34 | 0 | 0 | 0.967 | 0.815 | 0 | 0.366 | 1.009 | 3.98 | 0 | 2.809 | 0 | -5.891 | 0 | 0 | 12.585 |  | 1 | Unknown | 2023-01-30 13:58
WO-23003-206-A-1B | 0 | -17.296 | 5.148 | -1.224 | 0.394 | 0 | 1064.734 | 0.356 | 11.594 | 296.595 | 20.732 | 2969.033 | 0 | 0 | 0.282 | 0 | 0.347 | 0 | 0 | 1.012 | 0.794 | 0 | 0.362 | 1.199 | 4.89 | 0 | 2.78 | 0 | -6.146 | 1.669 | 0 | 0 |  | 1 | Unknown | 2023-01-30 13:30
WO-23003-205-A-1B | 0 | 12.562 | 105.739 | 0.794 | 0.449 | 0 | 641.965 | 0.35 | 44.219 | 640.608 | 74.136 | 4923.452 | 0 | 0 | 0.233 | 0 | 0.343 | 0 | 0 | 2.017 | 1.411 | 0 | 0 | 1.626 | 7.698 | 0 | 3.51 | 0 | -4.735 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-30 13:03
WO-23003-204-A-1B | 0 | -11.839 | 27.602 | -0.253 | 0.485 | 0 | 614.722 | 0.313 | 44.764 | 643.996 | 73.237 | 5298.356 | 0 | 0 | 0.303 | 0 | 0.315 | 0 | 0 | 2.904 | 1.758 | 0 | 0 | 1.389 | 9.339 | 0 | 2.867 | 0 | -4.387 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-30 12:36
LRB-22-reprep-inj2 | 0 | -26.773 | 7.198 | 0.367 | 0.419 | 0 | 1845.449 | 0.212 | 0 | 0.3 | 0.799 | 3.082 | 0 | 0 | 0.411 | 0 | 0.404 | 0 | 0 | 0.041 | 0.494 | 0 | 0.435 | 0.678 | 2.24 | 0 | 2.98 | 0 | -5.507 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-30 11:42
LRB-22-reprep | 0 | -19.826 | 6.092 | 1.198 | 0.465 | 0 | 1516.417 | 0 | 0 | 0.273 | 0.85 | 3.365 | 0 | 0 | 0.378 | 0 | 0.417 | 0 | 0 | 0.042 | 0.435 | 0 | 0.439 | 0.68 | 2.334 | 0 | 2.759 | 0 | -5.898 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-30 11:15
217288-d4-reprep | 0 | -9.074 | 10.978 | -1.272 | 0.339 | 0 | -596.574 | 0.593 | 4342.461 | 0.77 | 0.858 | 4.749 | 0 | 0 | 0.029 | 0 | 0.577 | 0 | 0 | 3.487 | 0.768 | 0 | 0.407 | 0.641 | 0.739 | 0 | 1.391 | 0 | -4.475 | 0 | 0 | 0 |  | 4 | Unknown | 2023-01-30 15:56
217289-d4-reprep | 0 | -13.822 | 12.459 | -1.583 | 0.339 | 0 | -537.227 | 0 | 4355.518 | 0.624 | 0.848 | 2.443 | 0 | 0 | 0.202 | 0 | 0.368 | 0 | 0 | 3.221 | 0.642 | 0 | 0.343 | 0.723 | 0.655 | 0 | 1.237 | 0 | -4.853 | 0 | 0 | 0 |  | 4 | Unknown | 2023-01-30 16:23
217295-d1000-LD-reprep | 0 | -47.866 | -0.431 | -1.582 | 0 | 0 | 1063.903 | 0.177 | 0 | 26.189 | 1.598 | 1222.979 | 0 | 0 | 0.126 | 0 | 1.087 | 0 | 0 | -0.202 | 0.308 | 0 | 0.28 | 2.205 | 0.299 | 0 | 0.725 | 0 | -5.991 | 0 | 0 | 0 |  | 1000 | Unknown | 2023-01-30 23:09
217296-d1000-LFSM-200-reprep | 0 | 145.29 | 189.732 | 202.065 | 404.612 | 0 | 1364.617 | 199.356 | 171.849 | 209.69 | 164.478 | 1611.097 | 0 | 0 | 177.964 | 166.141 | 190.796 | 0 | 0 | 182.061 | 189.823 | 0 | 184.164 | 144.268 | 180.753 | 0 | 177.374 | 0 | 265.776 | 0 | 0 | 0 |  | 1000 | Unknown | 2023-01-31 02:46
217315-d2-reprep | 0 | 4.342 | 16.17 | -1.427 | 0.34 | 0 | -2029.531 | 0 | 0 | 0.515 | 0.802 | 1.534 | 0 | 0 | 0.108 | 137.147 | 0.228 | 0 | 0 | 1.512 | 0 | 0 | 0.321 | 0.741 | 0.537 | 0 | 1.215 | 0 | -4.692 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-30 16:50
217335-d2-LD-reprep | 0 | -0.844 | 14.757 | -0.376 | 0 | 0 | -2349.483 | 0 | 0 | 1.315 | 0.725 | 53.394 | 0 | 0 | 0.006 | 430.994 | 1.008 | 0 | 0 | -0.13 | 0 | 0 | 0.328 | 0.762 | 0.574 | 0 | 1.18 | 0 | -5.296 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-30 23:36
217336-d2-LFSM-200-reprep | 0 | 262.827 | 229.683 | 208.752 | 216.508 | 0 | -3057.123 | 185.585 | 175.682 | 310.298 | 149.657 | 211.498 | 0 | 0 | 185.41 | 606.641 | 194.566 | 0 | 0 | 180.656 | 184.866 | 0 | 187.674 | 202.297 | 175.24 | 0 | 239.752 | 0 | 475.374 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-31 03:13
217356-d2-reprep | 0 | -9.612 | 12.982 | 0.302 | 0.333 | 0 | -1843.624 | 0.299 | 0 | 0.256 | 0.792 | 1.099 | 0 | 0 | 0.195 | 0 | 4.268 | 0 | 0 | -0.193 | 0.477 | 0 | 0.326 | 0.753 | 0.546 | 0 | 1.376 | 0 | -5.28 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-30 17:17
217357-d2-reprep | 0 | 27.503 | 12.796 | 6.657 | 0.338 | 0 | -2048.512 | 0.835 | 0 | 0.222 | 0.82 | 1.143 | 0 | 0 | 0.173 | 0 | 48.327 | 0 | 0 | -0.174 | 0 | 0 | 0.929 | 0.68 | 0.607 | 0 | 1.21 | 0 | -4.977 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-30 17:44
217358-d2-reprep | 0 | 3.484 | 13.334 | 12.64 | 0 | 0 | -1817.146 | 0 | 0 | 0.686 | 0.917 | 0.749 | 0 | 0 | 0.188 | 0 | 2575.277 | 0 | 0 | 0 | 0 | 0 | 1.195 | 1.03 | 0.733 | 0 | 1.123 | 0 | -5.422 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-30 18:12
217359-d2-LD | 0 | -10.065 | 18.448 | 6.068 | 0 | 0 | -1977.266 | 0 | 0 | 0.176 | 0.759 | 0.25 | 0 | 0 | 0.479 | 0 | 8748.882 | 0 | 0 | -0.061 | 0 | 0 | 2.774 | 0.772 | 0.968 | 0 | 0.704 | 0 | -5.043 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-31 01:25
217359-d2-reprep | 0 | -11.12 | 14.511 | 11.553 | 0 | 0 | -1783.186 | 0 | 0 | 0.343 | 0.815 | 0.671 | 0 | 0 | 0.688 | 0 | 9903.264 | 0 | 0 | 0.399 | 0 | 0 | 3.011 | 1.398 | 1.217 | 0 | 1.23 | 0 | -5.261 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-30 18:39
217360-d2-LFSM-200 | 0 | 132.551 | 225.235 | 210.292 | 218.738 | 0 | -2191.951 | 187.722 | 168.197 | 267.17 | 144.38 | 160.884 | 0 | 0 | 178.944 | 182.595 | 384.137 | 0 | 0 | 182.089 | 186.158 | 0 | 185.543 | 185.019 | 185.148 | 0 | 182.929 | 0 | 336.899 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-31 04:34
217360-d2-reprep | 0 | 6.774 | 17.627 | -0.119 | 0 | 0 | -2344.388 | 0 | 0 | 0.313 | 0.8 | 0.63 | 0 | 0 | 0.221 | 0 | 378.052 | 0 | 0 | 0.348 | 0 | 0 | 0.345 | 0.666 | 0.642 | 0 | 1.125 | 0 | -4.772 | 2.774 | 0 | 0 |  | 2 | Unknown | 2023-01-30 19:06
217375-d2-LD-reprep | 0 | 30.848 | 190.929 | 3.309 | 0.335 | 0 | -2467.426 | 0 | 0 | 0.233 | 0.772 | 1.407 | 0 | 0 | 0.205 | 0 | 28.506 | 0 | 0 | -0.006 | 0.931 | 0 | 0.367 | 0.646 | 0.848 | 0 | 1.16 | 0 | -5.024 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-31 00:04
217376-d2-LFSM-200-reprep | 0 | 131.198 | 322.854 | 204.968 | 212.858 | 0 | -2446.857 | 173.032 | 158.711 | 211.603 | 146.373 | 153.34 | 0 | 0 | 174.657 | 196.844 | 266.746 | 0 | 0 | 170.416 | 173.811 | 0 | 177.327 | 175.577 | 166.325 | 0 | 230.508 | 0 | 384.556 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-31 03:40
217393-d4-reprep | 0 | -18.662 | 12.629 | 5.858 | 0 | 0 | -693.821 | 0 | 0 | 0.055 | 0.771 | 0.47 | 0 | 0 | 0.256 | 0 | 1061.939 | 0 | 0 | 1.123 | 0.775 | 0 | 2.949 | 0.697 | 3.747 | 0 | 1.235 | 0 | -5.328 | 0 | 0 | 13.117 |  | 4 | Unknown | 2023-01-30 19:33
217394-d4-reprep | 0 | -16.32 | 11.577 | 4.122 | 0 | 0 | -612.677 | 0.511 | 0 | 0.105 | 0.752 | 0.294 | 0 | 0 | 0.327 | 0 | 1414.578 | 0 | 0 | 1.642 | 1.218 | 0 | 4.034 | 0.749 | 4.579 | 0 | 1.236 | 0 | -5.208 | 0 | 0 | 0 |  | 4 | Unknown | 2023-01-30 20:00
217395-d4-LD-reprep | 0 | -18.905 | 9.055 | 1.839 | 0 | 0 | -729.492 | 0 | 0 | 0.302 | 0.717 | 0.313 | 1164.357 | 0 | 0.254 | 0 | 2607.572 | 0 | 0 | 2.363 | 1.523 | 0 | 8.406 | 0.779 | 8.776 | 0 | 1.381 | 0 | -5.074 | 0 | 0 | 0 |  | 4 | Unknown | 2023-01-31 00:31
217395-d4-reprep | 0 | -26.093 | 5.597 | 3.305 | 0 | 0 | -685.422 | 0.132 | 0 | 0.028 | 0.747 | 0.241 | 0 | 0 | 0.287 | 0 | 2714.735 | 0 | 0 | 2.849 | 1.534 | 0 | 8.354 | 0.741 | 9.032 | 0 | 1.16 | 0 | -5.257 | 0 | 0 | 0 |  | 4 | Unknown | 2023-01-30 21:21
217396-d200-LFSM-200 | 0 | 133.178 | 192.08 | 248.904 | 228.127 | 0 | 1291.449 | 175.395 | 162.554 | 165.612 | 144.704 | 154.142 | 0 | 0 | 160.84 | 153.37 | 4040.399 | 0 | 0 | 171.237 | 175.651 | 0 | 178.6 | 189.747 | 168.361 | 0 | 221.107 | 0 | 249.6 | 0 | 0 | 0 |  | 200 | Unknown | 2023-01-31 04:07
217399-d2-reprep | 0 | -3.205 | 5.743 | -0.074 | 0.332 | 0 | -1948.401 | 0 | 11.758 | 0.21 | 0 | 0.353 | 0 | 0 | 0.101 | 0 | 4.311 | 0 | 0 | 1.515 | 0 | 0 | 0.312 | 0.64 | 0.679 | 0 | 1.382 | 0 | -4.933 | -0.95 | 0 | 0 |  | 2 | Unknown | 2023-01-30 22:15
217400-d2-reprep | 0 | -14.742 | 2.805 | -0.786 | 0 | 0 | -2338.926 | 0 | 11.796 | 0.423 | 0 | 0.213 | 0 | 0 | 0.136 | 0 | 1.945 | 0 | 0 | 0.695 | 0 | 0 | 0.321 | 0.667 | 0.707 | 0 | 1.275 | 0 | -4.006 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-30 22:42
217415-d2-LD-reprep | 0 | 14.772 | 7.025 | -2.034 | 0 | 0 | -2375.359 | 0.57 | 121.83 | 0.572 | 0 | 0.179 | 0 | 0 | 0.021 | 0 | 10.557 | 0 | 0 | 2.022 | 0 | 0 | 0.339 | 0.658 | 0.48 | 0 | 0.698 | 0 | -4.645 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-31 02:19
217416-d2-LFSM-200-reprep | 0 | 230.751 | 209.798 | 211.635 | 218.234 | 0 | -2292.378 | 183.534 | 312.093 | 254.925 | 152.563 | 161.847 | 0 | 0 | 180.684 | 199.788 | 195.064 | 0 | 0 | 181.643 | 184.672 | 0 | 187.492 | 208.736 | 186.591 | 0 | 183.094 | 0 | 369.985 | 0 | 0 | 0 |  | 2 | Unknown | 2023-01-31 05:01
CCC-H-12-reinject | 0 | 319.537 | 380.522 | 398.231 | 516.567 | 0 | 827.953 | 364.623 | 333.02 | 348.29 | 324.299 | 326.958 | 0 | 0 | 337.682 | 305.952 | 353.248 | 0 | 0 | 350.683 | 349.782 | 0 | 354.686 | 246.665 | 351.819 | 0 | 325.504 | 0 | 536.228 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-31 07:17
LFB-H-11-reinject | 0 | 156.333 | 231.008 | 191.795 | 284.856 | 0 | 1952.581 | 225.808 | 199.836 | 220.341 | 195.122 | 194.608 | 0 | 0 | 17.623 | 194.643 | 213.202 | 0 | 0 | 209.699 | 219.041 | 0 | 224.483 | 16.591 | 205.408 | 0 | 27.003 | 0 | 305.009 | 530.887 | 0 | 0 |  | 1 | Unknown | 2023-01-31 06:50
QCS-2-Reinject | 0 | 160.3 | 228.079 | 220.864 | 255.969 | 0 | 1773.213 | 211.41 | 192.951 | 198.851 | 176.719 | 175.82 | 0 | 0 | 203.422 | 174.903 | 205.073 | 0 | 0 | 197.138 | 200.785 | 0 | 202.109 | 158.331 | 189.964 | 0 | 290.941 | 0 | 289.231 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-31 06:23
LRB-22-reprep-inj2 | 0 | -44.147 | 5.235 | -0.648 | 0.721 | 0 | 1478.318 | 0.24 | 0 | 0.395 | 0.899 | 0.384 | 0 | 0 | 0.653 | 0 | 1.636 | 0 | 0 | 0.203 | 0.315 | 0 | 0.537 | 1.116 | 0.764 | 0 | 1.633 | 0 | -6.141 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-31 05:56
LRB-22-reprep | 0 | -45.825 | 8.435 | 5.787 | 11.563 | 0 | 889.628 | 0 | 0 | 5.062 | 3.048 | 6.458 | 0 | 0 | 8.185 | 0 | 9.672 | 0 | 0 | 1.213 | 0.383 | 0 | 5.822 | 9.12 | 3.603 | 0 | 9.906 | 0 | -5.673 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-31 05:29
WO-23003-57-A-1 | 0 | 54.949 | 16.549 | -1.681 | 0 | 0 | -2289.125 | 0 | 1.166 | 0.59 | 0 | 0.509 | 0 | 0 | 0.42 | 0 | 4.897 | 0 | 0 | 0.591 | 0 | 0 | 0.365 | 0.678 | 0.337 | 0 | 0.698 | 0 | -3.46 | 0 | 0 | 0 |  | 1 | Unknown | 2023-01-20 06:46

## Sheet3
Batch Folder Path: | D:\TraceFinderData\Projects\PFAS Online SPE\2023-01-18\2023-01-18
Batch Name: | 2023-01-18
Last Acquisition Date: | 2023-01-31 07:17
User Name: | Quantiva
Sample Count: | 365
Area - N/F
Data File | 11-Cl-PF3OUDS | 4-2:FTS | 6-2:FTS | 8-2:FTS | 9-Cl-PF3ONS | D3-N-Et-FOSA | D3-N-Me-FOSA | FBSA | FHxA | HFPO-DA | L-PFBS | L-PFHpS | L-PFHxS | L-PFOS | M2-4:2-FTS | M2-6:2-FTS | M2-8:2-FTS | M3PFBS | M3PFHxS | M4PFHpA | M5PFHxA | M5PFPeA | M6PFDA | M8FOSA | M8PFOA | M8PFOS | M9PFNA | MPFBA | MPFDoA | MPFUDA | MPHFPO-DA | NaDONA | N-Et-FOSA | N-Me-FOSA | PFBA | PFDA | PFDoA | PFDS | PFHpA | PFHxA | PFHxDA | PFNA | PFNS | PFOA | PFODA | PFOSA | PFPeA | PFPeS | PFTeDA | PFTrDA | PFUdA | All Flags | Sample Type
AB-230117-02_Cal_Blank_1 | N/F | 112516 | 227843 | 17023 | 994323 | N/F | 3288127 | N/F | 654123 | N/F | N/F | 49573 | 26032 | 89073 | 212284 | 1097131 | 1252931 | 791404 | 3461600 | 15471327 | 6168538 | 1139560 | 30301110 | 4635033 | 27779178 | 4273192 | 30410864 | 632991 | N/F | N/F | 999100 | N/F | N/F | 14483 | N/F | 375299 | N/F | N/F | 261105 | 16507 | 93148 | 330634 | 5902 | 321565 | 11786 | 9268 | N/F | 14859 | 72946 | N/F | 307619 |  | Unknown
AB-230117-02_Cal_Std_1_1 | N/F | 77777 | 77386 | 28590 | 440258 | N/F | 5220775 | N/F | 396975 | 9558 | N/F | 23909 | 17310 | 51290 | 205380 | 956924 | 1148416 | 474992 | 3721715 | 17841442 | 7254807 | 563429 | 40439150 | 4445809 | 31153664 | 4544352 | 35852585 | 374515 | N/F | N/F | 1412490 | N/F | N/F | 33229 | N/F | 265543 | N/F | N/F | 254669 | 50035 | 129187 | 211226 | 23099 | 247234 | 13991 | 19796 | N/F | 15209 | 215094 | N/F | 216278 |  | Cal Std
AB-230117-02_Cal_Std_2_1 | N/F | 62203 | 56912 | 18006 | 745795 | N/F | 5001560 | N/F | 342400 | 16470 | 6034 | 46984 | 33917 | 60367 | 195991 | 926769 | 975062 | 914458 | 3812773 | 17947361 | 6943986 | 1179073 | 39003692 | 4183962 | 31781766 | 4444601 | 35137464 | 334728 | N/F | N/F | 1300088 | N/F | 1045 | 51901 | N/F | 388301 | N/F | 1636 | 189621 | 71334 | 139081 | 317888 | 32714 | 352822 | 18587 | 37787 | N/F | 34689 | 393070 | N/F | 375224 |  | Cal Std
AB-230117-02_Cal_Std_3_1 | N/F | 69177 | 52924 | 20348 | 1595719 | N/F | 5140054 | 43 | 456698 | 35668 | 16020 | 81831 | 68453 | 98708 | 208317 | 818327 | 1052832 | 964959 | 3747583 | 17691989 | 7322397 | 1125545 | 41799297 | 4081775 | 27679054 | 4174088 | 34506802 | 433302 | 2234 | N/F | 1271882 | N/F | N/F | 86354 | 4306 | 831686 | N/F | N/F | 341420 | 139410 | 147347 | 598183 | 87517 | 568100 | 17990 | 74404 | N/F | 66633 | 742431 | N/F | 701872 |  | Cal Std
AB-230117-02_Cal_Std_4_1 | N/F | 75233 | 87011 | 55683 | 4341376 | N/F | 5174106 | N/F | 409949 | 77901 | 31394 | 234223 | 159856 | 211795 | 196635 | 922189 | 1089840 | 775858 | 3751819 | 17911014 | 6862671 | 1142105 | 42979199 | 4355222 | 31867700 | 4570958 | 37211667 | 389333 | N/F | N/F | 1286073 | N/F | 804 | 202411 | 16664 | 1982722 | N/F | N/F | 857101 | 310164 | 160570 | 1577968 | 209710 | 1486141 | 8975 | 201087 | N/F | 162255 | 1770992 | N/F | 1792559 |  | Cal Std
AB-230117-02_Cal_Std_5_1 | N/F | 81494 | 115586 | 98530 | 8341543 | N/F | 4897010 | N/F | 399852 | 152272 | 55155 | 402582 | 338372 | 365677 | 173345 | 809384 | 1033500 | 536479 | 3716010 | 17529330 | 6974866 | 536464 | 41768495 | 4131540 | 27530639 | 4177449 | 33993178 | 384293 | N/F | N/F | 1283661 | N/F | 1149 | 385491 | 31872 | 3774582 | N/F | N/F | 1624702 | 617567 | 164135 | 2913228 | 414950 | 2515035 | 13909 | 374702 | N/F | 308080 | 3464585 | N/F | 3481031 |  | Cal Std
AB-230117-02_Cal_Std_6_1 | N/F | 101858 | 204888 | 182076 | 15888415 | N/F | 4680611 | N/F | 434158 | 276434 | 132726 | 899511 | 651156 | 699407 | 180813 | 871544 | 1003720 | 715733 | 3596018 | 16645776 | 6780447 | 1051047 | 38998738 | 3817485 | 30647728 | 3900993 | 32905329 | 411914 | N/F | N/F | 1184376 | N/F | 733 | 700485 | 76673 | 7152575 | N/F | N/F | 3179923 | 1182559 | 162361 | 5766346 | 773530 | 5612352 | 20213 | 726718 | N/F | 609580 | 6777654 | N/F | 6703160 |  | Cal Std
AB-230117-02_Cal_Std_7_1 | N/F | 144976 | 418191 | 426365 | 39370726 | N/F | 3935057 | N/F | 287137 | 614770 | 373027 | 2202332 | 1559870 | 1805281 | 162804 | 772185 | 861916 | 737784 | 3292381 | 14061854 | 5695371 | 955118 | 33433459 | 3825231 | 27053869 | 4019358 | 32561100 | 266916 | N/F | N/F | 1006545 | N/F | N/F | 1531077 | 136686 | 15974735 | N/F | N/F | 6969311 | 2652843 | 178718 | 14508535 | 1821237 | 12973330 | 30320 | 1857538 | N/F | 1324925 | 16176710 | N/F | 14081290 |  | Cal Std
AB-230117-02_Cal_Std_8_1 | N/F | 252728 | 620784 | 677193 | 71997932 | N/F | 3222112 | N/F | 367743 | 1636057 | 958314 | 4199129 | 3215157 | 3317470 | 193926 | 623863 | 676473 | 943092 | 3257272 | 15570319 | 7138944 | 1073867 | 32156115 | 3658943 | 27803507 | 3648036 | 31218635 | 355650 | N/F | N/F | 1298685 | N/F | N/F | 2529592 | 317397 | 30451155 | N/F | N/F | 15394166 | 6805908 | 192354 | 27303504 | 3418214 | 26008307 | 50615 | 3743469 | N/F | 3246458 | 34014521 | N/F | 27246902 |  | Cal Std
AB-230117-02_Cal_Std_9_1 | N/F | 372077 | 905320 | 851712 | 116193964 | N/F | 2442342 | N/F | 323358 | 3247807 | 1701817 | 6846874 | 5460470 | 5632878 | 170163 | 430793 | 483992 | 823849 | 2689261 | 13298661 | 6821542 | 1041605 | 26016671 | 3254875 | 23663020 | 2894577 | 25709605 | 307127 | N/F | N/F | 1263633 | N/F | 608 | 3901024 | 557472 | 49179179 | N/F | N/F | 26176244 | 13033490 | 238950 | 45911037 | 5400832 | 44738056 | 84652 | 6346673 | N/F | 5856515 | 56919922 | N/F | 44285709 |  | Cal Std
AB-230117-02_Cal_Std_10_1 | N/F | 805127 | 1433828 | 1275925 | 197461920 | N/F | 1467537 | N/F | 375270 | 7667671 | 4123554 | 11567701 | 9750392 | 9590556 | 146615 | 267204 | 320707 | 735204 | 1848502 | 8850581 | 5884437 | 1086232 | 16098522 | 2612917 | 16194891 | 1890722 | 16620897 | 355318 | N/F | N/F | 1166714 | 2538 | N/F | 6112365 | 1880324 | 76750243 | N/F | 820 | 44673947 | 29261490 | 329132 | 75456240 | 8544439 | 79267874 | 159352 | 12970999 | N/F | 11150786 | 119694688 | N/F | 69996361 |  | Cal Std
AB-230117-02_Cal_Std_11_1 | N/F | 1332529 | 2118795 | 2077052 | 272702448 | N/F | 990370 | N/F | 419402 | 12061892 | 7476913 | 16347014 | 13118827 | 11968729 | 123995 | 217203 | 190259 | 743175 | 1182015 | 5514718 | 4663536 | 1027056 | 9992543 | 1895890 | 10669119 | 1117397 | 10205299 | 399743 | 1447 | N/F | 930811 | 4193 | N/F | 8231722 | 4061475 | 95537350 | N/F | N/F | 55812284 | 46123697 | 575935 | 91743660 | 13093151 | 101445144 | 354082 | 20481959 | N/F | 15936312 | 164908377 | N/F | 85329497 |  | Cal Std
AB-230117-02_Cal_Std_12_1 | N/F | 2052249 | 3752510 | 3400229 | 322467204 | N/F | 756279 | N/F | 366638 | 17411401 | 13623368 | 20315876 | 19347546 | 16014447 | 88970 | 177876 | 168650 | 624962 | 801246 | 3271930 | 3204749 | 893305 | 5309212 | 1428471 | 6088887 | 717869 | 5450072 | 344196 | N/F | N/F | 650094 | 9030 | N/F | 12733947 | 8016201 | 104205973 | N/F | N/F | 68699541 | 66607272 | 888421 | 100296778 | 13710573 | 122200751 | 645550 | 31534699 | N/F | 22410258 | 245054542 | N/F | 93843873 |  | Cal Std
AB-230117-02_Cal_Std_13_1 | N/F | 3596951 | 7130291 | 6678346 | 412145160 | N/F | 509906 | N/F | 325335 | 20974427 | 22784388 | 26960346 | 27672585 | 23051221 | 70362 | 158377 | 145781 | 398687 | 462298 | 1495614 | 1609916 | 697777 | 2537051 | 987007 | 2671403 | 409631 | 2633403 | 307430 | N/F | 2123 | 337049 | 19283 | N/F | 20401024 | 17446659 | 125687496 | N/F | N/F | 78667431 | 78848802 | 1634231 | 123270974 | 17814751 | 135061825 | 1390553 | 54196293 | N/F | 29004789 | 351506807 | N/F | 117265456 |  | Cal Std
AB-230117-02_CCC-H_1 | N/F | 388495 | 810157 | 875305 | 109538928 | N/F | 2325725 | N/F | 384587 | 3128817 | 1503395 | 6509411 | 5212732 | 4718647 | 164895 | 412475 | 452597 | 782284 | 2601336 | 12865665 | 6532921 | 994672 | 24555753 | 2877663 | 22963695 | 2474341 | 23069723 | 360810 | N/F | N/F | 1247833 | N/F | N/F | 3695792 | 659799 | 45540584 | N/F | N/F | 24994117 | 12482335 | 259446 | 40492486 | 5165347 | 43778576 | 123621 | 5456105 | N/F | 5576200 | 59517656 | N/F | 44937897 |  | Unknown
AB-230117-02_CCC-H_2 | N/F | 290696 | 705630 | 708675 | 115648404 | N/F | 2394529 | N/F | 344961 | 2808570 | 1426258 | 5823821 | 4682335 | 4794788 | 135186 | 353676 | 415796 | 687121 | 2292144 | 11308462 | 5777486 | 832964 | 25113502 | 2757946 | 20239051 | 2438598 | 22720073 | 329585 | N/F | N/F | 1089784 | 1845 | N/F | 3933474 | 623594 | 46895626 | 1910 | N/F | 22701593 | 10898005 | 224647 | 40696373 | 4750761 | 38501227 | 77710 | 5425710 | N/F | 4845984 | 56699300 | N/F | 41864943 |  | Unknown
AB-230117-02_CCC-H_3 | N/F | 184226 | 663586 | 646724 | 117127535 | N/F | 2284633 | N/F | 369891 | 1911978 | 1328501 | 5577277 | 4419230 | 4852225 | 87394 | 321704 | 394601 | 681707 | 2126517 | 10854643 | 3462214 | 810462 | 25012779 | 2736437 | 20023795 | 2428572 | 22761914 | 355150 | N/F | N/F | 747444 | N/F | 2337 | 3773025 | 718988 | 48248028 | N/F | 432 | 21440626 | 6484508 | 242485 | 40397962 | 4890617 | 37734645 | 82046 | 5301462 | N/F | 3457511 | 57544884 | N/F | 42088094 |  | Unknown
AB-230117-02_IBC_1 | N/F | 230782 | 560563 | 569511 | 73666088 | N/F | 3281939 | N/F | 360258 | 1529370 | 678693 | 4019409 | 3050938 | 3175699 | 172133 | 531069 | 615180 | 694321 | 3145626 | 15020003 | 6831085 | 1028940 | 33994344 | 3566642 | 26669533 | 3525478 | 30178374 | 343074 | N/F | N/F | 1265968 | N/F | N/F | 2526377 | 316501 | 32587756 | N/F | N/F | 14728164 | 6462707 | 291855 | 27403201 | 3219645 | 25442507 | 192708 | 3228987 | N/F | 2935547 | 30909836 | N/F | 29909284 |  | Unknown
AB-230117-02_LD_1 | N/F | 68371 | 117144 | 7141 | 29129 | N/F | 3197003 | N/F | 101563 | N/F | 30078 | 9431 | 6717 | 64906 | 153391 | 1339623 | 1434346 | 532312 | 3556895 | 16094248 | 5676603 | 514038 | 32960813 | 3685056 | 27273438 | 4080404 | 27937670 | 92502 | N/F | N/F | 969578 | N/F | 1076 | 10774 | 294730 | 219451 | N/F | N/F | 117539 | 29312 | 100593 | 69529 | 2296 | 78651 | 10933 | 1986 | N/F | 5293 | 49425 | 1341 | 25363 |  | Unknown
AB-230117-02_LD_2 | N/F | 44930 | N/F | 7948 | 22047 | N/F | 3993193 | N/F | 326052 | N/F | 8284 | 3367635 | 98029 | 39020108 | 163341 | 481989 | 659851 | 757165 | 3240989 | 16001726 | 6420086 | 926430 | 36730462 | 1661483 | 25510276 | 897677 | 9308318 | 314101 | N/F | N/F | 1139792 | N/F | 427 | 14243 | N/F | 72927 | N/F | 612 | 71803 | N/F | 138662 | N/F | 194050 | 89979 | 15455 | 502 | N/F | 6648 | 67062 | N/F | 64493 |  | Unknown
AB-230117-02_LFB-H_1 | N/F | 64541 | 37800 | 13404 | 1799905 | N/F | 3518377 | N/F | 354509 | 1382273 | 584790 | 4239384 | 3568705 | 3433927 | 177110 | 600556 | 639819 | 625103 | 3203708 | 15779197 | 6609916 | 878562 | 31292462 | 3487758 | 27298231 | 3380199 | 29701401 | 344493 | N/F | N/F | 1214377 | N/F | N/F | 61885 | 345397 | 29431932 | N/F | N/F | 17083687 | 6375282 | 30094960 | 28135449 | 3095157 | 27093950 | 18320672 | 55987 | N/F | 3137992 | 31971867 | N/F | 30976023 |  | Unknown
AB-230117-02_LFB-H_2 | N/F | 202154 | 467443 | 550623 | 65960133 | 1567 | 3045895 | N/F | 348908 | 1429351 | 773186 | 3368892 | 2665792 | 2818829 | 160348 | 447603 | 565583 | 786185 | 2698290 | 13657522 | 6097995 | 917574 | 31196287 | 3049218 | 23586649 | 3069555 | 27131872 | 332746 | N/F | N/F | 1141890 | N/F | N/F | 2458413 | 305169 | 29265704 | N/F | N/F | 13348964 | 5776510 | 176979 | 24412799 | 2962277 | 22252299 | 39068 | 2951827 | N/F | 2610078 | 29727846 | N/F | 25811068 |  | Unknown
AB-230117-02_LFB-H_3 | N/F | 205785 | 425956 | 464244 | 63153442 | N/F | 3211706 | N/F | 346938 | 1389553 | 741120 | 3409855 | 2585846 | 8274632 | 146474 | 418177 | 519769 | 728784 | 2651172 | 13461380 | 6134568 | 892572 | 31838461 | 2822516 | 23679978 | 2536330 | 25011867 | 336965 | N/F | N/F | 1098817 | 637 | 405 | 2541470 | 311284 | 29095176 | N/F | 310 | 13351595 | 5849304 | 185020 | 21905938 | 2862308 | 22469808 | 43689 | 2627944 | N/F | 2638704 | 30888735 | N/F | 26999380 |  | Unknown
AB-230117-02_LFSM-H_1 | N/F | 106459 | 653915 | 891040 | 65925722 | N/F | 2511692 | N/F | 36312 | 496934 | 288538 | 3061909 | 1699067 | 2803874 | 54038 | 582642 | 795963 | 252528 | 1754914 | 6334983 | 2070882 | 212919 | 21592786 | 2657785 | 14564917 | 2837556 | 18266669 | 32194 | 968 | N/F | 393725 | N/F | 942 | 2258626 | 101860 | 22519566 | N/F | N/F | 6573200 | 2031886 | 194132 | 17673841 | 3081657 | 14708479 | 53550 | 2881515 | N/F | 1082646 | 28006220 | N/F | 20600453 |  | Unknown
AB-230117-02_LFSM-H_2 | N/F | 163842 | 353499 | 399778 | 47160721 | N/F | 2493852 | N/F | 381508 | 1406750 | 733658 | 5913930 | 2771860 | 36369609 | 154709 | 363257 | 432832 | 703762 | 2642276 | 13506293 | 6218454 | 883837 | 27198676 | 1672447 | 21889536 | 871340 | 8742054 | 365992 | N/F | N/F | 1111637 | N/F | N/F | 2138552 | 338236 | 27524513 | N/F | 710 | 13128119 | 5875916 | 197645 | 8179450 | 2893418 | 21489183 | 56183 | 1597583 | N/F | 2549931 | 36114997 | N/F | 26850794 |  | Unknown
AB-230117-02_LRB_2 | N/F | 55961 | 30946 | 10821 | 59342 | N/F | 4273045 | N/F | 423362 | N/F | N/F | N/F | N/F | 21794 | 176728 | 727569 | 797187 | 709896 | 3550835 | 17311626 | 6426584 | 1025419 | 38121486 | 3879353 | 30052213 | 4082623 | 34535528 | 404157 | N/F | N/F | 1239312 | N/F | N/F | 15670 | N/F | 38159832 | N/F | N/F | 37212 | N/F | 150348 | 51138 | 3680 | 92068 | 18988 | 3493 | N/F | 7297 | 109276 | N/F | 63530 |  | Unknown
AB-230117-02_LRB_2_20230119014718 | N/F | 56734 | 31701 | 14418 | 40828 | N/F | 4314714 | N/F | 302056 | N/F | N/F | N/F | N/F | 17229 | 176285 | 677491 | 821716 | 579252 | 3186044 | 16097250 | 6983145 | 938344 | 38345817 | 3522166 | 26557316 | 3692244 | 32251309 | 287597 | N/F | N/F | 1250979 | N/F | N/F | 17857 | N/F | 1388027 | N/F | N/F | 55166 | 11706 | 112364 | 38070 | N/F | 82074 | 17409 | 3579 | N/F | 8962 | 95883 | N/F | 67715 |  | Unknown
AB-230117-02_LRB_3 | N/F | 51265 | 28095 | 5237 | 49408 | N/F | 4290546 | N/F | 345033 | N/F | N/F | 30759 | N/F | 1465948 | 162304 | 688457 | 832141 | 608685 | 3316109 | 16259220 | 6519835 | 847660 | 39060283 | 3656214 | 28054916 | 3830699 | 32548149 | 329839 | 2263 | N/F | 1180544 | N/F | N/F | 13301 | N/F | 191061 | N/F | N/F | 55310 | N/F | 140710 | 41604 | 4706 | 84850 | 13717 | 4050 | N/F | 7259 | 85868 | N/F | 51669 |  | Unknown
AB-230117-02_LRB_4 | N/F | 54344 | 34495 | 8715 | 15601 | N/F | 4532799 | N/F | 343234 | N/F | N/F | N/F | N/F | 29077 | 153937 | 687636 | 922012 | 783700 | 3160405 | 16377241 | 6208058 | 914384 | 40439399 | 3664301 | 27216054 | 3910568 | 33682350 | 331054 | N/F | N/F | 1167946 | N/F | 619 | 13268 | N/F | 69039 | N/F | 514 | 19679 | 8406 | 133018 | 29358 | N/F | 70823 | 12142 | 1718 | N/F | 6783 | 72418 | N/F | 46718 |  | Unknown
AB-230117-02_LRB_4_20230119142526 | N/F | 50431 | 30205 | 8966 | 16343 | N/F | 4392534 | 91 | 268763 | 1498 | N/F | N/F | N/F | 23100 | 101270 | 659458 | 799942 | 813269 | 3118516 | 15967860 | 4316238 | 715444 | 39118334 | 3439577 | 27292063 | 3707401 | 32347881 | 255481 | N/F | N/F | 727112 | N/F | 888 | 13530 | N/F | 71640 | N/F | N/F | N/F | N/F | 112853 | 31262 | N/F | 64581 | 7393 | 1877 | N/F | 6911 | 64201 | N/F | 34556 |  | Unknown
AB-230117-02_LRB_5 | N/F | 54361 | 31299 | 10647 | 42549 | N/F | 4453469 | N/F | 347304 | N/F | N/F | 3225 | 2629 | 21627 | 175622 | 648371 | 832802 | 790561 | 3170322 | 15967631 | 6538012 | 912161 | 39665652 | 3524230 | 27224949 | 3827007 | 33056389 | 333530 | N/F | N/F | 1158540 | N/F | 1968 | 16406 | N/F | 94746 | 1630 | 538 | 33302 | N/F | 142625 | 42538 | 1670 | 79172 | 15109 | 1414 | N/F | 6427 | 68715 | N/F | 61327 |  | Unknown
AB-230117-02_LRB_5_20230119195023 | N/F | 56464 | 30565 | 8425 | 11897 | N/F | 4190921 | N/F | 294495 | N/F | 780 | N/F | 1706 | 18524 | 148766 | 611427 | 750065 | 617862 | 2983135 | 15041867 | 5896299 | 785669 | 37000585 | 3226949 | 26102836 | 3543112 | 30826818 | 284660 | N/F | N/F | 1129294 | N/F | N/F | 13784 | N/F | 63331 | N/F | 415 | 30661 | N/F | 121710 | 27751 | 1469 | 68182 | 7490 | 1695 | N/F | 8470 | 61281 | N/F | 29983 |  | Unknown
AB-230117-02_LRB_5_20230119220550 | N/F | 54408 | 53104 | 42883 | 5083783 | N/F | 386841 | 372 | 13216 | N/F | N/F | 399579 | 109528 | 8757986 | 13230 | 32311 | 95278 | 16713 | 145690 | 494601 | 117238 | 25355 | 2805805 | 216278 | 1308521 | 219481 | 1786131 | 12456 | N/F | N/F | 23414 | N/F | 411 | 299800 | N/F | 2412917 | N/F | 491 | 360025 | 23164 | 37271 | 1402788 | 296905 | 1050877 | 13912 | 169002 | N/F | 42328 | 4805192 | N/F | 2912854 |  | Unknown
AB-230117-02_LRB_6 | N/F | 50416 | 33221 | 7862 | N/F | N/F | 3844791 | N/F | 269118 | N/F | 529 | N/F | N/F | 20912 | 166676 | 678945 | 722657 | 491110 | 3338896 | 16246570 | 6487216 | 695913 | 34146175 | 3494909 | 28081546 | 3831279 | 33107061 | 258969 | N/F | N/F | 1091249 | N/F | 1090 | 13093 | N/F | 59812 | N/F | 641 | 23849 | 10461 | 149184 | 19557 | 1532 | 72284 | 15125 | 1187 | N/F | 9062 | 62503 | N/F | 92673 |  | Unknown
AB-230117-02_LRB_6_20230120033053 | N/F | 47992 | 23756 | 7840 | 10749 | N/F | 4041035 | N/F | 302395 | N/F | 1800 | 1129 | N/F | 20292 | 147615 | 575148 | 666458 | 526049 | 2312291 | 11270072 | 5772719 | 781754 | 36778822 | 3150299 | 23284126 | 3361080 | 29128442 | 288321 | N/F | N/F | 764375 | N/F | 619 | 13082 | N/F | 65001 | 1013 | 289 | 39044 | N/F | 108033 | 19685 | 3813 | 53686 | 10681 | 1146 | N/F | 8132 | 58746 | N/F | 80134 |  | Unknown
AB-230117-02_LRB-1 | N/F | 58960 | 47604 | 12235 | 1957729 | N/F | 4192864 | N/F | 400447 | 9154 | N/F | 97712 | 57814 | 103376 | 182974 | 816006 | 842732 | 881418 | 3663879 | 17592203 | 7132049 | 1052182 | 37221036 | 4031444 | 30802350 | 4205808 | 35806446 | 377315 | N/F | N/F | 1295992 | N/F | N/F | 77019 | N/F | 851531 | N/F | N/F | 189075 | N/F | 143615 | 740342 | 89295 | 567140 | 20595 | 116280 | N/F | 30481 | 1150536 | N/F | 907810 |  | Unknown
AB-230117-02_QCS_1 | N/F | 251510 | 624834 | 746610 | 74162754 | N/F | 3780004 | N/F | 391464 | 1793397 | 876621 | 4480499 | 3426289 | 3142146 | 167272 | 536181 | 680083 | 761275 | 3056412 | 14912965 | 6912932 | 1012352 | 34623193 | 3224640 | 26340525 | 2988900 | 27591477 | 368604 | N/F | N/F | 1279834 | N/F | N/F | 103721 | 435977 | 36327856 | N/F | N/F | 16714995 | 7299392 | 5181116 | 26847701 | 315261 | 27798052 | 2377420 | 94137 | N/F | 3424892 | 4705810 | N/F | 35946324 |  | Unknown
WO-23003-01-A-1 | N/F | 65796 | 107908 | 5466 | 299302 | N/F | 3536696 | 196 | 75724 | N/F | 21768 | 15634 | 6652 | 45987 | 74836 | 741764 | 1288195 | 336384 | 1862793 | 7333015 | 2617430 | 354799 | 32301641 | 3661954 | 18071944 | 3844692 | 25697042 | 64238 | N/F | N/F | 497588 | N/F | N/F | 17383 | 152786 | 234212 | N/F | N/F | 48368 | 14580 | 157601 | 142326 | 15416 | 115520 | 50915 | 16797 | N/F | 7144 | 269861 | N/F | 127646 |  | Unknown
WO-23003-02-A-1 | N/F | 57507 | 109341 | 3563 | 195902 | N/F | 3585473 | N/F | 44961 | 8119 | 17379 | 14261 | 7300 | 45543 | 71662 | 992753 | 1519457 | 259309 | 2189835 | 7792307 | 2457350 | 243464 | 31419146 | 3730910 | 19370109 | 4064659 | 26032384 | 39407 | N/F | N/F | 447622 | N/F | N/F | 16779 | 100099 | 218265 | N/F | N/F | 175529 | N/F | 170343 | 110571 | 9222 | 108590 | 32218 | 13210 | N/F | 4550 | 170801 | N/F | 101520 |  | Unknown
WO-23003-03-A-1 | N/F | 69589 | 121081 | 5736 | 158435 | N/F | 3133212 | 209 | 108334 | N/F | 25958 | 11824 | 9613 | 38303 | 93068 | 1188539 | 1262236 | 435784 | 2930056 | 11311334 | 3443525 | 452675 | 28375547 | 3803107 | 24292510 | 4041821 | 28062070 | 96340 | N/F | N/F | 662710 | N/F | N/F | 13724 | 247185 | 157865 | N/F | N/F | 41113 | 14961 | 146593 | 105724 | 7756 | 120842 | 24973 | 10802 | N/F | 6671 | 133604 | N/F | 79790 |  | Unknown
WO-23003-04-A-1 | N/F | 43812 | 75305 | 17617 | 128345 | N/F | 4254250 | 59 | 11107 | N/F | N/F | 7713 | N/F | 44631 | 35619 | 512831 | 1167231 | 170395 | 1635815 | 5611845 | 1372283 | 91096 | 34514886 | 3756604 | 16332245 | 3964193 | 26011891 | 8462 | N/F | N/F | 231614 | N/F | N/F | 16000 | 20365 | 141643 | N/F | 3254 | 45874 | N/F | 151759 | 81655 | 6751 | 74199 | 27864 | 8000 | N/F | N/F | 133785 | N/F | 115397 |  | Unknown
WO-23003-05-A-1 | N/F | 50987 | 89994 | 6603 | 117141 | N/F | 4486300 | 73 | 21630 | N/F | N/F | 9579 | N/F | 38253 | 39714 | 509584 | 1154501 | 207211 | 1528397 | 5613681 | 1572217 | 131664 | 33872567 | 3799640 | 15172803 | 3859187 | 24593314 | 16128 | N/F | N/F | 279961 | N/F | N/F | 15430 | 45680 | 184023 | N/F | N/F | N/F | N/F | 165355 | 73980 | 6302 | 62779 | 25266 | 6052 | N/F | N/F | 137509 | N/F | 81909 |  | Unknown
WO-23003-06-A-1 | N/F | 43741 | 62735 | 7431 | 96830 | N/F | 4336593 | 69 | 14109 | N/F | N/F | N/F | N/F | 169716 | 32115 | 378761 | 933750 | 120550 | 1284902 | 4746972 | 1375153 | 88385 | 32033462 | 3469485 | 13088314 | 3521650 | 22472696 | 10215 | N/F | N/F | 259336 | N/F | N/F | 18337 | 26178 | 123003 | N/F | N/F | 147410 | N/F | 165982 | 70451 | 4882 | 52937 | 27343 | 6165 | N/F | N/F | 127715 | N/F | 100027 |  | Unknown
WO-23003-07-A-1 | N/F | 48295 | 34614 | 2721 | 63683 | N/F | 2520477 | N/F | 354715 | N/F | N/F | 5652 | 5192 | 21800 | 140279 | 749409 | 338303 | 816905 | 3383417 | 16427420 | 5638969 | 882621 | 11240896 | 3790390 | 29457912 | 3844275 | 33917331 | 330783 | N/F | N/F | 1141879 | N/F | N/F | 12872 | N/F | 214903398 | N/F | 1357 | 80700 | N/F | 140611 | 267639 | N/F | 170804 | 18222 | 5815 | N/F | 6296 | 137247 | N/F | 141251 |  | Unknown
WO-23003-08-A-1 | N/F | 47927 | 33300 | 2770 | 64409 | N/F | 2428873 | 103 | 360537 | N/F | N/F | 3671 | N/F | 11942 | 157557 | 688493 | 334223 | 763057 | 3440710 | 16639864 | 6072207 | 881338 | 10905532 | 3705502 | 29295926 | 3787128 | 33862932 | 342805 | N/F | N/F | 1182503 | N/F | N/F | 12715 | N/F | 230428836 | N/F | N/F | 45691 | N/F | 141393 | 267119 | N/F | 164238 | 18160 | 4882 | N/F | 6986 | 126941 | N/F | 145666 |  | Unknown
WO-23003-09-A-1 | N/F | 49083 | 33527 | 2323 | 50138 | N/F | 2549553 | N/F | 303536 | N/F | N/F | N/F | N/F | 15050 | 94552 | 719347 | 328498 | 605447 | 3398277 | 16488995 | 4196726 | 893275 | 11603574 | 3764838 | 29795410 | 3893301 | 34190517 | 285021 | N/F | N/F | 1180678 | N/F | N/F | 13868 | N/F | 223715045 | N/F | N/F | N/F | N/F | 136364 | 248559 | N/F | 179090 | 17019 | 5458 | N/F | 7220 | 120479 | N/F | 113077 |  | Unknown
WO-23003-10-A-1 | N/F | 50471 | 34306 | 8061 | 41350 | N/F | 2340514 | N/F | 323447 | N/F | N/F | 3442 | N/F | 10838 | 168284 | 696669 | 345999 | 677764 | 3351815 | 16583246 | 6539975 | 953275 | 10743050 | 3753494 | 29682382 | 3876799 | 34149551 | 307870 | N/F | N/F | 1147085 | N/F | 1824 | 12702 | N/F | 233506924 | N/F | N/F | 45690 | N/F | 120048 | 268489 | N/F | 163514 | 14613 | 3917 | N/F | 6689 | 115593 | N/F | 121748 |  | Unknown
WO-23003-11-A-1 | N/F | 42616 | 32074 | 3365 | 25439 | 559 | 2369921 | N/F | 311200 | 2024 | N/F | N/F | N/F | 11571 | 164966 | 687477 | 324313 | 706630 | 3316637 | 16000175 | 6441159 | 919025 | 10981126 | 3563273 | 28043019 | 3669322 | 31943220 | 295229 | N/F | N/F | 1151204 | N/F | 668 | 13396 | N/F | 223283789 | N/F | N/F | 53477 | N/F | 132337 | 238939 | N/F | 141948 | 17057 | 3520 | N/F | 8289 | 100480 | N/F | 111030 |  | Unknown
WO-23003-12-A-1 | N/F | 42091 | 28967 | 3737 | 29679 | N/F | 2294500 | N/F | 301075 | N/F | N/F | N/F | N/F | 14806 | 150148 | 673725 | 305797 | 626626 | 3308669 | 15773920 | 5594863 | 857055 | 11063233 | 3653343 | 28505799 | 3817697 | 32764869 | 286308 | N/F | N/F | 1048456 | N/F | N/F | 13317 | N/F | 224464407 | N/F | N/F | 48064 | N/F | 133109 | 244732 | N/F | 158345 | 14890 | 1804 | N/F | 8666 | 89670 | N/F | 96990 |  | Unknown
WO-23003-13-A-1 | N/F | 43638 | 32734 | 7334 | 35447 | 1393 | 2163174 | N/F | 363035 | 5604 | N/F | N/F | N/F | 12109 | 139041 | 673442 | 329550 | 671594 | 3307942 | 16148977 | 6350247 | 952339 | 10872151 | 3625980 | 28549100 | 3731629 | 33171763 | 343592 | N/F | N/F | 1136431 | N/F | 539 | 15762 | N/F | 226866512 | N/F | N/F | 46799 | N/F | 140940 | 255481 | N/F | 158734 | 14500 | 2263 | N/F | 7965 | 90516 | 1641 | 98975 |  | Unknown
WO-23003-14-A-1 | N/F | 43518 | 67742 | 2645 | 43181 | N/F | 3611110 | N/F | 148587 | N/F | 9669455 | 4795 | N/F | 25927 | 85808 | 755620 | 1037845 | 328067 | 2563857 | 10024593 | 3410265 | 536991 | 28911726 | 3648749 | 22823195 | 4013959 | 29279951 | 133833 | N/F | N/F | 642898 | N/F | 543 | 10232 | N/F | 27441137 | N/F | N/F | 243519 | N/F | 148426 | 55545 | N/F | 77732 | 17219 | 4064 | N/F | 4945 | 72695 | 1202 | 47053 |  | Unknown
WO-23003-15-A-1 | N/F | 50488 | 77536 | 9035 | 51950 | N/F | 4284119 | 11 | 158463 | N/F | 7477856 | N/F | 4705 | 31336 | 51210 | 821682 | 1209287 | 260209 | 2694573 | 10253063 | 1935784 | 472421 | 35402869 | 3935916 | 23280241 | 4034252 | 28846629 | 145471 | N/F | N/F | 424411 | N/F | 1537 | 9214 | N/F | 1347931 | N/F | 1059 | 239356 | 7597 | 130395 | 52657 | 3613 | 76144 | 15357 | 5852 | N/F | 5100 | 79576 | N/F | 47214 |  | Unknown
WO-23003-16-A-1 | N/F | 49171 | 57691 | 4049 | 41181 | 866 | 4053882 | N/F | 149121 | 2814 | 10418045 | N/F | 4340 | 29156 | 95565 | 688850 | 963586 | 355383 | 2613035 | 10449870 | 3583419 | 582988 | 36763017 | 3664319 | 23060882 | 3988486 | 30758028 | 137413 | N/F | N/F | 667889 | N/F | 1090 | 9336 | N/F | 738697 | N/F | N/F | 237363 | 13172 | 137599 | 58558 | N/F | 71212 | 16602 | 2702 | N/F | 6390 | 76879 | N/F | 38695 |  | Unknown
WO-23003-17-A-1 | N/F | 44935 | 53768 | 3764 | 47168 | N/F | 4181060 | N/F | 140211 | 3227 | 7372536 | 4367 | N/F | 30895 | 85972 | 600102 | 985196 | 255377 | 2527701 | 9952404 | 3321449 | 557827 | 36557273 | 3707816 | 22483101 | 3934598 | 29735618 | 127206 | 1770 | N/F | 606825 | N/F | 876 | 15804 | N/F | 425209 | N/F | N/F | 349739 | 13324 | 140809 | 51665 | 2420 | 82250 | 15119 | 2822 | N/F | 4787 | 78512 | N/F | 55923 |  | Unknown
WO-23003-18-A-1 | N/F | 44412 | 50524 | 5463 | 66113 | N/F | 3963927 | 33 | 131267 | 1124 | 9216444 | 6421 | 4810 | 28392 | 79378 | 603041 | 903702 | 314019 | 2335472 | 9200001 | 3151576 | 491141 | 33091451 | 3640022 | 21654419 | 3888248 | 29123978 | 121050 | N/F | N/F | 593985 | N/F | N/F | 17236 | N/F | 311487 | N/F | 988 | 211979 | 14054 | 146903 | 55797 | 4165 | 91159 | 16662 | 3347 | N/F | 5694 | 80714 | N/F | 56478 |  | Unknown
WO-23003-19-A-1 | N/F | 43746 | 43498 | 8165 | 59804 | N/F | 3750251 | 44 | 152910 | 1452 | 9987341 | 4420 | N/F | 23465 | 68440 | 581157 | 881655 | 342972 | 2232321 | 8418990 | 2525868 | 515345 | 33150001 | 3153772 | 20665717 | 3207515 | 23736871 | 143465 | N/F | N/F | 398253 | N/F | 521 | 15694 | N/F | 263279 | N/F | N/F | 188438 | N/F | 145125 | 52226 | 2809 | 86068 | 11573 | 3874 | N/F | 5496 | 81834 | N/F | 57224 |  | Unknown
WO-23003-20-A-1 | N/F | 43971 | N/F | 8485 | 63665 | N/F | 4135379 | N/F | 340495 | 1482 | N/F | 700120 | 20829 | 17373617 | 88669 | 587467 | 788053 | 675868 | 2914086 | 13904021 | 3341284 | 867519 | 37950471 | 2694033 | 25587257 | 2077513 | 20617946 | 326795 | N/F | N/F | 717575 | N/F | 577 | 14368 | N/F | 261759 | N/F | 704 | 60458 | N/F | 133315 | 41672 | 49465 | 86517 | 13812 | 3077 | N/F | 9211 | 79742 | N/F | 46286 |  | Unknown
WO-23003-21-A-1 | N/F | 47950 | 25410 | 26859 | 1102457 | N/F | 4317833 | 16 | 297697 | 7756 | N/F | 717244 | 40440 | 16610191 | 161171 | 630119 | 702089 | 697557 | 3199409 | 15733599 | 6276252 | 848524 | 36473897 | 2657982 | 26868769 | 2063733 | 21316953 | 280253 | N/F | N/F | 1114067 | N/F | 1641 | 65712 | N/F | 564430 | N/F | N/F | 133024 | 32040 | 129744 | 240900 | 98327 | 275025 | 15171 | 38417 | N/F | 13118 | 1136749 | N/F | 628358 |  | Unknown
WO-23003-22-A-1 | N/F | 49355 | N/F | 3305 | 158329 | N/F | 4272559 | N/F | 352440 | N/F | N/F | 674958 | 26295 | 17180750 | 160768 | 608605 | 765125 | 542265 | 3311045 | 15852635 | 6347969 | 779924 | 39365969 | 2811605 | 27368674 | 2181366 | 22130353 | 330967 | N/F | N/F | 1142107 | N/F | 708 | 20458 | N/F | 159650 | N/F | 507 | 162961 | N/F | 139097 | 51909 | 48687 | 111781 | 13196 | 6978 | N/F | 6193 | 151903 | N/F | 123218 |  | Unknown
WO-23003-23-A-1 | N/F | 45767 | N/F | 3416 | 89441 | N/F | 4266222 | N/F | 297433 | N/F | N/F | 758815 | 24961 | 18257538 | 149797 | 606454 | 747767 | 771498 | 3325718 | 15908187 | 5864427 | 927045 | 38399615 | 2745882 | 27119147 | 2083832 | 21230023 | 283492 | N/F | N/F | 1078946 | N/F | N/F | 19889 | N/F | 137975 | N/F | N/F | 66073 | 16595 | 126890 | 38347 | 52148 | 94379 | 7793 | 5156 | N/F | 7196 | 107058 | N/F | 92865 |  | Unknown
WO-23003-24-A-1 | N/F | 49612 | N/F | 11700 | 138693 | N/F | 4339451 | N/F | 335674 | 3332 | N/F | 705692 | 25732 | 16942544 | 141269 | 618268 | 781383 | 614004 | 3195824 | 15654005 | 5411334 | 742559 | 39298517 | 2745079 | 27217484 | 2133264 | 21992191 | 324427 | N/F | N/F | 1056532 | N/F | N/F | 21819 | N/F | 178672 | N/F | N/F | 110676 | 23606 | 133624 | 64395 | 52225 | 138106 | 6586 | 2727 | N/F | 9794 | 110755 | N/F | 100418 |  | Unknown
WO-23003-25-A-1 | N/F | 48362 | N/F | 5958 | 57738 | 1436 | 4136842 | N/F | 341051 | 5694 | N/F | 747746 | 24402 | 17817531 | 138646 | 615971 | 739224 | 666989 | 3109176 | 15498868 | 5929044 | 839167 | 38685040 | 2633953 | 27087305 | 2030032 | 21199897 | 329385 | N/F | N/F | 1115481 | N/F | 1427 | 15838 | N/F | 112373 | N/F | 936 | 69816 | 14295 | 129043 | 33550 | 50724 | 90408 | 6455 | 3987 | N/F | 8445 | 91358 | N/F | 50022 |  | Unknown
WO-23003-26-A-1 | N/F | 47043 | 75395 | 12474 | 35016 | N/F | 4151047 | 11 | 77551 | 2696 | N/F | 24618 | N/F | 1302804 | 72995 | 628312 | 1074200 | 284220 | 1958657 | 7380959 | 2503721 | 281548 | 35019415 | 3291704 | 19606438 | 3615594 | 27032188 | 66085 | N/F | N/F | 428529 | N/F | 797 | 14839 | N/F | 86557 | N/F | 1728 | 28615 | N/F | 148468 | 37364 | 6487 | 49233 | 16982 | 2233 | N/F | 5535 | 88954 | N/F | 47911 |  | Unknown
WO-23003-27-A-1 | N/F | 45955 | 80330 | 2765 | 22161 | 1315 | 3696884 | N/F | 95593 | N/F | 7669 | 5138 | N/F | 63738 | 80555 | 702314 | 1031436 | 387147 | 2251039 | 8737588 | 3127946 | 439822 | 30702952 | 3268564 | 20932697 | 3558184 | 26742961 | 85744 | N/F | N/F | 567057 | N/F | 811 | 15020 | N/F | 71634 | N/F | N/F | 25997 | N/F | 153972 | 33304 | N/F | 60311 | 18927 | 2408 | N/F | 4991 | 74753 | N/F | 36075 |  | Unknown
WO-23003-28-A-1 | N/F | 47319 | 72939 | 7895 | 28283 | 831 | 4388725 | N/F | 105965 | N/F | N/F | N/F | N/F | 39357 | 72749 | 625978 | 1095247 | 253068 | 2098475 | 8486416 | 2979528 | 359506 | 35387265 | 3484410 | 20636117 | 3857353 | 28505227 | 96008 | N/F | N/F | 537621 | N/F | 536 | 14078 | N/F | 79545 | N/F | N/F | N/F | N/F | 167136 | 40031 | N/F | 52291 | 16838 | 2658 | N/F | 3941 | 77592 | N/F | 50555 |  | Unknown
WO-23003-29-A-1 | N/F | 48192 | 74390 | 7003 | 25405 | N/F | 4339688 | N/F | 93626 | N/F | 4268 | 3534 | N/F | 29226 | 70345 | 604662 | 1058647 | 369494 | 2128875 | 8489643 | 3070656 | 434821 | 37248270 | 3560874 | 21057908 | 3906672 | 29615862 | 85066 | N/F | N/F | 567838 | N/F | 842 | 15242 | N/F | 76769 | N/F | 1277 | N/F | N/F | 162793 | 31516 | N/F | 45629 | 19572 | 3470 | N/F | N/F | 82294 | N/F | 38834 |  | Unknown
WO-23003-30-A-1 | N/F | 46986 | 95805 | 6666 | 18446 | N/F | 3488537 | N/F | 78313 | N/F | N/F | 4006 | N/F | 27904 | 76623 | 718493 | 1351081 | 341668 | 1990198 | 7796759 | 2835254 | 345601 | 31784162 | 3347975 | 19285062 | 3740955 | 26277998 | 70563 | N/F | N/F | 501333 | N/F | 1650 | 13568 | N/F | 77662 | N/F | 2077 | 32004 | N/F | 155475 | 33648 | N/F | 43204 | 20094 | 2066 | N/F | 6217 | 58636 | N/F | 31747 |  | Unknown
WO-23003-31-A-1 | N/F | 41051 | 62020 | 8497 | 23951 | N/F | 4007131 | N/F | 64876 | N/F | N/F | 3812 | N/F | 27304 | 40293 | 499043 | 982381 | 303279 | 1609450 | 6478285 | 1478468 | 342226 | 34027840 | 3276460 | 17080853 | 3446208 | 26018884 | 55125 | N/F | N/F | 321171 | N/F | N/F | 11182 | N/F | 161815 | N/F | 1555 | N/F | N/F | 152400 | 54723 | N/F | 72921 | 18467 | 1555 | N/F | N/F | 72433 | N/F | 42700 |  | Unknown
WO-23003-32-A-1 | N/F | 43566 | 62290 | 9533 | N/F | N/F | 4149553 | N/F | 48926 | 4129 | N/F | N/F | N/F | 19716 | 35717 | 515555 | 924501 | 286371 | 1665290 | 6449193 | 1450216 | 279869 | 34280801 | 3299448 | 17533167 | 3615937 | 26624290 | 40994 | N/F | N/F | 308123 | N/F | 468 | 11859 | N/F | 67106 | N/F | 1529 | N/F | N/F | 159263 | 31107 | N/F | 47961 | 16157 | 1704 | N/F | 3997 | 77397 | N/F | 37681 |  | Unknown
WO-23003-33-A-1 | N/F | 50652 | 68198 | 6301 | 18624 | N/F | 3962194 | N/F | 101078 | N/F | 12845 | N/F | N/F | 28493 | 72932 | 581987 | 1038054 | 258808 | 1698006 | 6165550 | 2886147 | 402420 | 35053863 | 3385576 | 17870923 | 3737407 | 27006693 | 91680 | N/F | N/F | 298250 | N/F | 851 | 14531 | N/F | 118380 | N/F | 2116 | N/F | 15754 | 156572 | 49308 | 1149 | 45379 | 17591 | 1013 | N/F | 4910 | 70366 | N/F | 40658 |  | Unknown
WO-23003-34-A-1 | N/F | 55701 | 73332 | 2272 | 21697 | 976 | 4398414 | N/F | 88734 | N/F | 19423 | N/F | 3349 | 28882 | 75409 | 601659 | 1074942 | 364625 | 2128378 | 8202520 | 2851048 | 392547 | 35959334 | 3563049 | 19697415 | 3886488 | 28660834 | 80214 | N/F | N/F | 512736 | N/F | 343 | 15347 | N/F | 119557 | N/F | 1563 | N/F | 13650 | 158794 | 56686 | 2955 | 62662 | 16905 | 2022 | N/F | 5165 | 65209 | 782 | 44286 |  | Unknown
WO-23003-35-A-1 | N/F | 56116 | 64529 | 8460 | 35736 | N/F | 4094428 | N/F | 93227 | N/F | 15314 | 5628 | 2872 | 33852 | 71425 | 613725 | 1069268 | 327495 | 2086886 | 8144241 | 2912434 | 427565 | 34775503 | 3470249 | 19886142 | 3905300 | 27913901 | 82771 | N/F | N/F | 522159 | N/F | 1048 | 13443 | N/F | 123600 | N/F | 586 | N/F | 15398 | 139678 | 54257 | 1031 | 63969 | 18788 | 1952 | N/F | 3700 | 62761 | N/F | 47520 |  | Unknown
WO-23003-36-A-1 | N/F | 55867 | 76179 | 11548 | 53867 | N/F | 3952433 | N/F | 99796 | 3295 | 14467 | 4361 | N/F | 37813 | 77111 | 616474 | 979009 | 316645 | 2130115 | 8054042 | 2949618 | 405961 | 35097046 | 3437932 | 20278508 | 3891160 | 28228487 | 89538 | N/F | N/F | 531139 | N/F | 1108 | 15884 | N/F | 133835 | N/F | 2380 | N/F | 12572 | 154499 | 64466 | 2980 | 67750 | 14246 | 2018 | N/F | 4612 | 66504 | N/F | 52595 |  | Unknown
WO-23003-37-A-1 | N/F | 57811 | 75071 | 10092 | 16721 | N/F | 4063769 | N/F | 85279 | N/F | 14330 | 2194 | 3103 | 31464 | 69150 | 642254 | 1041707 | 310697 | 2033498 | 8257018 | 2863062 | 409691 | 35420039 | 3390190 | 20464234 | 3927748 | 28772017 | 74620 | N/F | N/F | 539847 | N/F | 789 | 10336 | N/F | 102752 | N/F | 1257 | N/F | 13340 | 156566 | 50557 | N/F | 50862 | 9302 | 1839 | N/F | 4378 | 65102 | N/F | 60302 |  | Unknown
WO-23003-38-A-1 | N/F | 54134 | 73991 | 11735 | N/F | N/F | 4555539 | N/F | 79803 | N/F | 20974 | N/F | N/F | 32675 | 75176 | 628734 | 1043859 | 391407 | 2210565 | 8389168 | 3071048 | 359730 | 36217472 | 3516625 | 20761933 | 3900660 | 28472937 | 69658 | N/F | N/F | 561070 | N/F | 923 | 14579 | N/F | 116001 | N/F | 3613 | N/F | 13364 | 150389 | 45000 | 943 | 60789 | 16146 | 1285 | N/F | 6822 | 61177 | N/F | 34317 |  | Unknown
WO-23003-39-A-1 | N/F | 39109 | 59357 | 7161 | 7334 | N/F | 4414300 | N/F | 41010 | 1169 | N/F | N/F | N/F | 16466 | 46005 | 454739 | 947926 | 240789 | 1597783 | 6058998 | 2020304 | 192611 | 34449016 | 3348800 | 16282203 | 3623185 | 26067846 | 34161 | N/F | N/F | 370251 | N/F | 699 | 11671 | 23418 | 54740 | N/F | 1469 | 36041 | N/F | 156860 | 25182 | N/F | 45159 | 15043 | 1228 | N/F | 6802 | 58453 | 1767 | 54284 |  | Unknown
WO-23003-40-A-1 | N/F | 46421 | 76078 | 7864 | 10320 | 1447 | 4242833 | N/F | 84392 | 4218 | N/F | 3079 | 1056 | 18256 | 74118 | 614001 | 1055597 | 345027 | 1961904 | 7642129 | 2778990 | 363812 | 35276162 | 3457853 | 19481074 | 3727989 | 27583554 | 74297 | N/F | N/F | 501418 | N/F | N/F | 9656 | 53973 | 92167 | N/F | 1352 | 33817 | N/F | 151640 | 44556 | 2396 | 47023 | 15521 | 1662 | N/F | 7943 | 58565 | N/F | 31095 |  | Unknown
WO-23003-41-A-1 | N/F | 44324 | 64123 | 4275 | 201107 | N/F | 4126951 | 43 | 68046 | 1368 | 1688 | 11675 | 4642 | 215971 | 62536 | 553192 | 986747 | 227095 | 1921842 | 7275314 | 2619758 | 199945 | 34946438 | 3265805 | 19253007 | 3646302 | 27213616 | 56781 | N/F | N/F | 380418 | N/F | 529 | 20667 | 46579 | 118567 | N/F | 1950 | N/F | N/F | 135571 | 84415 | 11095 | 78678 | 17042 | 11585 | N/F | 3398 | 181102 | N/F | 158606 |  | Unknown
WO-23003-42-A-1 | N/F | 45763 | 67744 | 6776 | 83621 | N/F | 4256173 | N/F | 52936 | N/F | 611 | 8999 | 1647 | 89479 | 56900 | 554002 | 1028146 | 230489 | 1909302 | 7420564 | 2180358 | 170913 | 35680754 | 3365511 | 18875295 | 3761177 | 27039117 | 44191 | N/F | N/F | 458326 | N/F | 1645 | 16244 | 32743 | 79834 | N/F | 3424 | 41632 | N/F | 154161 | 46633 | 4827 | 57263 | 17726 | 4123 | N/F | 5221 | 120821 | N/F | 66154 |  | Unknown
WO-23003-43-A-1 | N/F | 45564 | 86141 | 12144 | 46197 | 800 | 3676012 | N/F | 68772 | N/F | N/F | 3430 | N/F | 52639 | 63584 | 697835 | 1274802 | 251081 | 1777769 | 6778221 | 2502751 | 283729 | 31060291 | 3209005 | 17773173 | 3622146 | 23914508 | 60196 | N/F | N/F | 380678 | N/F | 622 | 15435 | 36081 | 63126 | N/F | 775 | 33571 | N/F | 154736 | 35726 | 3673 | 45018 | 17786 | 4738 | N/F | 2783 | 82558 | N/F | 43371 |  | Unknown
WO-23003-44-A-1 | N/F | 42943 | 73876 | 8519 | 35579 | N/F | 4308759 | N/F | 83828 | 3579 | 1186 | 10949 | 1664 | 43808 | 41401 | 616172 | 1072920 | 220833 | 1963128 | 7289874 | 1667776 | 297415 | 34820074 | 3462454 | 19266918 | 3978394 | 28279691 | 74464 | N/F | N/F | 369180 | N/F | 527 | 17021 | 44078 | 63836 | N/F | 1439 | 44996 | N/F | 163609 | 41841 | 7854 | 48880 | 18136 | 3386 | N/F | 5229 | 81794 | N/F | 47572 |  | Unknown
WO-23003-45-A-1 | N/F | 48009 | 71859 | 7964 | 27283 | N/F | 4114888 | N/F | 95556 | N/F | 1101 | 3653 | 1196 | 32178 | 45960 | 636114 | 1129784 | 249583 | 1938531 | 7549690 | 1785092 | 314632 | 34854193 | 3299750 | 19560443 | 3725257 | 27438800 | 83660 | N/F | N/F | 365309 | N/F | 2221 | 13036 | 57078 | 65361 | N/F | 1373 | 39919 | N/F | 157638 | 34562 | 2599 | 54190 | 15607 | 3008 | N/F | 6995 | 70265 | N/F | 51484 |  | Unknown
WO-23003-46-A-1 | N/F | 42412 | 67452 | 7471 | 21039 | N/F | 4092494 | N/F | 52334 | N/F | 699 | 10038 | 1949 | 30785 | 59513 | 512861 | 1035197 | 285681 | 1376805 | 5202503 | 2429265 | 269594 | 33672580 | 3362456 | 15617674 | 3768171 | 25851385 | 43989 | 2153 | N/F | 349917 | N/F | 577 | 13150 | 27788 | 76662 | N/F | 1268 | 79381 | N/F | 150209 | 39877 | 5044 | 37553 | 16410 | 2272 | N/F | 5573 | 65802 | N/F | 78259 |  | Unknown
WO-23003-47-A-1 | N/F | 43149 | 84231 | 8732 | 16995 | N/F | 3519822 | N/F | 12270 | 1935 | N/F | 8997 | N/F | 28747 | 39976 | 537631 | 1240942 | 145000 | 1394350 | 5164156 | 1573853 | 105018 | 27574645 | 3071519 | 13784367 | 3462368 | 21632529 | 9210 | N/F | N/F | 272427 | N/F | 836 | 9231 | 13594 | 44933 | N/F | 3441 | N/F | N/F | 151614 | 26142 | 7727 | 40113 | 18203 | 1935 | N/F | 5943 | 57540 | N/F | 57896 |  | Unknown
WO-23003-48-A-1 | N/F | 44490 | 68944 | 16408 | 14675 | N/F | 4057883 | N/F | 77761 | 1309 | N/F | 4413 | N/F | 23642 | 71458 | 552834 | 1067960 | 211177 | 1847476 | 7218451 | 2722086 | 279897 | 35375815 | 3331087 | 18954867 | 3623848 | 27459965 | 66177 | N/F | N/F | 508707 | N/F | 608 | 13702 | 92251 | 60539 | N/F | 2410 | 39450 | N/F | 155965 | 34194 | 2232 | 60940 | 17795 | 2121 | N/F | 4873 | 65038 | N/F | 32268 |  | Unknown
WO-23003-49-A-1 | N/F | 45440 | 69846 | 11764 | 12246 | N/F | 4365610 | N/F | 80239 | N/F | 604 | 4822 | N/F | 24863 | 61597 | 562730 | 1027755 | 262748 | 2007731 | 7804912 | 2492440 | 293788 | 35076715 | 3497404 | 19278484 | 3894264 | 27906762 | 69942 | N/F | N/F | 508233 | N/F | 853 | 18972 | 96428 | 58067 | N/F | 2735 | 42395 | N/F | 160287 | 37108 | 5652 | 92464 | 17459 | 1482 | N/F | 7923 | 62868 | N/F | 34595 |  | Unknown
WO-23003-50-A-1 | N/F | 41272 | 49482 | 8676 | 9397 | N/F | 3780408 | N/F | 12436 | 1648 | 717 | 6687 | N/F | 20453 | 30948 | 311454 | 903869 | 99474 | 976903 | 3450546 | 1119243 | 74449 | 31442073 | 2799112 | 11365017 | 3133749 | 21242150 | 10422 | N/F | N/F | 129790 | N/F | 358 | 12004 | 12271 | 42925 | N/F | 3012 | 123558 | N/F | 172006 | 28891 | 2526 | N/F | 16268 | 1490 | N/F | 4219 | 61006 | N/F | 71844 |  | Unknown
WO-23003-51-A-1 | N/F | 37022 | 70003 | 6243 | 10007 | N/F | 4260043 | N/F | 46503 | N/F | 2625 | 4114 | 1035 | 23900 | 33086 | 533822 | 999562 | 214580 | 1641215 | 5885332 | 1164454 | 198381 | 35034929 | 3251221 | 17883015 | 3685757 | 27104782 | 40187 | N/F | N/F | 241504 | N/F | 1126 | 12868 | 111026 | 53206 | N/F | 2855 | N/F | N/F | 144794 | 28919 | 3069 | 47213 | 16239 | 1311 | N/F | 7262 | 55768 | 1989 | 31753 |  | Unknown
WO-23003-52-A-1 | N/F | 48599 | 89469 | 7591 | N/F | N/F | 3840625 | N/F | 95963 | 2574 | 782 | 4014 | N/F | 22862 | 49516 | 826098 | 1304489 | 259214 | 1771089 | 6694704 | 2021364 | 315274 | 32019259 | 3373915 | 18689352 | 3788554 | 25366446 | 86481 | N/F | N/F | 283266 | N/F | 702 | 12335 | 197869 | 39182 | N/F | 2809 | 54895 | N/F | 153150 | 29211 | 3924 | 41853 | 16293 | 1583 | N/F | 5921 | 56386 | N/F | 30337 |  | Unknown
WO-23003-53-A-1 | N/F | 57614 | 82439 | 7102 | 12684 | N/F | 3960738 | N/F | 99614 | N/F | 2890 | 3396 | 1949 | 22735 | 84926 | 779655 | 1110517 | 265947 | 2580651 | 9918375 | 3350863 | 363825 | 36530098 | 3546546 | 23159173 | 4157094 | 30016345 | 91104 | N/F | N/F | 600538 | N/F | 1325 | 14017 | 211358 | 79214 | 1261 | 1669 | 53650 | N/F | 150534 | 37246 | 3932 | 58841 | 15715 | 1009 | N/F | 6091 | 54718 | N/F | 31713 |  | Unknown
WO-23003-54-A-1 | N/F | 56719 | 66687 | 7620 | 15714 | N/F | 4157633 | N/F | 58787 | N/F | 10113 | 9878 | 3252 | 32759 | 65779 | 478726 | 990564 | 203735 | 1747890 | 6889137 | 2518745 | 201832 | 34325965 | 3304252 | 17113731 | 3725653 | 26555731 | 48294 | N/F | N/F | 472734 | N/F | 297 | 14380 | 119279 | 106312 | N/F | 2580 | 44382 | N/F | 155584 | 50756 | 5470 | 52950 | 18487 | 1343 | N/F | 7942 | 58021 | 643 | 47420 |  | Unknown
WO-23003-55-A-1 | N/F | 66565 | 100783 | 8454 | N/F | N/F | 3459638 | N/F | 61589 | N/F | 9199 | 9855 | 3801 | 39249 | 75580 | 975446 | 1315699 | 239124 | 2842774 | 10997251 | 3250452 | 192177 | 33577065 | 3316461 | 24470312 | 3959883 | 29039247 | 53430 | N/F | N/F | 566303 | N/F | 965 | 8696 | 135014 | 220252 | N/F | 3898 | N/F | 9003 | 140997 | 60068 | 7075 | 82578 | 14156 | 1592 | N/F | 9512 | 45201 | N/F | 24966 |  | Unknown
WO-23003-56-A-1 | N/F | 54754 | 73528 | 8224 | 5865 | N/F | 4113574 | N/F | 63414 | N/F | 8047 | 8540 | 2823 | 34131 | 59720 | 572132 | 1147381 | 205437 | 1620637 | 6430738 | 2287421 | 192928 | 31568213 | 3257242 | 16203550 | 3684997 | 24908344 | 53390 | N/F | N/F | 423742 | N/F | 398 | 11529 | 128107 | 91395 | 1293 | 2102 | 147501 | N/F | 159067 | 54031 | 3330 | 54759 | 18767 | 1086 | N/F | 4709 | 52545 | N/F | 30819 |  | Unknown
AB-230117-02_LFB-H_3-inj2 | N/F | 153612 | 336859 | 441324 | 64145861 | N/F | 2973007 | N/F | 277484 | 1392582 | 444397 | 2602453 | 2442836 | 2341203 | 142349 | 319073 | 429616 | 416290 | 2467529 | 12552801 | 5737650 | 445474 | 26770409 | 2583228 | 18087861 | 2470966 | 22319355 | 263890 | N/F | N/F | 1113353 | N/F | 524 | 2349773 | 250719 | 25916085 | 1556 | 16 | 12462757 | 5446794 | 96421 | 20118678 | 2563266 | 17472107 | 16286 | 2519809 | N/F | 2513773 | 26173116 | N/F | 21948918 |  | Unknown
WO-23003-14-A-1B | N/F | 49056 | 66242 | 2687 | N/F | N/F | 3545364 | N/F | 109291 | N/F | 4968388 | 8272 | N/F | 26025 | 78563 | 650627 | 1133240 | 191303 | 2367076 | 9367266 | 3228103 | 287674 | 31673484 | 3146910 | 20850445 | 3522701 | 26497504 | 98056 | N/F | N/F | 586884 | N/F | 809 | 9845 | N/F | 309995 | N/F | 2110 | 204581 | 9414 | 117321 | 35442 | 3137 | 64863 | 6165 | 863 | N/F | 8676 | 46595 | N/F | 30554 |  | Unknown
AB-230117-02_LRB_7-inj3 | N/F | 45259 | 64188 | 3199 | N/F | N/F | 4019182 | N/F | 61785 | 4347 | 640 | 1448 | 981 | 17671 | 66358 | 520203 | 898276 | 224057 | 1984692 | 7730305 | 2671578 | 195172 | 35105550 | 3364711 | 18860380 | 3754570 | 26903633 | 50728 | N/F | N/F | 489201 | N/F | 184 | 12504 | 43509 | 704461 | N/F | 1948 | 52028 | N/F | 153196 | 28650 | 1382 | 42219 | 16243 | 1202 | N/F | 5747 | 51230 | N/F | 64362 |  | Unknown
AB-230117-02_LRB_7-inj2 | N/F | 49481 | 25399 | 6637 | N/F | N/F | 3792511 | N/F | 294182 | N/F | N/F | 1636 | 1690 | 15279 | 151355 | 644639 | 600235 | 443639 | 3248161 | 15789821 | 5967825 | 638026 | 34617381 | 3312026 | 27653250 | 3378218 | 28828488 | 281986 | N/F | N/F | 1175727 | N/F | 354 | 10268 | N/F | 26116379 | N/F | 736 | 58705 | N/F | 147078 | 32025 | 3845 | 61756 | 5755 | 1022 | N/F | 8558 | 58492 | 1031 | 76009 |  | Unknown
WO-23003-60-A-1-inj2 | N/F | 39460 | 50916 | 6845 | 8313 | N/F | 2737052 | N/F | 11721 | N/F | 439 | 6500 | N/F | 14152 | 36022 | 355569 | 261560 | 111850 | 1565684 | 5657666 | 1522380 | 70603 | 8884746 | 3159818 | 15651949 | 3340961 | 24657370 | 9744 | N/F | N/F | 283109 | N/F | 939 | 13536 | N/F | 182259961 | N/F | 1852 | 103308 | N/F | 154496 | 309617 | 3985 | 135906 | 14498 | 2030 | N/F | 4940 | 58218 | N/F | 79961 |  | Unknown
WO-23003-59-A-1-inj2 | N/F | 45600 | 74101 | 7090 | 23647 | N/F | 3849593 | N/F | 77239 | 727 | 359 | 11099 | 2061 | 20188 | 59574 | 589685 | 578477 | 243432 | 2211458 | 8545054 | 2796228 | 334263 | 23508883 | 3402480 | 20857792 | 3757713 | 27380388 | 69366 | 1082 | N/F | 527354 | N/F | 481 | 14291 | N/F | 72956531 | N/F | 2623 | 147373 | N/F | 158265 | 94449 | 2401 | 80114 | 14989 | 1414 | N/F | 5093 | 62121 | N/F | 77224 |  | Unknown
WO-23003-58-A-1-inj2 | N/F | 52409 | 64405 | 11183 | 25051 | N/F | 3714525 | N/F | 99426 | N/F | 1963 | 7705 | 2369 | 24383 | 56758 | 503818 | 770937 | 318973 | 1823727 | 7288617 | 2343061 | 390620 | 31407933 | 3289026 | 17636975 | 3520584 | 24650996 | 89357 | N/F | N/F | 502433 | N/F | 465 | 10819 | N/F | 9415704 | N/F | 2068 | 182375 | N/F | 140119 | 132649 | 6437 | 56876 | 12543 | 1382 | N/F | 6510 | 59929 | N/F | 74948 |  | Unknown
WO-23003-26-A-1B | N/F | 48819 | 62804 | 6138 | 11128 | N/F | 2013496 | N/F | 47323 | N/F | 6118 | 5196 | N/F | 17993 | 45197 | 420948 | 572414 | 221861 | 1686468 | 6428241 | 2061350 | 200735 | 23123741 | 3339024 | 15830622 | 3128580 | 22726578 | 42089 | N/F | N/F | 386664 | N/F | 542 | 6958 | N/F | 54466 | N/F | 1116 | 18970 | N/F | 98818 | 34021 | 2722 | 34926 | 5226 | 2137 | N/F | 5806 | 45595 | 1022 | 32033 |  | Unknown
WO-23003-124-A-1 | N/F | 34719 | 23467 | 2656 | 24746 | N/F | 2632325 | N/F | 365396 | N/F | N/F | 2323 | 2910 | 6312 | 155082 | 608072 | 295106 | 782216 | 3413219 | 17596767 | 6545217 | 831639 | 11648092 | 4063813 | 31540850 | 4028406 | 37306194 | 344771 | N/F | N/F | 1147141 | 652 | 704 | 14827 | N/F | 242967382 | N/F | 313 | 46743 | 22561 | 173594 | 423630 | 1429 | 232683 | 16609 | 2749 | N/F | 8323 | 76541 | N/F | 97115 |  | Unknown
WO-23003-123-A-1 | N/F | 33859 | 26268 | 3872 | 30665 | N/F | 2702609 | N/F | 352448 | N/F | N/F | 6324 | 2455 | 8045 | 154965 | 627453 | 296199 | 787922 | 3447890 | 17733282 | 6652997 | 833105 | 11760718 | 3950551 | 31781255 | 3914703 | 36705423 | 341203 | N/F | N/F | 1191848 | N/F | 843 | 16371 | N/F | 237258758 | N/F | 109 | 49348 | 26372 | 160672 | 628802 | 3488 | 253914 | 14883 | 3018 | N/F | 7769 | 87789 | 2079 | 110449 |  | Unknown
WO-23003-122-A-1 | N/F | 34926 | 25576 | 5114 | 54560 | N/F | 2629092 | 97 | 331367 | N/F | N/F | 5816 | 2669 | 9783 | 150716 | 584792 | 301961 | 799705 | 3260898 | 16984551 | 6513227 | 832140 | 11794822 | 3872784 | 30676932 | 3824035 | 35822718 | 317514 | N/F | N/F | 1172222 | N/F | 498 | 15828 | N/F | 234371957 | N/F | 186 | 57374 | 29611 | 147921 | 530815 | 5325 | 253879 | 12853 | 4910 | N/F | 7891 | 97264 | N/F | 108730 |  | Unknown
WO-23003-121-A-1 | N/F | 37841 | 42574 | 7427 | 142716 | N/F | 2967155 | 81 | 180324 | N/F | N/F | 14785 | 4044 | 14266 | 91775 | 664466 | 398625 | 477397 | 2724072 | 11914775 | 3912310 | 528181 | 14897412 | 3875690 | 26466602 | 4032880 | 33793745 | 164724 | N/F | N/F | 701356 | N/F | 1066 | 16931 | N/F | 189787097 | N/F | 1519 | 224636 | 32985 | 151013 | 1300624 | 10081 | 1172758 | 17474 | 9578 | N/F | 7204 | 163404 | N/F | 122769 |  | Unknown
AB-230117-02_LFSM-H_6 | N/F | 116843 | 447190 | 348575 | 63586790 | N/F | 2300881 | N/F | 154827 | 695183 | 393240 | 3387062 | 2144937 | 2924427 | 85858 | 445574 | 355081 | 413834 | 2373337 | 10288661 | 3534758 | 480307 | 13096971 | 3308097 | 22728676 | 3280289 | 28306027 | 142739 | N/F | 2102 | 634016 | N/F | 935 | 1742611 | 124972 | 182605113 | N/F | 811 | 9490044 | 3090430 | 215638 | 25074130 | 1588159 | 20943236 | 61252 | 3224244 | N/F | 1361388 | 32788478 | 2029 | 25847585 |  | Unknown
AB-230117-02_CCC-H_7 | N/F | 297783 | 582678 | 666213 | 117420968 | N/F | 2713124 | N/F | 370853 | 2750476 | 1375403 | 5371324 | 4389541 | 4617408 | 150132 | 311929 | 396768 | 749296 | 2371745 | 13252045 | 6157485 | 842192 | 29369183 | 3158942 | 23190512 | 2628619 | 26586794 | 353402 | N/F | N/F | 1158056 | N/F | 505 | 4022216 | 599798 | 51655975 | N/F | 62 | 24261914 | 10865858 | 277502 | 43120094 | 4914207 | 40425414 | 105154 | 5673620 | N/F | 4563799 | 60471286 | 941 | 43702308 |  | Unknown
AB-230117-02_LFB-H_7 | N/F | 178898 | 385187 | 466921 | 66276186 | N/F | 3442716 | N/F | 373431 | 1379527 | 720895 | 3232929 | 2575639 | 2638110 | 143965 | 399806 | 491510 | 795193 | 2802953 | 15058626 | 6414093 | 854628 | 32994516 | 3450724 | 26765152 | 3134189 | 30536985 | 356495 | N/F | N/F | 1194388 | 837 | 591 | 2529429 | 304992 | 49165790 | N/F | N/F | 13936591 | 5800047 | 218334 | 25093229 | 2732141 | 23728345 | 53999 | 2991377 | N/F | 2379622 | 32310047 | N/F | 27928437 |  | Unknown
AB-230117-02_LD_6_inj2 | N/F | 36978 | 28093 | 9371 | 4134 | N/F | 2609383 | N/F | 353754 | N/F | N/F | 1943 | N/F | 6663 | 165073 | 598949 | 285219 | 796687 | 3246144 | 16951110 | 6648637 | 833625 | 11439467 | 3711137 | 29696384 | 3770630 | 34722557 | 340517 | N/F | N/F | 1197901 | N/F | 387 | 14745 | N/F | 232336403 | 848 | 250 | 43967 | 27665 | 132306 | 479246 | 1940 | 226398 | 16229 | 672 | N/F | 7973 | 57041 | 969 | 90152 |  | Unknown
AB-230117-02_LD_6 | 847 | 38832 | 28390 | 9040 | 8416 | N/F | 2411193 | N/F | 367440 | N/F | N/F | 2537 | 2330 | 7979 | 164561 | 650784 | 294197 | 816555 | 3345917 | 17809044 | 6657933 | 850469 | 11938961 | 3844232 | 31273925 | 3868396 | 37076510 | 346314 | N/F | N/F | 1177883 | N/F | 670 | 17452 | N/F | 229640168 | N/F | N/F | 41437 | 28298 | 174758 | 479994 | 991 | 236508 | 8348 | 1608 | N/F | 9415 | 59994 | 708 | 80397 |  | Unknown
AB-230117-02_LRB_13 | N/F | 39223 | 30139 | 4025 | 1156 | N/F | 4975935 | N/F | 339699 | 1861 | 615 | 1563 | 1444 | 10457 | 154524 | 632889 | 829400 | 808357 | 3308258 | 17309829 | 6512473 | 817461 | 42552876 | 3892402 | 30286976 | 3878390 | 36636843 | 327227 | N/F | N/F | 1190410 | N/F | 451 | 14385 | N/F | 14418173 | N/F | 148 | N/F | 23313 | 168585 | 47596 | 1473 | 116935 | 10597 | 839 | N/F | 7862 | 56021 | 1148 | 34851 |  | Unknown
WO-23003-120-A-1 | 4227 | 33365 | 56665 | 2462 | 3391 | N/F | 3665778 | N/F | 150847 | 2596 | N/F | 3116 | 1298 | N/F | 97770 | 623493 | 554026 | 434712 | 2644378 | 11163978 | 3851120 | 490883 | 21611490 | 4158510 | 26321859 | 4314884 | 34541785 | 139198 | N/F | N/F | 680859 | N/F | N/F | 15909 | N/F | 140835849 | N/F | 1090 | 110370 | 24236 | 175452 | 579885 | 1974 | 604169 | 18459 | 1225 | N/F | 8136 | 59547 | 875 | 35916 |  | Unknown
WO-23003-119-A-1 | N/F | 35232 | 80836 | 4076 | N/F | N/F | 3815725 | N/F | 165575 | N/F | 4337 | 5482 | 2072 | 8747 | 91723 | 787661 | 660417 | 470814 | 2850510 | 11838608 | 3887057 | 522272 | 22798424 | 4144631 | 27604625 | 4353409 | 34960984 | 152503 | N/F | N/F | 705098 | N/F | 566 | 12328 | N/F | 121342878 | N/F | 1701 | 118487 | 19743 | 181045 | 445774 | 1925 | 543012 | 13728 | 1443 | N/F | 6455 | 53594 | N/F | 31511 |  | Unknown
WO-23003-118-A-1 | N/F | 36558 | 113908 | 5997 | N/F | 1879 | 2944013 | N/F | 174270 | N/F | 7131 | 2261 | 1421 | 8843 | 97907 | 636039 | 406147 | 455736 | 2874973 | 12188667 | 3944874 | 546398 | 13732261 | 3973276 | 27047440 | 4072631 | 34571765 | 157920 | N/F | N/F | 702467 | N/F | 336 | 14545 | N/F | 192387453 | N/F | 2143 | 237454 | 33780 | 179666 | 1489565 | 3363 | 1680672 | 16713 | 864 | N/F | 8077 | 48850 | N/F | 36295 |  | Unknown
WO-23003-117-A-1 | N/F | 34058 | 150084 | 3806 | 2854 | N/F | 3106564 | N/F | 149272 | 1501 | N/F | 3312 | 2373 | 6234 | 90027 | 584326 | 361377 | 434364 | 2638588 | 11170020 | 3649594 | 476715 | 15207107 | 4032701 | 25776455 | 4130614 | 34091479 | 137545 | N/F | N/F | 656751 | N/F | 255 | 14239 | N/F | 187287220 | N/F | 2432 | 181558 | 28654 | 188295 | 1284611 | 1840 | 1427064 | 8949 | 1193 | N/F | 6576 | 56056 | 2294 | 42028 |  | Unknown
WO-23003-116-A-1 | N/F | 33501 | 54993 | 3443 | N/F | 1536 | 3078558 | N/F | 154023 | N/F | 6289 | 3421 | 1195 | 10234 | 90926 | 546205 | 372813 | 415987 | 2579445 | 10893207 | 3620587 | 469612 | 15028164 | 3931505 | 24439633 | 3932868 | 33143414 | 140200 | 1047 | N/F | 654291 | N/F | 888 | 13192 | N/F | 182537121 | N/F | 1339 | 189508 | 30260 | 179065 | 1334154 | 1910 | 1442508 | 16605 | 832 | N/F | 6269 | 51820 | N/F | 38360 |  | Unknown
WO-23003-115-A-1 | N/F | 40632 | 565499 | 7063 | N/F | N/F | 3828036 | N/F | 99766 | 1472 | N/F | 3483 | 2258 | 11296 | 71519 | 711567 | 771773 | 354215 | 2318518 | 9357415 | 3003752 | 360502 | 25004515 | 4027729 | 23193579 | 4320818 | 30846423 | 88662 | N/F | N/F | 550322 | N/F | 699 | 11555 | N/F | 87544147 | N/F | 1338 | 137835 | 19960 | 170679 | 256899 | 1519 | 292195 | 16274 | 1236 | N/F | 8630 | 53099 | N/F | 35809 |  | Unknown
WO-23003-114-A-1 | N/F | 35881 | 194083 | 10685 | N/F | 817 | 4118739 | N/F | 86142 | N/F | 3135 | 3714 | 1135 | 7999 | 71092 | 686954 | 733569 | 347830 | 2270014 | 9114971 | 3000690 | 325864 | 27456944 | 4035660 | 22992151 | 4320860 | 31365991 | 74455 | N/F | N/F | 529838 | N/F | 750 | 12838 | N/F | 83814010 | 798 | 3098 | 134764 | 14932 | 183311 | 214016 | 3595 | 251009 | 10678 | 1379 | N/F | 5911 | 58376 | N/F | 32140 |  | Unknown
WO-23003-113-A-1 | N/F | 37314 | 122428 | 5545 | N/F | N/F | 3982422 | N/F | 79559 | N/F | 3122 | 8169 | 907 | 9220 | 74839 | 598344 | 668272 | 353182 | 2249019 | 9107594 | 2967227 | 317505 | 27159198 | 4075045 | 22864066 | 4327492 | 31830157 | 70741 | N/F | N/F | 516099 | N/F | 457 | 13358 | N/F | 85481441 | 772 | 2557 | 47735 | 10873 | 181245 | 209128 | 2319 | 261159 | 17243 | 1749 | N/F | 5160 | 55355 | 1275 | 42318 |  | Unknown
WO-23003-112-A-1 | N/F | 70441 | 250477 | 10900 | 4911 | N/F | 3895384 | N/F | 86547 | N/F | 6939 | 6500 | N/F | 11981 | 79594 | 631806 | 737993 | 364283 | 2237851 | 9200671 | 2902462 | 315301 | 27511818 | 3949798 | 22209911 | 4242494 | 30748381 | 74317 | N/F | N/F | 528078 | N/F | 525 | 9908 | N/F | 72351283 | N/F | 3826 | 160297 | 16035 | 184078 | 230012 | 3458 | 328561 | 18257 | 653 | N/F | 6592 | 52834 | 558 | 35163 |  | Unknown
WO-23003-111-A-1 | N/F | 38230 | 209062 | 6649 | N/F | N/F | 4168487 | N/F | 80890 | N/F | 6314 | 7501 | N/F | 18695 | 70307 | 567611 | 626483 | 328678 | 2170294 | 8671592 | 2876144 | 300553 | 25777832 | 4035736 | 21485097 | 4174581 | 31212515 | 67943 | N/F | N/F | 508806 | N/F | 761 | 12426 | N/F | 93632999 | 1372 | 2846 | 173499 | 18897 | 172782 | 325423 | 2667 | 475100 | 18296 | 509 | N/F | 6804 | 55001 | 353 | 32276 |  | Unknown
AB-230117-02_LRB_12_20230122013852 | N/F | 43278 | 28783 | 8315 | 7181 | 949 | 4847639 | N/F | 377613 | N/F | 868 | 3751 | N/F | 9251 | 164550 | 594945 | 788881 | 846668 | 3298787 | 17428215 | 7028879 | 823081 | 41966592 | 3878621 | 29379328 | 3753277 | 34883484 | 362305 | N/F | N/F | 1260290 | N/F | 1033 | 15503 | N/F | 209488 | N/F | 283 | 34050 | 25539 | 141780 | 29832 | 2265 | 100780 | 14247 | 2113 | N/F | 11646 | 56499 | 892 | 42838 |  | Unknown
AB-230117-02_LRB_12 | N/F | 38740 | 37547 | 10631 | 5742 | 973 | 4991287 | N/F | 385085 | N/F | N/F | 2692 | 2351 | 8989 | 167954 | 615136 | 828284 | 837112 | 3497300 | 18450364 | 6972498 | 906039 | 43582855 | 4160630 | 31493553 | 4072598 | 37572020 | 368063 | N/F | N/F | 1263690 | N/F | 1106 | 12646 | N/F | 5345236 | 1143 | 142 | 45094 | 26748 | 182322 | 36176 | 3370 | 125989 | 16040 | 1845 | N/F | 8378 | 64305 | 634 | 44187 |  | Unknown
WO-23003-110-A-1 | N/F | 43680 | 715911 | 5291 | 7681 | 1252 | 4099605 | N/F | 97891 | N/F | 8501 | 4984 | N/F | 9908 | 75773 | 637488 | 758319 | 358982 | 2298698 | 9223789 | 3090429 | 368565 | 25868952 | 4103483 | 23004724 | 4257264 | 31172627 | 85601 | N/F | N/F | 543250 | N/F | 1171 | 12800 | N/F | 84482675 | N/F | 3721 | 117786 | 26617 | 170608 | 290283 | 2375 | 435245 | 17323 | 1409 | N/F | 6373 | 53982 | N/F | 37531 |  | Unknown
WO-23003-109-A-1 | N/F | 40040 | 239885 | 5605 | 7635 | N/F | 4985822 | N/F | 93207 | N/F | 668 | 6382 | N/F | 13058 | 79160 | 643334 | 975287 | 370341 | 2499747 | 9934193 | 3179144 | 340439 | 37906152 | 4197930 | 24204764 | 4340752 | 33272051 | 80475 | N/F | N/F | 559441 | N/F | 604 | 17143 | N/F | 25547220 | N/F | 3055 | 18772 | 7946 | 156557 | 63519 | 3060 | 87969 | 16360 | 1136 | N/F | 4909 | 56399 | 959 | 32256 |  | Unknown
WO-23003-108-A-1 | 860 | 36287 | 232806 | 5270 | 11835 | N/F | 4542665 | N/F | 51190 | N/F | N/F | 2838 | 844 | 7886 | 52594 | 313278 | 784434 | 243079 | 1222131 | 5647825 | 2020116 | 203289 | 31394130 | 3482061 | 12475249 | 3400362 | 22459853 | 40644 | N/F | N/F | 371179 | N/F | 523 | 14322 | N/F | 18147756 | 1254 | 1897 | 20826 | N/F | 185834 | 40563 | 977 | 40973 | 19976 | 1497 | N/F | 4549 | 67471 | 703 | 38022 |  | Unknown
WO-23003-107-A-1 | N/F | 33231 | 120040 | 4816 | 10833 | N/F | 4689467 | N/F | 90926 | N/F | N/F | 7406 | N/F | 10396 | 66208 | 481019 | 896330 | 336779 | 1848875 | 7520697 | 2685215 | 316411 | 34808537 | 3900449 | 18844557 | 3963746 | 28511479 | 79995 | N/F | N/F | 474385 | N/F | 385 | 11824 | N/F | 24168389 | N/F | 1819 | N/F | N/F | 168219 | 48890 | 4632 | 77260 | 16205 | 1688 | N/F | 4305 | 56138 | 530 | 49791 |  | Unknown
WO-23003-106-A-1 | N/F | 32879 | 89069 | 5159 | 15182 | N/F | 4357248 | N/F | 71490 | 569 | N/F | 5922 | 372 | 8156 | 57502 | 334537 | 799107 | 280040 | 1306281 | 6007469 | 2339066 | 275072 | 30720915 | 3533365 | 13398584 | 3342309 | 22560018 | 61625 | N/F | N/F | 421655 | N/F | 489 | 17387 | N/F | 20692211 | 513 | 1854 | N/F | N/F | 181635 | 43934 | 2945 | 59856 | 17930 | 1591 | N/F | 3069 | 65717 | N/F | 40865 |  | Unknown
WO-23003-105-A-1 | N/F | 46142 | 280017 | 7320 | 20119 | N/F | 4462092 | N/F | 55168 | 1053 | N/F | 4046 | 1045 | 11564 | 60753 | 446202 | 930307 | 275513 | 1697435 | 6886482 | 2374984 | 222694 | 34342726 | 3838787 | 17543190 | 3832481 | 27261056 | 44312 | N/F | N/F | 432451 | N/F | 362 | 14072 | N/F | 21561961 | N/F | 2633 | 77629 | N/F | 180589 | 75802 | 2528 | 85172 | 17366 | 2376 | N/F | 4300 | 64144 | N/F | 47004 |  | Unknown
WO-23003-104-A-1 | N/F | 45733 | 213630 | 5806 | 30557 | N/F | 4696497 | 11 | 133262 | N/F | N/F | 5730 | 2265 | 13158 | 92673 | 800180 | 1098460 | 447858 | 2776678 | 11384213 | 3711085 | 449765 | 37655788 | 4359688 | 26677085 | 4577412 | 34076493 | 119110 | N/F | N/F | 697535 | N/F | 611 | 14555 | N/F | 23769062 | 1829 | 3487 | 45834 | N/F | 185682 | 82689 | 4371 | 115313 | 14412 | 3353 | N/F | 6267 | 67633 | N/F | 46289 |  | Unknown
WO-23003-103-A-1 | N/F | 35556 | 256387 | 3120 | 39351 | N/F | 4931396 | N/F | 77484 | N/F | N/F | 3433 | N/F | 9752 | 59322 | 352121 | 869022 | 308437 | 1552503 | 6640754 | 2466184 | 272827 | 35768513 | 3790642 | 15504438 | 3727708 | 26318356 | 62960 | N/F | N/F | 448814 | N/F | 421 | 17022 | N/F | 11714912 | N/F | 1921 | 84857 | 11615 | 193991 | 52557 | 3800 | 66237 | 19818 | 3512 | N/F | 1750 | 83871 | N/F | 58593 |  | Unknown
WO-23003-102-A-1 | N/F | 36606 | 197697 | 5621 | 61956 | N/F | 4239856 | N/F | 72988 | 2040 | N/F | 8518 | N/F | 11663 | 51890 | 223738 | 639834 | 274722 | 1056082 | 5394730 | 2251411 | 264739 | 27775775 | 3097088 | 10581129 | 2695615 | 18015535 | 60077 | N/F | N/F | 415107 | N/F | 178 | 14311 | N/F | 8827538 | N/F | 2023 | 22438 | 11244 | 190345 | 47656 | 4991 | 46784 | 20984 | 5433 | N/F | 2560 | 128142 | N/F | 58196 |  | Unknown
WO-23003-101-A-1 | N/F | 48825 | 455527 | 7983 | 163899 | N/F | 4895635 | N/F | 85771 | N/F | N/F | 10107 | 3462 | 18546 | 73288 | 563161 | 1062014 | 376523 | 2309340 | 9069846 | 3044917 | 331168 | 38768749 | 4222824 | 22389687 | 4216694 | 31868378 | 71332 | N/F | N/F | 558194 | N/F | 281 | 18870 | N/F | 5374162 | 741 | 2114 | 33927 | 13383 | 171646 | 81852 | 7653 | 86243 | 19448 | 11464 | N/F | 4379 | 177624 | N/F | 106679 |  | Unknown
AB-230117-02_LFSM-H_5 | N/F | 91830 | 944386 | 685679 | 76529547 | 1662 | 3418497 | N/F | 102282 | 826078 | 434520 | 3679856 | 2569068 | 3188418 | 101507 | 551719 | 745732 | 455406 | 2828511 | 12308621 | 4003771 | 386917 | 28982045 | 3485923 | 24178040 | 3427596 | 26412031 | 89635 | N/F | N/F | 727272 | N/F | 913 | 2697595 | 89176 | 31734667 | N/F | 1082 | 11313624 | 3510215 | 198344 | 22793915 | 3349018 | 21979592 | 48432 | 3459348 | N/F | 1923500 | 32850160 | N/F | 24640187 |  | Unknown
AB-230117-02_CCC-H_6 | N/F | 281850 | 562778 | 654485 | 113141419 | N/F | 2793348 | N/F | 395495 | 2745124 | 1414539 | 5030896 | 4276671 | 4300009 | 152035 | 306839 | 350063 | 810488 | 2380427 | 12735101 | 6477511 | 865033 | 27124276 | 3102225 | 22052524 | 2540641 | 24084407 | 379586 | N/F | 2242 | 1219465 | N/F | 442 | 3982013 | 636652 | 45429877 | N/F | 125 | 22407500 | 10915888 | 233820 | 37961336 | 4726407 | 36661472 | 78086 | 5307399 | N/F | 4498921 | 55792339 | N/F | 39491287 |  | Unknown
AB-230117-02_LFB-H_6 | N/F | 187695 | 371368 | 459032 | 68015432 | N/F | 3536832 | N/F | 380146 | 1429983 | 714174 | 3140950 | 2475938 | 2613609 | 162044 | 409177 | 503164 | 822027 | 2775932 | 15094220 | 6712510 | 851740 | 33569510 | 3413768 | 25321860 | 3062121 | 28650436 | 363560 | N/F | N/F | 1249401 | 1903 | 443 | 2563681 | 313365 | 28981468 | N/F | 567 | 13619043 | 5845666 | 195156 | 23399281 | 2934172 | 21953767 | 48517 | 2943589 | N/F | 2483996 | 30205275 | 1357 | 25864201 |  | Unknown
AB-230117-02_LD_5_inj2 | N/F | 41217 | 443693 | 6992 | N/F | N/F | 4440593 | N/F | 78523 | N/F | 622 | 5586 | 1135 | 11568 | 76315 | 631259 | 946446 | 333710 | 2467575 | 10065037 | 3003137 | 299169 | 34578186 | 3844573 | 23295803 | 3998089 | 30311159 | 69999 | N/F | N/F | 543923 | N/F | 205 | 11594 | N/F | 14938611 | N/F | 2488 | 36249 | 11088 | 126724 | 48313 | 3249 | 92906 | 13116 | 1313 | N/F | 6943 | 50838 | N/F | 32846 |  | Unknown
AB-230117-02_LD_5 | N/F | 45285 | 451029 | 8268 | 5574 | 1799 | 4809439 | N/F | 94342 | N/F | N/F | 2824 | 1035 | 12246 | 76969 | 631388 | 1004335 | 380794 | 2654414 | 10624228 | 3291411 | 345139 | 38110657 | 4207699 | 25128152 | 4291168 | 32939070 | 82474 | N/F | N/F | 584499 | N/F | 387 | 11042 | N/F | 15494357 | N/F | 2393 | 24924 | 12665 | 169770 | 58472 | 1657 | 101600 | 16301 | 1159 | N/F | 7123 | 56945 | N/F | 40642 |  | Unknown
AB-230117-02_LRB_11 | N/F | 39580 | 28935 | 5754 | N/F | N/F | 5075547 | 9 | 382255 | 2634 | N/F | 1216 | 996 | 9996 | 156160 | 611546 | 844320 | 852826 | 3472176 | 17655766 | 6981552 | 840649 | 43350177 | 4214839 | 30665008 | 4082172 | 35547285 | 366041 | N/F | N/F | 1260144 | N/F | 853 | 15474 | N/F | 59282 | N/F | 246 | N/F | 25286 | 168981 | 28598 | 1503 | 107613 | 16949 | 892 | N/F | 8715 | 57099 | 1341 | 37846 |  | Unknown
WO-23003-100-A-1 | N/F | 34285 | 68719 | 2342 | N/F | N/F | 4918862 | N/F | 86077 | 1043 | 2126 | 5398 | 475 | 13296 | 70157 | 588867 | 1192737 | 344729 | 2109992 | 8525143 | 2963581 | 341636 | 39237589 | 4295001 | 21267105 | 4328504 | 31068711 | 75533 | N/F | N/F | 539778 | N/F | 299 | 15997 | 171522 | 45966 | N/F | 2465 | 28073 | N/F | 183621 | 31065 | 1032 | 40050 | 19725 | 801 | N/F | 5706 | 53858 | N/F | 36250 |  | Unknown
WO-23003-99-A-1 | N/F | 36064 | 71744 | 4917 | N/F | N/F | 4845368 | N/F | 52013 | N/F | N/F | 4464 | 1310 | 9538 | 61737 | 582119 | 1171392 | 296170 | 1936346 | 7270327 | 2419840 | 207284 | 38353517 | 4116694 | 19214059 | 4178947 | 28842466 | 41441 | N/F | N/F | 435231 | N/F | 432 | 12334 | 102641 | 37652 | 511 | 2894 | N/F | N/F | 175311 | 28743 | 645 | 39903 | 17864 | 835 | N/F | 6974 | 57370 | N/F | 42938 |  | Unknown
WO-23003-98-A-1 | 1496 | 37144 | 80204 | 7376 | N/F | N/F | 4980257 | N/F | 105663 | N/F | N/F | 6394 | 1205 | 13269 | 78960 | 648369 | 1183660 | 373117 | 2241053 | 9166810 | 3129094 | 364852 | 40201443 | 4346142 | 22766224 | 4420815 | 31947734 | 90738 | N/F | N/F | 576514 | N/F | 528 | 13129 | 202295 | 44451 | N/F | 4278 | 97968 | N/F | 177790 | 35644 | 3469 | 50887 | 15463 | 1483 | N/F | 5425 | 54845 | N/F | 28652 |  | Unknown
WO-23003-97-A-1 | N/F | 46464 | 106242 | 8331 | N/F | N/F | 4151628 | N/F | 99665 | 1590 | N/F | 7861 | 1018 | 15028 | 79190 | 848698 | 1477455 | 388074 | 2284322 | 8722568 | 3111777 | 341121 | 35663824 | 4266521 | 21661622 | 4408049 | 28777567 | 85366 | N/F | N/F | 528734 | N/F | 1155 | 11628 | 180250 | 102495 | N/F | 1201 | 38245 | N/F | 179280 | 34869 | 1302 | 50608 | 18943 | 743 | N/F | 7139 | 48575 | N/F | 30897 |  | Unknown
WO-23003-96-A-1 | N/F | 35884 | 89358 | 10665 | N/F | N/F | 4768265 | N/F | 104361 | N/F | N/F | 5795 | 798 | 11436 | 74530 | 720088 | 1309270 | 369718 | 2335994 | 9139782 | 3137650 | 357902 | 38779746 | 4329298 | 22640971 | 4407111 | 30472169 | 86582 | N/F | N/F | 570404 | N/F | 442 | 11591 | 189649 | 34754 | N/F | 4211 | N/F | N/F | 187492 | 27887 | 2250 | 38915 | 17942 | 1247 | N/F | 6496 | 59496 | N/F | 31061 |  | Unknown
WO-23003-95-A-1 | N/F | 40749 | 74634 | 6125 | 4408 | N/F | 5066551 | N/F | 102039 | N/F | N/F | 6433 | N/F | 11058 | 81800 | 675641 | 1242406 | 414155 | 2438136 | 9559093 | 3290070 | 363695 | 41590416 | 4582143 | 23651394 | 4621082 | 33004661 | 87444 | N/F | N/F | 591866 | N/F | 783 | 12604 | 191933 | 42827 | N/F | 3180 | N/F | N/F | 183387 | 27130 | 2327 | 46472 | 19709 | 1789 | N/F | 5387 | 58706 | 947 | 35657 |  | Unknown
WO-23003-94-A-1 | N/F | 43345 | 80482 | 13923 | N/F | N/F | 4827694 | N/F | 106735 | N/F | N/F | 6371 | N/F | 12891 | 85205 | 678765 | 1228343 | 400040 | 2432299 | 9572668 | 3204689 | 363698 | 40658864 | 4502627 | 23272982 | 4484651 | 32934546 | 90503 | N/F | N/F | 587438 | N/F | 943 | 15081 | 198645 | 64108 | N/F | 3087 | N/F | N/F | 193293 | 39343 | 3144 | 49186 | 18426 | 1316 | N/F | 7048 | 56215 | N/F | 35806 |  | Unknown
WO-23003-93-A-1 | N/F | 42493 | 88702 | 7486 | N/F | N/F | 4175637 | N/F | 73928 | N/F | 3847 | 3285 | N/F | 12455 | 70402 | 753665 | 1373606 | 337582 | 2216247 | 8701125 | 2874110 | 290619 | 35550674 | 4301119 | 21726871 | 4303332 | 29449513 | 62051 | N/F | N/F | 531128 | N/F | 952 | 12714 | 129541 | 59010 | N/F | 4083 | 19470 | N/F | 166811 | 40509 | 1810 | 46660 | 16372 | 1014 | N/F | 5669 | 50074 | N/F | 32908 |  | Unknown
WO-23003-92-A-1 | N/F | 43261 | 124911 | 4837 | N/F | N/F | 3595092 | N/F | 128538 | N/F | N/F | 5439 | 1286 | 15492 | 97673 | 1345341 | 1560244 | 415217 | 2653892 | 10153929 | 3722523 | 453875 | 33403984 | 4198498 | 21748898 | 4256224 | 26672539 | 111912 | N/F | N/F | 631015 | N/F | 967 | 12315 | 238715 | 30828 | N/F | 3641 | N/F | N/F | 176967 | 30731 | 4085 | 50573 | 18497 | 959 | N/F | 7252 | 49122 | N/F | 22674 |  | Unknown
WO-23003-91-A-1 | N/F | 43025 | 89302 | 2803 | 5926 | N/F | 4271476 | N/F | 59936 | N/F | N/F | 7516 | 877 | 12267 | 62875 | 645389 | 1288519 | 289060 | 2073458 | 7833762 | 2551234 | 230993 | 34928638 | 4071211 | 20445424 | 4097531 | 28354687 | 48649 | N/F | N/F | 458185 | N/F | 533 | 13800 | 103651 | 72298 | N/F | 4485 | N/F | N/F | 159478 | 29295 | 1478 | 54443 | 16935 | 1012 | N/F | 6350 | 49190 | 824 | 28777 |  | Unknown
AB-230117-02_LRB_10_20230121123333 | N/F | 42578 | 31864 | 2894 | 6406 | N/F | 4680891 | N/F | 408529 | 510 | 2535 | 2529 | 1752 | 13682 | 164204 | 589105 | 782192 | 821316 | 3334539 | 17091618 | 7168889 | 900066 | 40242698 | 3990148 | 28633631 | 3766065 | 33684803 | 393815 | N/F | N/F | 1316014 | N/F | 1073 | 13436 | N/F | 66161 | N/F | 253 | 27407 | 29237 | 141648 | 26538 | 1637 | 103485 | 15035 | 2062 | N/F | 10445 | 55137 | N/F | 34718 |  | Unknown
AB-230117-02_LRB_10 | N/F | 40587 | 31551 | 1856 | 7023 | N/F | 5022080 | N/F | 407220 | N/F | 1116 | 2831 | 1579 | 10750 | 157638 | 608177 | 961506 | 854107 | 3525650 | 17720729 | 7266040 | 904756 | 43561553 | 4294162 | 30607554 | 4033211 | 35895538 | 390942 | N/F | N/F | 1318397 | N/F | 1255 | 15687 | N/F | 63479 | 1474 | 335 | 26599 | 27529 | 177426 | 27848 | 2212 | 102101 | 17952 | 1667 | N/F | 8722 | 59876 | N/F | 35599 |  | Unknown
WO-23003-90-A-1 | N/F | 45562 | 92628 | 5498 | 7545 | 639 | 3948810 | N/F | 82212 | N/F | 2574 | 7608 | 1461 | 13421 | 65637 | 734020 | 1371313 | 309238 | 2158117 | 8119001 | 2768161 | 274352 | 33684003 | 4120983 | 20567756 | 4188114 | 28199051 | 69073 | N/F | N/F | 489232 | N/F | 777 | 12909 | 139472 | 63734 | N/F | 4223 | 22583 | N/F | 159001 | 29679 | 1620 | 51089 | 17386 | 1384 | N/F | 6129 | 50952 | N/F | 28581 |  | Unknown
WO-23003-89-A-1 | N/F | 38279 | 79213 | 6754 | 9134 | N/F | 4735152 | N/F | 95993 | N/F | 366 | 6199 | 997 | 11757 | 71005 | 655632 | 1145179 | 350298 | 2267966 | 8629936 | 2981804 | 347262 | 36900501 | 4205412 | 21841201 | 4240476 | 29903150 | 82000 | N/F | N/F | 557271 | N/F | 1240 | 11416 | 188674 | 37302 | N/F | 2515 | N/F | N/F | 168105 | 27389 | 3025 | 47765 | 13276 | 1868 | N/F | 7912 | 58170 | 2823 | 44396 |  | Unknown
WO-23003-88-A-1 | N/F | 38829 | 93066 | 6348 | 9758 | N/F | 4439384 | N/F | 111624 | N/F | N/F | 8092 | N/F | 12012 | 78322 | 692529 | 1217157 | 390901 | 2367506 | 9157525 | 3178984 | 402264 | 36457724 | 4151183 | 21706626 | 4184832 | 29585549 | 98078 | N/F | N/F | 573552 | N/F | 1446 | 11422 | 207475 | 42418 | N/F | 2416 | 23064 | N/F | 171792 | 34064 | 2155 | 47793 | 18540 | 1670 | N/F | 5986 | 55510 | N/F | 37627 |  | Unknown
WO-23003-87-A-1 | N/F | 38982 | 81028 | 5381 | 13552 | N/F | 4182229 | N/F | 36844 | N/F | 908 | 3888 | 723 | 13506 | 47644 | 525993 | 1137560 | 220206 | 1712563 | 6429706 | 2003607 | 161238 | 33573338 | 3842692 | 16981543 | 3880549 | 25806511 | 30259 | N/F | N/F | 359063 | N/F | 605 | 12969 | 68038 | 38381 | N/F | 3327 | 32970 | N/F | 171086 | 32815 | 3037 | 43472 | 21940 | 887 | N/F | 3761 | 54945 | N/F | 32862 |  | Unknown
WO-23003-86-A-1 | N/F | 36997 | 85805 | 6907 | 19279 | N/F | 4297233 | N/F | 57143 | 2350 | 2282 | 7341 | 857 | 11112 | 59857 | 592529 | 1147029 | 285832 | 2010931 | 7751944 | 2547631 | 218688 | 34794228 | 4089669 | 19859541 | 4037136 | 28078988 | 45454 | N/F | N/F | 431507 | N/F | 542 | 12506 | 109116 | 44493 | N/F | 3708 | 26148 | N/F | 177086 | 43880 | 2822 | 49940 | 20467 | 1523 | N/F | 5194 | 62960 | N/F | 36296 |  | Unknown
WO-23003-85-A-1 | N/F | 35400 | 107009 | 5109 | 24433 | N/F | 3786609 | N/F | 110322 | N/F | N/F | 7072 | 879 | 13240 | 76890 | 867893 | 1368510 | 375658 | 2413172 | 9236864 | 3211135 | 383660 | 32886770 | 4063122 | 22201556 | 4087810 | 28031744 | 95808 | N/F | N/F | 567741 | N/F | 728 | 12402 | 222215 | 54985 | N/F | 3042 | N/F | N/F | 168629 | 39050 | 3061 | 51360 | 18493 | 1941 | N/F | 8424 | 60896 | N/F | 33879 |  | Unknown
WO-23003-84-A-1 | N/F | 39589 | 110374 | 10330 | 28875 | N/F | 3557205 | 10 | 111901 | N/F | N/F | 8774 | 1460 | 14276 | 77508 | 965799 | 1367011 | 394791 | 2350619 | 9039349 | 3076516 | 367253 | 31050434 | 3961576 | 21416800 | 4017463 | 26361109 | 97362 | N/F | N/F | 559854 | N/F | 996 | 15772 | 218437 | 37155 | 1331 | 2157 | 63947 | N/F | 154301 | 37946 | 2742 | 51385 | 16948 | 3031 | N/F | 6114 | 61943 | N/F | 35713 |  | Unknown
WO-23003-83-A-1 | N/F | 42154 | 78304 | 6997 | 47761 | N/F | 4306540 | N/F | 101614 | N/F | 1250 | 9175 | 1667 | 13920 | 71904 | 646235 | 1134941 | 355826 | 2218956 | 8502791 | 2958244 | 349169 | 35975599 | 4163662 | 20622987 | 4111584 | 29319967 | 88731 | N/F | N/F | 550934 | N/F | 729 | 13990 | 152419 | 72851 | N/F | 3392 | N/F | N/F | 152148 | 44847 | 7460 | 54402 | 10742 | 4131 | N/F | 5542 | 76961 | N/F | 43246 |  | Unknown
WO-23003-82-A-1 | N/F | 36669 | 88482 | 8673 | 83049 | N/F | 2760858 | N/F | 64066 | 4693 | 1896 | 5763 | 2257 | 12964 | 44051 | 395691 | 716333 | 200694 | 1288854 | 5089784 | 1697629 | 203267 | 21307863 | 2508618 | 12461196 | 2465134 | 17761291 | 54739 | N/F | N/F | 314867 | N/F | 715 | 11261 | 183129 | 67303 | N/F | 3263 | N/F | N/F | 106950 | 50199 | 4814 | 61051 | 14145 | 7615 | N/F | 7228 | 87487 | N/F | 81124 |  | Unknown
WO-23003-81-A-1 | N/F | 40275 | 73708 | 6795 | 186171 | N/F | 4580446 | N/F | 108350 | N/F | 397 | 17604 | 4134 | 18571 | 79644 | 645820 | 1086927 | 360637 | 2247196 | 8751644 | 3165164 | 379155 | 37402474 | 4128737 | 21767220 | 4168274 | 29882331 | 93811 | N/F | N/F | 559477 | N/F | 968 | 17215 | 165211 | 107977 | N/F | 2923 | N/F | N/F | 147214 | 72777 | 13146 | 75513 | 21049 | 10288 | N/F | 8551 | 163694 | N/F | 95327 |  | Unknown
AB-230117-02_LFSM-H_4 | N/F | 75644 | 517165 | 696449 | 77253824 | N/F | 2984644 | N/F | 101765 | 701671 | 362620 | 3491492 | 2048727 | 3134915 | 69799 | 445395 | 615184 | 355626 | 2072653 | 8030274 | 3032232 | 346767 | 26066773 | 3306119 | 18871381 | 3159809 | 23253872 | 87624 | N/F | N/F | 549664 | N/F | 536 | 2536250 | 90678 | 27264006 | N/F | 2408 | 8042465 | 2785199 | 214580 | 21990528 | 3441522 | 18596951 | 62122 | 3574153 | N/F | 1257305 | 33567077 | N/F | 23155319 |  | Unknown
AB-230117-02_CCC-H_5 | N/F | 307698 | 613751 | 693538 | 121297746 | N/F | 2810612 | N/F | 459649 | 3142889 | 1511279 | 5522622 | 4467624 | 4792732 | 156257 | 341482 | 360765 | 857556 | 2362328 | 12581303 | 6803754 | 913930 | 26265725 | 3326480 | 21883588 | 2634839 | 23985052 | 442018 | N/F | N/F | 1310115 | 2635 | 1057 | 4099183 | 753193 | 45782815 | N/F | 240 | 22937199 | 11996790 | 304999 | 39602258 | 4809239 | 38569334 | 147176 | 6003753 | N/F | 4867710 | 58760997 | 671 | 38797049 |  | Unknown
AB-230117-02_LFB-H_5 | N/F | 188235 | 401534 | 497652 | 70350332 | N/F | 3540640 | N/F | 416079 | 1523645 | 759132 | 3285952 | 2613512 | 2737006 | 157270 | 397222 | 509752 | 830324 | 2845482 | 14855764 | 6832898 | 921954 | 32549183 | 3636391 | 25545598 | 3116264 | 28676707 | 399255 | N/F | N/F | 1328738 | 2247 | 986 | 2630197 | 344044 | 28781019 | 1013 | 216 | 13759351 | 6077429 | 193120 | 24041448 | 3117654 | 22857644 | 57124 | 3311879 | N/F | 2608302 | 31969395 | N/F | 25155507 |  | Unknown
AB-230117-02_LD_4_inj2 | N/F | 38291 | 74865 | 8665 | 3860 | N/F | 3324112 | N/F | 23978 | 2523 | 850 | 7807 | 1143 | 15194 | 39773 | 496866 | 919464 | 196260 | 1650355 | 6058693 | 1760888 | 119907 | 29138161 | 3441877 | 15897027 | 3415913 | 23339370 | 20917 | N/F | N/F | 312969 | N/F | 687 | 11237 | N/F | 555027 | N/F | 1755 | 34412 | N/F | 117461 | 23791 | 3700 | 51995 | 14236 | 940 | N/F | 8810 | 44239 | N/F | 75607 |  | Unknown
AB-230117-02_LD_4 | N/F | 36635 | 75221 | 8325 | N/F | N/F | 3259106 | N/F | 45206 | 4075 | N/F | 6757 | 1372 | 13468 | 51515 | 545968 | 985466 | 246215 | 1778462 | 6602078 | 2132653 | 175981 | 31087726 | 3496045 | 16990993 | 3622521 | 24656530 | 37954 | N/F | N/F | 380776 | N/F | 525 | 15239 | N/F | 557700 | N/F | 736 | 40485 | N/F | 147048 | 28322 | 1769 | 46727 | 15425 | 996 | N/F | 4340 | 47646 | N/F | 64070 |  | Unknown
AB-230117-02_LRB_9 | N/F | 43806 | 29790 | 7965 | N/F | N/F | 4743394 | N/F | 403384 | N/F | 743 | 2590 | 273 | 12941 | 162256 | 635067 | 915237 | 837006 | 3396911 | 17294441 | 7073193 | 831223 | 41477591 | 4294577 | 30238560 | 4029941 | 34633235 | 386958 | N/F | N/F | 1298730 | N/F | 475 | 14251 | N/F | 66408 | N/F | 522 | 38512 | 26942 | 158299 | 21751 | 1762 | 106449 | 16536 | 1060 | N/F | 9778 | 57279 | N/F | 33980 |  | Unknown
WO-23003-80-A-1 | N/F | 38102 | 83962 | 6354 | N/F | N/F | 4162855 | N/F | 54489 | N/F | 2790 | 3643 | 1152 | 10896 | 53066 | 565658 | 1227347 | 273658 | 1872359 | 7082416 | 2327100 | 199419 | 33311920 | 4033154 | 18043143 | 3979224 | 26625541 | 43395 | N/F | N/F | 423898 | N/F | 442 | 11964 | 84914 | 43224 | N/F | 4304 | 83278 | N/F | 171128 | 28667 | 1138 | 45498 | 17280 | 509 | N/F | 5769 | 51618 | N/F | 28337 |  | Unknown
WO-23003-79-A-1 | N/F | 41053 | 91339 | 7649 | N/F | N/F | 3771658 | N/F | 88679 | N/F | 1071 | 5038 | N/F | 13143 | 67436 | 751633 | 1300578 | 352862 | 2187823 | 8366060 | 2837187 | 297518 | 32949161 | 4090681 | 20666124 | 4033920 | 27348729 | 75986 | N/F | N/F | 527901 | N/F | 461 | 11075 | 143335 | 67083 | 1379 | 3411 | 79953 | N/F | 156368 | 36709 | 4103 | 57485 | 20199 | 1336 | N/F | 6648 | 46381 | N/F | 33233 |  | Unknown
WO-23003-78-A-1 | N/F | 37839 | 92682 | 6297 | N/F | N/F | 4191131 | N/F | 103394 | N/F | 400 | 5682 | N/F | 10643 | 75309 | 712237 | 1226427 | 354479 | 2259321 | 8830643 | 3047045 | 359893 | 33969588 | 4121518 | 20974092 | 4070156 | 28502611 | 87338 | N/F | N/F | 566963 | N/F | 608 | 11364 | 165687 | 49528 | N/F | 4524 | N/F | N/F | 154032 | 27406 | 3208 | 51312 | 19459 | 1226 | N/F | 7058 | 46479 | 1256 | 33700 |  | Unknown
WO-23003-77-A-1 | 1413 | 42162 | 77711 | 6835 | N/F | N/F | 4296232 | N/F | 113873 | 2539 | 717 | 3389 | 1898 | 11366 | 80098 | 702413 | 1067492 | 367850 | 2404844 | 9290518 | 3165726 | 372623 | 36194142 | 4131071 | 22688275 | 4163590 | 30051573 | 98954 | N/F | N/F | 593294 | N/F | 1148 | 11626 | 115120 | 72356 | N/F | 2461 | 32505 | N/F | 166724 | 40591 | 567 | 57090 | 18195 | 1652 | N/F | 7300 | 51679 | N/F | 35343 |  | Unknown
WO-23003-76-A-1 | N/F | 38848 | 79715 | 5004 | N/F | N/F | 4211035 | N/F | 98821 | N/F | 1382 | 7911 | 1351 | 15022 | 76075 | 667277 | 1152153 | 354298 | 2199167 | 8550553 | 3045691 | 325144 | 35596451 | 4144216 | 21155629 | 4064344 | 28576413 | 82684 | N/F | N/F | 553255 | N/F | 534 | 16033 | 91014 | 308167 | N/F | 3242 | N/F | N/F | 167889 | 33876 | 2995 | 49988 | 20067 | 848 | N/F | 7468 | 53338 | 1745 | 35940 |  | Unknown
WO-23003-75-A-1 | N/F | 47130 | 94980 | 6951 | N/F | N/F | 3049833 | N/F | 83507 | N/F | 1746 | 6246 | 1540 | 19837 | 68871 | 817940 | 1148460 | 346097 | 2037652 | 7754786 | 2759560 | 273566 | 28726809 | 3853319 | 19069664 | 3906238 | 25054025 | 70217 | N/F | N/F | 513406 | N/F | 295 | 10872 | N/F | 6579318 | N/F | 3479 | 149587 | N/F | 164657 | 48309 | 3851 | 52961 | 15749 | 777 | N/F | 6346 | 51633 | N/F | 32150 |  | Unknown
WO-23003-74-A-1 | N/F | 39192 | 74579 | 11921 | N/F | N/F | 3733174 | N/F | 108401 | N/F | N/F | 7967 | 1164 | 16075 | 77935 | 656935 | 1016146 | 363087 | 2248390 | 8660413 | 3016217 | 350115 | 33163671 | 4007846 | 20917351 | 3985448 | 28137426 | 94132 | N/F | N/F | 555461 | N/F | 1639 | 11095 | N/F | 5436768 | N/F | 2706 | N/F | N/F | 157426 | 34329 | 2751 | 48518 | 17290 | 1112 | N/F | 5268 | 47775 | N/F | 30740 |  | Unknown
WO-23003-73-A-1 | N/F | 40538 | 80964 | 10823 | 6483 | N/F | 3540604 | N/F | 82134 | N/F | 961 | 8188 | N/F | 16111 | 66858 | 603927 | 1090156 | 323190 | 1979874 | 7786665 | 2800246 | 283428 | 32476129 | 3990103 | 19586057 | 3936148 | 27116037 | 68866 | N/F | N/F | 484900 | N/F | 664 | 10970 | N/F | 2614769 | N/F | 3400 | 46548 | N/F | 159792 | 34021 | 5179 | 67060 | 15320 | 1019 | N/F | 6723 | 52244 | N/F | 34513 |  | Unknown
WO-23003-72-A-1 | N/F | 42212 | 104299 | 4758 | 7837 | N/F | 2812552 | N/F | 79615 | N/F | 531 | 7273 | 1441 | 17728 | 67336 | 880284 | 1262845 | 316272 | 2063857 | 7724853 | 2682500 | 262593 | 28298413 | 3708329 | 18645532 | 3652557 | 23551651 | 67634 | N/F | N/F | 494171 | N/F | 1638 | 9340 | N/F | 2684556 | N/F | 1121 | N/F | N/F | 142634 | 24156 | 4869 | 62260 | 14989 | 1076 | N/F | 7553 | 45004 | N/F | 24157 |  | Unknown
WO-23003-71-A-1 | N/F | 37151 | 79580 | 8125 | 6186 | N/F | 3506798 | N/F | 101254 | 1943 | N/F | 8846 | N/F | 12175 | 66904 | 664588 | 1024692 | 342097 | 2109933 | 8452869 | 2904367 | 300213 | 32070510 | 3912493 | 20358612 | 3924923 | 27434685 | 86431 | N/F | N/F | 521621 | N/F | 504 | 11737 | N/F | 2288947 | N/F | 2991 | N/F | N/F | 143354 | 29705 | 2638 | 49139 | 15822 | 2415 | N/F | 4605 | 52205 | N/F | 29638 |  | Unknown
AB-230117-02_LRB_8_20230120232821 | N/F | 45788 | 39184 | 3904 | 7974 | N/F | 3837556 | N/F | 358646 | N/F | 319 | 6277 | 2086 | 15015 | 146321 | 689894 | 800103 | 736670 | 3110107 | 15238182 | 6418043 | 686043 | 35232704 | 3788237 | 27135711 | 3791022 | 31335940 | 346903 | N/F | N/F | 1219569 | N/F | 1127 | 10662 | N/F | 46877 | 1030 | 427 | 22577 | N/F | 101059 | 26590 | 4695 | 76192 | 8706 | 1605 | N/F | 12442 | 50214 | N/F | 62278 |  | Unknown
AB-230117-02_LRB_8 | N/F | 45470 | 34712 | 5413 | 8942 | N/F | 3966551 | N/F | 359720 | 2419 | 681 | 6551 | 282 | 17682 | 156982 | 708782 | 865522 | 734076 | 3310690 | 15916496 | 6387645 | 745656 | 38150621 | 4183761 | 28461204 | 4128230 | 32962886 | 344404 | N/F | N/F | 1155057 | N/F | N/F | 13526 | N/F | 149199 | N/F | 229 | N/F | N/F | 134576 | 28417 | 4863 | 72203 | 14261 | 2052 | N/F | 11799 | 57896 | N/F | 53253 |  | Unknown
WO-23003-70-A-1 | N/F | 45173 | 72011 | 12687 | 8167 | N/F | 3518196 | N/F | 75845 | N/F | 2424 | 4034 | N/F | 17050 | 68661 | 588174 | 1028540 | 325383 | 2085017 | 8036131 | 2863392 | 270631 | 33314099 | 3892897 | 19678201 | 3802834 | 27121773 | 64949 | N/F | N/F | 523354 | N/F | 720 | 12194 | N/F | 2329145 | N/F | 2906 | 29568 | N/F | 147873 | 39412 | 3658 | 50952 | 16524 | 1961 | N/F | 8013 | 51197 | N/F | 62907 |  | Unknown
WO-23003-69-A-1 | N/F | 36117 | 69499 | 8439 | 10090 | N/F | 3645925 | N/F | 81183 | N/F | 584 | 8235 | 1010 | 16050 | 67118 | 571184 | 1024676 | 323597 | 2013235 | 7864338 | 2797238 | 237319 | 33335946 | 4046453 | 19684253 | 3994878 | 27044723 | 68747 | N/F | N/F | 505162 | N/F | 365 | 13067 | N/F | 1565467 | N/F | 2623 | 43899 | N/F | 150976 | 27313 | 5815 | 66575 | 15929 | 1963 | N/F | 5963 | 54231 | N/F | 72081 |  | Unknown
WO-23003-68-A-1 | N/F | 38072 | 104729 | 15040 | 11230 | N/F | 3056028 | N/F | 83911 | N/F | 1055 | 8405 | 1693 | 18271 | 69700 | 873192 | 1237500 | 313311 | 2117441 | 8067058 | 2893781 | 245906 | 29097804 | 3866376 | 19475658 | 3818392 | 24328483 | 72278 | N/F | N/F | 540784 | N/F | 275 | 10673 | N/F | 1459317 | N/F | 1571 | 137373 | N/F | 142698 | 33296 | 3668 | 52101 | 17221 | 2009 | N/F | 6306 | 51242 | N/F | 53900 |  | Unknown
WO-23003-67-A-1 | N/F | 37516 | 73767 | 5970 | 15940 | N/F | 3715637 | N/F | 84558 | N/F | 955 | 9135 | 1651 | 16398 | 69783 | 637269 | 1065555 | 315898 | 2198727 | 8506885 | 2982161 | 267423 | 33824164 | 4103182 | 20463965 | 4021916 | 27985442 | 72354 | N/F | N/F | 568400 | N/F | 626 | 12580 | N/F | 949598 | N/F | 2480 | 41186 | N/F | 152874 | 31621 | 2777 | 52641 | 16755 | 2806 | N/F | 7134 | 56979 | 768 | 44873 |  | Unknown
WO-23003-66-A-1 | N/F | 39261 | 84573 | 7686 | 18202 | 1046 | 4340512 | N/F | 94881 | N/F | 571 | 2572 | N/F | 18658 | 73805 | 740523 | 1110991 | 301071 | 2196066 | 8600488 | 3053264 | 283648 | 29837440 | 4199735 | 21091039 | 4144308 | 26677106 | 83558 | N/F | N/F | 575404 | N/F | 618 | 12831 | N/F | 828275 | N/F | 3520 | 27989 | N/F | 146301 | 30854 | 4815 | 51724 | 18136 | 3568 | N/F | 8616 | 52431 | N/F | 28461 |  | Unknown
WO-23003-65-A-1 | 1570 | 39124 | 83669 | 10955 | 23224 | N/F | 3266137 | N/F | 48721 | N/F | 1089 | 8340 | N/F | 17030 | 52388 | 681552 | 1135874 | 226076 | 1923390 | 7144751 | 2397431 | 172613 | 30651184 | 3933942 | 18142797 | 3902899 | 25240218 | 39898 | N/F | N/F | 449024 | N/F | 894 | 13849 | N/F | 1337561 | N/F | 678 | 44829 | 2651 | 153439 | 28808 | 6979 | 42589 | 18439 | 2667 | N/F | 6484 | 55407 | N/F | 30505 |  | Unknown
WO-23003-64-A-1 | N/F | 38138 | 64613 | 9837 | 30519 | N/F | 3540348 | N/F | 44213 | N/F | 791 | 10668 | 971 | 17396 | 49076 | 508028 | 926326 | 201717 | 1729750 | 6608632 | 2264569 | 140572 | 31846879 | 3905064 | 16931737 | 3776287 | 25722526 | 34007 | N/F | N/F | 423399 | N/F | 430 | 14653 | N/F | 1880229 | N/F | 2422 | 41689 | N/F | 158777 | 45387 | 7226 | 45122 | 16065 | 4579 | N/F | 5469 | 66451 | N/F | 74762 |  | Unknown
WO-23003-63-A-1 | N/F | 36004 | 67696 | 8968 | 45451 | N/F | 3829429 | N/F | 56945 | 1382 | 681 | 9425 | 1753 | 19475 | 60563 | 589679 | 1024970 | 245238 | 1889404 | 7221355 | 2596358 | 176897 | 33426157 | 4088866 | 19240240 | 4026837 | 27207511 | 47697 | N/F | N/F | 501018 | N/F | 392 | 15904 | N/F | 1522078 | N/F | 3484 | 51325 | N/F | 163453 | 42634 | 8271 | 52239 | 17863 | 4691 | N/F | 4038 | 76247 | N/F | 86294 |  | Unknown
WO-23003-62-A-1 | N/F | 37306 | 74484 | 8885 | 70423 | N/F | 3752867 | N/F | 69091 | N/F | 1788 | 9919 | 3103 | 20923 | 70215 | 625665 | 1079636 | 244470 | 2105867 | 8268549 | 2937539 | 211321 | 33327527 | 4027168 | 20501940 | 4020798 | 28012206 | 58574 | N/F | N/F | 533107 | N/F | 716 | 14347 | N/F | 980633 | N/F | 3715 | 30162 | N/F | 144824 | 46212 | 3461 | 66276 | 19679 | 5206 | N/F | 7169 | 86053 | N/F | 64381 |  | Unknown
WO-23003-61-A-1 | N/F | 38310 | 83062 | 16143 | 169893 | N/F | 3716879 | N/F | 49089 | N/F | 626 | 16747 | N/F | 24713 | 61495 | 614972 | 1086930 | 184116 | 1857673 | 7382427 | 2553636 | 173143 | 31165442 | 3882644 | 18165095 | 3818796 | 25807274 | 39650 | 1514 | N/F | 473310 | N/F | 445 | 24553 | N/F | 659319 | N/F | 1839 | 68951 | N/F | 130772 | 69055 | 14259 | 66110 | 15462 | 11978 | N/F | 6549 | 143976 | N/F | 108778 |  | Unknown
AB-230117-02_LFSM-H_3 | N/F | 56850 | 456647 | 643062 | 73707881 | N/F | 2726498 | N/F | 41829 | 578319 | 192546 | 3387854 | 1799709 | 3073156 | 55229 | 406774 | 600001 | 210395 | 1830236 | 7067342 | 2497078 | 206173 | 25558148 | 3377565 | 17167760 | 3212403 | 22029527 | 34135 | N/F | N/F | 468608 | N/F | 730 | 2287862 | 63012 | 25536709 | N/F | 3124 | 6921542 | 2362256 | 212301 | 20199923 | 3363898 | 16377660 | 60341 | 3573590 | N/F | 1098719 | 31569909 | N/F | 21796468 |  | Unknown
AB-230117-02_CCC-H_4 | N/F | 325555 | 670045 | 740961 | 120230961 | N/F | 2248001 | N/F | 258513 | 2951692 | 896575 | 5733532 | 4509367 | 4832420 | 135482 | 362711 | 364696 | 463618 | 2194755 | 11188051 | 5962507 | 545200 | 22953023 | 3114180 | 20043337 | 2555118 | 21827783 | 245197 | N/F | N/F | 1173296 | N/F | 353 | 3584230 | 451380 | 42633436 | N/F | 279 | 21725133 | 11117525 | 223974 | 38195355 | 4999985 | 37579983 | 72389 | 5846053 | N/F | 4734852 | 53125089 | N/F | 35077246 |  | Unknown
AB-230117-02_LFB-H_4 | N/F | 144101 | 454007 | 551792 | 68280380 | N/F | 2773144 | N/F | 252306 | 1504306 | 456291 | 3376106 | 2663538 | 2752725 | 142217 | 440656 | 502842 | 443262 | 2745742 | 13343664 | 6098605 | 582983 | 28066312 | 3332796 | 23171554 | 2951750 | 25512501 | 238833 | N/F | N/F | 1181283 | 904 | 232 | 2231474 | 208346 | 26451828 | N/F | 305 | 13253636 | 5693054 | 163076 | 23028556 | 2992807 | 21527843 | 38387 | 3138970 | N/F | 2572635 | 29132558 | N/F | 22970324 |  | Unknown
AB-230117-02_LD_3 | N/F | 40607 | 64993 | 4225 | 6501 | N/F | 3379100 | N/F | 37987 | N/F | 532 | 10460 | 1302 | 15109 | 58006 | 484238 | 841083 | 195206 | 1797806 | 6986338 | 2303037 | 178149 | 28505998 | 3264048 | 17701996 | 3312750 | 24386723 | 33217 | N/F | N/F | 398970 | N/F | 846 | 8487 | 25399 | 244830 | N/F | 1872 | 45214 | N/F | 57604 | 35649 | 1033 | 47772 | 9545 | 1098 | N/F | 8481 | 42795 | N/F | 30066 |  | Unknown
AB-230117-02_LRB_7_20230120154805 | 1263 | 15209 | 35579 | 2036 | 10032 | N/F | 192510 | N/F | 5418 | N/F | 530 | 4203 | N/F | 14687 | 4709 | 19612 | 44665 | 6978 | 78334 | 250544 | 78411 | 7836 | 1533140 | 146825 | 663256 | 154930 | 1096848 | 5123 | N/F | N/F | 12758 | N/F | 2152 | 2483 | N/F | 580373 | N/F | 26 | N/F | N/F | 18119 | 23992 | 1936 | 47828 | 6610 | 1567 | N/F | 13677 | 20833 | 1062 | 27134 |  | Unknown
AB-230117-02_LRB_7 | 2038 | 53970 | 31236 | 5536 | 11304 | N/F | 3537374 | N/F | 253227 | N/F | 386 | 5817 | 1297 | 18694 | 145586 | 611887 | 735305 | 501100 | 1875354 | 13421045 | 6075374 | 445401 | 34977386 | 3432771 | 26555056 | 3619684 | 30394983 | 240978 | N/F | N/F | 1135717 | N/F | 2126 | 11251 | N/F | 4665299 | N/F | 174 | 36282 | N/F | 62873 | 30176 | 4439 | 63583 | 5504 | 1968 | N/F | 12392 | 53584 | N/F | 82383 |  | Unknown
WO-23003-60-A-1 | N/F | 40697 | 45819 | 7707 | 17520 | N/F | 2298348 | N/F | 5085 | N/F | 904 | 7987 | 1318 | 11980 | 26657 | 372629 | 275521 | 99288 | 1576405 | 5395806 | 1307468 | 55113 | 9685815 | 2699768 | 15734824 | 2805663 | 21786661 | 3676 | N/F | N/F | 229596 | N/F | 818 | 10683 | N/F | 202156810 | N/F | 2020 | 108396 | N/F | 54824 | 288133 | 4453 | 152927 | 8818 | 1370 | N/F | 4011 | 55533 | 950 | 94826 |  | Unknown
WO-23003-59-A-1 | N/F | 43976 | 68020 | 3092 | 35296 | N/F | 2815865 | 46 | 45359 | N/F | N/F | 4105 | 2083 | 15373 | 52962 | 564596 | 532379 | 267109 | 2051000 | 7946611 | 2062473 | 143849 | 20147347 | 2954149 | 18992116 | 3414709 | 25958296 | 35406 | N/F | N/F | 404907 | N/F | 492 | 11901 | N/F | 65216052 | N/F | 2958 | 28261 | N/F | 69472 | 90322 | 4792 | 76702 | 11863 | 4632 | N/F | 9008 | 71072 | N/F | 92753 |  | Unknown
WO-23003-58-A-1 | N/F | 53117 | 66607 | 12998 | 60583 | N/F | 2967392 | 39 | 50012 | N/F | N/F | 12661 | 2596 | 21062 | 58400 | 505157 | 733475 | 191825 | 1777737 | 6905818 | 2336952 | 230470 | 27314869 | 2849235 | 16830488 | 3252455 | 24137122 | 43031 | N/F | N/F | 419100 | N/F | 543 | 12625 | N/F | 8133209 | N/F | 3131 | 54950 | N/F | 63211 | 137152 | 9962 | 58849 | 11751 | 4166 | N/F | 6244 | 86593 | N/F | 100105 |  | Unknown
WO-23003-138-A-1 | N/F | 22640 | 71618 | 4963 | N/F | N/F | 5714048 | N/F | 113300 | N/F | 90863 | 5267 | 4079 | 10255 | 81370 | 959218 | 1543897 | 379201 | 2401875 | 9433895 | 3349023 | 365920 | 35416341 | 4321380 | 22679273 | 4161902 | 28595325 | 96973 | N/F | N/F | 583191 | N/F | 602 | 14789 | N/F | 66502 | N/F | 3461 | 39268 | N/F | 159969 | 25178 | 1400 | 41085 | 15305 | 1720 | N/F | 6585 | 48708 | N/F | 28011 |  | Unknown
WO-23003-137-A-1 | N/F | 26465 | 77300 | 8334 | N/F | N/F | 5683525 | N/F | 151711 | N/F | 79951 | 4894 | N/F | 8338 | 89840 | 955129 | 1479764 | 456434 | 2499664 | 9881285 | 3548669 | 431510 | 35718603 | 4543633 | 22937395 | 4280646 | 29179604 | 132259 | 1113 | N/F | 629946 | N/F | 419 | 14611 | N/F | 70238 | N/F | 2918 | N/F | N/F | 165375 | 22172 | 921 | 54106 | 17720 | 1285 | N/F | 6417 | 44177 | N/F | 32367 |  | Unknown
WO-23003-136-A-1 | N/F | 22127 | 70840 | 8207 | N/F | N/F | 5377131 | N/F | 14069 | 1707 | 32484 | 4723 | 85 | 7985 | 35638 | 557647 | 1336808 | 156718 | 1386247 | 4975276 | 1406570 | 77642 | 30705787 | 3865038 | 13518015 | 3699614 | 22565939 | 10602 | N/F | N/F | 255224 | N/F | 517 | 14024 | N/F | 68033 | N/F | 1991 | N/F | N/F | 167303 | 16462 | 1316 | 27299 | 17368 | 741 | N/F | 3780 | 51889 | N/F | 25694 |  | Unknown
WO-23003-135-A-1 | N/F | 24477 | 71095 | 5038 | N/F | N/F | 4832413 | N/F | 119113 | N/F | 68330 | 2864 | 704 | 11655 | 72970 | 663416 | 1431970 | 351604 | 1568439 | 6554848 | 2814769 | 355541 | 30085287 | 3942866 | 14745212 | 3590842 | 22759139 | 103340 | 896 | N/F | 508344 | N/F | 165 | 10799 | N/F | 67432 | N/F | 2608 | 48168 | N/F | 165981 | 25240 | 1028 | 41601 | 14571 | 972 | N/F | 5583 | 46396 | 1138 | 21265 |  | Unknown
WO-23003-134-A-1 | N/F | 24527 | 88605 | 4900 | N/F | N/F | 5533954 | N/F | 116705 | 624 | 81449 | 5063 | 1024 | 8854 | 73249 | 829767 | 1748752 | 365967 | 2043914 | 7994462 | 3030173 | 367619 | 34656130 | 4726989 | 19418592 | 4282981 | 26244060 | 98231 | N/F | N/F | 543832 | N/F | 281 | 15000 | N/F | 72886 | 352 | 3441 | 19754 | N/F | 170367 | 20337 | 3026 | 50856 | 14721 | 1043 | N/F | 8274 | 47341 | N/F | 34518 |  | Unknown
WO-23003-133-A-1 | N/F | 3457 | 78269 | 9854 | N/F | 2045 | 4892994 | N/F | 73186 | 3078 | 52812 | 4138 | N/F | 9430 | 60825 | 582022 | 1242523 | 274414 | 1214269 | 5539638 | 2242924 | 245807 | 27027843 | 3543979 | 12101965 | 3200615 | 19143175 | 61954 | N/F | N/F | 406838 | N/F | 559 | 11276 | N/F | 71217 | 726 | 2926 | N/F | N/F | 161791 | 17824 | 741 | 31440 | 16882 | 1176 | N/F | 1346 | 44741 | 1353 | 31651 |  | Unknown
WO-23003-132-A-1 | N/F | 24412 | 72305 | 5768 | N/F | N/F | 5513665 | N/F | 38092 | N/F | 21350 | 3411 | N/F | 8120 | 48919 | 590867 | 1312655 | 237812 | 1478371 | 6017671 | 2055522 | 158852 | 33892710 | 4002897 | 15305524 | 3707534 | 24372583 | 31432 | N/F | N/F | 362798 | N/F | 327 | 13049 | N/F | 130371 | N/F | 2874 | N/F | N/F | 182212 | 155153 | 2668 | 96747 | 18465 | 1034 | N/F | 5428 | 48194 | N/F | 42759 |  | Unknown
WO-23003-131-A-1 | N/F | N/F | 58576 | 6062 | N/F | N/F | 6015082 | N/F | 111380 | 1054 | 27311 | 6142 | N/F | 7067 | 68496 | 655395 | 1271623 | 347480 | 1809424 | 7525095 | 2844506 | 341053 | 34747393 | 4159293 | 18207144 | 3956493 | 26549325 | 94356 | N/F | N/F | 530790 | N/F | 810 | 11666 | N/F | 129982 | N/F | 3308 | 22775 | N/F | 157890 | 26148 | 1692 | 47243 | 15283 | 786 | N/F | 3864 | 45513 | N/F | 27020 |  | Unknown
AB-230117-02_LRB_14_20230123152733 | 1050 | 26295 | 33956 | 1614 | N/F | N/F | 4987787 | N/F | 534411 | N/F | N/F | 1923 | 1453 | 8398 | 185531 | 654226 | 738541 | 984231 | 3160520 | 16995537 | 7990089 | 1061146 | 40930026 | 4078276 | 28640931 | 3665930 | 33465281 | 515255 | N/F | N/F | 1427915 | N/F | 415 | 13041 | N/F | 136180 | N/F | N/F | 27874 | 32432 | 123799 | 27972 | 1310 | 104436 | 14037 | 1130 | N/F | 7346 | 56038 | N/F | 32790 |  | Unknown
AB-230117-02_LRB_14 | N/F | 25334 | 33745 | 3472 | N/F | N/F | 5546533 | N/F | 552859 | N/F | 1337 | 3994 | 1962 | 8087 | 189896 | 728109 | 934100 | 970248 | 3645393 | 18338564 | 7939528 | 1092118 | 44684916 | 4437394 | 31459571 | 4127495 | 36416657 | 523742 | 1572 | N/F | 1423067 | N/F | 966 | 17661 | N/F | 152220 | N/F | 62 | 57116 | 30452 | 171266 | 25896 | 2350 | 100379 | 17426 | 1708 | N/F | 8293 | 53273 | 897 | 34496 |  | Unknown
WO-23003-130-A-1 | N/F | 24966 | 72673 | 6479 | N/F | N/F | 5013821 | N/F | 48031 | 1162 | 19435 | 2603 | N/F | 8552 | 56984 | 759322 | 1376809 | 263797 | 1830424 | 6867673 | 2212515 | 195110 | 32804320 | 4018962 | 17821794 | 3863138 | 25163894 | 37183 | N/F | N/F | 392233 | N/F | 313 | 15547 | N/F | 127522 | N/F | 1383 | 44304 | N/F | 166665 | 23579 | 2074 | 56465 | 15340 | 750 | N/F | 6646 | 45548 | N/F | 24257 |  | Unknown
WO-23003-129-A-1 | N/F | 23507 | 66874 | 7996 | N/F | N/F | 5134153 | N/F | 75829 | N/F | 18985 | 3631 | 637 | 11667 | 58825 | 467697 | 1141750 | 277829 | 1148923 | 5448219 | 2274633 | 257650 | 28034179 | 3583773 | 11565885 | 3174631 | 18874520 | 60467 | N/F | N/F | 408787 | N/F | 826 | 10124 | N/F | 161030 | 1329 | 2917 | 21013 | N/F | 164716 | 21990 | 1899 | 34916 | 16158 | 1122 | N/F | 2010 | 50498 | 1068 | 31022 |  | Unknown
WO-23003-128-A-1 | N/F | 18812 | 45827 | 5026 | N/F | N/F | 4521838 | N/F | 55745 | 1461 | 11438 | 4285 | 531 | 6493 | 50172 | 269409 | 704893 | 237811 | 934843 | 4814475 | 1982972 | 202431 | 23495238 | 2804990 | 9105334 | 2260065 | 14623735 | 45411 | N/F | N/F | 350118 | N/F | 556 | 12780 | N/F | 182487 | 1140 | 1431 | 19144 | N/F | 168441 | 17179 | 1111 | 33477 | 18839 | 1100 | N/F | 1166 | 54024 | 721 | 36177 |  | Unknown
WO-23003-127-A-1 | N/F | 21473 | 57751 | 7152 | N/F | N/F | 5124173 | N/F | 96709 | N/F | 14519 | 2172 | 1078 | 7406 | 60889 | 372386 | 977979 | 289169 | 1117778 | 5581262 | 2451430 | 307105 | 28731300 | 3410322 | 11281829 | 2945807 | 19103573 | 79591 | N/F | N/F | 446521 | N/F | 232 | 10830 | N/F | 267880 | N/F | 1747 | N/F | N/F | 163671 | 24903 | 1790 | 34051 | 17623 | 690 | N/F | 5109 | 51149 | N/F | 25144 |  | Unknown
WO-23003-126-A-1 | N/F | 18768 | 55411 | 5999 | N/F | N/F | 4323259 | N/F | 36613 | N/F | 9455 | 4677 | 661 | 7609 | 38435 | 272297 | 734369 | 184292 | 846387 | 4204367 | 1591971 | 142551 | 19273236 | 2465688 | 8130357 | 1976140 | 12752173 | 30509 | N/F | N/F | 292620 | N/F | 362 | 11958 | N/F | 447069 | N/F | 1539 | N/F | N/F | 156849 | 18785 | 1683 | N/F | 19381 | 1552 | N/F | 1334 | 57323 | 1523 | 24456 |  | Unknown
WO-23003-125-A-1 | N/F | 23650 | 78437 | 10344 | 3747 | N/F | 4651750 | N/F | 40013 | N/F | 16372 | 5374 | N/F | 10272 | 56069 | 799273 | 1335919 | 256722 | 1690241 | 6368431 | 2125778 | 159849 | 28262592 | 3845229 | 15625123 | 3598788 | 22372578 | 32552 | N/F | N/F | 367291 | N/F | 489 | 10809 | N/F | 6511456 | N/F | 2964 | N/F | N/F | 149126 | 33089 | 1446 | 51902 | 16838 | 1273 | N/F | 6853 | 50670 | N/F | 23675 |  | Unknown
WO-23003-121-A-1B | N/F | 23009 | 50176 | 6308 | 6891 | N/F | 3437692 | N/F | 207188 | 1268 | N/F | 3327 | N/F | 8109 | 92009 | 697381 | 390908 | 507402 | 2728083 | 11420592 | 4054371 | 548877 | 13605851 | 3981506 | 24419468 | 3747443 | 30006898 | 191988 | N/F | N/F | 745087 | N/F | 266 | 12075 | N/F | 169342682 | 1139 | 1767 | 178871 | 29376 | 104584 | 1086943 | 3317 | 1059285 | 10819 | 1245 | N/F | 9294 | 48739 | 784 | 42559 |  | Unknown
WO-23003-111-A-1B | N/F | 22987 | 215961 | 5057 | 8746 | N/F | 4416627 | N/F | 56235 | N/F | 5967 | 6483 | 1554 | 18040 | 58277 | 573944 | 598968 | 285769 | 2043807 | 7810096 | 2481005 | 228257 | 24339869 | 3970597 | 19681184 | 3769791 | 27636877 | 52232 | N/F | N/F | 445432 | N/F | 747 | 12013 | N/F | 88156947 | N/F | 1702 | 106086 | 21224 | 77852 | 291450 | 1371 | 460268 | 8147 | 2138 | N/F | 7769 | 55371 | 1239 | 35912 |  | Unknown
AB-230117-02_LRB_13_20230123105625 | N/F | 48 | 31003 | 2695 | 12976 | N/F | 4969519 | N/F | 500247 | 1324 | N/F | 4883 | 2644 | 8344 | 169339 | 618027 | 739289 | 899201 | 3126958 | 16592397 | 7254703 | 977390 | 37983588 | 3999587 | 27657980 | 3550476 | 32002182 | 477597 | N/F | N/F | 1377450 | N/F | 613 | 13969 | N/F | 622893 | N/F | 394 | 38693 | 28051 | 103917 | 27822 | 3275 | 105403 | 9754 | 1106 | N/F | 9059 | 60145 | 547 | 41851 |  | Unknown
AB-230117-02_LRB_12_20230123102918 | N/F | 51924 | 39787 | 1712 | 11402 | N/F | 38195 | N/F | 9141 | N/F | N/F | 1854 | 1260 | 8123 | 2499 | 5368 | 6186 | 7032 | 22189 | 107309 | 50717 | 10281 | 273976 | 25458 | 176137 | 24467 | 226469 | 7694 | N/F | N/F | 9823 | N/F | 661 | 700 | N/F | 1267000 | 1504 | 72 | 21364 | N/F | 11025 | 27721 | 3603 | 43623 | 6560 | 892 | N/F | 9575 | 29345 | 614 | 28837 | C | Unknown
AB-230117-02_QCS_2 | N/F | 204253 | 441359 | 603646 | 81574396 | N/F | 4420463 | 368 | 422987 | 1618969 | 947855 | 3470683 | 2709675 | 3151001 | 150030 | 363559 | 496802 | 802983 | 2333913 | 11682604 | 6179804 | 951336 | 29133809 | 3171892 | 20001675 | 2572196 | 22790999 | 401060 | N/F | 2142 | 1130478 | 1272 | 754 | 457901 | 425937 | 31453513 | N/F | N/F | 13269483 | 6529999 | 217306 | 24509136 | 429956 | 21781698 | 63781 | 378738 | 1042298 | 2843790 | 7168617 | N/F | 30110560 |  | Unknown
AB-230117-02_CCC-H_12 | 482 | 289264 | 554433 | 603771 | 109505664 | 1561 | 3107982 | N/F | 417924 | 2739489 | 1648536 | 4821065 | 3840892 | 8282863 | 135023 | 282032 | 290018 | 796744 | 1940837 | 9825469 | 5657886 | 970177 | 22029282 | 2691385 | 17254104 | 1870799 | 17173406 | 397349 | N/F | N/F | 1095778 | N/F | 314 | 4711087 | 752635 | 40232440 | N/F | 296 | 19466371 | 10779089 | 223959 | 29681762 | 4255560 | 32324222 | 80667 | 5001092 | 1884785 | 4503255 | 47603263 | 1165 | 33908553 |  | Unknown
WO-23003-212-A-1 | N/F | 19419 | N/F | 3222 | 5466 | 1194 | 5070144 | N/F | 477435 | 1700 | 8131 | 3080891 | 94447 | 31515941 | 164522 | 411196 | 580215 | 937794 | 3311931 | 16646158 | 7037727 | 1087278 | 37972771 | 1891630 | 26368237 | 818293 | 7639842 | 454164 | N/F | N/F | 1295538 | N/F | 300 | 13453 | N/F | 55514 | 1046 | 319 | 50215 | 27137 | 145798 | N/F | 194759 | 171241 | 14140 | 413 | N/F | 8545 | 52427 | 1310 | 29615 |  | Unknown
AB-230117-02_LRB_22 | N/F | 20776 | 21580 | 3009 | 5542 | N/F | 5414775 | N/F | 486154 | 1201 | N/F | 196503 | 7930 | 6682665 | 171590 | 576160 | 666014 | 956316 | 3315056 | 16895262 | 7312818 | 1124938 | 41046510 | 3648200 | 28520880 | 2760987 | 27771783 | 462212 | N/F | N/F | 1317309 | N/F | 519 | 13221 | N/F | 56623 | N/F | N/F | 30966 | 26495 | 145642 | 11911 | 17080 | 103592 | 6599 | 1402 | N/F | 8452 | 60896 | 728 | 34669 |  | Unknown
WO-23003-211-A-1 | N/F | 18241 | N/F | 3545 | 6295 | N/F | 5074136 | N/F | 476536 | N/F | 11928 | 3254453 | 95711 | 32536957 | 166175 | 415258 | 567747 | 925945 | 3336733 | 16561750 | 7069050 | 1128506 | 38808812 | 1924035 | 26980041 | 851737 | 7497363 | 453256 | N/F | N/F | 1336035 | N/F | 83 | 14156 | N/F | 53221 | N/F | 491 | 50261 | 31206 | 160001 | N/F | 203630 | 165269 | 14465 | 1071 | N/F | 5640 | 54935 | 1134 | 32574 |  | Unknown
WO-23003-210-A-1 | 1268 | 20087 | N/F | 2656 | 6746 | N/F | 4789669 | N/F | 438982 | 1066 | 12992 | 4235526 | 156049 | 32563042 | 159055 | 375087 | 561719 | 883799 | 3225027 | 16267255 | 6958152 | 1056920 | 37852841 | 1807117 | 25194154 | 765246 | 6993795 | 417165 | N/F | N/F | 1243565 | N/F | 496 | 12246 | N/F | 56277 | N/F | N/F | 55223 | 25410 | 156773 | N/F | 195119 | 207478 | 12164 | 542 | N/F | 6156 | 55722 | 405 | 39044 |  | Unknown
WO-23003-209-A-1 | N/F | 16521 | 55666 | 4033 | 6840 | N/F | 4838962 | N/F | 206665 | 1093 | 282685 | 17813585 | 2359020 | 44410429 | 91140 | 123396 | 606109 | 511993 | 2358730 | 10502171 | 3913333 | 622713 | 34154426 | 1368453 | 11487087 | 490738 | 3580538 | 189791 | N/F | N/F | 695284 | N/F | 287 | 13231 | N/F | 54296 | N/F | 241 | 504604 | 86410 | 145035 | N/F | 63903 | 1463923 | 10388 | 738 | 10985 | 20916 | 50729 | 1283 | 39549 |  | Unknown
WO-23003-208-A-1 | N/F | 15930 | 69997 | 3365 | 8842 | N/F | 4611754 | N/F | 150711 | 1534 | 237609 | 17234114 | 2163204 | 44379838 | 82548 | 114085 | 595119 | 463464 | 2194316 | 9255556 | 3484935 | 516429 | 34058358 | 1430752 | 11019572 | 534977 | 3671026 | 137549 | N/F | N/F | 641950 | N/F | 324 | 12584 | N/F | 45890 | N/F | 501 | 455636 | 84758 | 166427 | N/F | 61337 | 1375569 | 13978 | 491 | N/F | 19779 | 52614 | 1616 | 30435 |  | Unknown
WO-23003-207-A-1 | N/F | 17046 | 102659 | 6844 | 7940 | N/F | 4351378 | N/F | 149658 | N/F | 253584 | 17174020 | 2223808 | 45579388 | 82367 | 113003 | 600897 | 452265 | 2228845 | 9434078 | 3493807 | 522053 | 33979020 | 1387819 | 10652995 | 526947 | 3472443 | 133800 | N/F | N/F | 613027 | N/F | 344 | 13622 | N/F | 45362 | N/F | 1301 | 459228 | 78642 | 169542 | N/F | 61520 | 1423802 | 15579 | 977 | N/F | 15266 | 55974 | 722 | 40990 |  | Unknown
WO-23003-206-A-1 | N/F | 15853 | 94194 | 2851 | 11424 | 1393 | 4380456 | N/F | 171368 | 1987 | 279986 | 20102645 | 2573993 | 52157495 | 88220 | 112528 | 579466 | 503077 | 2217176 | 9988571 | 3682234 | 597765 | 31615078 | 1158911 | 9117678 | 396341 | 2571564 | 157626 | 1160 | N/F | 674686 | N/F | 457 | 13305 | N/F | 50005 | N/F | 486 | 591642 | 95677 | 142309 | N/F | 84171 | 1609566 | 12972 | 841 | N/F | 19419 | 54321 | 1124 | 41815 |  | Unknown
WO-23003-204-A-1 | N/F | 19091 | 232622 | 8059 | 14706 | N/F | 4457283 | N/F | 54720 | 1285 | 352169 | 21427404 | 3471206 | 49172875 | 55157 | 86133 | 896529 | 289048 | 1553806 | 6530989 | 2327833 | 244962 | 24559170 | 1094035 | 6360974 | 409960 | 2250114 | 44987 | N/F | N/F | 400465 | N/F | 289 | 9865 | N/F | 39040 | N/F | 944 | 373127 | 60363 | 147629 | N/F | 58207 | 941185 | 10345 | 628 | N/F | 29248 | 52330 | 1052 | 38913 |  | Unknown
WO-23003-203-A-1 | N/F | 15339 | 83037 | 4319 | 21224 | N/F | 4408325 | N/F | 41609 | 2572 | 297646 | 21399204 | 3196260 | 49723305 | 45012 | 81385 | 914523 | 247272 | 1473104 | 6136837 | 2082824 | 190796 | 24116360 | 1042863 | 5989206 | 363290 | 2065247 | 36289 | N/F | N/F | 361261 | N/F | 565 | 12044 | N/F | 44092 | 1092 | 437 | 448061 | 68990 | 151079 | N/F | 61176 | 1188738 | 9120 | 3290 | N/F | 23298 | 57431 | N/F | 36132 |  | Unknown
WO-23003-202-A-1 | N/F | 3310 | 40998 | 5869 | 52014 | N/F | 5499305 | 26 | 93189 | N/F | 304201 | 11494993 | 1379644 | 32216989 | 68950 | 190759 | 1033057 | 348026 | 1808625 | 7524275 | 2742275 | 354772 | 31378239 | 1780291 | 13140881 | 854523 | 6111114 | 80081 | N/F | N/F | 481883 | N/F | 387 | 15766 | N/F | 63395 | N/F | 1015 | 256110 | 55309 | 153111 | N/F | 31367 | 738795 | 13702 | 1891 | N/F | 19085 | 73874 | N/F | 56083 |  | Unknown
WO-23003-201-A-1 | N/F | 19192 | 34513 | 4056 | 123587 | N/F | 5353869 | N/F | 79493 | 4089 | 306194 | 11732335 | 1307838 | 32425999 | 62493 | 187388 | 989081 | 344747 | 1756825 | 7393348 | 2646422 | 334233 | 30882941 | 1754895 | 12833795 | 824947 | 6282357 | 65972 | 1654 | N/F | 454091 | N/F | 509 | 19440 | N/F | 93716 | 1277 | 1143 | 260033 | 48550 | 132485 | 20799 | 42662 | 728548 | 12476 | 5597 | N/F | 17364 | 124888 | 1086 | 112740 |  | Unknown
AB-230117-02_LFSM-H_10 | N/F | 70953 | 204025 | 630315 | 57277270 | N/F | 4038662 | N/F | 103082 | 615482 | 688288 | 12265304 | 2920422 | 31792958 | 69780 | 178413 | 615366 | 383532 | 1727810 | 6980925 | 2898810 | 392898 | 22408400 | 1452498 | 12232708 | 831195 | 5712771 | 88543 | N/F | N/F | 525551 | N/F | 349 | 3130516 | 94035 | 21394431 | N/F | 1773 | 6873541 | 2715295 | 205138 | 5076872 | 2779753 | 11647483 | 60066 | 1700508 | 360721 | 1145766 | 28175259 | 2351 | 19999415 |  | Unknown
AB-230117-02_CCC-H_11 | N/F | 268222 | 513983 | 529486 | 112284316 | N/F | 3566690 | N/F | 483672 | 2857702 | 1774053 | 4539835 | 3798387 | 4060758 | 140379 | 263608 | 279671 | 906117 | 2011273 | 10596408 | 6258063 | 1091221 | 22917705 | 3086199 | 17894682 | 2175740 | 19829498 | 460334 | N/F | N/F | 1221504 | N/F | 491 | 4986840 | 784632 | 40559193 | 907 | 50 | 19343685 | 11049893 | 231831 | 32857948 | 3977206 | 31793768 | 74378 | 5518390 | 2004953 | 4345397 | 49147562 | N/F | 37518346 |  | Unknown
AB-230117-02_LFB-H_11 | N/F | 170024 | 371829 | 420457 | 64756087 | N/F | 4883870 | N/F | 442155 | 1391869 | 863970 | 3366270 | 2279066 | 5924030 | 149140 | 357736 | 428283 | 839595 | 2303348 | 12023295 | 6137861 | 982929 | 26941243 | 3014461 | 20602763 | 2379547 | 21417114 | 409354 | N/F | N/F | 1116067 | N/F | 1025 | 3695826 | 364703 | 24627744 | N/F | 143 | 11601376 | 5689178 | 168253 | 18369737 | 2356765 | 18917556 | 36302 | 2918014 | 926579 | 2537970 | 26231482 | 783 | 22540694 |  | Unknown
AB-230117-02_LD_10 | N/F | 18272 | 30213 | 12258 | N/F | N/F | 4836993 | N/F | 13709 | 1457 | 119322 | 9963744 | 832213 | 32349720 | 32101 | 136780 | 936647 | 153826 | 1171678 | 4465287 | 1269400 | 92056 | 26862747 | 1353161 | 9817354 | 825097 | 5515946 | 11231 | N/F | N/F | 219738 | N/F | 560 | 11785 | N/F | 37934 | 2735 | 1936 | 159365 | 21506 | 146587 | N/F | 27535 | 536224 | 11910 | 1299 | N/F | 11112 | 45455 | 823 | 34180 |  | Unknown
AB-230117-02_LRB_21_20230129012547 | 448 | 19044 | 29776 | 4928 | N/F | N/F | 4661309 | N/F | 368862 | 348 | 920 | 41098 | 2931 | 153907 | 147842 | 474214 | 601928 | 869335 | 2577534 | 14108098 | 6373522 | 929339 | 32716061 | 3282778 | 23046320 | 2926896 | 26813713 | 354881 | N/F | N/F | 1185440 | 862 | 203 | 11024 | N/F | 54481 | 1646 | N/F | 33752 | 22000 | 101044 | 17796 | 1172 | 87678 | 4398 | 733 | N/F | 9934 | 41623 | N/F | 33136 |  | Unknown
AB-230117-02_LRB_21 | N/F | 22881 | 25148 | 8435 | N/F | N/F | 5480285 | N/F | 459951 | N/F | N/F | 598753 | 33889 | 4144237 | 167395 | 582188 | 734722 | 904832 | 3249198 | 17016118 | 7134991 | 1115417 | 42486830 | 3895662 | 28451576 | 3086215 | 29645595 | 437337 | N/F | N/F | 1252977 | N/F | 172 | 11208 | N/F | 63670 | 967 | 55 | 30678 | 31970 | 152410 | 19148 | 2979 | 114533 | 6702 | 707 | N/F | 7995 | 53048 | 559 | 33275 |  | Unknown
WO-23003-200-A-1 | N/F | 15533 | 20406 | 2924 | 2840 | N/F | 5430894 | N/F | 91589 | N/F | 305820 | 11327929 | 1305896 | 32182845 | 66944 | 171366 | 853529 | 366120 | 1842376 | 7731473 | 2804146 | 357210 | 34422132 | 1947721 | 14015330 | 896167 | 7062188 | 75617 | N/F | N/F | 489517 | N/F | 215 | 12377 | N/F | 37273 | 948 | 403 | 321363 | 63442 | 158357 | N/F | 27306 | 956748 | 15421 | 192 | N/F | 19489 | 46140 | N/F | 28762 |  | Unknown
WO-23003-199-A-1 | 604 | 19323 | 26449 | 5155 | N/F | 1006 | 5919865 | N/F | 117446 | 2140 | 380486 | 8530076 | 1156529 | 24121155 | 76961 | 324737 | 1088913 | 401948 | 2049307 | 8530274 | 3228719 | 462322 | 33610321 | 2375162 | 16534708 | 1399210 | 10973119 | 103257 | N/F | N/F | 569404 | N/F | 339 | 10843 | N/F | 43221 | N/F | 1241 | 232757 | 49058 | 165054 | 7386 | 14390 | 564085 | 13400 | 670 | N/F | 19712 | 45117 | 1115 | 31587 |  | Unknown
WO-23003-198-A-1 | 1236 | 16277 | 31790 | 5987 | N/F | 962 | 5025130 | N/F | 29450 | 1209 | 208593 | 12431155 | 1529477 | 33122706 | 39771 | 161211 | 1077263 | 200546 | 1312846 | 5297880 | 1671541 | 130518 | 27215197 | 1603065 | 9768776 | 821874 | 5348640 | 22061 | N/F | 3670 | 280313 | N/F | 632 | 9453 | N/F | 33696 | 1215 | 1133 | 224934 | 36742 | 168458 | 8460 | 24677 | 724754 | 15985 | 274 | N/F | 16796 | 42085 | 1744 | 22934 |  | Unknown
WO-23003-197-A-1 | 1526 | 15549 | 34712 | 6383 | N/F | 821 | 5331614 | N/F | 80004 | 3691 | 366496 | 14037953 | 2120324 | 33654683 | 65794 | 133657 | 806799 | 356339 | 1753471 | 7263200 | 2828714 | 328986 | 33394941 | 1835139 | 11764041 | 818531 | 6191372 | 68455 | N/F | N/F | 480511 | N/F | 274 | 12520 | N/F | 41671 | N/F | 1091 | 288228 | 53340 | 169628 | N/F | 26434 | 817122 | 15311 | 206 | N/F | 20853 | 47597 | N/F | 27613 |  | Unknown
WO-23003-196-A-1 | N/F | 16346 | 46070 | 6034 | 3682 | N/F | 5840777 | N/F | 39933 | 2787 | 62666 | 992694 | 86302 | 7464271 | 48782 | 380038 | 879971 | 250865 | 1581333 | 6057717 | 2018624 | 201306 | 34051812 | 3422189 | 15600340 | 2732450 | 20086012 | 35397 | N/F | N/F | 346963 | N/F | 366 | 11800 | N/F | 33805 | N/F | 2977 | 96976 | 10320 | 163803 | 15138 | 4401 | 74703 | 14399 | 800 | N/F | 5753 | 50055 | 1180 | 41679 |  | Unknown
WO-23003-195-A-1 | N/F | 17344 | 43781 | 5562 | N/F | 1386 | 5883393 | N/F | 77675 | 875 | 80332 | 1123101 | 100803 | 7419197 | 60371 | 467396 | 951624 | 345615 | 1913407 | 7573851 | 2680389 | 325336 | 35999627 | 3446401 | 18682395 | 2779978 | 21631474 | 64274 | N/F | N/F | 483473 | N/F | 184 | 12991 | N/F | 38050 | N/F | 2028 | 44059 | 14517 | 170354 | N/F | 3804 | 94285 | 14539 | 1121 | N/F | 4900 | 44098 | 768 | 34210 |  | Unknown
WO-23003-194-A-1 | N/F | 19255 | 45450 | 7060 | 3255 | N/F | 6166172 | N/F | 97597 | 3990 | 105819 | 1192217 | 107192 | 7836621 | 69391 | 538330 | 1092333 | 383637 | 2008907 | 8124924 | 2968645 | 362662 | 35050320 | 3469696 | 19273079 | 2755911 | 21045555 | 81652 | N/F | N/F | 509360 | N/F | 395 | 11051 | N/F | 38492 | 1613 | 1421 | 63898 | 11957 | 160526 | 17871 | 4215 | 101817 | 13435 | 665 | N/F | 4061 | 45384 | 740 | 29874 |  | Unknown
WO-23003-193-A-1 | N/F | 18269 | 43431 | 4115 | N/F | N/F | 5875707 | N/F | 77905 | 578 | 72626 | 842913 | 77144 | 6135963 | 62469 | 509278 | 967140 | 353054 | 1916812 | 7536871 | 2720983 | 328437 | 36116897 | 3554282 | 18691242 | 2918698 | 22666720 | 64886 | N/F | N/F | 481457 | N/F | 530 | 13755 | N/F | 31500 | N/F | 1700 | 48981 | 10963 | 168995 | 17257 | 4506 | 73375 | 16234 | 897 | N/F | 4577 | 48748 | N/F | 23420 |  | Unknown
WO-23003-192-A-1 | N/F | 17008 | 40866 | 7763 | N/F | N/F | 5956394 | N/F | 81339 | N/F | 169397 | 1071079 | 129371 | 5888696 | 61431 | 464571 | 1038094 | 358394 | 1938047 | 7685883 | 2784636 | 330002 | 35628495 | 3587671 | 18266146 | 2910447 | 22132711 | 67594 | N/F | N/F | 491857 | N/F | 380 | 10311 | N/F | 33166 | N/F | 1869 | 61063 | 15170 | 181493 | 23894 | 3653 | 95026 | 14357 | 1484 | N/F | 6308 | 44811 | N/F | 40344 |  | Unknown
WO-23003-191-A-1 | N/F | 15805 | 45192 | 5134 | N/F | 1019 | 5305642 | N/F | 20738 | 2244 | 81713 | 854251 | 91272 | 5402853 | 34592 | 375770 | 979005 | 182248 | 1279172 | 4983417 | 1507161 | 116649 | 31907415 | 3289420 | 13583557 | 2775798 | 19573895 | 15693 | N/F | N/F | 258299 | N/F | 786 | 10671 | N/F | 28528 | 2216 | 1477 | 91502 | 11261 | 168909 | 18004 | 2251 | 67846 | 16098 | 790 | N/F | 3713 | 51335 | 952 | 35571 |  | Unknown
AB-230117-02_LRB_20_20230128200047 | N/F | 21860 | 24698 | 1993 | 7853 | N/F | 5313269 | N/F | 452410 | 1315 | N/F | 4327 | 1558 | 14661 | 167592 | 588226 | 680975 | 968936 | 3149172 | 16405925 | 7324266 | 1109714 | 39648533 | 4020201 | 28240850 | 3481649 | 31973298 | 437246 | N/F | N/F | 1374949 | N/F | 218 | 15117 | N/F | 57752 | N/F | 243 | 45513 | 29633 | 152191 | 21965 | 1083 | 91655 | 12865 | 1027 | N/F | 8236 | 54712 | N/F | 30553 |  | Unknown
AB-230117-02_LRB_20 | 883 | 18912 | 30241 | 2231 | 7216 | N/F | 5681547 | N/F | 461591 | N/F | N/F | 37585 | 5775 | 251705 | 169333 | 586673 | 727335 | 957130 | 3272710 | 17037185 | 7502454 | 1143863 | 41119870 | 4158859 | 28880873 | 3544377 | 32546918 | 439978 | N/F | N/F | 1357430 | N/F | 432 | 14621 | N/F | 64856 | 807 | 237 | 36366 | 27426 | 169109 | 22722 | 2129 | 97222 | 13852 | 1899 | N/F | 8333 | 58129 | 746 | 40901 |  | Unknown
WO-23003-190-A-1 | N/F | 17753 | 43830 | 5521 | 9986 | N/F | 5312904 | N/F | 48964 | N/F | 147146 | 1107758 | 128322 | 6160498 | 51364 | 423956 | 940632 | 293861 | 1684141 | 6778091 | 2251107 | 229402 | 33587846 | 3588782 | 17055709 | 2896586 | 21148930 | 39881 | N/F | N/F | 394071 | N/F | 349 | 12122 | N/F | 65791 | N/F | 2595 | 46970 | 17026 | 171524 | 23060 | 4369 | 106894 | 14640 | 1391 | N/F | 5463 | 52839 | 568 | 35349 |  | Unknown
WO-23003-189-A-1 | N/F | 18329 | 51966 | 2645 | 9870 | N/F | 5798846 | N/F | 100311 | 2484 | 202440 | 1502818 | 190058 | 7520299 | 66707 | 573473 | 1151973 | 382234 | 2040648 | 8009665 | 2905790 | 379687 | 34075321 | 3460972 | 18794074 | 2918972 | 21142994 | 85580 | N/F | N/F | 512194 | N/F | 226 | 11627 | N/F | 65867 | N/F | 2347 | 112033 | 23597 | 167181 | 37239 | 5788 | 107755 | 14500 | 1383 | N/F | 13440 | 45746 | 2041 | 27851 |  | Unknown
WO-23003-188-A-1 | N/F | 17461 | 53949 | 5467 | 9911 | N/F | 5461830 | N/F | 82363 | N/F | 11997 | 107694 | 8947 | 1246448 | 62805 | 537514 | 956384 | 351928 | 1994692 | 7913775 | 2742557 | 339425 | 36804811 | 4020680 | 19244736 | 3607841 | 26529526 | 69449 | N/F | N/F | 484668 | N/F | 364 | 12425 | N/F | 39256 | 1138 | 2753 | 70396 | N/F | 170338 | 24493 | 1327 | 47615 | 16149 | 1639 | N/F | 2516 | 57931 | N/F | 46490 |  | Unknown
WO-23003-187-A-1 | N/F | 16312 | 49524 | 7741 | 10080 | 1442 | 5811317 | N/F | 82884 | N/F | 7269 | 91078 | 8413 | 1089742 | 68118 | 524589 | 999776 | 355746 | 1924367 | 7665300 | 2762067 | 344564 | 36518213 | 3997412 | 18966931 | 3637597 | 26285689 | 67936 | N/F | N/F | 474748 | N/F | 539 | 11240 | N/F | 43855 | N/F | 2566 | 18145 | N/F | 162875 | 17240 | 6008 | 40138 | 14766 | 1954 | N/F | 5789 | 54127 | N/F | 39627 |  | Unknown
WO-23003-186-A-1 | N/F | 20291 | 56233 | 2637 | 22171 | N/F | 5875287 | N/F | 89518 | 746 | 8764 | 94955 | 5730 | 1107808 | 72680 | 612699 | 1063101 | 361916 | 2032877 | 8345047 | 2949424 | 367737 | 35994281 | 3993570 | 19807118 | 3634040 | 25912912 | 74934 | N/F | N/F | 529083 | N/F | 395 | 12949 | N/F | 42108 | N/F | 1796 | 16895 | N/F | 171658 | 19701 | 3031 | 51814 | 15437 | 2015 | N/F | 7334 | 54059 | 1871 | 34912 |  | Unknown
WO-23003-185-A-1 | N/F | 19636 | 53015 | 5749 | 20589 | N/F | 5742435 | N/F | 90247 | 1589 | 26003 | 78216 | 8677 | 758587 | 71130 | 575329 | 1005008 | 380600 | 1990488 | 8085328 | 2904735 | 378090 | 35863933 | 4113732 | 19844099 | 3640898 | 26407794 | 77847 | N/F | N/F | 506374 | N/F | 341 | 15714 | N/F | 45017 | N/F | 1768 | N/F | 5834 | 161258 | 33578 | 3663 | 52530 | 15543 | 2521 | N/F | 8199 | 65089 | 871 | 51651 |  | Unknown
WO-23003-184-A-1 | N/F | 21220 | 54208 | 5408 | 26340 | N/F | 5555022 | N/F | 89785 | 2103 | 29062 | 98810 | 12248 | 861360 | 67579 | 614094 | 983345 | 403204 | 2344458 | 9300913 | 3094685 | 379749 | 35514802 | 4197163 | 21936335 | 3691711 | 27768915 | 76146 | N/F | N/F | 537967 | N/F | 165 | 14145 | N/F | 45911 | 1748 | 2047 | N/F | N/F | 162701 | 26832 | 2284 | 62005 | 14646 | 4062 | N/F | 8487 | 61122 | N/F | 42122 |  | Unknown
WO-23003-183-A-1 | N/F | 18983 | 53834 | 6023 | 40390 | 1044 | 5873467 | N/F | 96783 | 1902 | 31949 | 95078 | 11393 | 905281 | 70533 | 532112 | 1022261 | 379271 | 1920242 | 7749766 | 2915834 | 379708 | 36206525 | 4082820 | 19124889 | 3641152 | 26049925 | 82736 | 741 | N/F | 520642 | N/F | 73 | 11767 | N/F | 60280 | N/F | 2378 | N/F | N/F | 171342 | 24340 | 7859 | 51851 | 13878 | 4523 | N/F | 5706 | 72525 | N/F | 47103 |  | Unknown
WO-23003-182-A-1 | N/F | 19964 | 54736 | 4336 | 70544 | N/F | 5591705 | N/F | 87411 | 1022 | N/F | 29536 | 3616 | 299928 | 70023 | 607631 | 1037320 | 376204 | 2053163 | 8094669 | 2889862 | 345390 | 35678436 | 4108076 | 19655372 | 3699468 | 25915073 | 73792 | N/F | N/F | 498288 | N/F | 956 | 16708 | N/F | 57922 | N/F | 1758 | 21207 | N/F | 160905 | 35686 | 3837 | 60630 | 14609 | 5292 | N/F | 3763 | 84746 | N/F | 57621 |  | Unknown
WO-23003-181-A-1 | N/F | 19135 | 48323 | 2369 | 166338 | N/F | 5675313 | 38 | 92596 | 2822 | N/F | 30185 | 4652 | 321537 | 66458 | 546008 | 946983 | 387068 | 2055317 | 8126718 | 2959952 | 373765 | 35900244 | 3985933 | 19786204 | 3623051 | 26881852 | 79367 | N/F | N/F | 520554 | N/F | 568 | 25276 | N/F | 95090 | 1456 | 1605 | N/F | N/F | 155041 | 53566 | 7224 | 75533 | 16701 | 10481 | N/F | 5263 | 158450 | N/F | 99874 |  | Unknown
AB-230117-02_LFSM-H_9 | N/F | 32384 | 267992 | 448033 | 66906671 | N/F | 3856330 | N/F | 17742 | 284813 | 166334 | 2410899 | 1153267 | 2745748 | 30225 | 243236 | 438694 | 165991 | 1255344 | 4647554 | 1389125 | 100715 | 25494037 | 3190817 | 12329974 | 2649459 | 18646924 | 14244 | N/F | N/F | 242322 | N/F | 183 | 3088604 | 13110 | 23688789 | 701 | 1972 | 4399134 | 1239693 | 240014 | 16456350 | 2940918 | 11241430 | 62525 | 3180136 | 88281 | 659251 | 32990263 | N/F | 23187317 |  | Unknown
AB-230117-02_CCC-H_10 | N/F | 290172 | 523343 | 526720 | 113452906 | N/F | 4211550 | N/F | 519751 | 2910868 | 1806060 | 4429739 | 3951852 | 4203307 | 153719 | 289518 | 301436 | 952125 | 2100146 | 11053440 | 6632334 | 1163934 | 23743526 | 3304165 | 18727431 | 2348444 | 20658508 | 493383 | N/F | N/F | 1279916 | N/F | 207 | 5748985 | 819684 | 39821091 | 871 | N/F | 19088423 | 11549089 | 256409 | 33231224 | 4163821 | 31886679 | 89069 | 5653111 | 2099382 | 4562258 | 53680169 | N/F | 37459508 |  | Unknown
AB-230117-02_LFB-H_10 | N/F | 164708 | 338197 | 371151 | 65671209 | N/F | 4243422 | N/F | 497557 | 1516092 | 909153 | 2857946 | 2190806 | 2212065 | 158320 | 363146 | 417744 | 985252 | 2433931 | 13098976 | 7015382 | 1154004 | 29407136 | 3513343 | 22259117 | 2675897 | 24459732 | 474655 | N/F | N/F | 1319412 | N/F | 302 | 2927973 | 390545 | 24982872 | 696 | 39 | 11813877 | 5923381 | 204849 | 19498312 | 2520733 | 19062464 | 45627 | 2994708 | 1035459 | 2469719 | 28294699 | 1207 | 23566734 |  | Unknown
AB-230117-02_LD_9 | N/F | 17897 | 53592 | 7994 | N/F | N/F | 5428884 | N/F | 57753 | 1662 | N/F | 23938 | N/F | 306434 | 56838 | 568906 | 957577 | 317353 | 1834832 | 6972579 | 2457924 | 264953 | 33893145 | 4223106 | 18049984 | 3726813 | 25240310 | 49183 | N/F | N/F | 423819 | 1153 | 310 | 14358 | N/F | 41624 | 601 | 1514 | N/F | N/F | 152562 | 20532 | 3695 | 42711 | 7181 | 1031 | N/F | 4257 | 46272 | 645 | 29160 |  | Unknown
AB-230117-02_LRB_19_20230128122014 | N/F | 20459 | 28787 | 4756 | N/F | N/F | 5586287 | N/F | 494139 | N/F | N/F | 2365 | 1417 | 4320 | 172302 | 598899 | 703044 | 979270 | 3210571 | 16560517 | 7657476 | 1200068 | 39049072 | 4146167 | 27607743 | 3510414 | 31221498 | 475072 | N/F | N/F | 1364602 | N/F | 361 | 12275 | N/F | 57681 | N/F | 87 | 28202 | 30422 | 142486 | 26466 | 2099 | 91748 | 6970 | 850 | N/F | 8529 | 48293 | N/F | 35267 |  | Unknown
AB-230117-02_LRB_19 | N/F | 19613 | 29780 | 4606 | N/F | 668 | 5404909 | N/F | 484961 | 1366 | N/F | 2748 | 897 | 20673 | 159754 | 613166 | 881002 | 975650 | 3402141 | 16798319 | 7363776 | 1179640 | 40849895 | 4040698 | 28767903 | 3416155 | 32008275 | 462613 | N/F | N/F | 1333367 | N/F | 48 | 12497 | N/F | 59145 | N/F | 461 | 37173 | 32143 | 177734 | 23030 | 1270 | 96571 | 13930 | 1280 | N/F | 9718 | 46409 | N/F | 28132 |  | Unknown
WO-23003-180-A-1 | N/F | 20313 | 70873 | 5818 | N/F | N/F | 4942557 | N/F | 51502 | N/F | N/F | 20820 | N/F | 276511 | 53407 | 468558 | 1209231 | 266808 | 1270011 | 5428612 | 2117294 | 227330 | 26608435 | 3586635 | 12516657 | 3073816 | 18800561 | 41414 | N/F | N/F | 351343 | N/F | 338 | 9519 | N/F | 31590 | N/F | 1490 | N/F | N/F | 168473 | 16426 | 1484 | 30908 | 14922 | 1588 | N/F | 3428 | 38074 | 857 | 28712 |  | Unknown
WO-23003-179-A-1 | N/F | 13037 | 44232 | 6409 | N/F | N/F | 3958668 | N/F | 50030 | 1029 | N/F | 6214 | N/F | 81028 | 41144 | 176462 | 497361 | 235710 | 797813 | 4088557 | 1727550 | 209111 | 19752521 | 2471074 | 7635678 | 1761842 | 11761625 | 38951 | N/F | N/F | 306279 | N/F | 360 | 11686 | N/F | 25500 | 1291 | 1754 | 58461 | N/F | 164177 | N/F | 1599 | 20186 | 18132 | 589 | N/F | 1162 | 49869 | 966 | 24049 |  | Unknown
WO-23003-178-A-1 | N/F | 14936 | 57427 | 10542 | N/F | N/F | 5124041 | N/F | 61954 | 435 | N/F | 11370 | 1528 | 145760 | 50033 | 369986 | 961501 | 287900 | 1080273 | 5047723 | 2095557 | 244182 | 26034693 | 3434464 | 10682658 | 2808552 | 17221198 | 49961 | 1353 | 1450 | 378797 | N/F | 694 | 10818 | N/F | 28199 | 544 | 1793 | 65457 | N/F | 162124 | 13111 | 942 | 17024 | 18038 | 823 | N/F | 1599 | 39643 | 808 | 29092 |  | Unknown
WO-23003-177-A-1 | N/F | 17204 | 57413 | 8704 | N/F | N/F | 5755279 | N/F | 58342 | N/F | N/F | 19165 | 2039 | 233231 | 54269 | 467704 | 946314 | 304495 | 1708034 | 6642716 | 2401343 | 249877 | 35158130 | 4133434 | 16951043 | 3653857 | 24435144 | 47934 | N/F | N/F | 412953 | N/F | 228 | 14686 | N/F | 44867 | N/F | 2174 | 34101 | N/F | 171980 | 21597 | 2029 | 44959 | 14592 | 718 | N/F | 3868 | 53026 | 485 | 33476 |  | Unknown
WO-23003-176-A-1 | N/F | 15472 | 49559 | 2410 | N/F | N/F | 5743114 | N/F | 77044 | 3672 | N/F | 9862 | N/F | 109010 | 56337 | 453338 | 912424 | 334777 | 1633243 | 6725419 | 2623095 | 329025 | 34980436 | 4075895 | 16601625 | 3523101 | 24917770 | 62970 | N/F | N/F | 475018 | N/F | 180 | 11734 | N/F | 33908 | N/F | 1737 | N/F | N/F | 169438 | 17931 | 2979 | 43738 | 15268 | 611 | N/F | 4074 | 50657 | 987 | 42125 |  | Unknown
WO-23003-175-A-1 | N/F | 18904 | 94111 | 5666 | N/F | 1315 | 5819420 | 43 | 90208 | 1055 | N/F | 11159 | 1581 | 111070 | 66209 | 517187 | 987646 | 369722 | 1999689 | 7824120 | 2959853 | 363998 | 35315191 | 4183782 | 19082738 | 3626223 | 25711792 | 74350 | N/F | N/F | 511880 | N/F | 373 | 14006 | N/F | 35492 | N/F | 2312 | 11451 | N/F | 154626 | 22955 | 2801 | 46917 | 12396 | 1045 | N/F | 5371 | 49377 | N/F | 33902 |  | Unknown
WO-23003-174-A-1 | N/F | 20565 | 181962 | 3794 | N/F | N/F | 5982507 | N/F | 72251 | 2472 | N/F | 10316 | 2055 | 103769 | 61088 | 493560 | 987552 | 351613 | 1943904 | 7512317 | 2663445 | 312279 | 35919221 | 4338605 | 18862711 | 3765508 | 25850470 | 59001 | N/F | N/F | 468510 | N/F | 462 | 10536 | N/F | 42895 | N/F | 2285 | 16682 | N/F | 163282 | 21217 | 2378 | 46328 | 13249 | 1406 | N/F | 5654 | 54091 | N/F | 33864 |  | Unknown
WO-23003-173-A-1 | N/F | 21769 | 483769 | 6252 | N/F | N/F | 5668810 | N/F | 19145 | 508 | N/F | 8296 | 1649 | 104701 | 33451 | 317791 | 883639 | 167549 | 1318094 | 4821696 | 1406071 | 99326 | 33014680 | 3907669 | 13489980 | 3306529 | 21850161 | 14126 | N/F | N/F | 258711 | N/F | 298 | 12253 | N/F | 45647 | 902 | 1670 | 63934 | N/F | 158053 | 21392 | 2801 | 34207 | 14080 | 660 | N/F | 3935 | 51009 | 833 | 37383 |  | Unknown
WO-23003-172-A-1 | N/F | 20481 | 132575 | 9242 | N/F | N/F | 6051473 | N/F | 74304 | 2307 | N/F | 15109 | 1152 | 439001 | 70487 | 615949 | 1187780 | 372477 | 2023323 | 7784524 | 2828001 | 324177 | 33821093 | 4304216 | 19258794 | 3705559 | 25363940 | 61777 | N/F | N/F | 493198 | N/F | 644 | 13679 | N/F | 49182 | 1415 | 2554 | 34009 | N/F | 165134 | 23035 | 2734 | 31121 | 13503 | 687 | N/F | 6400 | 44404 | 1471 | 31103 |  | Unknown
WO-23003-171-A-1 | N/F | 20613 | 82471 | 7442 | N/F | N/F | 5805586 | N/F | 111450 | N/F | N/F | 15558 | N/F | 436563 | 80798 | 704555 | 1313062 | 417810 | 2132079 | 8466266 | 3174860 | 440752 | 32821029 | 4227071 | 20015689 | 3679650 | 24823141 | 97618 | N/F | N/F | 558713 | N/F | 983 | 9877 | N/F | 46321 | N/F | 3687 | N/F | N/F | 145423 | 20636 | 1852 | 53941 | 13538 | 802 | N/F | 8338 | 44892 | 1256 | 29169 |  | Unknown
AB-230117-02_LRB_18_20230128065511 | N/F | 19947 | 32473 | 3654 | 6477 | N/F | 5589208 | N/F | 502462 | N/F | N/F | 1231 | 1412 | 2376 | 170856 | 620570 | 720985 | 1016943 | 3250951 | 16438593 | 7600592 | 1207701 | 38512113 | 4161213 | 27888958 | 3514841 | 30466100 | 480642 | 1323 | N/F | 1380572 | N/F | 33 | 10016 | N/F | 61101 | N/F | 137 | 31803 | 25851 | 128273 | 20447 | 1103 | 91758 | 10133 | 1213 | N/F | 9227 | 52721 | N/F | 37121 |  | Unknown
AB-230117-02_LRB_18 | N/F | 22623 | 29539 | 4078 | 5787 | N/F | 5854183 | N/F | 489437 | N/F | N/F | 1988 | 1200 | 8296 | 172865 | 602302 | 753951 | 999773 | 3231707 | 16830928 | 7666312 | 1214027 | 40339726 | 4371654 | 28983106 | 3593951 | 31732010 | 465299 | N/F | N/F | 1407079 | N/F | 1099 | 12160 | N/F | 59669 | 881 | 198 | N/F | 26257 | 156584 | 20160 | 1767 | 96280 | 12830 | 1129 | N/F | 9391 | 58781 | 888 | 32785 |  | Unknown
WO-23003-170-A-1 | N/F | 16960 | 105726 | 8968 | 8496 | N/F | 5917529 | N/F | 58044 | 2020 | N/F | 7900 | 117 | 79463 | 57206 | 512977 | 978720 | 321024 | 1918476 | 7331982 | 2560727 | 269131 | 35655967 | 4409533 | 18979309 | 3701497 | 25869636 | 48476 | 3108 | N/F | 441216 | N/F | 201 | 14523 | N/F | 41185 | N/F | 2113 | 30190 | N/F | 157365 | 18150 | 3464 | 44012 | 12588 | 1473 | N/F | 5799 | 48169 | 862 | 33062 |  | Unknown
WO-23003-169-A-1 | N/F | 19089 | 117183 | 3395 | 8208 | N/F | 5645860 | N/F | 16061 | 705 | N/F | 9547 | 2200 | 102070 | 29215 | 457559 | 1085844 | 151875 | 1289295 | 4800114 | 1260188 | 87573 | 31308682 | 4052739 | 13380648 | 3482955 | 22045140 | 12005 | N/F | N/F | 230313 | N/F | 415 | 10630 | N/F | 41679 | 1004 | 1413 | 12200 | N/F | 147720 | 18126 | 4124 | 34607 | 13523 | 1592 | N/F | 1864 | 50188 | 717 | 25370 |  | Unknown
WO-23003-168-A-1 | N/F | 52389 | 282364 | 8779 | 8653 | N/F | 3206745 | N/F | 124553 | 2839 | 17220004 | 4564 | N/F | 19896 | 107894 | 1204166 | 1483243 | 338226 | 2389505 | 8665500 | 3560271 | 477327 | 30011092 | 4283816 | 19229820 | 3746978 | 22941001 | 107661 | N/F | N/F | 612076 | N/F | 433 | 9604 | N/F | 96993 | N/F | 3078 | 402077 | 15243 | 145602 | 46946 | 1276 | 62983 | 8701 | 1348 | N/F | 6217 | 42678 | 1122 | 16648 |  | Unknown
WO-23003-167-A-1 | N/F | 56219 | 969237 | 8778 | 12161 | N/F | 3076346 | N/F | 80943 | N/F | 15089772 | 6156 | 2911 | 20072 | 93333 | 916137 | 1450492 | 283386 | 2079503 | 7744759 | 2920180 | 336671 | 29415023 | 4147141 | 17747007 | 3630087 | 22371425 | 66326 | N/F | N/F | 515902 | N/F | 302 | 7574 | N/F | 104484 | N/F | 3714 | 371567 | 23801 | 146247 | 48884 | 2077 | 56837 | 14462 | 2577 | N/F | 7677 | 40186 | 936 | 15278 |  | Unknown
WO-23003-166-A-1 | N/F | 41102 | 559975 | 7125 | 14664 | 935 | 4871720 | N/F | 152127 | N/F | 18877176 | 8835 | 4833 | 20268 | 106965 | 996926 | 1396697 | 362297 | 2539693 | 9476008 | 3810986 | 563524 | 31878795 | 4264002 | 21890359 | 3901813 | 25407205 | 136720 | N/F | N/F | 681771 | N/F | 189 | 10612 | N/F | 100409 | N/F | 2168 | 453809 | 22010 | 158691 | 48274 | 3629 | 59592 | 14832 | 2415 | N/F | 6983 | 49384 | N/F | 24545 |  | Unknown
WO-23003-165-A-1 | N/F | 33724 | 82375 | 8942 | 17715 | 1363 | 5041041 | N/F | 85476 | 3530 | 14718012 | 9303 | 2541 | 20635 | 74201 | 890067 | 1408951 | 284546 | 2077889 | 7691472 | 2977046 | 357214 | 31033433 | 4129061 | 18802474 | 3702482 | 23474981 | 69809 | N/F | N/F | 520027 | N/F | 233 | 12497 | N/F | 77862 | N/F | 1829 | 371353 | 11265 | 165660 | 42184 | 4983 | 82315 | 15342 | 2908 | N/F | 6463 | 50118 | N/F | 32573 |  | Unknown
WO-23003-164-A-1 | N/F | 24282 | 77223 | 6590 | 24362 | N/F | 4415184 | N/F | 38712 | 2239 | 3876149 | 4691 | 1819 | 7928 | 52071 | 835361 | 1384845 | 222523 | 1648049 | 6145104 | 2117720 | 188997 | 28647728 | 3911273 | 15546125 | 3467978 | 20805999 | 33077 | N/F | N/F | 361499 | N/F | 651 | 11223 | N/F | 44639 | 843 | 2610 | 57158 | N/F | 161888 | 23630 | 2578 | 49552 | 10304 | 2873 | N/F | 5238 | 53165 | N/F | 42032 |  | Unknown
WO-23003-163-A-1 | N/F | 17339 | 65770 | 5194 | 39764 | N/F | 5348519 | 298 | 18531 | 2813 | 2732124 | 7087 | 1326 | 7003 | 32663 | 436918 | 1218578 | 168104 | 1319432 | 4843303 | 1435424 | 105682 | 30799128 | 3829879 | 13691119 | 3292263 | 21426804 | 16334 | N/F | N/F | 245372 | N/F | 355 | 15498 | N/F | 47181 | 874 | 1218 | N/F | N/F | 168267 | 24993 | 6727 | 41189 | 16016 | 2765 | N/F | 2306 | 70557 | N/F | 43443 |  | Unknown
WO-23003-162-A-1 | N/F | 21825 | 66990 | 7666 | 62632 | N/F | 5232667 | N/F | 61833 | 2326 | 5285361 | 9612 | 1860 | 10652 | 62408 | 753630 | 1325620 | 295461 | 1993288 | 7401610 | 2617183 | 290613 | 32557192 | 4343109 | 18864513 | 3892407 | 24824767 | 53537 | N/F | N/F | 465424 | N/F | 418 | 12234 | N/F | 103152 | 1237 | 3603 | 100563 | N/F | 167683 | 42326 | 4663 | 57217 | 15538 | 6911 | N/F | 5193 | 76178 | N/F | 55215 |  | Unknown
WO-23003-161-A-1 | N/F | 22787 | 76440 | 3586 | 152694 | N/F | 5088481 | 81 | 85480 | 2682 | 9288137 | 13156 | 3464 | 12946 | 80928 | 965444 | 1378803 | 320634 | 2091007 | 7897875 | 3033931 | 353953 | 30116754 | 4135400 | 18538399 | 3656138 | 23473716 | 69693 | N/F | N/F | 541326 | N/F | 118 | 22271 | N/F | 101707 | N/F | 2191 | 110312 | N/F | 149565 | 50707 | 9207 | 65287 | 18017 | 10847 | N/F | 4824 | 137458 | N/F | 83798 |  | Unknown
AB-230117-02_LFSM-H_8 | N/F | 74011 | 639663 | 840003 | 63919630 | N/F | 3612855 | N/F | 80115 | 582113 | 9507546 | 2988031 | 1650643 | 2593001 | 70695 | 614889 | 842819 | 314744 | 1896632 | 6956418 | 2890177 | 338215 | 21881749 | 3362595 | 15740897 | 2939093 | 18646299 | 67281 | N/F | N/F | 497745 | N/F | 630 | 2793355 | 65112 | 20508600 | N/F | 2162 | 6422955 | 2449174 | 221852 | 16017906 | 2846015 | 14060570 | 60033 | 3259519 | 296262 | 1153197 | 30150560 | N/F | 20674873 |  | Unknown
AB-230117-02_CCC-H_9 | N/F | 306484 | 565672 | 574419 | 114967475 | N/F | 3303441 | N/F | 530533 | 3108414 | 1879967 | 4782937 | 3958558 | 4171529 | 160801 | 299670 | 311580 | 997179 | 2169590 | 11147682 | 7007399 | 1220309 | 24904792 | 3456287 | 19631655 | 2417403 | 21703755 | 502223 | 2366 | N/F | 1384656 | N/F | 499 | 4606806 | 837726 | 42220888 | 1012 | 26 | 19688885 | 12187522 | 253031 | 34274579 | 4348208 | 33047919 | 79735 | 5797848 | 2165743 | 4761844 | 53697048 | N/F | 39196538 |  | Unknown
AB-230117-02_LFB-H_9 | N/F | 225283 | 415022 | 539995 | 71579149 | N/F | 4175842 | N/F | 575944 | 1944739 | 1170389 | 3159892 | 2554066 | 2666954 | 193231 | 422993 | 531671 | 1168510 | 2731349 | 14133441 | 8495625 | 1408230 | 32099934 | 3846696 | 24747417 | 3098356 | 26967994 | 546872 | N/F | N/F | 1664664 | N/F | 712 | 2957691 | 479426 | 27930035 | N/F | 177 | 12982365 | 7655089 | 206678 | 22393229 | 2947254 | 22038961 | 46676 | 3429393 | 1295998 | 2963244 | 31311722 | 886 | 26967713 |  | Unknown
AB-230117-02_LD_8 | 1611 | 25954 | 70832 | 8331 | 10559 | 912 | 4748298 | N/F | 46632 | N/F | 4681698 | 3760 | N/F | 5032 | 64320 | 701205 | 1287837 | 244540 | 1713812 | 6467909 | 2240979 | 220661 | 31064596 | 4233622 | 16793031 | 3731604 | 23159551 | 41236 | N/F | N/F | 387823 | N/F | 364 | 11429 | N/F | 76846 | 740 | 1760 | 55817 | N/F | 144366 | 30724 | 1657 | 47573 | 13061 | 1612 | N/F | 3577 | 45043 | N/F | 26800 |  | Unknown
AB-230117-02_LRB_17_20230127231451 | N/F | 24826 | 41324 | 3212 | N/F | N/F | 5462917 | N/F | 490612 | N/F | N/F | 2966 | 2972 | 5781 | 206021 | 788146 | 894505 | 1036892 | 3197912 | 16600735 | 8001582 | 1224859 | 38781650 | 4352404 | 28252225 | 3737957 | 32070431 | 472207 | N/F | N/F | 1408963 | N/F | 1094 | 15323 | N/F | 65644 | N/F | 135 | 60173 | 34258 | 132417 | 20214 | 1007 | 109993 | 12076 | 781 | N/F | 10364 | 49486 | 879 | 26536 |  | Unknown
AB-230117-02_LRB_17 | N/F | 27496 | 48338 | 5217 | N/F | N/F | 5912452 | N/F | 509488 | N/F | N/F | 4445 | 1791 | 6345 | 207061 | 872961 | 1025191 | 1050190 | 3379366 | 16961166 | 7902759 | 1252637 | 40057113 | 4798619 | 29850546 | 4215275 | 32933569 | 484708 | N/F | N/F | 1439985 | N/F | 547 | 11256 | N/F | 66179 | N/F | 318 | 36308 | 29797 | 177830 | 24722 | 2891 | 110538 | 13454 | 1097 | N/F | 7832 | 52510 | 880 | 32420 |  | Unknown
WO-23003-160-A-1 | N/F | 24192 | 97428 | 6833 | N/F | N/F | 4869433 | N/F | 69393 | N/F | 8007913 | 6435 | 1656 | 6165 | 75084 | 1114633 | 1656464 | 282179 | 1772186 | 6610913 | 2716039 | 292414 | 30127225 | 4210310 | 15260217 | 3469802 | 20522506 | 54321 | N/F | N/F | 451646 | N/F | 353 | 8607 | N/F | 74133 | 1499 | 2498 | 101063 | 23874 | 179169 | 20829 | 1123 | 39208 | 13684 | 1718 | N/F | 4194 | 45985 | 510 | 24830 |  | Unknown
WO-23003-159-A-1 | N/F | 18199 | 54811 | 8685 | N/F | N/F | 5813320 | N/F | 47925 | N/F | 7473240 | 5898 | N/F | 6418 | 58561 | 510795 | 1121185 | 260875 | 1670059 | 6774462 | 2368282 | 236671 | 35695339 | 4425844 | 17583008 | 3888346 | 25763344 | 38069 | N/F | N/F | 410528 | N/F | 661 | 14204 | N/F | 62563 | 891 | 1700 | 118306 | N/F | 178283 | 27650 | 3048 | 39182 | 13153 | 1343 | N/F | 4567 | 50806 | N/F | 29580 |  | Unknown
WO-23003-158-A-1 | N/F | 18669 | 54091 | 4822 | N/F | N/F | 5839522 | N/F | 108189 | 2999 | 3044505 | 6001 | 821 | 4161 | 75476 | 624740 | 1187744 | 386892 | 1987225 | 8038532 | 3121432 | 425090 | 36461695 | 4599970 | 19224864 | 3959551 | 26801827 | 92849 | N/F | N/F | 556363 | 1826 | 467 | 13345 | N/F | 318373 | N/F | 2067 | 85261 | N/F | 189222 | 21466 | 1523 | 44288 | 14856 | 946 | N/F | 5690 | 49863 | 1703 | 29593 |  | Unknown
WO-23003-157-A-1 | N/F | 16778 | 70621 | 3441 | N/F | N/F | 4730005 | N/F | 23847 | 638 | 1244427 | 4039 | N/F | 8379 | 36092 | 598311 | 1308848 | 165777 | 1257699 | 4863416 | 1513836 | 113209 | 27568727 | 3801571 | 12335038 | 3244749 | 19264744 | 15191 | N/F | N/F | 255885 | N/F | 339 | 9467 | N/F | 35938 | 1593 | 1734 | 10591 | N/F | 172878 | 17168 | 1255 | 32881 | 16553 | 1275 | N/F | 2634 | 46457 | 924 | 23984 |  | Unknown
WO-23003-156-A-1 | N/F | 8465 | 73909 | 5430 | N/F | N/F | 4607077 | N/F | 96035 | N/F | 2858671 | 8114 | N/F | 7770 | 84509 | 933303 | 1433286 | 377701 | 2226147 | 8307874 | 3178765 | 404392 | 31947500 | 4308448 | 19713169 | 3762476 | 24511254 | 82072 | N/F | N/F | 554647 | N/F | 487 | 10384 | N/F | 41384 | N/F | 2522 | 74313 | N/F | 162504 | 18913 | 1191 | 50223 | 12842 | 927 | N/F | 6944 | 42391 | 1730 | 25579 |  | Unknown
WO-23003-155-A-1 | N/F | 18984 | 56806 | 5907 | N/F | 1434 | 5745681 | N/F | 80135 | 3658 | 4609996 | 5391 | 1190 | N/F | 65083 | 540690 | 1078858 | 311653 | 1783363 | 7339596 | 2785854 | 334619 | 36089654 | 4463528 | 17934940 | 3803417 | 26608371 | 64013 | N/F | N/F | 488040 | 1636 | 456 | 12988 | N/F | 49169 | N/F | 2308 | 34952 | N/F | 177305 | 34513 | 1786 | 42917 | 17140 | 1505 | N/F | 6405 | 48169 | 1180 | 35323 |  | Unknown
WO-23003-154-A-1 | N/F | 19292 | 69099 | 5652 | N/F | N/F | 5591321 | N/F | 77779 | 3250 | 3987003 | 4813 | 1992 | 5138 | 66763 | 587517 | 1297606 | 310399 | 1614797 | 6881998 | 2742671 | 324559 | 33684806 | 4330562 | 16827970 | 3772794 | 24393253 | 63488 | N/F | N/F | 471730 | N/F | 219 | 11812 | N/F | 50777 | N/F | 2314 | N/F | N/F | 174011 | 173823 | 1430 | 46819 | 16439 | 1069 | N/F | 2600 | 47230 | 1098 | 26725 |  | Unknown
WO-23003-153-A-1 | N/F | 15489 | 48111 | 3590 | 4898 | N/F | 4749892 | N/F | 55849 | N/F | 2042070 | 5116 | N/F | 4158 | 50120 | 274975 | 837140 | 256724 | 969843 | 4818959 | 2031301 | 230281 | 25804341 | 3521984 | 9743480 | 2604457 | 15843986 | 45525 | N/F | N/F | 358758 | N/F | 452 | 11485 | N/F | 52909 | 628 | 1828 | N/F | N/F | 183103 | 18557 | 1791 | 24967 | 18237 | 1113 | N/F | 1894 | 55324 | 1332 | 29217 |  | Unknown
WO-23003-152-A-1 | N/F | 15366 | 46886 | 5006 | N/F | 1069 | 5371058 | N/F | 41256 | 2505 | 720672 | 4646 | N/F | 3685 | 46744 | 354280 | 912526 | 235400 | 1233668 | 5391073 | 1952820 | 189287 | 33483385 | 3767259 | 13012060 | 3345514 | 22355921 | 35316 | N/F | N/F | 337422 | N/F | 15 | 11140 | N/F | 46667 | N/F | 2135 | N/F | N/F | 180310 | 17818 | 4413 | 35985 | 16314 | 1304 | N/F | 3974 | 55874 | N/F | 28553 |  | Unknown
WO-23003-151-A-1 | N/F | 19257 | 47544 | 6173 | 24246 | N/F | 5270755 | N/F | 79245 | N/F | 921625 | 2622 | N/F | 9689 | 61429 | 391388 | 871021 | 328737 | 1307811 | 6057087 | 2494156 | 331070 | 34276127 | 3530487 | 13904521 | 3304266 | 23442267 | 64985 | N/F | N/F | 439616 | N/F | 247 | 14239 | N/F | 176764 | N/F | 1533 | N/F | N/F | 181496 | 24927 | 1650 | 50373 | 17180 | 1808 | N/F | 2479 | 66138 | N/F | 38360 |  | Unknown
AB-230117-02_LRB_16_20230127174957 | N/F | 25617 | 25258 | 3181 | 6149 | N/F | 4772346 | N/F | 329369 | N/F | N/F | 1611 | 1748 | 7824 | 152551 | 505547 | 581921 | 823464 | 2838192 | 15611997 | 6549995 | 848175 | 38439401 | 3360589 | 26693963 | 3350159 | 31526459 | 319065 | N/F | N/F | 1158028 | N/F | 787 | 12880 | N/F | 84008 | N/F | 784 | 25142 | 23480 | 132719 | 25155 | 3709 | 102717 | 9155 | 1412 | N/F | 7397 | 57002 | 739 | 38475 |  | Unknown
AB-230117-02_LRB_16 | N/F | 34633 | 26579 | 3279 | 5744 | N/F | 5131704 | N/F | 371747 | 5594 | N/F | 2333 | 589 | 7384 | 154118 | 579847 | 729806 | 785930 | 3153054 | 16541082 | 6366458 | 807509 | 43038656 | 3538721 | 28695723 | 3752352 | 34424377 | 356839 | N/F | N/F | 1105469 | N/F | 370 | 13662 | N/F | 82227 | N/F | 175 | 33639 | 22574 | 169831 | 28595 | 1780 | 95879 | 16268 | 1527 | N/F | 10960 | 61780 | 1178 | 47942 |  | Unknown
WO-23003-150-A-1 | N/F | 25999 | 48550 | 4610 | 6739 | N/F | 5402224 | N/F | 41987 | N/F | 772102 | 3593 | 2229 | 12904 | 54961 | 476900 | 1045410 | 226839 | 1624270 | 6436548 | 2039599 | 175331 | 37504622 | 3583959 | 17346607 | 3851752 | 27896754 | 31830 | N/F | N/F | 344916 | N/F | 205 | 17239 | N/F | 67162 | N/F | 2774 | N/F | N/F | 168347 | 29764 | 2597 | 50465 | 18144 | 961 | N/F | 3584 | 55651 | N/F | 31969 |  | Unknown
WO-23003-149-A-1 | N/F | 31756 | 74306 | 7934 | N/F | N/F | 4995140 | N/F | 42190 | 2371 | 714003 | 7170 | 455 | N/F | 53871 | 729774 | 1462005 | 233992 | 1703351 | 6632741 | 2007245 | 173026 | 33375145 | 3332020 | 17551789 | 3793251 | 26469640 | 33126 | N/F | N/F | 340903 | N/F | 451 | 12412 | N/F | 62182 | N/F | 2708 | 22728 | N/F | 165988 | 30888 | 4574 | 34272 | 15205 | 1579 | N/F | 5554 | 44561 | N/F | 33342 |  | Unknown
WO-23003-148-A-1 | N/F | 33242 | 71676 | 4193 | 7807 | N/F | 4922542 | N/F | 52706 | 3259 | 822590 | 6784 | N/F | 7467 | 60120 | 741259 | 1422064 | 242650 | 1794748 | 6992145 | 2224399 | 215368 | 33859088 | 3427961 | 18475522 | 3756816 | 26373254 | 42946 | N/F | N/F | 376745 | 1045 | 176 | 11789 | N/F | 62022 | 1517 | 549 | 26655 | N/F | 165293 | 32288 | 1776 | 49955 | 15657 | 1175 | N/F | 5834 | 50330 | 988 | 26221 |  | Unknown
WO-23003-147-A-1 | N/F | 30924 | 71512 | 4472 | 10219 | N/F | 5142605 | N/F | 64852 | N/F | 921139 | 8907 | N/F | 8118 | 65484 | 824335 | 1416512 | 273863 | 1906985 | 7518003 | 2453173 | 242382 | 35745303 | 3448670 | 19525177 | 3832871 | 27348352 | 49981 | N/F | N/F | 429523 | N/F | 174 | 13148 | N/F | 66734 | N/F | 3297 | N/F | N/F | 170796 | 29791 | 2720 | 60509 | 16224 | 1115 | N/F | 4259 | 55962 | 2079 | 26723 |  | Unknown
WO-23003-146-A-1 | N/F | 41385 | 82883 | 9759 | 13771 | N/F | 4132950 | N/F | 73875 | N/F | 1275663 | 7409 | 1396 | 9969 | 68778 | 1053999 | 1545661 | 295815 | 2056685 | 7941116 | 2652987 | 273558 | 32156990 | 3404050 | 20046846 | 3906282 | 25700237 | 61502 | N/F | N/F | 443958 | N/F | 778 | 10733 | N/F | 66596 | N/F | 2056 | 86259 | N/F | 160014 | 27436 | 3431 | 68942 | 14760 | 1095 | N/F | 4417 | 46385 | N/F | 36219 |  | Unknown
WO-23003-145-A-1 | N/F | 33084 | 75418 | 18477 | 18471 | 977 | 4784068 | N/F | 85705 | N/F | 1189624 | 8559 | 2345 | 9466 | 72646 | 815428 | 1493470 | 317013 | 2031906 | 8174231 | 2684127 | 285286 | 34635365 | 3507343 | 20234548 | 3890902 | 27672780 | 72268 | N/F | N/F | 458870 | N/F | 39 | 13539 | N/F | 73664 | N/F | 1682 | N/F | N/F | 168834 | 30509 | 2001 | 56555 | 15471 | 2486 | N/F | 6089 | 63483 | N/F | 33928 |  | Unknown
WO-23003-144-A-1 | N/F | 32643 | 50119 | 9115 | 26776 | N/F | 5449959 | 14 | 71134 | N/F | 220632 | 6891 | N/F | 10689 | 64226 | 512297 | 997226 | 297135 | 1800955 | 7520719 | 2550021 | 259757 | 40148666 | 3600614 | 19453429 | 4008874 | 30179345 | 56921 | N/F | 1296 | 430289 | N/F | 989 | 16926 | N/F | 94491 | N/F | 1947 | 33772 | N/F | 177620 | 34045 | 4223 | 57860 | 18222 | 3894 | N/F | 7134 | 70804 | N/F | 48991 |  | Unknown
WO-23003-143-A-1 | N/F | 30072 | 49826 | 3394 | 40609 | N/F | 5507041 | N/F | 67125 | 3961 | 175736 | 9048 | 1431 | 11692 | 63220 | 535788 | 1024663 | 295991 | 1912046 | 7728237 | 2412260 | 258218 | 40160262 | 3553445 | 19532034 | 3855879 | 29438075 | 54078 | N/F | N/F | 411477 | N/F | 835 | 13077 | N/F | 135017 | N/F | 2482 | 9602 | N/F | 184242 | 36691 | 4147 | 56924 | 17961 | 2552 | N/F | 5635 | 78639 | N/F | 51888 |  | Unknown
WO-23003-142-A-1 | N/F | 40119 | 105636 | 5778 | 69571 | N/F | 4303242 | N/F | 53273 | N/F | 238490 | 9350 | N/F | 9341 | 68546 | 1213708 | 1857096 | 269680 | 1970255 | 7378311 | 2269744 | 226409 | 33259350 | 3742151 | 18317393 | 3757708 | 23460827 | 43086 | 1484 | N/F | 393507 | N/F | 724 | 15336 | N/F | 152427 | N/F | 2769 | 41147 | 16232 | 173756 | 42580 | 9146 | 55478 | 16163 | 4753 | N/F | 4451 | 87557 | N/F | 51764 |  | Unknown
WO-23003-141-A-1 | N/F | 33862 | 73455 | 7331 | 202024 | N/F | 4012351 | N/F | 15985 | N/F | 98896 | 13572 | N/F | 13210 | 36234 | 609994 | 1365222 | 149370 | 1267875 | 4997696 | 1333669 | 75888 | 29934648 | 2967222 | 13563890 | 3333604 | 22187134 | 12317 | N/F | N/F | 225755 | N/F | 376 | 15725 | N/F | 137878 | N/F | 3093 | 60250 | N/F | 159210 | 64726 | 9620 | 66169 | 17581 | 8538 | N/F | 2978 | 121292 | N/F | 68336 |  | Unknown
AB-230117-02_LFSM-H_7 | N/F | 77033 | 574253 | 845422 | 64946442 | N/F | 3055770 | N/F | 35356 | 364126 | 356101 | 2931477 | 1413249 | 2739312 | 46555 | 564917 | 817266 | 219237 | 1618655 | 6120918 | 1886454 | 151982 | 24540149 | 2659010 | 15199777 | 3058586 | 20777930 | 29375 | N/F | N/F | 326061 | 1576 | 600 | 2370424 | 30329 | 22725019 | N/F | 1579 | 5600255 | 1626107 | 207792 | 18101405 | 3019053 | 13862555 | 63331 | 2684124 | N/F | 828444 | 25706787 | N/F | 18496056 |  | Unknown
AB-230117-02_CCC-H_8 | N/F | 283634 | 540834 | 571652 | 112230191 | N/F | 3094138 | N/F | 397317 | 2619326 | 1402414 | 4929560 | 4048473 | 4314371 | 142821 | 294559 | 299509 | 793697 | 2157964 | 12328264 | 6025403 | 809826 | 28531073 | 2863416 | 21121392 | 2383178 | 24578455 | 380122 | N/F | 841 | 1095340 | 1343 | 531 | 4369011 | 641387 | 48461870 | N/F | 158 | 22191139 | 10407726 | 256513 | 40138544 | 4324048 | 36438473 | 90069 | 5011331 | N/F | 4411997 | 55940087 | N/F | 44911171 |  | Unknown
AB-230117-02_LFB-H_8 | N/F | 188692 | 370761 | 477303 | 65860546 | N/F | 4175234 | N/F | 393779 | 1257647 | 702162 | 2974567 | 2374829 | 2524445 | 160540 | 401129 | 522896 | 766022 | 2647643 | 14439857 | 6265642 | 844149 | 35223208 | 3148685 | 24434056 | 3035404 | 29190429 | 373163 | N/F | N/F | 1103741 | N/F | 352 | 2898271 | 304173 | 29820113 | 1037 | 508 | 12869591 | 5444796 | 220698 | 23541569 | 2822259 | 21261067 | 48319 | 2692329 | N/F | 2391425 | 30845357 | N/F | 26687083 |  | Unknown
AB-230117-02_LD_7 | N/F | 37694 | 99384 | 4779 | 7757 | N/F | 4265080 | N/F | 25222 | N/F | 151098 | 6204 | N/F | N/F | 44592 | 919188 | 1734024 | 171470 | 1522140 | 5446739 | 1583020 | 103016 | 30179486 | 3487801 | 14327453 | 3377280 | 20794686 | 17781 | N/F | N/F | 258314 | N/F | 434 | 10583 | N/F | 179349 | 1198 | 629 | 20106 | 13626 | 150196 | 23302 | 2912 | 40808 | 13323 | 1198 | N/F | 2389 | 47071 | N/F | 34956 |  | Unknown
AB-230117-02_LRB_15_20230127100939 | N/F | 36561 | 29706 | 3350 | N/F | 900 | 4501277 | N/F | 311627 | 4308 | N/F | 2372 | N/F | 7481 | 156236 | 493514 | 607056 | 725349 | 2627401 | 14095219 | 5947417 | 712231 | 35095576 | 2922875 | 23444026 | 3163740 | 29548719 | 302709 | N/F | N/F | 1054039 | N/F | 270 | 13232 | N/F | 223796 | N/F | 319 | N/F | 25794 | 119012 | 28448 | 2146 | 79462 | 13375 | 1174 | N/F | 6263 | 48352 | 496 | 37811 |  | Unknown
AB-230117-02_LRB_15 | N/F | 35677 | 24942 | 2977 | 8069 | N/F | 5368014 | N/F | 369245 | 2553 | N/F | 677 | N/F | 7071 | 169071 | 603724 | 806236 | 776399 | 3163272 | 16402500 | 6363611 | 833703 | 42072343 | 3527219 | 28246415 | 3750497 | 33904567 | 353297 | N/F | N/F | 1122844 | 968 | 664 | 14291 | N/F | 295261 | N/F | 580 | 28963 | 27899 | 165345 | 33895 | 2643 | 106634 | 17126 | 1247 | N/F | 8153 | 59646 | N/F | 41605 |  | Unknown
WO-23003-140-A-1 | N/F | 26785 | 58847 | 3020 | N/F | 919 | 5629488 | N/F | 60590 | 2213 | 239678 | 7245 | N/F | 7782 | 58816 | 612534 | 1176296 | 269934 | 1776710 | 7167007 | 2241844 | 228036 | 36342071 | 3473055 | 18299776 | 3732851 | 27809792 | 50264 | N/F | 2062 | 386869 | N/F | 111 | 12385 | N/F | 360866 | N/F | 2617 | N/F | N/F | 164935 | 31433 | 2528 | 46615 | 13623 | 1394 | N/F | 4089 | 55989 | 955 | 33577 |  | Unknown
WO-23003-139-A-1B | N/F | 29012 | 51109 | 3933 | 13611 | N/F | 5394826 | N/F | 56577 | 6464 | 211700 | 7344 | 1874 | 25846 | 55362 | 526887 | 918806 | 247705 | 1842386 | 7243909 | 2194892 | 203238 | 35046837 | 3279201 | 18277425 | 3553271 | 27124738 | 47777 | N/F | N/F | 393619 | N/F | 654 | 12674 | N/F | 1076143 | N/F | 2592 | N/F | N/F | 126606 | 31596 | 6511 | 48481 | 14010 | 1439 | N/F | 5891 | 55714 | N/F | 51714 |  | Unknown
WO-23003-139-A-1 | N/F | 31677 | 78099 | 2887 | 56842 | N/F | 5594901 | N/F | 88805 | 9909 | 165071 | 4129 | N/F | 23235 | 55825 | 434235 | 818506 | 200900 | 1893044 | 7669561 | 2565629 | 223068 | 31886772 | 3793322 | 18606048 | 3544326 | 26342001 | 75063 | N/F | N/F | 501825 | N/F | 296 | 17674 | N/F | 20941273 | 1125 | 665 | 33655 | N/F | 133668 | 58013 | 7003 | 78779 | 7423 | 5750 | N/F | 11430 | 73011 | 745 | 64482 |  | Unknown
AB-230117-02_LRB_12_20230123100212 | N/F | 69108 | 38289 | 1611 | 44362 | 526 | 403553 | N/F | 14699 | 4833 | N/F | 3768 | 3206 | 8803 | 3058 | 24669 | 54053 | 6697 | 124329 | 448804 | 100284 | 8211 | 2431795 | 293562 | 1287286 | 265578 | 2143667 | 12916 | 1743 | N/F | 18818 | N/F | 112 | 4649 | N/F | 46587122 | N/F | N/F | 38972 | N/F | 29102 | 59619 | 4323 | 72330 | 8497 | 6410 | N/F | 12964 | 37454 | 1177 | 48108 | C | Unknown
WO-23003-203-A-1B | 1217 | 47274 | 26042 | 10615 | 21734 | N/F | 5035644 | N/F | 520821 | 1827 | 54088 | 4246293 | 312484 | 17162755 | 189706 | 375335 | 719115 | 965643 | 3062842 | 16186626 | 7261125 | 1091433 | 38456781 | 1705287 | 25456947 | 1662963 | 17509722 | 500203 | N/F | N/F | 1342376 | N/F | 46 | 11755 | N/F | 73952 | N/F | N/F | 141859 | 30330 | 152455 | 18826 | 9494 | 589947 | 15896 | 20671 | N/F | 12169 | 62247 | 796 | 51875 |  | Unknown
WO-23003-208-A-1B | N/F | 50776 | 30813 | 5820 | 6824 | 1010 | 5449984 | N/F | 468018 | 5653 | 53138 | 4509247 | 323889 | 17762006 | 173990 | 381752 | 744333 | 936372 | 3246724 | 16632516 | 7138306 | 1048049 | 39806308 | 1792928 | 26469318 | 1709606 | 18251157 | 449514 | N/F | N/F | 1264936 | N/F | 78 | 12573 | N/F | 64591 | N/F | 512 | 137623 | 25757 | 145826 | 16946 | 9986 | 532428 | 15794 | 18041 | N/F | 8482 | 53437 | 1376 | 41182 |  | Unknown
WO-23003-207-A-1B | 1199 | 63726 | 43149 | 11019 | 12344 | N/F | 5229263 | N/F | 488640 | N/F | 57334 | 4736971 | 339119 | 18623889 | 169531 | 369012 | 714456 | 955947 | 3170272 | 16481551 | 7203777 | 1064852 | 39965297 | 1736488 | 26394783 | 1638880 | 17579970 | 465880 | N/F | 1658 | 1261198 | N/F | 125 | 12554 | N/F | 69410 | N/F | 8 | 136627 | 28185 | 170601 | 22137 | 10276 | 544293 | 15528 | 20264 | N/F | 7930 | 60092 | 1346 | 35236 |  | Unknown
WO-23003-206-A-1B | N/F | 46003 | 28949 | 4921 | 9961 | N/F | 5277615 | N/F | 517899 | 2866 | 61629 | 6158966 | 391191 | 22867936 | 180495 | 328000 | 692831 | 982724 | 3234755 | 17195748 | 7396823 | 1065695 | 38969513 | 1474619 | 25605917 | 1336272 | 14106716 | 495795 | 1481 | N/F | 1330125 | N/F | 352 | 13542 | N/F | 69182 | 1332 | 210 | 146637 | 28185 | 170843 | 17544 | 14533 | 644270 | 17091 | 16978 | N/F | 5924 | 59071 | 917 | 50468 |  | Unknown
WO-23003-205-A-1B | N/F | 50932 | 131176 | 14087 | 13320 | N/F | 5034118 | N/F | 315303 | 1835 | 163867 | 11344895 | 1223008 | 29055816 | 124021 | 231073 | 844695 | 646357 | 2759228 | 13341459 | 4933200 | 775866 | 35260483 | 1129694 | 18925067 | 1005913 | 9179316 | 295207 | N/F | N/F | 865937 | N/F | 173 | 11906 | N/F | 61895 | 999 | 882 | 183530 | 33918 | 149655 | N/F | 22286 | 741288 | 14662 | 17489 | N/F | 12026 | 48120 | 481 | 35573 |  | Unknown
WO-23003-204-A-1B | N/F | 37706 | 43674 | 10397 | 16622 | 1286 | 5304768 | N/F | 330524 | 1732 | 171614 | 11563289 | 1224814 | 30353372 | 133075 | 223653 | 887806 | 668503 | 2797550 | 13671055 | 5203349 | 807996 | 36292475 | 1261782 | 19003081 | 973200 | 9018587 | 317187 | N/F | N/F | 918572 | N/F | 127 | 14053 | N/F | 58739 | 833 | 544 | 251246 | 44746 | 168461 | N/F | 17724 | 899930 | 14707 | 15121 | N/F | 14517 | 46377 | 1058 | 36795 |  | Unknown
AB-230117-02_LRB_22B_20230130113632 | N/F | 38167 | 47778 | 8019 | 28191 | 534 | 4117270 | N/F | 586108 | 1678 | N/F | 6515 | 1812 | 54856 | 185799 | 487393 | 547399 | 1013485 | 2515105 | 13811734 | 7210561 | 1106935 | 30900382 | 1775117 | 23303576 | 2794056 | 25322946 | 572685 | N/F | N/F | 1356743 | N/F | 249 | 12724 | N/F | 63395 | N/F | 126 | 47965 | 16745 | 108165 | 40019 | 1732 | 278282 | 13739 | 22361 | N/F | 11880 | 49133 | 807 | 59138 |  | Unknown
AB-230117-02_LRB_22B | N/F | 46681 | 48010 | 10821 | 46088 | N/F | 4861937 | N/F | 602141 | N/F | N/F | 7036 | 2998 | 65300 | 193159 | 517580 | 582015 | 1038861 | 2908422 | 15865842 | 7867451 | 1143671 | 35010831 | 2026559 | 26885937 | 3077562 | 28422575 | 587046 | N/F | N/F | 1501253 | N/F | 191 | 14361 | N/F | 74148 | N/F | 231 | 55189 | 15968 | 126285 | 45443 | 2004 | 333638 | 16239 | 23096 | N/F | 8553 | 72411 | 611 | 64068 |  | Unknown
217288-d4-reprep | N/F | 25116 | 90144 | 8774 | 5972 | 924 | 5841501 | N/F | 173166 | 2389 | 10483422 | 14464 | 2776 | 106021 | 84346 | 776789 | 1276571 | 393326 | 2579520 | 10548765 | 3761331 | 546421 | 33642603 | 2810267 | 24106931 | 3658609 | 27874542 | 157464 | N/F | N/F | 650768 | N/F | 270 | 8955 | N/F | 97569 | N/F | 1022 | 225873 | 13843 | 149657 | 40414 | 1117 | 107384 | 16273 | 11139 | N/F | 8231 | 46063 | 1332 | 27299 |  | Unknown
217289-d4-reprep | 984 | 25491 | 102093 | 6685 | 6408 | N/F | 6090198 | N/F | 183091 | N/F | 11259640 | 12745 | 2776 | 63669 | 93373 | 829328 | 1237734 | 421121 | 2727214 | 11279977 | 3904277 | 606199 | 37427258 | 3215678 | 25204572 | 3967468 | 28154739 | 168253 | N/F | N/F | 695262 | N/F | 95 | 13646 | N/F | 70233 | N/F | 1819 | 225909 | 11936 | 165897 | 32560 | 3128 | 101759 | 18200 | 10049 | N/F | 7395 | 51852 | 927 | 39689 |  | Unknown
217295-d1000-LD-reprep | 696 | 20883 | 37335 | 4391 | N/F | N/F | 6272970 | N/F | 665964 | 1827 | N/F | 593721 | 19486 | 13785228 | 219005 | 605430 | 812553 | 1318561 | 3518786 | 18814017 | 9513170 | 1431177 | 42625232 | 3040714 | 30393908 | 1986992 | 21408323 | 641965 | N/F | N/F | 1798305 | N/F | 685 | 12102 | N/F | 230246 | N/F | 371 | 41469 | 13264 | 171099 | 18479 | 41942 | 68628 | 16383 | 1046 | N/F | 9763 | 52855 | N/F | 44600 |  | Unknown
217296-d1000-LFSM-200-reprep | 1991 | 199052 | 330968 | 416262 | 66925428 | N/F | 4567584 | N/F | 576390 | 1843876 | 1114612 | 3209818 | 2358648 | 14344332 | 180419 | 341907 | 428412 | 1113075 | 2384164 | 13926006 | 7752940 | 1181855 | 28143191 | 2543040 | 21920330 | 1563970 | 16594272 | 556150 | N/F | N/F | 1449487 | N/F | 261 | 3316007 | 503731 | 26331193 | N/F | 348 | 13244952 | 7308955 | 199059 | 14122256 | 2465503 | 19783496 | 48641 | 2442402 | 1153771 | 2706163 | 27604180 | 684 | 24509778 |  | Unknown
217315-d2-reprep | N/F | 20725 | 71577 | 6322 | 6503 | N/F | 5917056 | N/F | 63208 | N/F | N/F | 7349 | 1369 | 40863 | 56351 | 508385 | 1029380 | 315346 | 1848991 | 7099238 | 2492778 | 267490 | 35013426 | 2894322 | 19174601 | 3730780 | 25701650 | 52733 | N/F | N/F | 398883 | N/F | 547 | 10980 | 39223 | 41662 | 934 | 2365 | 79003 | N/F | 156992 | 27076 | 3305 | 66067 | 14281 | 8693 | N/F | 5990 | 50405 | 1053 | 52289 |  | Unknown
217335-d2-LD-reprep | 1252 | 21156 | 71679 | 10688 | N/F | N/F | 5853497 | N/F | 49101 | N/F | N/F | 17196 | 523 | 1083104 | 62092 | 534687 | 960509 | 305429 | 1888781 | 7420324 | 2367644 | 225857 | 33019176 | 2852319 | 19509902 | 3588274 | 25665879 | 39079 | N/F | N/F | 398711 | N/F | 637 | 8444 | 93208 | 165617 | 1780 | 3306 | 19139 | N/F | 171100 | 27821 | 3533 | 70810 | 10500 | 8026 | N/F | 4155 | 46380 | 830 | 32456 |  | Unknown
217336-d2-LFSM-200-reprep | N/F | 45298 | 331633 | 530673 | 59872634 | N/F | 4068641 | N/F | 12256 | 238309 | 138888 | 2291264 | 1034837 | 3114858 | 26388 | 286281 | 528893 | 135649 | 1150272 | 4144848 | 1164473 | 79394 | 21729644 | 1999570 | 11821768 | 2616570 | 16744072 | 9796 | N/F | N/F | 201239 | N/F | 302 | 3077124 | 32977 | 20732272 | 803 | 1314 | 3911811 | 1069112 | 209474 | 14521447 | 2675781 | 10344228 | 55875 | 2598330 | 73737 | 583376 | 26119912 | 981 | 20291816 |  | Unknown
217356-d2-reprep | N/F | 20069 | 82053 | 15334 | 3814 | N/F | 5870232 | 8 | 76680 | 875 | N/F | 4803 | 1412 | 31725 | 68036 | 653299 | 1069240 | 362507 | 2077732 | 8096076 | 2734706 | 308952 | 35078474 | 2845001 | 21287780 | 3725669 | 26439273 | 64045 | N/F | N/F | 488862 | N/F | 559 | 12986 | N/F | 736551 | 946 | 4141 | 18245 | 6113 | 151352 | 28435 | 3570 | 74310 | 17170 | 11041 | N/F | 4985 | 45274 | 613 | 34940 |  | Unknown
217357-d2-reprep | N/F | 32052 | 75356 | 44566 | 5801 | N/F | 5244048 | N/F | 67015 | 2362 | N/F | 4165 | 1693 | 32992 | 65595 | 604243 | 1001574 | 337621 | 1990548 | 7773840 | 2671781 | 296061 | 33193268 | 2849331 | 19919468 | 3765206 | 26011235 | 54807 | N/F | N/F | 453180 | N/F | 230 | 11121 | N/F | 7868034 | N/F | 3284 | 18296 | N/F | 164534 | 100436 | 1901 | 75649 | 15712 | 8492 | N/F | 5553 | 42301 | N/F | 41112 |  | Unknown
217358-d2-reprep | N/F | 23677 | 72136 | 24997 | N/F | N/F | 4729105 | N/F | 79584 | N/F | N/F | 10085 | 2848 | 24194 | 65171 | 566785 | 342994 | 358746 | 1989970 | 7845774 | 2800127 | 333002 | 11191331 | 2870189 | 20363549 | 3693069 | 26309173 | 66583 | N/F | N/F | 489764 | N/F | 798 | 10327 | N/F | 141319505 | 621 | 4149 | N/F | N/F | 163871 | 134025 | 3024 | 90072 | 15353 | 7193 | N/F | 4477 | 41984 | 1018 | 69786 |  | Unknown
217359-d2-LD | 971 | 17870 | 66847 | 8947 | N/F | N/F | 3414232 | N/F | 65162 | N/F | N/F | 3112 | 837 | 12195 | 61072 | 440804 | 214548 | 310530 | 1734908 | 7024500 | 2495623 | 280112 | 5267128 | 3279024 | 17329609 | 3248796 | 23961101 | 52260 | N/F | N/F | 423000 | N/F | 461 | 11501 | N/F | 225954508 | N/F | 2149 | 20654 | N/F | 141407 | 296928 | 595 | 96978 | 17439 | 744 | N/F | 4925 | 47573 | 1027 | 134731 |  | Unknown
217359-d2-reprep | N/F | 18373 | 65981 | 17265 | N/F | N/F | 3297853 | N/F | 80302 | N/F | N/F | 5443 | 1556 | 22622 | 63996 | 496533 | 254931 | 371808 | 1898308 | 7682663 | 2777479 | 336158 | 5216648 | 2893285 | 19783128 | 3700322 | 27355959 | 65076 | N/F | N/F | 490242 | N/F | 443 | 13916 | N/F | 253317030 | N/F | 2051 | 40984 | N/F | 162467 | 369067 | 2577 | 135285 | 18232 | 8939 | N/F | 5175 | 47985 | 1078 | 153944 |  | Unknown
217360-d2-LFSM-200 | N/F | 51737 | 386323 | 499128 | 65617849 | N/F | 4506121 | N/F | 50386 | 464645 | 293350 | 2807931 | 1420577 | 2570668 | 49901 | 339709 | 493857 | 299353 | 1637116 | 6241613 | 2189551 | 228744 | 21867262 | 2969746 | 15306684 | 2838369 | 20095160 | 41412 | N/F | N/F | 387900 | N/F | 774 | 3289355 | 41315 | 41189890 | N/F | 2566 | 5937270 | 2024299 | 209091 | 17229779 | 2461164 | 14150233 | 61962 | 2941876 | 218097 | 917692 | 28022948 | 937 | 22066640 |  | Unknown
217360-d2-reprep | 860 | 20584 | 86535 | 9732 | N/F | N/F | 5448112 | 70 | 47413 | N/F | N/F | 4708 | 1281 | 21560 | 54102 | 585753 | 788245 | 289713 | 1758120 | 6538550 | 2278169 | 216977 | 24984103 | 2875732 | 17853311 | 3663221 | 23347612 | 36267 | 967 | N/F | 381007 | N/F | 661 | 12630 | N/F | 46315496 | 791 | 2628 | 33153 | N/F | 155210 | 27133 | 1220 | 70856 | 13699 | 7236 | N/F | 5296 | 45437 | 730 | 42640 |  | Unknown
217375-d2-LD-reprep | N/F | 25384 | 390526 | 26969 | 3790 | N/F | 5411978 | N/F | 39026 | N/F | N/F | 3022 | 777 | 33629 | 50156 | 401069 | 942737 | 252033 | 1396354 | 5781199 | 2050127 | 199392 | 30356496 | 2682278 | 15040827 | 3283182 | 21663026 | 32562 | N/F | N/F | 353343 | N/F | 436 | 12187 | N/F | 4245281 | 560 | 3347 | 18647 | 9205 | 157693 | 27408 | 1103 | 75167 | 13059 | 7257 | N/F | 4041 | 47983 | N/F | 31763 |  | Unknown
217376-d2-LFSM-200-reprep | N/F | 65564 | 700419 | 558954 | 71558837 | N/F | 4773089 | N/F | 56985 | 567072 | 355833 | 3050272 | 1975287 | 2746035 | 63670 | 437084 | 567229 | 384978 | 2245190 | 9294029 | 2947540 | 272169 | 25146296 | 2431843 | 20267034 | 3180995 | 22022785 | 47681 | N/F | N/F | 513604 | N/F | 461 | 3400935 | 51364 | 32891947 | N/F | 4066 | 8276010 | 2544314 | 198470 | 18045981 | 2684779 | 16832591 | 52201 | 3037869 | 245830 | 1343814 | 26256051 | N/F | 21917322 |  | Unknown
217393-d4-reprep | N/F | 24587 | 100701 | 24049 | N/F | N/F | 5640495 | N/F | 183676 | N/F | N/F | 2823 | 1524 | 19025 | 99243 | 812677 | 590817 | 543343 | 2761566 | 11319101 | 4115075 | 582291 | 18289200 | 2968573 | 25726568 | 3815020 | 29222054 | 167116 | N/F | 1465 | 714264 | N/F | 1021 | 13864 | N/F | 95234437 | N/F | 1593 | 103059 | 15287 | 165174 | 385799 | 1237 | 500710 | 14230 | 9251 | N/F | 7239 | 46953 | N/F | 33570 |  | Unknown
217394-d4-reprep | N/F | 23761 | 78999 | 15420 | N/F | N/F | 5491444 | N/F | 179021 | 2159 | N/F | 3568 | 1171 | 15388 | 91400 | 664419 | 474956 | 526280 | 2660302 | 11534145 | 3908243 | 571202 | 16523189 | 3013250 | 26064711 | 3844497 | 30325157 | 164144 | N/F | N/F | 686028 | N/F | 386 | 15099 | N/F | 114609128 | N/F | 810 | 136162 | 23132 | 167818 | 552495 | 1636 | 615378 | 15394 | 9407 | N/F | 7571 | 44258 | N/F | 37063 |  | Unknown
217395-d4-LD-reprep | 838 | 22985 | 75246 | 8174 | N/F | N/F | 5095457 | N/F | 174845 | N/F | N/F | 7080 | 616 | 15865 | 93252 | 704026 | 377820 | 546714 | 2715983 | 11584406 | 3962948 | 566562 | 11471986 | 2855096 | 26488218 | 3864173 | 30315119 | 162917 | N/F | N/F | 686795 | 1317 | 681 | 12491 | N/F | 146680128 | 442 | 1469 | 180267 | 29448 | 165754 | 1165027 | 1345 | 1180013 | 14110 | 11149 | N/F | 8521 | 50782 | N/F | 38085 |  | Unknown
217395-d4-reprep | N/F | 20327 | 60736 | 10125 | N/F | N/F | 4851454 | N/F | 183884 | 515 | N/F | 2363 | 1137 | 14219 | 97274 | 671870 | 354150 | 528791 | 2784250 | 11700873 | 4108574 | 581720 | 11213632 | 3077095 | 26148752 | 3838100 | 30556450 | 167429 | N/F | N/F | 707509 | N/F | 235 | 12535 | N/F | 149269121 | 945 | 1328 | 211675 | 30759 | 163359 | 1167012 | 1060 | 1198362 | 15868 | 8321 | N/F | 7380 | 44033 | 745 | 40267 |  | Unknown
217396-d200-LFSM-200 | 768 | 206007 | 356184 | 237947 | 65496904 | N/F | 3821865 | N/F | 597045 | 1739541 | 1101227 | 2817359 | 2304057 | 2357196 | 198069 | 363751 | 199307 | 1163052 | 2649285 | 14478970 | 8146487 | 1222092 | 8403033 | 2527858 | 23886278 | 2716368 | 25767847 | 574760 | N/F | N/F | 1554299 | N/F | 195 | 2508156 | 479586 | 166477566 | N/F | 322 | 12954911 | 7106511 | 224430 | 21266355 | 970107 | 20081097 | 56253 | 3028653 | 1070801 | 2659873 | 30255459 | 651 | 24175109 |  | Unknown
217399-d2-reprep | 436 | 24780 | 79046 | 16361 | 2733 | N/F | 4784404 | N/F | 73560 | N/F | 23475 | 3918 | N/F | 14221 | 75456 | 867722 | 1302752 | 368700 | 1942365 | 7573260 | 2773002 | 304273 | 27407650 | 2628133 | 18007646 | 3282239 | 21440348 | 60835 | 1499 | N/F | 445606 | N/F | 560 | 8737 | N/F | 581372 | N/F | 3394 | 84409 | N/F | 133264 | 21631 | 900 | 74854 | 15246 | 10291 | N/F | 6211 | 34634 | N/F | 20985 |  | Unknown
217400-d2-reprep | 632 | 22957 | 81470 | 13331 | N/F | N/F | 5367379 | N/F | 68656 | N/F | 24767 | 9145 | N/F | 12985 | 85595 | 1056779 | 1451839 | 387634 | 2705573 | 10805928 | 3286410 | 308810 | 31039101 | 2832429 | 23929613 | 3662110 | 24811018 | 57369 | N/F | N/F | 543005 | N/F | 501 | 10580 | N/F | 298312 | N/F | 1818 | 74290 | N/F | 143622 | 26057 | 1539 | 102749 | 13292 | 9435 | N/F | 9733 | 35047 | N/F | 20493 |  | Unknown
217415-d2-LD-reprep | 808 | 25656 | 78343 | 4050 | N/F | N/F | 4155058 | N/F | 46103 | 1320 | 206464 | 7465 | N/F | 11063 | 60766 | 805962 | 1242937 | 291605 | 1720166 | 6871050 | 2263857 | 205106 | 26651054 | 3222867 | 16258879 | 3296585 | 20903479 | 37833 | N/F | N/F | 374812 | N/F | 443 | 6241 | N/F | 1381459 | N/F | 3889 | 94703 | N/F | 133017 | 23718 | 1170 | 51414 | 13381 | 628 | N/F | 5660 | 31225 | N/F | 16381 |  | Unknown
217416-d2-LFSM-200-reprep | N/F | 83241 | 788735 | 1003138 | 62131940 | N/F | 3782519 | N/F | 47699 | 456242 | 550463 | 2852845 | 1598869 | 2454321 | 53730 | 741540 | 986325 | 301510 | 1743168 | 6376900 | 2212125 | 223481 | 20335293 | 2868447 | 13558401 | 2693800 | 15469898 | 38456 | N/F | N/F | 389577 | N/F | 752 | 2787936 | 42059 | 19451561 | N/F | 2641 | 6051152 | 2028848 | 200643 | 13403418 | 2584366 | 12631617 | 60323 | 2844100 | 210802 | 1013276 | 25704393 | N/F | 17533860 |  | Unknown
AB-230117-02_CCC-H_12B | N/F | 264141 | 458136 | 485437 | 113179950 | N/F | 3427960 | N/F | 398600 | 2607896 | 1687348 | 4313477 | 3777218 | 3813315 | 131248 | 244035 | 255173 | 865786 | 1929331 | 10935415 | 5985086 | 907564 | 23410432 | 2706727 | 17713375 | 2071302 | 20159351 | 382165 | N/F | N/F | 1120831 | N/F | 240 | 4717789 | 644587 | 40550901 | N/F | 277 | 20000465 | 10397662 | 125859 | 33049742 | 3520165 | 31103722 | 19847 | 4778743 | 1641406 | 4193316 | 42349328 | N/F | 38429312 |  | Unknown
AB-230117-02_LFB-H_11_20230131064413 | N/F | 212962 | 385844 | 480726 | 72657344 | N/F | 4968740 | 973 | 616911 | 2088516 | 1278132 | 3202712 | 2659566 | 2642671 | 183445 | 331273 | 520881 | 1096484 | 2263951 | 13426920 | 7447833 | 1184394 | 28409833 | 1931638 | 22092585 | 2412550 | 23966959 | 590402 | 1534 | N/F | 1449455 | N/F | 485 | 363552 | 628744 | 29701807 | N/F | 35 | 14702218 | 8102199 | 226466 | 24864372 | 276486 | 22656358 | 66662 | 276542 | 1304200 | 3049501 | 5261060 | N/F | 33033199 |  | Unknown
AB-230117-02_QCS_2B | N/F | 215077 | 377764 | 370620 | 66660321 | N/F | 3917737 | N/F | 598270 | 1895423 | 1188016 | 2895607 | 2411660 | 2438116 | 182022 | 328270 | 349367 | 1055790 | 2267959 | 13391104 | 7455815 | 1167081 | 26824835 | 1872018 | 21990461 | 2463527 | 23190728 | 579051 | N/F | N/F | 1405041 | N/F | 611 | 3250291 | 552812 | 26975487 | 1918 | 129 | 13787266 | 7434792 | 200321 | 21660094 | 2580785 | 20857326 | 51320 | 2953398 | 1194124 | 2787749 | 30139689 | 417 | 25655151 |  | Unknown
AB-230117-02_LRB_22C_20230131055005 | N/F | 24277 | 47797 | 6246 | 138063 | 1479 | 5467133 | N/F | 653772 | 2285 | N/F | 10235 | 4227 | 14820 | 211547 | 539032 | 635159 | 1182442 | 3194747 | 17482559 | 8603870 | 1266632 | 39102975 | 2783438 | 28518683 | 3287098 | 31593070 | 629245 | N/F | N/F | 1611581 | N/F | 616 | 22288 | N/F | 316513 | N/F | 150 | 75425 | 12330 | 145128 | 64827 | 12603 | 130659 | 15483 | 14691 | N/F | 7181 | 135995 | 845 | 90870 |  | Unknown
AB-230117-02_LRB_22C | N/F | 23031 | 63576 | 31461 | 4232912 | N/F | 5989035 | 47 | 602175 | N/F | N/F | 114178 | 49182 | 137205 | 217265 | 611764 | 779377 | 1194069 | 3444078 | 18194704 | 8922549 | 1268080 | 40081974 | 2997945 | 30780928 | 3558349 | 32350728 | 580530 | N/F | N/F | 1595323 | N/F | 252 | 208061 | N/F | 1903816 | N/F | 719 | 174189 | 15791 | 186971 | 856870 | 207949 | 576952 | 25336 | 150619 | N/F | 12235 | 3503220 | 1184 | 2025680 |  | Unknown
WO-23003-57-A-1 | N/F | 40256 | 73763 | 4948 | N/F | N/F | 4214765 | N/F | 52953 | N/F | 296 | 7653 | N/F | 19400 | 63711 | 517283 | 1003054 | 169211 | 1718804 | 6773288 | 2450418 | 194259 | 34045107 | 3234054 | 16898084 | 3724831 | 25858895 | 44556 | N/F | N/F | 455739 | N/F | 303 | 13174 | N/F | 819985 | N/F | 1374 | 42898 | N/F | 164590 | 32501 | 1900 | 41419 | 18099 | 629 | N/F | 5073 | 61928 | 1558 | 37403 |  | Unknown

## Sheet4
Batch Folder Path: | D:\TraceFinderData\Projects\PFAS Online SPE\2023-01-18\2023-01-18
Batch Name: | 2023-01-18
Last Acquisition Date: | 2023-01-31 07:17
User Name: | Quantiva
Sample Count: | 365
Area - Zero
Data File | 11Cl-PF3OUdS | 4:2 FTS | 6:2 FTS | 8:2 FTS | 9Cl-PF3ONS | N-EtFOSA-D3 | N-MeFOSA-D3 | FBSA | FHxA | HFPO-DA | PFBS | PFHpS | PFHxS | PFOS | M2-4:2 FTS | M2-6:2 FTS | M2-8:2 FTS | M3PFBS | M3PFHxS | M4PFHpA | M5PFHxA | M5PFPeA | M6PFDA | M8FOSA | M8PFOA | M8PFOS | M9PFNA | MPFBA | MPFDoA | MPFUdA | MPHFPO-DA | NaDONA | N-EtFOSAA | N-MeFOSA | PFBA | PFDA | PFDoA | PFDS | PFHpA | PFHxA | PFHxDA | PFNA | PFNS | PFOA | PFODA | PFOSA | PFPeA | PFPeS | PFTeDA | PFTrDA | PFUdA | All Flags | Sample Type
AB-230117-02 Cal Blank 1 | 0 | 112516 | 227843 | 17023 | 994323 | 0 | 3288127 | 0 | 654123 | 0 | 0 | 49573 | 26032 | 89073 | 212284 | 1097131 | 1252931 | 791404 | 3461600 | 15471327 | 6168538 | 1139560 | 30301110 | 4635033 | 27779178 | 4273192 | 30410864 | 632991 | 0 | 0 | 999100 | 0 | 0 | 14483 | 0 | 375299 | 0 | 0 | 261105 | 16507 | 93148 | 330634 | 5902 | 321565 | 11786 | 9268 | 0 | 14859 | 72946 | 0 | 307619 |  | Unknown
AB-230117-02 Cal Std 1 1 | 0 | 77777 | 77386 | 28590 | 440258 | 0 | 5220775 | 0 | 396975 | 9558 | 0 | 23909 | 17310 | 51290 | 205380 | 956924 | 1148416 | 474992 | 3721715 | 17841442 | 7254807 | 563429 | 40439150 | 4445809 | 31153664 | 4544352 | 35852585 | 374515 | 0 | 0 | 1412490 | 0 | 0 | 33229 | 0 | 265543 | 0 | 0 | 254669 | 50035 | 129187 | 211226 | 23099 | 247234 | 13991 | 19796 | 0 | 15209 | 215094 | 0 | 216278 |  | Cal Std
AB-230117-02 Cal Std 2 1 | 0 | 62203 | 56912 | 18006 | 745795 | 0 | 5001560 | 0 | 342400 | 16470 | 6034 | 46984 | 33917 | 60367 | 195991 | 926769 | 975062 | 914458 | 3812773 | 17947361 | 6943986 | 1179073 | 39003692 | 4183962 | 31781766 | 4444601 | 35137464 | 334728 | 0 | 0 | 1300088 | 0 | 1045 | 51901 | 0 | 388301 | 0 | 1636 | 189621 | 71334 | 139081 | 317888 | 32714 | 352822 | 18587 | 37787 | 0 | 34689 | 393070 | 0 | 375224 |  | Cal Std
AB-230117-02 Cal Std 3 1 | 0 | 69177 | 52924 | 20348 | 1595719 | 0 | 5140054 | 43 | 456698 | 35668 | 16020 | 81831 | 68453 | 98708 | 208317 | 818327 | 1052832 | 964959 | 3747583 | 17691989 | 7322397 | 1125545 | 41799297 | 4081775 | 27679054 | 4174088 | 34506802 | 433302 | 2234 | 0 | 1271882 | 0 | 0 | 86354 | 4306 | 831686 | 0 | 0 | 341420 | 139410 | 147347 | 598183 | 87517 | 568100 | 17990 | 74404 | 0 | 66633 | 742431 | 0 | 701872 |  | Cal Std
AB-230117-02 Cal Std 4 1 | 0 | 75233 | 87011 | 55683 | 4341376 | 0 | 5174106 | 0 | 409949 | 77901 | 31394 | 234223 | 159856 | 211795 | 196635 | 922189 | 1089840 | 775858 | 3751819 | 17911014 | 6862671 | 1142105 | 42979199 | 4355222 | 31867700 | 4570958 | 37211667 | 389333 | 0 | 0 | 1286073 | 0 | 804 | 202411 | 16664 | 1982722 | 0 | 0 | 857101 | 310164 | 160570 | 1577968 | 209710 | 1486141 | 8975 | 201087 | 0 | 162255 | 1770992 | 0 | 1792559 |  | Cal Std
AB-230117-02 Cal Std 5 1 | 0 | 81494 | 115586 | 98530 | 8341543 | 0 | 4897010 | 0 | 399852 | 152272 | 55155 | 402582 | 338372 | 365677 | 173345 | 809384 | 1033500 | 536479 | 3716010 | 17529330 | 6974866 | 536464 | 41768495 | 4131540 | 27530639 | 4177449 | 33993178 | 384293 | 0 | 0 | 1283661 | 0 | 1149 | 385491 | 31872 | 3774582 | 0 | 0 | 1624702 | 617567 | 164135 | 2913228 | 414950 | 2515035 | 13909 | 374702 | 0 | 308080 | 3464585 | 0 | 3481031 |  | Cal Std
AB-230117-02 Cal Std 6 1 | 0 | 101858 | 204888 | 182076 | 15888415 | 0 | 4680611 | 0 | 434158 | 276434 | 132726 | 899511 | 651156 | 699407 | 180813 | 871544 | 1003720 | 715733 | 3596018 | 16645776 | 6780447 | 1051047 | 38998738 | 3817485 | 30647728 | 3900993 | 32905329 | 411914 | 0 | 0 | 1184376 | 0 | 733 | 700485 | 76673 | 7152575 | 0 | 0 | 3179923 | 1182559 | 162361 | 5766346 | 773530 | 5612352 | 20213 | 726718 | 0 | 609580 | 6777654 | 0 | 6703160 |  | Cal Std
AB-230117-02 Cal Std 7 1 | 0 | 144976 | 418191 | 426365 | 39370726 | 0 | 3935057 | 0 | 287137 | 614770 | 373027 | 2202332 | 1559870 | 1805281 | 162804 | 772185 | 861916 | 737784 | 3292381 | 14061854 | 5695371 | 955118 | 33433459 | 3825231 | 27053869 | 4019358 | 32561100 | 266916 | 0 | 0 | 1006545 | 0 | 0 | 1531077 | 136686 | 15974735 | 0 | 0 | 6969311 | 2652843 | 178718 | 14508535 | 1821237 | 12973330 | 30320 | 1857538 | 0 | 1324925 | 16176710 | 0 | 14081290 |  | Cal Std
AB-230117-02 Cal Std 8 1 | 0 | 252728 | 620784 | 677193 | 71997932 | 0 | 3222112 | 0 | 367743 | 1636057 | 958314 | 4199129 | 3215157 | 3317470 | 193926 | 623863 | 676473 | 943092 | 3257272 | 15570319 | 7138944 | 1073867 | 32156115 | 3658943 | 27803507 | 3648036 | 31218635 | 355650 | 0 | 0 | 1298685 | 0 | 0 | 2529592 | 317397 | 30451155 | 0 | 0 | 15394166 | 6805908 | 192354 | 27303504 | 3418214 | 26008307 | 50615 | 3743469 | 0 | 3246458 | 34014521 | 0 | 27246902 |  | Cal Std
AB-230117-02 Cal Std 9 1 | 0 | 372077 | 905320 | 851712 | 116193964 | 0 | 2442342 | 0 | 323358 | 3247807 | 1701817 | 6846874 | 5460470 | 5632878 | 170163 | 430793 | 483992 | 823849 | 2689261 | 13298661 | 6821542 | 1041605 | 26016671 | 3254875 | 23663020 | 2894577 | 25709605 | 307127 | 0 | 0 | 1263633 | 0 | 608 | 3901024 | 557472 | 49179179 | 0 | 0 | 26176244 | 13033490 | 238950 | 45911037 | 5400832 | 44738056 | 84652 | 6346673 | 0 | 5856515 | 56919922 | 0 | 44285709 |  | Cal Std
AB-230117-02 Cal Std 10 1 | 0 | 805127 | 1433828 | 1275925 | 197461920 | 0 | 1467537 | 0 | 375270 | 7667671 | 4123554 | 11567701 | 9750392 | 9590556 | 146615 | 267204 | 320707 | 735204 | 1848502 | 8850581 | 5884437 | 1086232 | 16098522 | 2612917 | 16194891 | 1890722 | 16620897 | 355318 | 0 | 0 | 1166714 | 2538 | 0 | 6112365 | 1880324 | 76750243 | 0 | 820 | 44673947 | 29261490 | 329132 | 75456240 | 8544439 | 79267874 | 159352 | 12970999 | 0 | 11150786 | 119694688 | 0 | 69996361 |  | Cal Std
AB-230117-02 Cal Std 11 1 | 0 | 1332529 | 2118795 | 2077052 | 272702448 | 0 | 990370 | 0 | 419402 | 12061892 | 7476913 | 16347014 | 13118827 | 11968729 | 123995 | 217203 | 190259 | 743175 | 1182015 | 5514718 | 4663536 | 1027056 | 9992543 | 1895890 | 10669119 | 1117397 | 10205299 | 399743 | 1447 | 0 | 930811 | 4193 | 0 | 8231722 | 4061475 | 95537350 | 0 | 0 | 55812284 | 46123697 | 575935 | 91743660 | 13093151 | 101445144 | 354082 | 20481959 | 0 | 15936312 | 164908377 | 0 | 85329497 |  | Cal Std
AB-230117-02 Cal Std 12 1 | 0 | 2052249 | 3752510 | 3400229 | 322467204 | 0 | 756279 | 0 | 366638 | 17411401 | 13623368 | 20315876 | 19347546 | 16014447 | 88970 | 177876 | 168650 | 624962 | 801246 | 3271930 | 3204749 | 893305 | 5309212 | 1428471 | 6088887 | 717869 | 5450072 | 344196 | 0 | 0 | 650094 | 9030 | 0 | 12733947 | 8016201 | 104205973 | 0 | 0 | 68699541 | 66607272 | 888421 | 100296778 | 13710573 | 122200751 | 645550 | 31534699 | 0 | 22410258 | 245054542 | 0 | 93843873 |  | Cal Std
AB-230117-02 Cal Std 13 1 | 0 | 3596951 | 7130291 | 6678346 | 412145160 | 0 | 509906 | 0 | 325335 | 20974427 | 22784388 | 26960346 | 27672585 | 23051221 | 70362 | 158377 | 145781 | 398687 | 462298 | 1495614 | 1609916 | 697777 | 2537051 | 987007 | 2671403 | 409631 | 2633403 | 307430 | 0 | 2123 | 337049 | 19283 | 0 | 20401024 | 17446659 | 125687496 | 0 | 0 | 78667431 | 78848802 | 1634231 | 123270974 | 17814751 | 135061825 | 1390553 | 54196293 | 0 | 29004789 | 351506807 | 0 | 117265456 |  | Cal Std
AB-230117-02 CCC-H 1 | 0 | 388495 | 810157 | 875305 | 109538928 | 0 | 2325725 | 0 | 384587 | 3128817 | 1503395 | 6509411 | 5212732 | 4718647 | 164895 | 412475 | 452597 | 782284 | 2601336 | 12865665 | 6532921 | 994672 | 24555753 | 2877663 | 22963695 | 2474341 | 23069723 | 360810 | 0 | 0 | 1247833 | 0 | 0 | 3695792 | 659799 | 45540584 | 0 | 0 | 24994117 | 12482335 | 259446 | 40492486 | 5165347 | 43778576 | 123621 | 5456105 | 0 | 5576200 | 59517656 | 0 | 44937897 |  | Unknown
AB-230117-02 CCC-H 2 | 0 | 290696 | 705630 | 708675 | 115648404 | 0 | 2394529 | 0 | 344961 | 2808570 | 1426258 | 5823821 | 4682335 | 4794788 | 135186 | 353676 | 415796 | 687121 | 2292144 | 11308462 | 5777486 | 832964 | 25113502 | 2757946 | 20239051 | 2438598 | 22720073 | 329585 | 0 | 0 | 1089784 | 1845 | 0 | 3933474 | 623594 | 46895626 | 1910 | 0 | 22701593 | 10898005 | 224647 | 40696373 | 4750761 | 38501227 | 77710 | 5425710 | 0 | 4845984 | 56699300 | 0 | 41864943 |  | Unknown
AB-230117-02 CCC-H 3 | 0 | 184226 | 663586 | 646724 | 117127535 | 0 | 2284633 | 0 | 369891 | 1911978 | 1328501 | 5577277 | 4419230 | 4852225 | 87394 | 321704 | 394601 | 681707 | 2126517 | 10854643 | 3462214 | 810462 | 25012779 | 2736437 | 20023795 | 2428572 | 22761914 | 355150 | 0 | 0 | 747444 | 0 | 2337 | 3773025 | 718988 | 48248028 | 0 | 432 | 21440626 | 6484508 | 242485 | 40397962 | 4890617 | 37734645 | 82046 | 5301462 | 0 | 3457511 | 57544884 | 0 | 42088094 |  | Unknown
AB-230117-02 IBC 1 | 0 | 230782 | 560563 | 569511 | 73666088 | 0 | 3281939 | 0 | 360258 | 1529370 | 678693 | 4019409 | 3050938 | 3175699 | 172133 | 531069 | 615180 | 694321 | 3145626 | 15020003 | 6831085 | 1028940 | 33994344 | 3566642 | 26669533 | 3525478 | 30178374 | 343074 | 0 | 0 | 1265968 | 0 | 0 | 2526377 | 316501 | 32587756 | 0 | 0 | 14728164 | 6462707 | 291855 | 27403201 | 3219645 | 25442507 | 192708 | 3228987 | 0 | 2935547 | 30909836 | 0 | 29909284 |  | Unknown
AB-230117-02 LD 1 | 0 | 68371 | 117144 | 7141 | 29129 | 0 | 3197003 | 0 | 101563 | 0 | 30078 | 9431 | 6717 | 64906 | 153391 | 1339623 | 1434346 | 532312 | 3556895 | 16094248 | 5676603 | 514038 | 32960813 | 3685056 | 27273438 | 4080404 | 27937670 | 92502 | 0 | 0 | 969578 | 0 | 1076 | 10774 | 294730 | 219451 | 0 | 0 | 117539 | 29312 | 100593 | 69529 | 2296 | 78651 | 10933 | 1986 | 0 | 5293 | 49425 | 1341 | 25363 |  | Unknown
AB-230117-02 LD 2 | 0 | 44930 | 0 | 7948 | 22047 | 0 | 3993193 | 0 | 326052 | 0 | 8284 | 3367635 | 98029 | 39020108 | 163341 | 481989 | 659851 | 757165 | 3240989 | 16001726 | 6420086 | 926430 | 36730462 | 1661483 | 25510276 | 897677 | 9308318 | 314101 | 0 | 0 | 1139792 | 0 | 427 | 14243 | 0 | 72927 | 0 | 612 | 71803 | 0 | 138662 | 0 | 194050 | 89979 | 15455 | 502 | 0 | 6648 | 67062 | 0 | 64493 |  | Unknown
AB-230117-02 LFB-H 1 | 0 | 64541 | 37800 | 13404 | 1799905 | 0 | 3518377 | 0 | 354509 | 1382273 | 584790 | 4239384 | 3568705 | 3433927 | 177110 | 600556 | 639819 | 625103 | 3203708 | 15779197 | 6609916 | 878562 | 31292462 | 3487758 | 27298231 | 3380199 | 29701401 | 344493 | 0 | 0 | 1214377 | 0 | 0 | 61885 | 345397 | 29431932 | 0 | 0 | 17083687 | 6375282 | 30094960 | 28135449 | 3095157 | 27093950 | 18320672 | 55987 | 0 | 3137992 | 31971867 | 0 | 30976023 |  | Unknown
AB-230117-02 LFB-H 2 | 0 | 202154 | 467443 | 550623 | 65960133 | 1567 | 3045895 | 0 | 348908 | 1429351 | 773186 | 3368892 | 2665792 | 2818829 | 160348 | 447603 | 565583 | 786185 | 2698290 | 13657522 | 6097995 | 917574 | 31196287 | 3049218 | 23586649 | 3069555 | 27131872 | 332746 | 0 | 0 | 1141890 | 0 | 0 | 2458413 | 305169 | 29265704 | 0 | 0 | 13348964 | 5776510 | 176979 | 24412799 | 2962277 | 22252299 | 39068 | 2951827 | 0 | 2610078 | 29727846 | 0 | 25811068 |  | Unknown
AB-230117-02 LFB-H 3 | 0 | 205785 | 425956 | 464244 | 63153442 | 0 | 3211706 | 0 | 346938 | 1389553 | 741120 | 3409855 | 2585846 | 8274632 | 146474 | 418177 | 519769 | 728784 | 2651172 | 13461380 | 6134568 | 892572 | 31838461 | 2822516 | 23679978 | 2536330 | 25011867 | 336965 | 0 | 0 | 1098817 | 637 | 405 | 2541470 | 311284 | 29095176 | 0 | 310 | 13351595 | 5849304 | 185020 | 21905938 | 2862308 | 22469808 | 43689 | 2627944 | 0 | 2638704 | 30888735 | 0 | 26999380 |  | Unknown
AB-230117-02 LFSM-H 1 | 0 | 106459 | 653915 | 891040 | 65925722 | 0 | 2511692 | 0 | 36312 | 496934 | 288538 | 3061909 | 1699067 | 2803874 | 54038 | 582642 | 795963 | 252528 | 1754914 | 6334983 | 2070882 | 212919 | 21592786 | 2657785 | 14564917 | 2837556 | 18266669 | 32194 | 968 | 0 | 393725 | 0 | 942 | 2258626 | 101860 | 22519566 | 0 | 0 | 6573200 | 2031886 | 194132 | 17673841 | 3081657 | 14708479 | 53550 | 2881515 | 0 | 1082646 | 28006220 | 0 | 20600453 |  | Unknown
AB-230117-02 LFSM-H 2 | 0 | 163842 | 353499 | 399778 | 47160721 | 0 | 2493852 | 0 | 381508 | 1406750 | 733658 | 5913930 | 2771860 | 36369609 | 154709 | 363257 | 432832 | 703762 | 2642276 | 13506293 | 6218454 | 883837 | 27198676 | 1672447 | 21889536 | 871340 | 8742054 | 365992 | 0 | 0 | 1111637 | 0 | 0 | 2138552 | 338236 | 27524513 | 0 | 710 | 13128119 | 5875916 | 197645 | 8179450 | 2893418 | 21489183 | 56183 | 1597583 | 0 | 2549931 | 36114997 | 0 | 26850794 |  | Unknown
LRB2-Inj1 | 0 | 55961 | 30946 | 10821 | 59342 | 0 | 4273045 | 0 | 423362 | 0 | 0 | 0 | 0 | 21794 | 176728 | 727569 | 797187 | 709896 | 3550835 | 17311626 | 6426584 | 1025419 | 38121486 | 3879353 | 30052213 | 4082623 | 34535528 | 404157 | 0 | 0 | 1239312 | 0 | 0 | 15670 | 0 | 38159832 | 0 | 0 | 37212 | 0 | 150348 | 51138 | 3680 | 92068 | 18988 | 3493 | 0 | 7297 | 109276 | 0 | 63530 |  | Unknown
AB-230117-02 LRB 2 | 0 | 56734 | 31701 | 14418 | 40828 | 0 | 4314714 | 0 | 302056 | 0 | 0 | 0 | 0 | 17229 | 176285 | 677491 | 821716 | 579252 | 3186044 | 16097250 | 6983145 | 938344 | 38345817 | 3522166 | 26557316 | 3692244 | 32251309 | 287597 | 0 | 0 | 1250979 | 0 | 0 | 17857 | 0 | 1388027 | 0 | 0 | 55166 | 11706 | 112364 | 38070 | 0 | 82074 | 17409 | 3579 | 0 | 8962 | 95883 | 0 | 67715 |  | Unknown
AB-230117-02 LRB 3 | 0 | 51265 | 28095 | 5237 | 49408 | 0 | 4290546 | 0 | 345033 | 0 | 0 | 30759 | 0 | 1465948 | 162304 | 688457 | 832141 | 608685 | 3316109 | 16259220 | 6519835 | 847660 | 39060283 | 3656214 | 28054916 | 3830699 | 32548149 | 329839 | 2263 | 0 | 1180544 | 0 | 0 | 13301 | 0 | 191061 | 0 | 0 | 55310 | 0 | 140710 | 41604 | 4706 | 84850 | 13717 | 4050 | 0 | 7259 | 85868 | 0 | 51669 |  | Unknown
LRB4-inj1 | 0 | 54344 | 34495 | 8715 | 15601 | 0 | 4532799 | 0 | 343234 | 0 | 0 | 0 | 0 | 29077 | 153937 | 687636 | 922012 | 783700 | 3160405 | 16377241 | 6208058 | 914384 | 40439399 | 3664301 | 27216054 | 3910568 | 33682350 | 331054 | 0 | 0 | 1167946 | 0 | 619 | 13268 | 0 | 69039 | 0 | 514 | 19679 | 8406 | 133018 | 29358 | 0 | 70823 | 12142 | 1718 | 0 | 6783 | 72418 | 0 | 46718 |  | Unknown
AB-230117-02 LRB 4 | 0 | 50431 | 30205 | 8966 | 16343 | 0 | 4392534 | 91 | 268763 | 1498 | 0 | 0 | 0 | 23100 | 101270 | 659458 | 799942 | 813269 | 3118516 | 15967860 | 4316238 | 715444 | 39118334 | 3439577 | 27292063 | 3707401 | 32347881 | 255481 | 0 | 0 | 727112 | 0 | 888 | 13530 | 0 | 71640 | 0 | 0 | 0 | 0 | 112853 | 31262 | 0 | 64581 | 7393 | 1877 | 0 | 6911 | 64201 | 0 | 34556 |  | Unknown
LRB5-inj1 | 0 | 54361 | 31299 | 10647 | 42549 | 0 | 4453469 | 0 | 347304 | 0 | 0 | 3225 | 2629 | 21627 | 175622 | 648371 | 832802 | 790561 | 3170322 | 15967631 | 6538012 | 912161 | 39665652 | 3524230 | 27224949 | 3827007 | 33056389 | 333530 | 0 | 0 | 1158540 | 0 | 1968 | 16406 | 0 | 94746 | 1630 | 538 | 33302 | 0 | 142625 | 42538 | 1670 | 79172 | 15109 | 1414 | 0 | 6427 | 68715 | 0 | 61327 |  | Unknown
AB-230117-02 LRB 5 | 0 | 56464 | 30565 | 8425 | 11897 | 0 | 4190921 | 0 | 294495 | 0 | 780 | 0 | 1706 | 18524 | 148766 | 611427 | 750065 | 617862 | 2983135 | 15041867 | 5896299 | 785669 | 37000585 | 3226949 | 26102836 | 3543112 | 30826818 | 284660 | 0 | 0 | 1129294 | 0 | 0 | 13784 | 0 | 63331 | 0 | 415 | 30661 | 0 | 121710 | 27751 | 1469 | 68182 | 7490 | 1695 | 0 | 8470 | 61281 | 0 | 29983 |  | Unknown
Rinse5 | 0 | 54408 | 53104 | 42883 | 5083783 | 0 | 386841 | 372 | 13216 | 0 | 0 | 399579 | 109528 | 8757986 | 13230 | 32311 | 95278 | 16713 | 145690 | 494601 | 117238 | 25355 | 2805805 | 216278 | 1308521 | 219481 | 1786131 | 12456 | 0 | 0 | 23414 | 0 | 411 | 299800 | 0 | 2412917 | 0 | 491 | 360025 | 23164 | 37271 | 1402788 | 296905 | 1050877 | 13912 | 169002 | 0 | 42328 | 4805192 | 0 | 2912854 |  | Unknown
LRB6-inj1 | 0 | 50416 | 33221 | 7862 | 0 | 0 | 3844791 | 0 | 269118 | 0 | 529 | 0 | 0 | 20912 | 166676 | 678945 | 722657 | 491110 | 3338896 | 16246570 | 6487216 | 695913 | 34146175 | 3494909 | 28081546 | 3831279 | 33107061 | 258969 | 0 | 0 | 1091249 | 0 | 1090 | 13093 | 0 | 59812 | 0 | 641 | 23849 | 10461 | 149184 | 19557 | 1532 | 72284 | 15125 | 1187 | 0 | 9062 | 62503 | 0 | 92673 |  | Unknown
AB-230117-02 LRB 6 | 0 | 47992 | 23756 | 7840 | 10749 | 0 | 4041035 | 0 | 302395 | 0 | 1800 | 1129 | 0 | 20292 | 147615 | 575148 | 666458 | 526049 | 2312291 | 11270072 | 5772719 | 781754 | 36778822 | 3150299 | 23284126 | 3361080 | 29128442 | 288321 | 0 | 0 | 764375 | 0 | 619 | 13082 | 0 | 65001 | 1013 | 289 | 39044 | 0 | 108033 | 19685 | 3813 | 53686 | 10681 | 1146 | 0 | 8132 | 58746 | 0 | 80134 |  | Unknown
AB-230117-02 LRB 1 | 0 | 58960 | 47604 | 12235 | 1957729 | 0 | 4192864 | 0 | 400447 | 9154 | 0 | 97712 | 57814 | 103376 | 182974 | 816006 | 842732 | 881418 | 3663879 | 17592203 | 7132049 | 1052182 | 37221036 | 4031444 | 30802350 | 4205808 | 35806446 | 377315 | 0 | 0 | 1295992 | 0 | 0 | 77019 | 0 | 851531 | 0 | 0 | 189075 | 0 | 143615 | 740342 | 89295 | 567140 | 20595 | 116280 | 0 | 30481 | 1150536 | 0 | 907810 |  | Unknown
AB-230117-02 QCS 1 | 0 | 251510 | 624834 | 746610 | 74162754 | 0 | 3780004 | 0 | 391464 | 1793397 | 876621 | 4480499 | 3426289 | 3142146 | 167272 | 536181 | 680083 | 761275 | 3056412 | 14912965 | 6912932 | 1012352 | 34623193 | 3224640 | 26340525 | 2988900 | 27591477 | 368604 | 0 | 0 | 1279834 | 0 | 0 | 103721 | 435977 | 36327856 | 0 | 0 | 16714995 | 7299392 | 5181116 | 26847701 | 315261 | 27798052 | 2377420 | 94137 | 0 | 3424892 | 4705810 | 0 | 35946324 |  | Unknown
WO-23003-01-A-1 | 0 | 65796 | 107908 | 5466 | 299302 | 0 | 3536696 | 196 | 75724 | 0 | 21768 | 15634 | 6652 | 45987 | 74836 | 741764 | 1288195 | 336384 | 1862793 | 7333015 | 2617430 | 354799 | 32301641 | 3661954 | 18071944 | 3844692 | 25697042 | 64238 | 0 | 0 | 497588 | 0 | 0 | 17383 | 152786 | 234212 | 0 | 0 | 48368 | 14580 | 157601 | 142326 | 15416 | 115520 | 50915 | 16797 | 0 | 7144 | 269861 | 0 | 127646 |  | Unknown
WO-23003-02-A-1 | 0 | 57507 | 109341 | 3563 | 195902 | 0 | 3585473 | 0 | 44961 | 8119 | 17379 | 14261 | 7300 | 45543 | 71662 | 992753 | 1519457 | 259309 | 2189835 | 7792307 | 2457350 | 243464 | 31419146 | 3730910 | 19370109 | 4064659 | 26032384 | 39407 | 0 | 0 | 447622 | 0 | 0 | 16779 | 100099 | 218265 | 0 | 0 | 175529 | 0 | 170343 | 110571 | 9222 | 108590 | 32218 | 13210 | 0 | 4550 | 170801 | 0 | 101520 |  | Unknown
WO-23003-03-A-1 | 0 | 69589 | 121081 | 5736 | 158435 | 0 | 3133212 | 209 | 108334 | 0 | 25958 | 11824 | 9613 | 38303 | 93068 | 1188539 | 1262236 | 435784 | 2930056 | 11311334 | 3443525 | 452675 | 28375547 | 3803107 | 24292510 | 4041821 | 28062070 | 96340 | 0 | 0 | 662710 | 0 | 0 | 13724 | 247185 | 157865 | 0 | 0 | 41113 | 14961 | 146593 | 105724 | 7756 | 120842 | 24973 | 10802 | 0 | 6671 | 133604 | 0 | 79790 |  | Unknown
WO-23003-04-A-1 | 0 | 43812 | 75305 | 17617 | 128345 | 0 | 4254250 | 59 | 11107 | 0 | 0 | 7713 | 0 | 44631 | 35619 | 512831 | 1167231 | 170395 | 1635815 | 5611845 | 1372283 | 91096 | 34514886 | 3756604 | 16332245 | 3964193 | 26011891 | 8462 | 0 | 0 | 231614 | 0 | 0 | 16000 | 20365 | 141643 | 0 | 3254 | 45874 | 0 | 151759 | 81655 | 6751 | 74199 | 27864 | 8000 | 0 | 0 | 133785 | 0 | 115397 |  | Unknown
WO-23003-05-A-1 | 0 | 50987 | 89994 | 6603 | 117141 | 0 | 4486300 | 73 | 21630 | 0 | 0 | 9579 | 0 | 38253 | 39714 | 509584 | 1154501 | 207211 | 1528397 | 5613681 | 1572217 | 131664 | 33872567 | 3799640 | 15172803 | 3859187 | 24593314 | 16128 | 0 | 0 | 279961 | 0 | 0 | 15430 | 45680 | 184023 | 0 | 0 | 0 | 0 | 165355 | 73980 | 6302 | 62779 | 25266 | 6052 | 0 | 0 | 137509 | 0 | 81909 |  | Unknown
WO-23003-06-A-1 | 0 | 43741 | 62735 | 7431 | 96830 | 0 | 4336593 | 69 | 14109 | 0 | 0 | 0 | 0 | 169716 | 32115 | 378761 | 933750 | 120550 | 1284902 | 4746972 | 1375153 | 88385 | 32033462 | 3469485 | 13088314 | 3521650 | 22472696 | 10215 | 0 | 0 | 259336 | 0 | 0 | 18337 | 26178 | 123003 | 0 | 0 | 147410 | 0 | 165982 | 70451 | 4882 | 52937 | 27343 | 6165 | 0 | 0 | 127715 | 0 | 100027 |  | Unknown
WO-23003-07-A-1 | 0 | 48295 | 34614 | 2721 | 63683 | 0 | 2520477 | 0 | 354715 | 0 | 0 | 5652 | 5192 | 21800 | 140279 | 749409 | 338303 | 816905 | 3383417 | 16427420 | 5638969 | 882621 | 11240896 | 3790390 | 29457912 | 3844275 | 33917331 | 330783 | 0 | 0 | 1141879 | 0 | 0 | 12872 | 0 | 214903398 | 0 | 1357 | 80700 | 0 | 140611 | 267639 | 0 | 170804 | 18222 | 5815 | 0 | 6296 | 137247 | 0 | 141251 |  | Unknown
WO-23003-08-A-1 | 0 | 47927 | 33300 | 2770 | 64409 | 0 | 2428873 | 103 | 360537 | 0 | 0 | 3671 | 0 | 11942 | 157557 | 688493 | 334223 | 763057 | 3440710 | 16639864 | 6072207 | 881338 | 10905532 | 3705502 | 29295926 | 3787128 | 33862932 | 342805 | 0 | 0 | 1182503 | 0 | 0 | 12715 | 0 | 230428836 | 0 | 0 | 45691 | 0 | 141393 | 267119 | 0 | 164238 | 18160 | 4882 | 0 | 6986 | 126941 | 0 | 145666 |  | Unknown
WO-23003-09-A-1 | 0 | 49083 | 33527 | 2323 | 50138 | 0 | 2549553 | 0 | 303536 | 0 | 0 | 0 | 0 | 15050 | 94552 | 719347 | 328498 | 605447 | 3398277 | 16488995 | 4196726 | 893275 | 11603574 | 3764838 | 29795410 | 3893301 | 34190517 | 285021 | 0 | 0 | 1180678 | 0 | 0 | 13868 | 0 | 223715045 | 0 | 0 | 0 | 0 | 136364 | 248559 | 0 | 179090 | 17019 | 5458 | 0 | 7220 | 120479 | 0 | 113077 |  | Unknown
WO-23003-10-A-1 | 0 | 50471 | 34306 | 8061 | 41350 | 0 | 2340514 | 0 | 323447 | 0 | 0 | 3442 | 0 | 10838 | 168284 | 696669 | 345999 | 677764 | 3351815 | 16583246 | 6539975 | 953275 | 10743050 | 3753494 | 29682382 | 3876799 | 34149551 | 307870 | 0 | 0 | 1147085 | 0 | 1824 | 12702 | 0 | 233506924 | 0 | 0 | 45690 | 0 | 120048 | 268489 | 0 | 163514 | 14613 | 3917 | 0 | 6689 | 115593 | 0 | 121748 |  | Unknown
WO-23003-11-A-1 | 0 | 42616 | 32074 | 3365 | 25439 | 559 | 2369921 | 0 | 311200 | 2024 | 0 | 0 | 0 | 11571 | 164966 | 687477 | 324313 | 706630 | 3316637 | 16000175 | 6441159 | 919025 | 10981126 | 3563273 | 28043019 | 3669322 | 31943220 | 295229 | 0 | 0 | 1151204 | 0 | 668 | 13396 | 0 | 223283789 | 0 | 0 | 53477 | 0 | 132337 | 238939 | 0 | 141948 | 17057 | 3520 | 0 | 8289 | 100480 | 0 | 111030 |  | Unknown
WO-23003-12-A-1 | 0 | 42091 | 28967 | 3737 | 29679 | 0 | 2294500 | 0 | 301075 | 0 | 0 | 0 | 0 | 14806 | 150148 | 673725 | 305797 | 626626 | 3308669 | 15773920 | 5594863 | 857055 | 11063233 | 3653343 | 28505799 | 3817697 | 32764869 | 286308 | 0 | 0 | 1048456 | 0 | 0 | 13317 | 0 | 224464407 | 0 | 0 | 48064 | 0 | 133109 | 244732 | 0 | 158345 | 14890 | 1804 | 0 | 8666 | 89670 | 0 | 96990 |  | Unknown
WO-23003-13-A-1 | 0 | 43638 | 32734 | 7334 | 35447 | 1393 | 2163174 | 0 | 363035 | 5604 | 0 | 0 | 0 | 12109 | 139041 | 673442 | 329550 | 671594 | 3307942 | 16148977 | 6350247 | 952339 | 10872151 | 3625980 | 28549100 | 3731629 | 33171763 | 343592 | 0 | 0 | 1136431 | 0 | 539 | 15762 | 0 | 226866512 | 0 | 0 | 46799 | 0 | 140940 | 255481 | 0 | 158734 | 14500 | 2263 | 0 | 7965 | 90516 | 1641 | 98975 |  | Unknown
WO-23003-14-A-1 | 0 | 43518 | 67742 | 2645 | 43181 | 0 | 3611110 | 0 | 148587 | 0 | 9669455 | 4795 | 0 | 25927 | 85808 | 755620 | 1037845 | 328067 | 2563857 | 10024593 | 3410265 | 536991 | 28911726 | 3648749 | 22823195 | 4013959 | 29279951 | 133833 | 0 | 0 | 642898 | 0 | 543 | 10232 | 0 | 27441137 | 0 | 0 | 243519 | 0 | 148426 | 55545 | 0 | 77732 | 17219 | 4064 | 0 | 4945 | 72695 | 1202 | 47053 |  | Unknown
WO-23003-15-A-1 | 0 | 50488 | 77536 | 9035 | 51950 | 0 | 4284119 | 11 | 158463 | 0 | 7477856 | 0 | 4705 | 31336 | 51210 | 821682 | 1209287 | 260209 | 2694573 | 10253063 | 1935784 | 472421 | 35402869 | 3935916 | 23280241 | 4034252 | 28846629 | 145471 | 0 | 0 | 424411 | 0 | 1537 | 9214 | 0 | 1347931 | 0 | 1059 | 239356 | 7597 | 130395 | 52657 | 3613 | 76144 | 15357 | 5852 | 0 | 5100 | 79576 | 0 | 47214 |  | Unknown
WO-23003-16-A-1 | 0 | 49171 | 57691 | 4049 | 41181 | 866 | 4053882 | 0 | 149121 | 2814 | 10418045 | 0 | 4340 | 29156 | 95565 | 688850 | 963586 | 355383 | 2613035 | 10449870 | 3583419 | 582988 | 36763017 | 3664319 | 23060882 | 3988486 | 30758028 | 137413 | 0 | 0 | 667889 | 0 | 1090 | 9336 | 0 | 738697 | 0 | 0 | 237363 | 13172 | 137599 | 58558 | 0 | 71212 | 16602 | 2702 | 0 | 6390 | 76879 | 0 | 38695 |  | Unknown
WO-23003-17-A-1 | 0 | 44935 | 53768 | 3764 | 47168 | 0 | 4181060 | 0 | 140211 | 3227 | 7372536 | 4367 | 0 | 30895 | 85972 | 600102 | 985196 | 255377 | 2527701 | 9952404 | 3321449 | 557827 | 36557273 | 3707816 | 22483101 | 3934598 | 29735618 | 127206 | 1770 | 0 | 606825 | 0 | 876 | 15804 | 0 | 425209 | 0 | 0 | 349739 | 13324 | 140809 | 51665 | 2420 | 82250 | 15119 | 2822 | 0 | 4787 | 78512 | 0 | 55923 |  | Unknown
WO-23003-18-A-1 | 0 | 44412 | 50524 | 5463 | 66113 | 0 | 3963927 | 33 | 131267 | 1124 | 9216444 | 6421 | 4810 | 28392 | 79378 | 603041 | 903702 | 314019 | 2335472 | 9200001 | 3151576 | 491141 | 33091451 | 3640022 | 21654419 | 3888248 | 29123978 | 121050 | 0 | 0 | 593985 | 0 | 0 | 17236 | 0 | 311487 | 0 | 988 | 211979 | 14054 | 146903 | 55797 | 4165 | 91159 | 16662 | 3347 | 0 | 5694 | 80714 | 0 | 56478 |  | Unknown
WO-23003-19-A-1 | 0 | 43746 | 43498 | 8165 | 59804 | 0 | 3750251 | 44 | 152910 | 1452 | 9987341 | 4420 | 0 | 23465 | 68440 | 581157 | 881655 | 342972 | 2232321 | 8418990 | 2525868 | 515345 | 33150001 | 3153772 | 20665717 | 3207515 | 23736871 | 143465 | 0 | 0 | 398253 | 0 | 521 | 15694 | 0 | 263279 | 0 | 0 | 188438 | 0 | 145125 | 52226 | 2809 | 86068 | 11573 | 3874 | 0 | 5496 | 81834 | 0 | 57224 |  | Unknown
WO-23003-20-A-1 | 0 | 43971 | 0 | 8485 | 63665 | 0 | 4135379 | 0 | 340495 | 1482 | 0 | 700120 | 20829 | 17373617 | 88669 | 587467 | 788053 | 675868 | 2914086 | 13904021 | 3341284 | 867519 | 37950471 | 2694033 | 25587257 | 2077513 | 20617946 | 326795 | 0 | 0 | 717575 | 0 | 577 | 14368 | 0 | 261759 | 0 | 704 | 60458 | 0 | 133315 | 41672 | 49465 | 86517 | 13812 | 3077 | 0 | 9211 | 79742 | 0 | 46286 |  | Unknown
WO-23003-21-A-1 | 0 | 47950 | 25410 | 26859 | 1102457 | 0 | 4317833 | 16 | 297697 | 7756 | 0 | 717244 | 40440 | 16610191 | 161171 | 630119 | 702089 | 697557 | 3199409 | 15733599 | 6276252 | 848524 | 36473897 | 2657982 | 26868769 | 2063733 | 21316953 | 280253 | 0 | 0 | 1114067 | 0 | 1641 | 65712 | 0 | 564430 | 0 | 0 | 133024 | 32040 | 129744 | 240900 | 98327 | 275025 | 15171 | 38417 | 0 | 13118 | 1136749 | 0 | 628358 |  | Unknown
WO-23003-22-A-1 | 0 | 49355 | 0 | 3305 | 158329 | 0 | 4272559 | 0 | 352440 | 0 | 0 | 674958 | 26295 | 17180750 | 160768 | 608605 | 765125 | 542265 | 3311045 | 15852635 | 6347969 | 779924 | 39365969 | 2811605 | 27368674 | 2181366 | 22130353 | 330967 | 0 | 0 | 1142107 | 0 | 708 | 20458 | 0 | 159650 | 0 | 507 | 162961 | 0 | 139097 | 51909 | 48687 | 111781 | 13196 | 6978 | 0 | 6193 | 151903 | 0 | 123218 |  | Unknown
WO-23003-23-A-1 | 0 | 45767 | 0 | 3416 | 89441 | 0 | 4266222 | 0 | 297433 | 0 | 0 | 758815 | 24961 | 18257538 | 149797 | 606454 | 747767 | 771498 | 3325718 | 15908187 | 5864427 | 927045 | 38399615 | 2745882 | 27119147 | 2083832 | 21230023 | 283492 | 0 | 0 | 1078946 | 0 | 0 | 19889 | 0 | 137975 | 0 | 0 | 66073 | 16595 | 126890 | 38347 | 52148 | 94379 | 7793 | 5156 | 0 | 7196 | 107058 | 0 | 92865 |  | Unknown
WO-23003-24-A-1 | 0 | 49612 | 0 | 11700 | 138693 | 0 | 4339451 | 0 | 335674 | 3332 | 0 | 705692 | 25732 | 16942544 | 141269 | 618268 | 781383 | 614004 | 3195824 | 15654005 | 5411334 | 742559 | 39298517 | 2745079 | 27217484 | 2133264 | 21992191 | 324427 | 0 | 0 | 1056532 | 0 | 0 | 21819 | 0 | 178672 | 0 | 0 | 110676 | 23606 | 133624 | 64395 | 52225 | 138106 | 6586 | 2727 | 0 | 9794 | 110755 | 0 | 100418 |  | Unknown
WO-23003-25-A-1 | 0 | 48362 | 0 | 5958 | 57738 | 1436 | 4136842 | 0 | 341051 | 5694 | 0 | 747746 | 24402 | 17817531 | 138646 | 615971 | 739224 | 666989 | 3109176 | 15498868 | 5929044 | 839167 | 38685040 | 2633953 | 27087305 | 2030032 | 21199897 | 329385 | 0 | 0 | 1115481 | 0 | 1427 | 15838 | 0 | 112373 | 0 | 936 | 69816 | 14295 | 129043 | 33550 | 50724 | 90408 | 6455 | 3987 | 0 | 8445 | 91358 | 0 | 50022 |  | Unknown
WO-23003-26-A-1 | 0 | 47043 | 75395 | 12474 | 35016 | 0 | 4151047 | 11 | 77551 | 2696 | 0 | 24618 | 0 | 1302804 | 72995 | 628312 | 1074200 | 284220 | 1958657 | 7380959 | 2503721 | 281548 | 35019415 | 3291704 | 19606438 | 3615594 | 27032188 | 66085 | 0 | 0 | 428529 | 0 | 797 | 14839 | 0 | 86557 | 0 | 1728 | 28615 | 0 | 148468 | 37364 | 6487 | 49233 | 16982 | 2233 | 0 | 5535 | 88954 | 0 | 47911 |  | Unknown
WO-23003-27-A-1 | 0 | 45955 | 80330 | 2765 | 22161 | 1315 | 3696884 | 0 | 95593 | 0 | 7669 | 5138 | 0 | 63738 | 80555 | 702314 | 1031436 | 387147 | 2251039 | 8737588 | 3127946 | 439822 | 30702952 | 3268564 | 20932697 | 3558184 | 26742961 | 85744 | 0 | 0 | 567057 | 0 | 811 | 15020 | 0 | 71634 | 0 | 0 | 25997 | 0 | 153972 | 33304 | 0 | 60311 | 18927 | 2408 | 0 | 4991 | 74753 | 0 | 36075 |  | Unknown
WO-23003-28-A-1 | 0 | 47319 | 72939 | 7895 | 28283 | 831 | 4388725 | 0 | 105965 | 0 | 0 | 0 | 0 | 39357 | 72749 | 625978 | 1095247 | 253068 | 2098475 | 8486416 | 2979528 | 359506 | 35387265 | 3484410 | 20636117 | 3857353 | 28505227 | 96008 | 0 | 0 | 537621 | 0 | 536 | 14078 | 0 | 79545 | 0 | 0 | 0 | 0 | 167136 | 40031 | 0 | 52291 | 16838 | 2658 | 0 | 3941 | 77592 | 0 | 50555 |  | Unknown
WO-23003-29-A-1 | 0 | 48192 | 74390 | 7003 | 25405 | 0 | 4339688 | 0 | 93626 | 0 | 4268 | 3534 | 0 | 29226 | 70345 | 604662 | 1058647 | 369494 | 2128875 | 8489643 | 3070656 | 434821 | 37248270 | 3560874 | 21057908 | 3906672 | 29615862 | 85066 | 0 | 0 | 567838 | 0 | 842 | 15242 | 0 | 76769 | 0 | 1277 | 0 | 0 | 162793 | 31516 | 0 | 45629 | 19572 | 3470 | 0 | 0 | 82294 | 0 | 38834 |  | Unknown
WO-23003-30-A-1 | 0 | 46986 | 95805 | 6666 | 18446 | 0 | 3488537 | 0 | 78313 | 0 | 0 | 4006 | 0 | 27904 | 76623 | 718493 | 1351081 | 341668 | 1990198 | 7796759 | 2835254 | 345601 | 31784162 | 3347975 | 19285062 | 3740955 | 26277998 | 70563 | 0 | 0 | 501333 | 0 | 1650 | 13568 | 0 | 77662 | 0 | 2077 | 32004 | 0 | 155475 | 33648 | 0 | 43204 | 20094 | 2066 | 0 | 6217 | 58636 | 0 | 31747 |  | Unknown
WO-23003-31-A-1 | 0 | 41051 | 62020 | 8497 | 23951 | 0 | 4007131 | 0 | 64876 | 0 | 0 | 3812 | 0 | 27304 | 40293 | 499043 | 982381 | 303279 | 1609450 | 6478285 | 1478468 | 342226 | 34027840 | 3276460 | 17080853 | 3446208 | 26018884 | 55125 | 0 | 0 | 321171 | 0 | 0 | 11182 | 0 | 161815 | 0 | 1555 | 0 | 0 | 152400 | 54723 | 0 | 72921 | 18467 | 1555 | 0 | 0 | 72433 | 0 | 42700 |  | Unknown
WO-23003-32-A-1 | 0 | 43566 | 62290 | 9533 | 0 | 0 | 4149553 | 0 | 48926 | 4129 | 0 | 0 | 0 | 19716 | 35717 | 515555 | 924501 | 286371 | 1665290 | 6449193 | 1450216 | 279869 | 34280801 | 3299448 | 17533167 | 3615937 | 26624290 | 40994 | 0 | 0 | 308123 | 0 | 468 | 11859 | 0 | 67106 | 0 | 1529 | 0 | 0 | 159263 | 31107 | 0 | 47961 | 16157 | 1704 | 0 | 3997 | 77397 | 0 | 37681 |  | Unknown
WO-23003-33-A-1 | 0 | 50652 | 68198 | 6301 | 18624 | 0 | 3962194 | 0 | 101078 | 0 | 12845 | 0 | 0 | 28493 | 72932 | 581987 | 1038054 | 258808 | 1698006 | 6165550 | 2886147 | 402420 | 35053863 | 3385576 | 17870923 | 3737407 | 27006693 | 91680 | 0 | 0 | 298250 | 0 | 851 | 14531 | 0 | 118380 | 0 | 2116 | 0 | 15754 | 156572 | 49308 | 1149 | 45379 | 17591 | 1013 | 0 | 4910 | 70366 | 0 | 40658 |  | Unknown
WO-23003-34-A-1 | 0 | 55701 | 73332 | 2272 | 21697 | 976 | 4398414 | 0 | 88734 | 0 | 19423 | 0 | 3349 | 28882 | 75409 | 601659 | 1074942 | 364625 | 2128378 | 8202520 | 2851048 | 392547 | 35959334 | 3563049 | 19697415 | 3886488 | 28660834 | 80214 | 0 | 0 | 512736 | 0 | 343 | 15347 | 0 | 119557 | 0 | 1563 | 0 | 13650 | 158794 | 56686 | 2955 | 62662 | 16905 | 2022 | 0 | 5165 | 65209 | 782 | 44286 |  | Unknown
WO-23003-35-A-1 | 0 | 56116 | 64529 | 8460 | 35736 | 0 | 4094428 | 0 | 93227 | 0 | 15314 | 5628 | 2872 | 33852 | 71425 | 613725 | 1069268 | 327495 | 2086886 | 8144241 | 2912434 | 427565 | 34775503 | 3470249 | 19886142 | 3905300 | 27913901 | 82771 | 0 | 0 | 522159 | 0 | 1048 | 13443 | 0 | 123600 | 0 | 586 | 0 | 15398 | 139678 | 54257 | 1031 | 63969 | 18788 | 1952 | 0 | 3700 | 62761 | 0 | 47520 |  | Unknown
WO-23003-36-A-1 | 0 | 55867 | 76179 | 11548 | 53867 | 0 | 3952433 | 0 | 99796 | 3295 | 14467 | 4361 | 0 | 37813 | 77111 | 616474 | 979009 | 316645 | 2130115 | 8054042 | 2949618 | 405961 | 35097046 | 3437932 | 20278508 | 3891160 | 28228487 | 89538 | 0 | 0 | 531139 | 0 | 1108 | 15884 | 0 | 133835 | 0 | 2380 | 0 | 12572 | 154499 | 64466 | 2980 | 67750 | 14246 | 2018 | 0 | 4612 | 66504 | 0 | 52595 |  | Unknown
WO-23003-37-A-1 | 0 | 57811 | 75071 | 10092 | 16721 | 0 | 4063769 | 0 | 85279 | 0 | 14330 | 2194 | 3103 | 31464 | 69150 | 642254 | 1041707 | 310697 | 2033498 | 8257018 | 2863062 | 409691 | 35420039 | 3390190 | 20464234 | 3927748 | 28772017 | 74620 | 0 | 0 | 539847 | 0 | 789 | 10336 | 0 | 102752 | 0 | 1257 | 0 | 13340 | 156566 | 50557 | 0 | 50862 | 9302 | 1839 | 0 | 4378 | 65102 | 0 | 60302 |  | Unknown
WO-23003-38-A-1 | 0 | 54134 | 73991 | 11735 | 0 | 0 | 4555539 | 0 | 79803 | 0 | 20974 | 0 | 0 | 32675 | 75176 | 628734 | 1043859 | 391407 | 2210565 | 8389168 | 3071048 | 359730 | 36217472 | 3516625 | 20761933 | 3900660 | 28472937 | 69658 | 0 | 0 | 561070 | 0 | 923 | 14579 | 0 | 116001 | 0 | 3613 | 0 | 13364 | 150389 | 45000 | 943 | 60789 | 16146 | 1285 | 0 | 6822 | 61177 | 0 | 34317 |  | Unknown
WO-23003-39-A-1 | 0 | 39109 | 59357 | 7161 | 7334 | 0 | 4414300 | 0 | 41010 | 1169 | 0 | 0 | 0 | 16466 | 46005 | 454739 | 947926 | 240789 | 1597783 | 6058998 | 2020304 | 192611 | 34449016 | 3348800 | 16282203 | 3623185 | 26067846 | 34161 | 0 | 0 | 370251 | 0 | 699 | 11671 | 23418 | 54740 | 0 | 1469 | 36041 | 0 | 156860 | 25182 | 0 | 45159 | 15043 | 1228 | 0 | 6802 | 58453 | 1767 | 54284 |  | Unknown
WO-23003-40-A-1 | 0 | 46421 | 76078 | 7864 | 10320 | 1447 | 4242833 | 0 | 84392 | 4218 | 0 | 3079 | 1056 | 18256 | 74118 | 614001 | 1055597 | 345027 | 1961904 | 7642129 | 2778990 | 363812 | 35276162 | 3457853 | 19481074 | 3727989 | 27583554 | 74297 | 0 | 0 | 501418 | 0 | 0 | 9656 | 53973 | 92167 | 0 | 1352 | 33817 | 0 | 151640 | 44556 | 2396 | 47023 | 15521 | 1662 | 0 | 7943 | 58565 | 0 | 31095 |  | Unknown
WO-23003-41-A-1 | 0 | 44324 | 64123 | 4275 | 201107 | 0 | 4126951 | 43 | 68046 | 1368 | 1688 | 11675 | 4642 | 215971 | 62536 | 553192 | 986747 | 227095 | 1921842 | 7275314 | 2619758 | 199945 | 34946438 | 3265805 | 19253007 | 3646302 | 27213616 | 56781 | 0 | 0 | 380418 | 0 | 529 | 20667 | 46579 | 118567 | 0 | 1950 | 0 | 0 | 135571 | 84415 | 11095 | 78678 | 17042 | 11585 | 0 | 3398 | 181102 | 0 | 158606 |  | Unknown
WO-23003-42-A-1 | 0 | 45763 | 67744 | 6776 | 83621 | 0 | 4256173 | 0 | 52936 | 0 | 611 | 8999 | 1647 | 89479 | 56900 | 554002 | 1028146 | 230489 | 1909302 | 7420564 | 2180358 | 170913 | 35680754 | 3365511 | 18875295 | 3761177 | 27039117 | 44191 | 0 | 0 | 458326 | 0 | 1645 | 16244 | 32743 | 79834 | 0 | 3424 | 41632 | 0 | 154161 | 46633 | 4827 | 57263 | 17726 | 4123 | 0 | 5221 | 120821 | 0 | 66154 |  | Unknown
WO-23003-43-A-1 | 0 | 45564 | 86141 | 12144 | 46197 | 800 | 3676012 | 0 | 68772 | 0 | 0 | 3430 | 0 | 52639 | 63584 | 697835 | 1274802 | 251081 | 1777769 | 6778221 | 2502751 | 283729 | 31060291 | 3209005 | 17773173 | 3622146 | 23914508 | 60196 | 0 | 0 | 380678 | 0 | 622 | 15435 | 36081 | 63126 | 0 | 775 | 33571 | 0 | 154736 | 35726 | 3673 | 45018 | 17786 | 4738 | 0 | 2783 | 82558 | 0 | 43371 |  | Unknown
WO-23003-44-A-1 | 0 | 42943 | 73876 | 8519 | 35579 | 0 | 4308759 | 0 | 83828 | 3579 | 1186 | 10949 | 1664 | 43808 | 41401 | 616172 | 1072920 | 220833 | 1963128 | 7289874 | 1667776 | 297415 | 34820074 | 3462454 | 19266918 | 3978394 | 28279691 | 74464 | 0 | 0 | 369180 | 0 | 527 | 17021 | 44078 | 63836 | 0 | 1439 | 44996 | 0 | 163609 | 41841 | 7854 | 48880 | 18136 | 3386 | 0 | 5229 | 81794 | 0 | 47572 |  | Unknown
WO-23003-45-A-1 | 0 | 48009 | 71859 | 7964 | 27283 | 0 | 4114888 | 0 | 95556 | 0 | 1101 | 3653 | 1196 | 32178 | 45960 | 636114 | 1129784 | 249583 | 1938531 | 7549690 | 1785092 | 314632 | 34854193 | 3299750 | 19560443 | 3725257 | 27438800 | 83660 | 0 | 0 | 365309 | 0 | 2221 | 13036 | 57078 | 65361 | 0 | 1373 | 39919 | 0 | 157638 | 34562 | 2599 | 54190 | 15607 | 3008 | 0 | 6995 | 70265 | 0 | 51484 |  | Unknown
WO-23003-46-A-1 | 0 | 42412 | 67452 | 7471 | 21039 | 0 | 4092494 | 0 | 52334 | 0 | 699 | 10038 | 1949 | 30785 | 59513 | 512861 | 1035197 | 285681 | 1376805 | 5202503 | 2429265 | 269594 | 33672580 | 3362456 | 15617674 | 3768171 | 25851385 | 43989 | 2153 | 0 | 349917 | 0 | 577 | 13150 | 27788 | 76662 | 0 | 1268 | 79381 | 0 | 150209 | 39877 | 5044 | 37553 | 16410 | 2272 | 0 | 5573 | 65802 | 0 | 78259 |  | Unknown
WO-23003-47-A-1 | 0 | 43149 | 84231 | 8732 | 16995 | 0 | 3519822 | 0 | 12270 | 1935 | 0 | 8997 | 0 | 28747 | 39976 | 537631 | 1240942 | 145000 | 1394350 | 5164156 | 1573853 | 105018 | 27574645 | 3071519 | 13784367 | 3462368 | 21632529 | 9210 | 0 | 0 | 272427 | 0 | 836 | 9231 | 13594 | 44933 | 0 | 3441 | 0 | 0 | 151614 | 26142 | 7727 | 40113 | 18203 | 1935 | 0 | 5943 | 57540 | 0 | 57896 |  | Unknown
WO-23003-48-A-1 | 0 | 44490 | 68944 | 16408 | 14675 | 0 | 4057883 | 0 | 77761 | 1309 | 0 | 4413 | 0 | 23642 | 71458 | 552834 | 1067960 | 211177 | 1847476 | 7218451 | 2722086 | 279897 | 35375815 | 3331087 | 18954867 | 3623848 | 27459965 | 66177 | 0 | 0 | 508707 | 0 | 608 | 13702 | 92251 | 60539 | 0 | 2410 | 39450 | 0 | 155965 | 34194 | 2232 | 60940 | 17795 | 2121 | 0 | 4873 | 65038 | 0 | 32268 |  | Unknown
WO-23003-49-A-1 | 0 | 45440 | 69846 | 11764 | 12246 | 0 | 4365610 | 0 | 80239 | 0 | 604 | 4822 | 0 | 24863 | 61597 | 562730 | 1027755 | 262748 | 2007731 | 7804912 | 2492440 | 293788 | 35076715 | 3497404 | 19278484 | 3894264 | 27906762 | 69942 | 0 | 0 | 508233 | 0 | 853 | 18972 | 96428 | 58067 | 0 | 2735 | 42395 | 0 | 160287 | 37108 | 5652 | 92464 | 17459 | 1482 | 0 | 7923 | 62868 | 0 | 34595 |  | Unknown
WO-23003-50-A-1 | 0 | 41272 | 49482 | 8676 | 9397 | 0 | 3780408 | 0 | 12436 | 1648 | 717 | 6687 | 0 | 20453 | 30948 | 311454 | 903869 | 99474 | 976903 | 3450546 | 1119243 | 74449 | 31442073 | 2799112 | 11365017 | 3133749 | 21242150 | 10422 | 0 | 0 | 129790 | 0 | 358 | 12004 | 12271 | 42925 | 0 | 3012 | 123558 | 0 | 172006 | 28891 | 2526 | 0 | 16268 | 1490 | 0 | 4219 | 61006 | 0 | 71844 |  | Unknown
WO-23003-51-A-1 | 0 | 37022 | 70003 | 6243 | 10007 | 0 | 4260043 | 0 | 46503 | 0 | 2625 | 4114 | 1035 | 23900 | 33086 | 533822 | 999562 | 214580 | 1641215 | 5885332 | 1164454 | 198381 | 35034929 | 3251221 | 17883015 | 3685757 | 27104782 | 40187 | 0 | 0 | 241504 | 0 | 1126 | 12868 | 111026 | 53206 | 0 | 2855 | 0 | 0 | 144794 | 28919 | 3069 | 47213 | 16239 | 1311 | 0 | 7262 | 55768 | 1989 | 31753 |  | Unknown
WO-23003-52-A-1 | 0 | 48599 | 89469 | 7591 | 0 | 0 | 3840625 | 0 | 95963 | 2574 | 782 | 4014 | 0 | 22862 | 49516 | 826098 | 1304489 | 259214 | 1771089 | 6694704 | 2021364 | 315274 | 32019259 | 3373915 | 18689352 | 3788554 | 25366446 | 86481 | 0 | 0 | 283266 | 0 | 702 | 12335 | 197869 | 39182 | 0 | 2809 | 54895 | 0 | 153150 | 29211 | 3924 | 41853 | 16293 | 1583 | 0 | 5921 | 56386 | 0 | 30337 |  | Unknown
WO-23003-53-A-1 | 0 | 57614 | 82439 | 7102 | 12684 | 0 | 3960738 | 0 | 99614 | 0 | 2890 | 3396 | 1949 | 22735 | 84926 | 779655 | 1110517 | 265947 | 2580651 | 9918375 | 3350863 | 363825 | 36530098 | 3546546 | 23159173 | 4157094 | 30016345 | 91104 | 0 | 0 | 600538 | 0 | 1325 | 14017 | 211358 | 79214 | 1261 | 1669 | 53650 | 0 | 150534 | 37246 | 3932 | 58841 | 15715 | 1009 | 0 | 6091 | 54718 | 0 | 31713 |  | Unknown
WO-23003-54-A-1 | 0 | 56719 | 66687 | 7620 | 15714 | 0 | 4157633 | 0 | 58787 | 0 | 10113 | 9878 | 3252 | 32759 | 65779 | 478726 | 990564 | 203735 | 1747890 | 6889137 | 2518745 | 201832 | 34325965 | 3304252 | 17113731 | 3725653 | 26555731 | 48294 | 0 | 0 | 472734 | 0 | 297 | 14380 | 119279 | 106312 | 0 | 2580 | 44382 | 0 | 155584 | 50756 | 5470 | 52950 | 18487 | 1343 | 0 | 7942 | 58021 | 643 | 47420 |  | Unknown
WO-23003-55-A-1 | 0 | 66565 | 100783 | 8454 | 0 | 0 | 3459638 | 0 | 61589 | 0 | 9199 | 9855 | 3801 | 39249 | 75580 | 975446 | 1315699 | 239124 | 2842774 | 10997251 | 3250452 | 192177 | 33577065 | 3316461 | 24470312 | 3959883 | 29039247 | 53430 | 0 | 0 | 566303 | 0 | 965 | 8696 | 135014 | 220252 | 0 | 3898 | 0 | 9003 | 140997 | 60068 | 7075 | 82578 | 14156 | 1592 | 0 | 9512 | 45201 | 0 | 24966 |  | Unknown
WO-23003-56-A-1 | 0 | 54754 | 73528 | 8224 | 5865 | 0 | 4113574 | 0 | 63414 | 0 | 8047 | 8540 | 2823 | 34131 | 59720 | 572132 | 1147381 | 205437 | 1620637 | 6430738 | 2287421 | 192928 | 31568213 | 3257242 | 16203550 | 3684997 | 24908344 | 53390 | 0 | 0 | 423742 | 0 | 398 | 11529 | 128107 | 91395 | 1293 | 2102 | 147501 | 0 | 159067 | 54031 | 3330 | 54759 | 18767 | 1086 | 0 | 4709 | 52545 | 0 | 30819 |  | Unknown
LFB-H-3-Inj2 | 0 | 153612 | 336859 | 441324 | 64145861 | 0 | 2973007 | 0 | 277484 | 1392582 | 444397 | 2602453 | 2442836 | 2341203 | 142349 | 319073 | 429616 | 416290 | 2467529 | 12552801 | 5737650 | 445474 | 26770409 | 2583228 | 18087861 | 2470966 | 22319355 | 263890 | 0 | 0 | 1113353 | 0 | 524 | 2349773 | 250719 | 25916085 | 1556 | 16 | 12462757 | 5446794 | 96421 | 20118678 | 2563266 | 17472107 | 16286 | 2519809 | 0 | 2513773 | 26173116 | 0 | 21948918 |  | Unknown
WO-23003-14-A-1B | 0 | 49056 | 66242 | 2687 | 0 | 0 | 3545364 | 0 | 109291 | 0 | 4968388 | 8272 | 0 | 26025 | 78563 | 650627 | 1133240 | 191303 | 2367076 | 9367266 | 3228103 | 287674 | 31673484 | 3146910 | 20850445 | 3522701 | 26497504 | 98056 | 0 | 0 | 586884 | 0 | 809 | 9845 | 0 | 309995 | 0 | 2110 | 204581 | 9414 | 117321 | 35442 | 3137 | 64863 | 6165 | 863 | 0 | 8676 | 46595 | 0 | 30554 |  | Unknown
LRB-7-inj3 | 0 | 45259 | 64188 | 3199 | 0 | 0 | 4019182 | 0 | 61785 | 4347 | 640 | 1448 | 981 | 17671 | 66358 | 520203 | 898276 | 224057 | 1984692 | 7730305 | 2671578 | 195172 | 35105550 | 3364711 | 18860380 | 3754570 | 26903633 | 50728 | 0 | 0 | 489201 | 0 | 184 | 12504 | 43509 | 704461 | 0 | 1948 | 52028 | 0 | 153196 | 28650 | 1382 | 42219 | 16243 | 1202 | 0 | 5747 | 51230 | 0 | 64362 |  | Unknown
LRB-7-inj2 | 0 | 49481 | 25399 | 6637 | 0 | 0 | 3792511 | 0 | 294182 | 0 | 0 | 1636 | 1690 | 15279 | 151355 | 644639 | 600235 | 443639 | 3248161 | 15789821 | 5967825 | 638026 | 34617381 | 3312026 | 27653250 | 3378218 | 28828488 | 281986 | 0 | 0 | 1175727 | 0 | 354 | 10268 | 0 | 26116379 | 0 | 736 | 58705 | 0 | 147078 | 32025 | 3845 | 61756 | 5755 | 1022 | 0 | 8558 | 58492 | 1031 | 76009 |  | Unknown
WO-23003-60-A-1B | 0 | 39460 | 50916 | 6845 | 8313 | 0 | 2737052 | 0 | 11721 | 0 | 439 | 6500 | 0 | 14152 | 36022 | 355569 | 261560 | 111850 | 1565684 | 5657666 | 1522380 | 70603 | 8884746 | 3159818 | 15651949 | 3340961 | 24657370 | 9744 | 0 | 0 | 283109 | 0 | 939 | 13536 | 0 | 182259961 | 0 | 1852 | 103308 | 0 | 154496 | 309617 | 3985 | 135906 | 14498 | 2030 | 0 | 4940 | 58218 | 0 | 79961 |  | Unknown
WO-23003-59-A-1B | 0 | 45600 | 74101 | 7090 | 23647 | 0 | 3849593 | 0 | 77239 | 727 | 359 | 11099 | 2061 | 20188 | 59574 | 589685 | 578477 | 243432 | 2211458 | 8545054 | 2796228 | 334263 | 23508883 | 3402480 | 20857792 | 3757713 | 27380388 | 69366 | 1082 | 0 | 527354 | 0 | 481 | 14291 | 0 | 72956531 | 0 | 2623 | 147373 | 0 | 158265 | 94449 | 2401 | 80114 | 14989 | 1414 | 0 | 5093 | 62121 | 0 | 77224 |  | Unknown
WO-23003-58-A-1B | 0 | 52409 | 64405 | 11183 | 25051 | 0 | 3714525 | 0 | 99426 | 0 | 1963 | 7705 | 2369 | 24383 | 56758 | 503818 | 770937 | 318973 | 1823727 | 7288617 | 2343061 | 390620 | 31407933 | 3289026 | 17636975 | 3520584 | 24650996 | 89357 | 0 | 0 | 502433 | 0 | 465 | 10819 | 0 | 9415704 | 0 | 2068 | 182375 | 0 | 140119 | 132649 | 6437 | 56876 | 12543 | 1382 | 0 | 6510 | 59929 | 0 | 74948 |  | Unknown
WO-23003-26-A-1B | 0 | 48819 | 62804 | 6138 | 11128 | 0 | 2013496 | 0 | 47323 | 0 | 6118 | 5196 | 0 | 17993 | 45197 | 420948 | 572414 | 221861 | 1686468 | 6428241 | 2061350 | 200735 | 23123741 | 3339024 | 15830622 | 3128580 | 22726578 | 42089 | 0 | 0 | 386664 | 0 | 542 | 6958 | 0 | 54466 | 0 | 1116 | 18970 | 0 | 98818 | 34021 | 2722 | 34926 | 5226 | 2137 | 0 | 5806 | 45595 | 1022 | 32033 |  | Unknown
WO-23003-124-A-1 | 0 | 34719 | 23467 | 2656 | 24746 | 0 | 2632325 | 0 | 365396 | 0 | 0 | 2323 | 2910 | 6312 | 155082 | 608072 | 295106 | 782216 | 3413219 | 17596767 | 6545217 | 831639 | 11648092 | 4063813 | 31540850 | 4028406 | 37306194 | 344771 | 0 | 0 | 1147141 | 652 | 704 | 14827 | 0 | 242967382 | 0 | 313 | 46743 | 22561 | 173594 | 423630 | 1429 | 232683 | 16609 | 2749 | 0 | 8323 | 76541 | 0 | 97115 |  | Unknown
WO-23003-123-A-1 | 0 | 33859 | 26268 | 3872 | 30665 | 0 | 2702609 | 0 | 352448 | 0 | 0 | 6324 | 2455 | 8045 | 154965 | 627453 | 296199 | 787922 | 3447890 | 17733282 | 6652997 | 833105 | 11760718 | 3950551 | 31781255 | 3914703 | 36705423 | 341203 | 0 | 0 | 1191848 | 0 | 843 | 16371 | 0 | 237258758 | 0 | 109 | 49348 | 26372 | 160672 | 628802 | 3488 | 253914 | 14883 | 3018 | 0 | 7769 | 87789 | 2079 | 110449 |  | Unknown
WO-23003-122-A-1 | 0 | 34926 | 25576 | 5114 | 54560 | 0 | 2629092 | 97 | 331367 | 0 | 0 | 5816 | 2669 | 9783 | 150716 | 584792 | 301961 | 799705 | 3260898 | 16984551 | 6513227 | 832140 | 11794822 | 3872784 | 30676932 | 3824035 | 35822718 | 317514 | 0 | 0 | 1172222 | 0 | 498 | 15828 | 0 | 234371957 | 0 | 186 | 57374 | 29611 | 147921 | 530815 | 5325 | 253879 | 12853 | 4910 | 0 | 7891 | 97264 | 0 | 108730 |  | Unknown
WO-23003-121-A-1 | 0 | 37841 | 42574 | 7427 | 142716 | 0 | 2967155 | 81 | 180324 | 0 | 0 | 14785 | 4044 | 14266 | 91775 | 664466 | 398625 | 477397 | 2724072 | 11914775 | 3912310 | 528181 | 14897412 | 3875690 | 26466602 | 4032880 | 33793745 | 164724 | 0 | 0 | 701356 | 0 | 1066 | 16931 | 0 | 189787097 | 0 | 1519 | 224636 | 32985 | 151013 | 1300624 | 10081 | 1172758 | 17474 | 9578 | 0 | 7204 | 163404 | 0 | 122769 |  | Unknown
AB-230117-02 LFSM-H 6 | 0 | 116843 | 447190 | 348575 | 63586790 | 0 | 2300881 | 0 | 154827 | 695183 | 393240 | 3387062 | 2144937 | 2924427 | 85858 | 445574 | 355081 | 413834 | 2373337 | 10288661 | 3534758 | 480307 | 13096971 | 3308097 | 22728676 | 3280289 | 28306027 | 142739 | 0 | 2102 | 634016 | 0 | 935 | 1742611 | 124972 | 182605113 | 0 | 811 | 9490044 | 3090430 | 215638 | 25074130 | 1588159 | 20943236 | 61252 | 3224244 | 0 | 1361388 | 32788478 | 2029 | 25847585 |  | Unknown
AB-230117-02 CCC-H 7 | 0 | 297783 | 582678 | 666213 | 117420968 | 0 | 2713124 | 0 | 370853 | 2750476 | 1375403 | 5371324 | 4389541 | 4617408 | 150132 | 311929 | 396768 | 749296 | 2371745 | 13252045 | 6157485 | 842192 | 29369183 | 3158942 | 23190512 | 2628619 | 26586794 | 353402 | 0 | 0 | 1158056 | 0 | 505 | 4022216 | 599798 | 51655975 | 0 | 62 | 24261914 | 10865858 | 277502 | 43120094 | 4914207 | 40425414 | 105154 | 5673620 | 0 | 4563799 | 60471286 | 941 | 43702308 |  | Unknown
AB-230117-02 LFB-H 7 | 0 | 178898 | 385187 | 466921 | 66276186 | 0 | 3442716 | 0 | 373431 | 1379527 | 720895 | 3232929 | 2575639 | 2638110 | 143965 | 399806 | 491510 | 795193 | 2802953 | 15058626 | 6414093 | 854628 | 32994516 | 3450724 | 26765152 | 3134189 | 30536985 | 356495 | 0 | 0 | 1194388 | 837 | 591 | 2529429 | 304992 | 49165790 | 0 | 0 | 13936591 | 5800047 | 218334 | 25093229 | 2732141 | 23728345 | 53999 | 2991377 | 0 | 2379622 | 32310047 | 0 | 27928437 |  | Unknown
AB-230117-02 LD 6 | 0 | 36978 | 28093 | 9371 | 4134 | 0 | 2609383 | 0 | 353754 | 0 | 0 | 1943 | 0 | 6663 | 165073 | 598949 | 285219 | 796687 | 3246144 | 16951110 | 6648637 | 833625 | 11439467 | 3711137 | 29696384 | 3770630 | 34722557 | 340517 | 0 | 0 | 1197901 | 0 | 387 | 14745 | 0 | 232336403 | 848 | 250 | 43967 | 27665 | 132306 | 479246 | 1940 | 226398 | 16229 | 672 | 0 | 7973 | 57041 | 969 | 90152 |  | Unknown
LD6-inj1 | 847 | 38832 | 28390 | 9040 | 8416 | 0 | 2411193 | 0 | 367440 | 0 | 0 | 2537 | 2330 | 7979 | 164561 | 650784 | 294197 | 816555 | 3345917 | 17809044 | 6657933 | 850469 | 11938961 | 3844232 | 31273925 | 3868396 | 37076510 | 346314 | 0 | 0 | 1177883 | 0 | 670 | 17452 | 0 | 229640168 | 0 | 0 | 41437 | 28298 | 174758 | 479994 | 991 | 236508 | 8348 | 1608 | 0 | 9415 | 59994 | 708 | 80397 |  | Unknown
AB-230117-02 LRB 13 | 0 | 39223 | 30139 | 4025 | 1156 | 0 | 4975935 | 0 | 339699 | 1861 | 615 | 1563 | 1444 | 10457 | 154524 | 632889 | 829400 | 808357 | 3308258 | 17309829 | 6512473 | 817461 | 42552876 | 3892402 | 30286976 | 3878390 | 36636843 | 327227 | 0 | 0 | 1190410 | 0 | 451 | 14385 | 0 | 14418173 | 0 | 148 | 0 | 23313 | 168585 | 47596 | 1473 | 116935 | 10597 | 839 | 0 | 7862 | 56021 | 1148 | 34851 |  | Unknown
WO-23003-120-A-1 | 4227 | 33365 | 56665 | 2462 | 3391 | 0 | 3665778 | 0 | 150847 | 2596 | 0 | 3116 | 1298 | 0 | 97770 | 623493 | 554026 | 434712 | 2644378 | 11163978 | 3851120 | 490883 | 21611490 | 4158510 | 26321859 | 4314884 | 34541785 | 139198 | 0 | 0 | 680859 | 0 | 0 | 15909 | 0 | 140835849 | 0 | 1090 | 110370 | 24236 | 175452 | 579885 | 1974 | 604169 | 18459 | 1225 | 0 | 8136 | 59547 | 875 | 35916 |  | Unknown
WO-23003-119-A-1 | 0 | 35232 | 80836 | 4076 | 0 | 0 | 3815725 | 0 | 165575 | 0 | 4337 | 5482 | 2072 | 8747 | 91723 | 787661 | 660417 | 470814 | 2850510 | 11838608 | 3887057 | 522272 | 22798424 | 4144631 | 27604625 | 4353409 | 34960984 | 152503 | 0 | 0 | 705098 | 0 | 566 | 12328 | 0 | 121342878 | 0 | 1701 | 118487 | 19743 | 181045 | 445774 | 1925 | 543012 | 13728 | 1443 | 0 | 6455 | 53594 | 0 | 31511 |  | Unknown
WO-23003-118-A-1 | 0 | 36558 | 113908 | 5997 | 0 | 1879 | 2944013 | 0 | 174270 | 0 | 7131 | 2261 | 1421 | 8843 | 97907 | 636039 | 406147 | 455736 | 2874973 | 12188667 | 3944874 | 546398 | 13732261 | 3973276 | 27047440 | 4072631 | 34571765 | 157920 | 0 | 0 | 702467 | 0 | 336 | 14545 | 0 | 192387453 | 0 | 2143 | 237454 | 33780 | 179666 | 1489565 | 3363 | 1680672 | 16713 | 864 | 0 | 8077 | 48850 | 0 | 36295 |  | Unknown
WO-23003-117-A-1 | 0 | 34058 | 150084 | 3806 | 2854 | 0 | 3106564 | 0 | 149272 | 1501 | 0 | 3312 | 2373 | 6234 | 90027 | 584326 | 361377 | 434364 | 2638588 | 11170020 | 3649594 | 476715 | 15207107 | 4032701 | 25776455 | 4130614 | 34091479 | 137545 | 0 | 0 | 656751 | 0 | 255 | 14239 | 0 | 187287220 | 0 | 2432 | 181558 | 28654 | 188295 | 1284611 | 1840 | 1427064 | 8949 | 1193 | 0 | 6576 | 56056 | 2294 | 42028 |  | Unknown
WO-23003-116-A-1 | 0 | 33501 | 54993 | 3443 | 0 | 1536 | 3078558 | 0 | 154023 | 0 | 6289 | 3421 | 1195 | 10234 | 90926 | 546205 | 372813 | 415987 | 2579445 | 10893207 | 3620587 | 469612 | 15028164 | 3931505 | 24439633 | 3932868 | 33143414 | 140200 | 1047 | 0 | 654291 | 0 | 888 | 13192 | 0 | 182537121 | 0 | 1339 | 189508 | 30260 | 179065 | 1334154 | 1910 | 1442508 | 16605 | 832 | 0 | 6269 | 51820 | 0 | 38360 |  | Unknown
WO-23003-115-A-1 | 0 | 40632 | 565499 | 7063 | 0 | 0 | 3828036 | 0 | 99766 | 1472 | 0 | 3483 | 2258 | 11296 | 71519 | 711567 | 771773 | 354215 | 2318518 | 9357415 | 3003752 | 360502 | 25004515 | 4027729 | 23193579 | 4320818 | 30846423 | 88662 | 0 | 0 | 550322 | 0 | 699 | 11555 | 0 | 87544147 | 0 | 1338 | 137835 | 19960 | 170679 | 256899 | 1519 | 292195 | 16274 | 1236 | 0 | 8630 | 53099 | 0 | 35809 |  | Unknown
WO-23003-114-A-1 | 0 | 35881 | 194083 | 10685 | 0 | 817 | 4118739 | 0 | 86142 | 0 | 3135 | 3714 | 1135 | 7999 | 71092 | 686954 | 733569 | 347830 | 2270014 | 9114971 | 3000690 | 325864 | 27456944 | 4035660 | 22992151 | 4320860 | 31365991 | 74455 | 0 | 0 | 529838 | 0 | 750 | 12838 | 0 | 83814010 | 798 | 3098 | 134764 | 14932 | 183311 | 214016 | 3595 | 251009 | 10678 | 1379 | 0 | 5911 | 58376 | 0 | 32140 |  | Unknown
WO-23003-113-A-1 | 0 | 37314 | 122428 | 5545 | 0 | 0 | 3982422 | 0 | 79559 | 0 | 3122 | 8169 | 907 | 9220 | 74839 | 598344 | 668272 | 353182 | 2249019 | 9107594 | 2967227 | 317505 | 27159198 | 4075045 | 22864066 | 4327492 | 31830157 | 70741 | 0 | 0 | 516099 | 0 | 457 | 13358 | 0 | 85481441 | 772 | 2557 | 47735 | 10873 | 181245 | 209128 | 2319 | 261159 | 17243 | 1749 | 0 | 5160 | 55355 | 1275 | 42318 |  | Unknown
WO-23003-112-A-1 | 0 | 70441 | 250477 | 10900 | 4911 | 0 | 3895384 | 0 | 86547 | 0 | 6939 | 6500 | 0 | 11981 | 79594 | 631806 | 737993 | 364283 | 2237851 | 9200671 | 2902462 | 315301 | 27511818 | 3949798 | 22209911 | 4242494 | 30748381 | 74317 | 0 | 0 | 528078 | 0 | 525 | 9908 | 0 | 72351283 | 0 | 3826 | 160297 | 16035 | 184078 | 230012 | 3458 | 328561 | 18257 | 653 | 0 | 6592 | 52834 | 558 | 35163 |  | Unknown
WO-23003-111-A-1 | 0 | 38230 | 209062 | 6649 | 0 | 0 | 4168487 | 0 | 80890 | 0 | 6314 | 7501 | 0 | 18695 | 70307 | 567611 | 626483 | 328678 | 2170294 | 8671592 | 2876144 | 300553 | 25777832 | 4035736 | 21485097 | 4174581 | 31212515 | 67943 | 0 | 0 | 508806 | 0 | 761 | 12426 | 0 | 93632999 | 1372 | 2846 | 173499 | 18897 | 172782 | 325423 | 2667 | 475100 | 18296 | 509 | 0 | 6804 | 55001 | 353 | 32276 |  | Unknown
AB-230117-02 LRB 12 | 0 | 43278 | 28783 | 8315 | 7181 | 949 | 4847639 | 0 | 377613 | 0 | 868 | 3751 | 0 | 9251 | 164550 | 594945 | 788881 | 846668 | 3298787 | 17428215 | 7028879 | 823081 | 41966592 | 3878621 | 29379328 | 3753277 | 34883484 | 362305 | 0 | 0 | 1260290 | 0 | 1033 | 15503 | 0 | 209488 | 0 | 283 | 34050 | 25539 | 141780 | 29832 | 2265 | 100780 | 14247 | 2113 | 0 | 11646 | 56499 | 892 | 42838 |  | Unknown
LRB-12-inj1 | 0 | 38740 | 37547 | 10631 | 5742 | 973 | 4991287 | 0 | 385085 | 0 | 0 | 2692 | 2351 | 8989 | 167954 | 615136 | 828284 | 837112 | 3497300 | 18450364 | 6972498 | 906039 | 43582855 | 4160630 | 31493553 | 4072598 | 37572020 | 368063 | 0 | 0 | 1263690 | 0 | 1106 | 12646 | 0 | 5345236 | 1143 | 142 | 45094 | 26748 | 182322 | 36176 | 3370 | 125989 | 16040 | 1845 | 0 | 8378 | 64305 | 634 | 44187 |  | Unknown
WO-23003-110-A-1 | 0 | 43680 | 715911 | 5291 | 7681 | 1252 | 4099605 | 0 | 97891 | 0 | 8501 | 4984 | 0 | 9908 | 75773 | 637488 | 758319 | 358982 | 2298698 | 9223789 | 3090429 | 368565 | 25868952 | 4103483 | 23004724 | 4257264 | 31172627 | 85601 | 0 | 0 | 543250 | 0 | 1171 | 12800 | 0 | 84482675 | 0 | 3721 | 117786 | 26617 | 170608 | 290283 | 2375 | 435245 | 17323 | 1409 | 0 | 6373 | 53982 | 0 | 37531 |  | Unknown
WO-23003-109-A-1 | 0 | 40040 | 239885 | 5605 | 7635 | 0 | 4985822 | 0 | 93207 | 0 | 668 | 6382 | 0 | 13058 | 79160 | 643334 | 975287 | 370341 | 2499747 | 9934193 | 3179144 | 340439 | 37906152 | 4197930 | 24204764 | 4340752 | 33272051 | 80475 | 0 | 0 | 559441 | 0 | 604 | 17143 | 0 | 25547220 | 0 | 3055 | 18772 | 7946 | 156557 | 63519 | 3060 | 87969 | 16360 | 1136 | 0 | 4909 | 56399 | 959 | 32256 |  | Unknown
WO-23003-108-A-1 | 860 | 36287 | 232806 | 5270 | 11835 | 0 | 4542665 | 0 | 51190 | 0 | 0 | 2838 | 844 | 7886 | 52594 | 313278 | 784434 | 243079 | 1222131 | 5647825 | 2020116 | 203289 | 31394130 | 3482061 | 12475249 | 3400362 | 22459853 | 40644 | 0 | 0 | 371179 | 0 | 523 | 14322 | 0 | 18147756 | 1254 | 1897 | 20826 | 0 | 185834 | 40563 | 977 | 40973 | 19976 | 1497 | 0 | 4549 | 67471 | 703 | 38022 |  | Unknown
WO-23003-107-A-1 | 0 | 33231 | 120040 | 4816 | 10833 | 0 | 4689467 | 0 | 90926 | 0 | 0 | 7406 | 0 | 10396 | 66208 | 481019 | 896330 | 336779 | 1848875 | 7520697 | 2685215 | 316411 | 34808537 | 3900449 | 18844557 | 3963746 | 28511479 | 79995 | 0 | 0 | 474385 | 0 | 385 | 11824 | 0 | 24168389 | 0 | 1819 | 0 | 0 | 168219 | 48890 | 4632 | 77260 | 16205 | 1688 | 0 | 4305 | 56138 | 530 | 49791 |  | Unknown
WO-23003-106-A-1 | 0 | 32879 | 89069 | 5159 | 15182 | 0 | 4357248 | 0 | 71490 | 569 | 0 | 5922 | 372 | 8156 | 57502 | 334537 | 799107 | 280040 | 1306281 | 6007469 | 2339066 | 275072 | 30720915 | 3533365 | 13398584 | 3342309 | 22560018 | 61625 | 0 | 0 | 421655 | 0 | 489 | 17387 | 0 | 20692211 | 513 | 1854 | 0 | 0 | 181635 | 43934 | 2945 | 59856 | 17930 | 1591 | 0 | 3069 | 65717 | 0 | 40865 |  | Unknown
WO-23003-105-A-1 | 0 | 46142 | 280017 | 7320 | 20119 | 0 | 4462092 | 0 | 55168 | 1053 | 0 | 4046 | 1045 | 11564 | 60753 | 446202 | 930307 | 275513 | 1697435 | 6886482 | 2374984 | 222694 | 34342726 | 3838787 | 17543190 | 3832481 | 27261056 | 44312 | 0 | 0 | 432451 | 0 | 362 | 14072 | 0 | 21561961 | 0 | 2633 | 77629 | 0 | 180589 | 75802 | 2528 | 85172 | 17366 | 2376 | 0 | 4300 | 64144 | 0 | 47004 |  | Unknown
WO-23003-104-A-1 | 0 | 45733 | 213630 | 5806 | 30557 | 0 | 4696497 | 11 | 133262 | 0 | 0 | 5730 | 2265 | 13158 | 92673 | 800180 | 1098460 | 447858 | 2776678 | 11384213 | 3711085 | 449765 | 37655788 | 4359688 | 26677085 | 4577412 | 34076493 | 119110 | 0 | 0 | 697535 | 0 | 611 | 14555 | 0 | 23769062 | 1829 | 3487 | 45834 | 0 | 185682 | 82689 | 4371 | 115313 | 14412 | 3353 | 0 | 6267 | 67633 | 0 | 46289 |  | Unknown
WO-23003-103-A-1 | 0 | 35556 | 256387 | 3120 | 39351 | 0 | 4931396 | 0 | 77484 | 0 | 0 | 3433 | 0 | 9752 | 59322 | 352121 | 869022 | 308437 | 1552503 | 6640754 | 2466184 | 272827 | 35768513 | 3790642 | 15504438 | 3727708 | 26318356 | 62960 | 0 | 0 | 448814 | 0 | 421 | 17022 | 0 | 11714912 | 0 | 1921 | 84857 | 11615 | 193991 | 52557 | 3800 | 66237 | 19818 | 3512 | 0 | 1750 | 83871 | 0 | 58593 |  | Unknown
WO-23003-102-A-1 | 0 | 36606 | 197697 | 5621 | 61956 | 0 | 4239856 | 0 | 72988 | 2040 | 0 | 8518 | 0 | 11663 | 51890 | 223738 | 639834 | 274722 | 1056082 | 5394730 | 2251411 | 264739 | 27775775 | 3097088 | 10581129 | 2695615 | 18015535 | 60077 | 0 | 0 | 415107 | 0 | 178 | 14311 | 0 | 8827538 | 0 | 2023 | 22438 | 11244 | 190345 | 47656 | 4991 | 46784 | 20984 | 5433 | 0 | 2560 | 128142 | 0 | 58196 |  | Unknown
WO-23003-101-A-1 | 0 | 48825 | 455527 | 7983 | 163899 | 0 | 4895635 | 0 | 85771 | 0 | 0 | 10107 | 3462 | 18546 | 73288 | 563161 | 1062014 | 376523 | 2309340 | 9069846 | 3044917 | 331168 | 38768749 | 4222824 | 22389687 | 4216694 | 31868378 | 71332 | 0 | 0 | 558194 | 0 | 281 | 18870 | 0 | 5374162 | 741 | 2114 | 33927 | 13383 | 171646 | 81852 | 7653 | 86243 | 19448 | 11464 | 0 | 4379 | 177624 | 0 | 106679 |  | Unknown
AB-230117-02 LFSM-H 5 | 0 | 91830 | 944386 | 685679 | 76529547 | 1662 | 3418497 | 0 | 102282 | 826078 | 434520 | 3679856 | 2569068 | 3188418 | 101507 | 551719 | 745732 | 455406 | 2828511 | 12308621 | 4003771 | 386917 | 28982045 | 3485923 | 24178040 | 3427596 | 26412031 | 89635 | 0 | 0 | 727272 | 0 | 913 | 2697595 | 89176 | 31734667 | 0 | 1082 | 11313624 | 3510215 | 198344 | 22793915 | 3349018 | 21979592 | 48432 | 3459348 | 0 | 1923500 | 32850160 | 0 | 24640187 |  | Unknown
AB-230117-02 CCC-H 6 | 0 | 281850 | 562778 | 654485 | 113141419 | 0 | 2793348 | 0 | 395495 | 2745124 | 1414539 | 5030896 | 4276671 | 4300009 | 152035 | 306839 | 350063 | 810488 | 2380427 | 12735101 | 6477511 | 865033 | 27124276 | 3102225 | 22052524 | 2540641 | 24084407 | 379586 | 0 | 2242 | 1219465 | 0 | 442 | 3982013 | 636652 | 45429877 | 0 | 125 | 22407500 | 10915888 | 233820 | 37961336 | 4726407 | 36661472 | 78086 | 5307399 | 0 | 4498921 | 55792339 | 0 | 39491287 |  | Unknown
AB-230117-02 LFB-H 6 | 0 | 187695 | 371368 | 459032 | 68015432 | 0 | 3536832 | 0 | 380146 | 1429983 | 714174 | 3140950 | 2475938 | 2613609 | 162044 | 409177 | 503164 | 822027 | 2775932 | 15094220 | 6712510 | 851740 | 33569510 | 3413768 | 25321860 | 3062121 | 28650436 | 363560 | 0 | 0 | 1249401 | 1903 | 443 | 2563681 | 313365 | 28981468 | 0 | 567 | 13619043 | 5845666 | 195156 | 23399281 | 2934172 | 21953767 | 48517 | 2943589 | 0 | 2483996 | 30205275 | 1357 | 25864201 |  | Unknown
AB-230117-02 LD 5 | 0 | 41217 | 443693 | 6992 | 0 | 0 | 4440593 | 0 | 78523 | 0 | 622 | 5586 | 1135 | 11568 | 76315 | 631259 | 946446 | 333710 | 2467575 | 10065037 | 3003137 | 299169 | 34578186 | 3844573 | 23295803 | 3998089 | 30311159 | 69999 | 0 | 0 | 543923 | 0 | 205 | 11594 | 0 | 14938611 | 0 | 2488 | 36249 | 11088 | 126724 | 48313 | 3249 | 92906 | 13116 | 1313 | 0 | 6943 | 50838 | 0 | 32846 |  | Unknown
LD-5-inj1 | 0 | 45285 | 451029 | 8268 | 5574 | 1799 | 4809439 | 0 | 94342 | 0 | 0 | 2824 | 1035 | 12246 | 76969 | 631388 | 1004335 | 380794 | 2654414 | 10624228 | 3291411 | 345139 | 38110657 | 4207699 | 25128152 | 4291168 | 32939070 | 82474 | 0 | 0 | 584499 | 0 | 387 | 11042 | 0 | 15494357 | 0 | 2393 | 24924 | 12665 | 169770 | 58472 | 1657 | 101600 | 16301 | 1159 | 0 | 7123 | 56945 | 0 | 40642 |  | Unknown
AB-230117-02 LRB 11 | 0 | 39580 | 28935 | 5754 | 0 | 0 | 5075547 | 9 | 382255 | 2634 | 0 | 1216 | 996 | 9996 | 156160 | 611546 | 844320 | 852826 | 3472176 | 17655766 | 6981552 | 840649 | 43350177 | 4214839 | 30665008 | 4082172 | 35547285 | 366041 | 0 | 0 | 1260144 | 0 | 853 | 15474 | 0 | 59282 | 0 | 246 | 0 | 25286 | 168981 | 28598 | 1503 | 107613 | 16949 | 892 | 0 | 8715 | 57099 | 1341 | 37846 |  | Unknown
WO-23003-100-A-1 | 0 | 34285 | 68719 | 2342 | 0 | 0 | 4918862 | 0 | 86077 | 1043 | 2126 | 5398 | 475 | 13296 | 70157 | 588867 | 1192737 | 344729 | 2109992 | 8525143 | 2963581 | 341636 | 39237589 | 4295001 | 21267105 | 4328504 | 31068711 | 75533 | 0 | 0 | 539778 | 0 | 299 | 15997 | 171522 | 45966 | 0 | 2465 | 28073 | 0 | 183621 | 31065 | 1032 | 40050 | 19725 | 801 | 0 | 5706 | 53858 | 0 | 36250 |  | Unknown
WO-23003-99-A-1 | 0 | 36064 | 71744 | 4917 | 0 | 0 | 4845368 | 0 | 52013 | 0 | 0 | 4464 | 1310 | 9538 | 61737 | 582119 | 1171392 | 296170 | 1936346 | 7270327 | 2419840 | 207284 | 38353517 | 4116694 | 19214059 | 4178947 | 28842466 | 41441 | 0 | 0 | 435231 | 0 | 432 | 12334 | 102641 | 37652 | 511 | 2894 | 0 | 0 | 175311 | 28743 | 645 | 39903 | 17864 | 835 | 0 | 6974 | 57370 | 0 | 42938 |  | Unknown
WO-23003-98-A-1 | 1496 | 37144 | 80204 | 7376 | 0 | 0 | 4980257 | 0 | 105663 | 0 | 0 | 6394 | 1205 | 13269 | 78960 | 648369 | 1183660 | 373117 | 2241053 | 9166810 | 3129094 | 364852 | 40201443 | 4346142 | 22766224 | 4420815 | 31947734 | 90738 | 0 | 0 | 576514 | 0 | 528 | 13129 | 202295 | 44451 | 0 | 4278 | 97968 | 0 | 177790 | 35644 | 3469 | 50887 | 15463 | 1483 | 0 | 5425 | 54845 | 0 | 28652 |  | Unknown
WO-23003-97-A-1 | 0 | 46464 | 106242 | 8331 | 0 | 0 | 4151628 | 0 | 99665 | 1590 | 0 | 7861 | 1018 | 15028 | 79190 | 848698 | 1477455 | 388074 | 2284322 | 8722568 | 3111777 | 341121 | 35663824 | 4266521 | 21661622 | 4408049 | 28777567 | 85366 | 0 | 0 | 528734 | 0 | 1155 | 11628 | 180250 | 102495 | 0 | 1201 | 38245 | 0 | 179280 | 34869 | 1302 | 50608 | 18943 | 743 | 0 | 7139 | 48575 | 0 | 30897 |  | Unknown
WO-23003-96-A-1 | 0 | 35884 | 89358 | 10665 | 0 | 0 | 4768265 | 0 | 104361 | 0 | 0 | 5795 | 798 | 11436 | 74530 | 720088 | 1309270 | 369718 | 2335994 | 9139782 | 3137650 | 357902 | 38779746 | 4329298 | 22640971 | 4407111 | 30472169 | 86582 | 0 | 0 | 570404 | 0 | 442 | 11591 | 189649 | 34754 | 0 | 4211 | 0 | 0 | 187492 | 27887 | 2250 | 38915 | 17942 | 1247 | 0 | 6496 | 59496 | 0 | 31061 |  | Unknown
WO-23003-95-A-1 | 0 | 40749 | 74634 | 6125 | 4408 | 0 | 5066551 | 0 | 102039 | 0 | 0 | 6433 | 0 | 11058 | 81800 | 675641 | 1242406 | 414155 | 2438136 | 9559093 | 3290070 | 363695 | 41590416 | 4582143 | 23651394 | 4621082 | 33004661 | 87444 | 0 | 0 | 591866 | 0 | 783 | 12604 | 191933 | 42827 | 0 | 3180 | 0 | 0 | 183387 | 27130 | 2327 | 46472 | 19709 | 1789 | 0 | 5387 | 58706 | 947 | 35657 |  | Unknown
WO-23003-94-A-1 | 0 | 43345 | 80482 | 13923 | 0 | 0 | 4827694 | 0 | 106735 | 0 | 0 | 6371 | 0 | 12891 | 85205 | 678765 | 1228343 | 400040 | 2432299 | 9572668 | 3204689 | 363698 | 40658864 | 4502627 | 23272982 | 4484651 | 32934546 | 90503 | 0 | 0 | 587438 | 0 | 943 | 15081 | 198645 | 64108 | 0 | 3087 | 0 | 0 | 193293 | 39343 | 3144 | 49186 | 18426 | 1316 | 0 | 7048 | 56215 | 0 | 35806 |  | Unknown
WO-23003-93-A-1 | 0 | 42493 | 88702 | 7486 | 0 | 0 | 4175637 | 0 | 73928 | 0 | 3847 | 3285 | 0 | 12455 | 70402 | 753665 | 1373606 | 337582 | 2216247 | 8701125 | 2874110 | 290619 | 35550674 | 4301119 | 21726871 | 4303332 | 29449513 | 62051 | 0 | 0 | 531128 | 0 | 952 | 12714 | 129541 | 59010 | 0 | 4083 | 19470 | 0 | 166811 | 40509 | 1810 | 46660 | 16372 | 1014 | 0 | 5669 | 50074 | 0 | 32908 |  | Unknown
WO-23003-92-A-1 | 0 | 43261 | 124911 | 4837 | 0 | 0 | 3595092 | 0 | 128538 | 0 | 0 | 5439 | 1286 | 15492 | 97673 | 1345341 | 1560244 | 415217 | 2653892 | 10153929 | 3722523 | 453875 | 33403984 | 4198498 | 21748898 | 4256224 | 26672539 | 111912 | 0 | 0 | 631015 | 0 | 967 | 12315 | 238715 | 30828 | 0 | 3641 | 0 | 0 | 176967 | 30731 | 4085 | 50573 | 18497 | 959 | 0 | 7252 | 49122 | 0 | 22674 |  | Unknown
WO-23003-91-A-1 | 0 | 43025 | 89302 | 2803 | 5926 | 0 | 4271476 | 0 | 59936 | 0 | 0 | 7516 | 877 | 12267 | 62875 | 645389 | 1288519 | 289060 | 2073458 | 7833762 | 2551234 | 230993 | 34928638 | 4071211 | 20445424 | 4097531 | 28354687 | 48649 | 0 | 0 | 458185 | 0 | 533 | 13800 | 103651 | 72298 | 0 | 4485 | 0 | 0 | 159478 | 29295 | 1478 | 54443 | 16935 | 1012 | 0 | 6350 | 49190 | 824 | 28777 |  | Unknown
AB-230117-02 LRB 10 | 0 | 42578 | 31864 | 2894 | 6406 | 0 | 4680891 | 0 | 408529 | 510 | 2535 | 2529 | 1752 | 13682 | 164204 | 589105 | 782192 | 821316 | 3334539 | 17091618 | 7168889 | 900066 | 40242698 | 3990148 | 28633631 | 3766065 | 33684803 | 393815 | 0 | 0 | 1316014 | 0 | 1073 | 13436 | 0 | 66161 | 0 | 253 | 27407 | 29237 | 141648 | 26538 | 1637 | 103485 | 15035 | 2062 | 0 | 10445 | 55137 | 0 | 34718 |  | Unknown
LRB-10-inj1 | 0 | 40587 | 31551 | 1856 | 7023 | 0 | 5022080 | 0 | 407220 | 0 | 1116 | 2831 | 1579 | 10750 | 157638 | 608177 | 961506 | 854107 | 3525650 | 17720729 | 7266040 | 904756 | 43561553 | 4294162 | 30607554 | 4033211 | 35895538 | 390942 | 0 | 0 | 1318397 | 0 | 1255 | 15687 | 0 | 63479 | 1474 | 335 | 26599 | 27529 | 177426 | 27848 | 2212 | 102101 | 17952 | 1667 | 0 | 8722 | 59876 | 0 | 35599 |  | Unknown
WO-23003-90-A-1 | 0 | 45562 | 92628 | 5498 | 7545 | 639 | 3948810 | 0 | 82212 | 0 | 2574 | 7608 | 1461 | 13421 | 65637 | 734020 | 1371313 | 309238 | 2158117 | 8119001 | 2768161 | 274352 | 33684003 | 4120983 | 20567756 | 4188114 | 28199051 | 69073 | 0 | 0 | 489232 | 0 | 777 | 12909 | 139472 | 63734 | 0 | 4223 | 22583 | 0 | 159001 | 29679 | 1620 | 51089 | 17386 | 1384 | 0 | 6129 | 50952 | 0 | 28581 |  | Unknown
WO-23003-89-A-1 | 0 | 38279 | 79213 | 6754 | 9134 | 0 | 4735152 | 0 | 95993 | 0 | 366 | 6199 | 997 | 11757 | 71005 | 655632 | 1145179 | 350298 | 2267966 | 8629936 | 2981804 | 347262 | 36900501 | 4205412 | 21841201 | 4240476 | 29903150 | 82000 | 0 | 0 | 557271 | 0 | 1240 | 11416 | 188674 | 37302 | 0 | 2515 | 0 | 0 | 168105 | 27389 | 3025 | 47765 | 13276 | 1868 | 0 | 7912 | 58170 | 2823 | 44396 |  | Unknown
WO-23003-88-A-1 | 0 | 38829 | 93066 | 6348 | 9758 | 0 | 4439384 | 0 | 111624 | 0 | 0 | 8092 | 0 | 12012 | 78322 | 692529 | 1217157 | 390901 | 2367506 | 9157525 | 3178984 | 402264 | 36457724 | 4151183 | 21706626 | 4184832 | 29585549 | 98078 | 0 | 0 | 573552 | 0 | 1446 | 11422 | 207475 | 42418 | 0 | 2416 | 23064 | 0 | 171792 | 34064 | 2155 | 47793 | 18540 | 1670 | 0 | 5986 | 55510 | 0 | 37627 |  | Unknown
WO-23003-87-A-1 | 0 | 38982 | 81028 | 5381 | 13552 | 0 | 4182229 | 0 | 36844 | 0 | 908 | 3888 | 723 | 13506 | 47644 | 525993 | 1137560 | 220206 | 1712563 | 6429706 | 2003607 | 161238 | 33573338 | 3842692 | 16981543 | 3880549 | 25806511 | 30259 | 0 | 0 | 359063 | 0 | 605 | 12969 | 68038 | 38381 | 0 | 3327 | 32970 | 0 | 171086 | 32815 | 3037 | 43472 | 21940 | 887 | 0 | 3761 | 54945 | 0 | 32862 |  | Unknown
WO-23003-86-A-1 | 0 | 36997 | 85805 | 6907 | 19279 | 0 | 4297233 | 0 | 57143 | 2350 | 2282 | 7341 | 857 | 11112 | 59857 | 592529 | 1147029 | 285832 | 2010931 | 7751944 | 2547631 | 218688 | 34794228 | 4089669 | 19859541 | 4037136 | 28078988 | 45454 | 0 | 0 | 431507 | 0 | 542 | 12506 | 109116 | 44493 | 0 | 3708 | 26148 | 0 | 177086 | 43880 | 2822 | 49940 | 20467 | 1523 | 0 | 5194 | 62960 | 0 | 36296 |  | Unknown
WO-23003-85-A-1 | 0 | 35400 | 107009 | 5109 | 24433 | 0 | 3786609 | 0 | 110322 | 0 | 0 | 7072 | 879 | 13240 | 76890 | 867893 | 1368510 | 375658 | 2413172 | 9236864 | 3211135 | 383660 | 32886770 | 4063122 | 22201556 | 4087810 | 28031744 | 95808 | 0 | 0 | 567741 | 0 | 728 | 12402 | 222215 | 54985 | 0 | 3042 | 0 | 0 | 168629 | 39050 | 3061 | 51360 | 18493 | 1941 | 0 | 8424 | 60896 | 0 | 33879 |  | Unknown
WO-23003-84-A-1 | 0 | 39589 | 110374 | 10330 | 28875 | 0 | 3557205 | 10 | 111901 | 0 | 0 | 8774 | 1460 | 14276 | 77508 | 965799 | 1367011 | 394791 | 2350619 | 9039349 | 3076516 | 367253 | 31050434 | 3961576 | 21416800 | 4017463 | 26361109 | 97362 | 0 | 0 | 559854 | 0 | 996 | 15772 | 218437 | 37155 | 1331 | 2157 | 63947 | 0 | 154301 | 37946 | 2742 | 51385 | 16948 | 3031 | 0 | 6114 | 61943 | 0 | 35713 |  | Unknown
WO-23003-83-A-1 | 0 | 42154 | 78304 | 6997 | 47761 | 0 | 4306540 | 0 | 101614 | 0 | 1250 | 9175 | 1667 | 13920 | 71904 | 646235 | 1134941 | 355826 | 2218956 | 8502791 | 2958244 | 349169 | 35975599 | 4163662 | 20622987 | 4111584 | 29319967 | 88731 | 0 | 0 | 550934 | 0 | 729 | 13990 | 152419 | 72851 | 0 | 3392 | 0 | 0 | 152148 | 44847 | 7460 | 54402 | 10742 | 4131 | 0 | 5542 | 76961 | 0 | 43246 |  | Unknown
WO-23003-82-A-1 | 0 | 36669 | 88482 | 8673 | 83049 | 0 | 2760858 | 0 | 64066 | 4693 | 1896 | 5763 | 2257 | 12964 | 44051 | 395691 | 716333 | 200694 | 1288854 | 5089784 | 1697629 | 203267 | 21307863 | 2508618 | 12461196 | 2465134 | 17761291 | 54739 | 0 | 0 | 314867 | 0 | 715 | 11261 | 183129 | 67303 | 0 | 3263 | 0 | 0 | 106950 | 50199 | 4814 | 61051 | 14145 | 7615 | 0 | 7228 | 87487 | 0 | 81124 |  | Unknown
WO-23003-81-A-1 | 0 | 40275 | 73708 | 6795 | 186171 | 0 | 4580446 | 0 | 108350 | 0 | 397 | 17604 | 4134 | 18571 | 79644 | 645820 | 1086927 | 360637 | 2247196 | 8751644 | 3165164 | 379155 | 37402474 | 4128737 | 21767220 | 4168274 | 29882331 | 93811 | 0 | 0 | 559477 | 0 | 968 | 17215 | 165211 | 107977 | 0 | 2923 | 0 | 0 | 147214 | 72777 | 13146 | 75513 | 21049 | 10288 | 0 | 8551 | 163694 | 0 | 95327 |  | Unknown
AB-230117-02 LFSM-H 4 | 0 | 75644 | 517165 | 696449 | 77253824 | 0 | 2984644 | 0 | 101765 | 701671 | 362620 | 3491492 | 2048727 | 3134915 | 69799 | 445395 | 615184 | 355626 | 2072653 | 8030274 | 3032232 | 346767 | 26066773 | 3306119 | 18871381 | 3159809 | 23253872 | 87624 | 0 | 0 | 549664 | 0 | 536 | 2536250 | 90678 | 27264006 | 0 | 2408 | 8042465 | 2785199 | 214580 | 21990528 | 3441522 | 18596951 | 62122 | 3574153 | 0 | 1257305 | 33567077 | 0 | 23155319 |  | Unknown
AB-230117-02 CCC-H 5 | 0 | 307698 | 613751 | 693538 | 121297746 | 0 | 2810612 | 0 | 459649 | 3142889 | 1511279 | 5522622 | 4467624 | 4792732 | 156257 | 341482 | 360765 | 857556 | 2362328 | 12581303 | 6803754 | 913930 | 26265725 | 3326480 | 21883588 | 2634839 | 23985052 | 442018 | 0 | 0 | 1310115 | 2635 | 1057 | 4099183 | 753193 | 45782815 | 0 | 240 | 22937199 | 11996790 | 304999 | 39602258 | 4809239 | 38569334 | 147176 | 6003753 | 0 | 4867710 | 58760997 | 671 | 38797049 |  | Unknown
AB-230117-02 LFB-H 5 | 0 | 188235 | 401534 | 497652 | 70350332 | 0 | 3540640 | 0 | 416079 | 1523645 | 759132 | 3285952 | 2613512 | 2737006 | 157270 | 397222 | 509752 | 830324 | 2845482 | 14855764 | 6832898 | 921954 | 32549183 | 3636391 | 25545598 | 3116264 | 28676707 | 399255 | 0 | 0 | 1328738 | 2247 | 986 | 2630197 | 344044 | 28781019 | 1013 | 216 | 13759351 | 6077429 | 193120 | 24041448 | 3117654 | 22857644 | 57124 | 3311879 | 0 | 2608302 | 31969395 | 0 | 25155507 |  | Unknown
AB-230117-02 LD 4 | 0 | 38291 | 74865 | 8665 | 3860 | 0 | 3324112 | 0 | 23978 | 2523 | 850 | 7807 | 1143 | 15194 | 39773 | 496866 | 919464 | 196260 | 1650355 | 6058693 | 1760888 | 119907 | 29138161 | 3441877 | 15897027 | 3415913 | 23339370 | 20917 | 0 | 0 | 312969 | 0 | 687 | 11237 | 0 | 555027 | 0 | 1755 | 34412 | 0 | 117461 | 23791 | 3700 | 51995 | 14236 | 940 | 0 | 8810 | 44239 | 0 | 75607 |  | Unknown
LD4-inj1 | 0 | 36635 | 75221 | 8325 | 0 | 0 | 3259106 | 0 | 45206 | 4075 | 0 | 6757 | 1372 | 13468 | 51515 | 545968 | 985466 | 246215 | 1778462 | 6602078 | 2132653 | 175981 | 31087726 | 3496045 | 16990993 | 3622521 | 24656530 | 37954 | 0 | 0 | 380776 | 0 | 525 | 15239 | 0 | 557700 | 0 | 736 | 40485 | 0 | 147048 | 28322 | 1769 | 46727 | 15425 | 996 | 0 | 4340 | 47646 | 0 | 64070 |  | Unknown
AB-230117-02 LRB 9 | 0 | 43806 | 29790 | 7965 | 0 | 0 | 4743394 | 0 | 403384 | 0 | 743 | 2590 | 273 | 12941 | 162256 | 635067 | 915237 | 837006 | 3396911 | 17294441 | 7073193 | 831223 | 41477591 | 4294577 | 30238560 | 4029941 | 34633235 | 386958 | 0 | 0 | 1298730 | 0 | 475 | 14251 | 0 | 66408 | 0 | 522 | 38512 | 26942 | 158299 | 21751 | 1762 | 106449 | 16536 | 1060 | 0 | 9778 | 57279 | 0 | 33980 |  | Unknown
WO-23003-80-A-1 | 0 | 38102 | 83962 | 6354 | 0 | 0 | 4162855 | 0 | 54489 | 0 | 2790 | 3643 | 1152 | 10896 | 53066 | 565658 | 1227347 | 273658 | 1872359 | 7082416 | 2327100 | 199419 | 33311920 | 4033154 | 18043143 | 3979224 | 26625541 | 43395 | 0 | 0 | 423898 | 0 | 442 | 11964 | 84914 | 43224 | 0 | 4304 | 83278 | 0 | 171128 | 28667 | 1138 | 45498 | 17280 | 509 | 0 | 5769 | 51618 | 0 | 28337 |  | Unknown
WO-23003-79-A-1 | 0 | 41053 | 91339 | 7649 | 0 | 0 | 3771658 | 0 | 88679 | 0 | 1071 | 5038 | 0 | 13143 | 67436 | 751633 | 1300578 | 352862 | 2187823 | 8366060 | 2837187 | 297518 | 32949161 | 4090681 | 20666124 | 4033920 | 27348729 | 75986 | 0 | 0 | 527901 | 0 | 461 | 11075 | 143335 | 67083 | 1379 | 3411 | 79953 | 0 | 156368 | 36709 | 4103 | 57485 | 20199 | 1336 | 0 | 6648 | 46381 | 0 | 33233 |  | Unknown
WO-23003-78-A-1 | 0 | 37839 | 92682 | 6297 | 0 | 0 | 4191131 | 0 | 103394 | 0 | 400 | 5682 | 0 | 10643 | 75309 | 712237 | 1226427 | 354479 | 2259321 | 8830643 | 3047045 | 359893 | 33969588 | 4121518 | 20974092 | 4070156 | 28502611 | 87338 | 0 | 0 | 566963 | 0 | 608 | 11364 | 165687 | 49528 | 0 | 4524 | 0 | 0 | 154032 | 27406 | 3208 | 51312 | 19459 | 1226 | 0 | 7058 | 46479 | 1256 | 33700 |  | Unknown
WO-23003-77-A-1 | 1413 | 42162 | 77711 | 6835 | 0 | 0 | 4296232 | 0 | 113873 | 2539 | 717 | 3389 | 1898 | 11366 | 80098 | 702413 | 1067492 | 367850 | 2404844 | 9290518 | 3165726 | 372623 | 36194142 | 4131071 | 22688275 | 4163590 | 30051573 | 98954 | 0 | 0 | 593294 | 0 | 1148 | 11626 | 115120 | 72356 | 0 | 2461 | 32505 | 0 | 166724 | 40591 | 567 | 57090 | 18195 | 1652 | 0 | 7300 | 51679 | 0 | 35343 |  | Unknown
WO-23003-76-A-1 | 0 | 38848 | 79715 | 5004 | 0 | 0 | 4211035 | 0 | 98821 | 0 | 1382 | 7911 | 1351 | 15022 | 76075 | 667277 | 1152153 | 354298 | 2199167 | 8550553 | 3045691 | 325144 | 35596451 | 4144216 | 21155629 | 4064344 | 28576413 | 82684 | 0 | 0 | 553255 | 0 | 534 | 16033 | 91014 | 308167 | 0 | 3242 | 0 | 0 | 167889 | 33876 | 2995 | 49988 | 20067 | 848 | 0 | 7468 | 53338 | 1745 | 35940 |  | Unknown
WO-23003-75-A-1 | 0 | 47130 | 94980 | 6951 | 0 | 0 | 3049833 | 0 | 83507 | 0 | 1746 | 6246 | 1540 | 19837 | 68871 | 817940 | 1148460 | 346097 | 2037652 | 7754786 | 2759560 | 273566 | 28726809 | 3853319 | 19069664 | 3906238 | 25054025 | 70217 | 0 | 0 | 513406 | 0 | 295 | 10872 | 0 | 6579318 | 0 | 3479 | 149587 | 0 | 164657 | 48309 | 3851 | 52961 | 15749 | 777 | 0 | 6346 | 51633 | 0 | 32150 |  | Unknown
WO-23003-74-A-1 | 0 | 39192 | 74579 | 11921 | 0 | 0 | 3733174 | 0 | 108401 | 0 | 0 | 7967 | 1164 | 16075 | 77935 | 656935 | 1016146 | 363087 | 2248390 | 8660413 | 3016217 | 350115 | 33163671 | 4007846 | 20917351 | 3985448 | 28137426 | 94132 | 0 | 0 | 555461 | 0 | 1639 | 11095 | 0 | 5436768 | 0 | 2706 | 0 | 0 | 157426 | 34329 | 2751 | 48518 | 17290 | 1112 | 0 | 5268 | 47775 | 0 | 30740 |  | Unknown
WO-23003-73-A-1 | 0 | 40538 | 80964 | 10823 | 6483 | 0 | 3540604 | 0 | 82134 | 0 | 961 | 8188 | 0 | 16111 | 66858 | 603927 | 1090156 | 323190 | 1979874 | 7786665 | 2800246 | 283428 | 32476129 | 3990103 | 19586057 | 3936148 | 27116037 | 68866 | 0 | 0 | 484900 | 0 | 664 | 10970 | 0 | 2614769 | 0 | 3400 | 46548 | 0 | 159792 | 34021 | 5179 | 67060 | 15320 | 1019 | 0 | 6723 | 52244 | 0 | 34513 |  | Unknown
WO-23003-72-A-1 | 0 | 42212 | 104299 | 4758 | 7837 | 0 | 2812552 | 0 | 79615 | 0 | 531 | 7273 | 1441 | 17728 | 67336 | 880284 | 1262845 | 316272 | 2063857 | 7724853 | 2682500 | 262593 | 28298413 | 3708329 | 18645532 | 3652557 | 23551651 | 67634 | 0 | 0 | 494171 | 0 | 1638 | 9340 | 0 | 2684556 | 0 | 1121 | 0 | 0 | 142634 | 24156 | 4869 | 62260 | 14989 | 1076 | 0 | 7553 | 45004 | 0 | 24157 |  | Unknown
WO-23003-71-A-1 | 0 | 37151 | 79580 | 8125 | 6186 | 0 | 3506798 | 0 | 101254 | 1943 | 0 | 8846 | 0 | 12175 | 66904 | 664588 | 1024692 | 342097 | 2109933 | 8452869 | 2904367 | 300213 | 32070510 | 3912493 | 20358612 | 3924923 | 27434685 | 86431 | 0 | 0 | 521621 | 0 | 504 | 11737 | 0 | 2288947 | 0 | 2991 | 0 | 0 | 143354 | 29705 | 2638 | 49139 | 15822 | 2415 | 0 | 4605 | 52205 | 0 | 29638 |  | Unknown
AB-230117-02 LRB 8 | 0 | 45788 | 39184 | 3904 | 7974 | 0 | 3837556 | 0 | 358646 | 0 | 319 | 6277 | 2086 | 15015 | 146321 | 689894 | 800103 | 736670 | 3110107 | 15238182 | 6418043 | 686043 | 35232704 | 3788237 | 27135711 | 3791022 | 31335940 | 346903 | 0 | 0 | 1219569 | 0 | 1127 | 10662 | 0 | 46877 | 1030 | 427 | 22577 | 0 | 101059 | 26590 | 4695 | 76192 | 8706 | 1605 | 0 | 12442 | 50214 | 0 | 62278 |  | Unknown
LRB-8-inj1 | 0 | 45470 | 34712 | 5413 | 8942 | 0 | 3966551 | 0 | 359720 | 2419 | 681 | 6551 | 282 | 17682 | 156982 | 708782 | 865522 | 734076 | 3310690 | 15916496 | 6387645 | 745656 | 38150621 | 4183761 | 28461204 | 4128230 | 32962886 | 344404 | 0 | 0 | 1155057 | 0 | 0 | 13526 | 0 | 149199 | 0 | 229 | 0 | 0 | 134576 | 28417 | 4863 | 72203 | 14261 | 2052 | 0 | 11799 | 57896 | 0 | 53253 |  | Unknown
WO-23003-70-A-1 | 0 | 45173 | 72011 | 12687 | 8167 | 0 | 3518196 | 0 | 75845 | 0 | 2424 | 4034 | 0 | 17050 | 68661 | 588174 | 1028540 | 325383 | 2085017 | 8036131 | 2863392 | 270631 | 33314099 | 3892897 | 19678201 | 3802834 | 27121773 | 64949 | 0 | 0 | 523354 | 0 | 720 | 12194 | 0 | 2329145 | 0 | 2906 | 29568 | 0 | 147873 | 39412 | 3658 | 50952 | 16524 | 1961 | 0 | 8013 | 51197 | 0 | 62907 |  | Unknown
WO-23003-69-A-1 | 0 | 36117 | 69499 | 8439 | 10090 | 0 | 3645925 | 0 | 81183 | 0 | 584 | 8235 | 1010 | 16050 | 67118 | 571184 | 1024676 | 323597 | 2013235 | 7864338 | 2797238 | 237319 | 33335946 | 4046453 | 19684253 | 3994878 | 27044723 | 68747 | 0 | 0 | 505162 | 0 | 365 | 13067 | 0 | 1565467 | 0 | 2623 | 43899 | 0 | 150976 | 27313 | 5815 | 66575 | 15929 | 1963 | 0 | 5963 | 54231 | 0 | 72081 |  | Unknown
WO-23003-68-A-1 | 0 | 38072 | 104729 | 15040 | 11230 | 0 | 3056028 | 0 | 83911 | 0 | 1055 | 8405 | 1693 | 18271 | 69700 | 873192 | 1237500 | 313311 | 2117441 | 8067058 | 2893781 | 245906 | 29097804 | 3866376 | 19475658 | 3818392 | 24328483 | 72278 | 0 | 0 | 540784 | 0 | 275 | 10673 | 0 | 1459317 | 0 | 1571 | 137373 | 0 | 142698 | 33296 | 3668 | 52101 | 17221 | 2009 | 0 | 6306 | 51242 | 0 | 53900 |  | Unknown
WO-23003-67-A-1 | 0 | 37516 | 73767 | 5970 | 15940 | 0 | 3715637 | 0 | 84558 | 0 | 955 | 9135 | 1651 | 16398 | 69783 | 637269 | 1065555 | 315898 | 2198727 | 8506885 | 2982161 | 267423 | 33824164 | 4103182 | 20463965 | 4021916 | 27985442 | 72354 | 0 | 0 | 568400 | 0 | 626 | 12580 | 0 | 949598 | 0 | 2480 | 41186 | 0 | 152874 | 31621 | 2777 | 52641 | 16755 | 2806 | 0 | 7134 | 56979 | 768 | 44873 |  | Unknown
WO-23003-66-A-1 | 0 | 39261 | 84573 | 7686 | 18202 | 1046 | 4340512 | 0 | 94881 | 0 | 571 | 2572 | 0 | 18658 | 73805 | 740523 | 1110991 | 301071 | 2196066 | 8600488 | 3053264 | 283648 | 29837440 | 4199735 | 21091039 | 4144308 | 26677106 | 83558 | 0 | 0 | 575404 | 0 | 618 | 12831 | 0 | 828275 | 0 | 3520 | 27989 | 0 | 146301 | 30854 | 4815 | 51724 | 18136 | 3568 | 0 | 8616 | 52431 | 0 | 28461 |  | Unknown
WO-23003-65-A-1 | 1570 | 39124 | 83669 | 10955 | 23224 | 0 | 3266137 | 0 | 48721 | 0 | 1089 | 8340 | 0 | 17030 | 52388 | 681552 | 1135874 | 226076 | 1923390 | 7144751 | 2397431 | 172613 | 30651184 | 3933942 | 18142797 | 3902899 | 25240218 | 39898 | 0 | 0 | 449024 | 0 | 894 | 13849 | 0 | 1337561 | 0 | 678 | 44829 | 2651 | 153439 | 28808 | 6979 | 42589 | 18439 | 2667 | 0 | 6484 | 55407 | 0 | 30505 |  | Unknown
WO-23003-64-A-1 | 0 | 38138 | 64613 | 9837 | 30519 | 0 | 3540348 | 0 | 44213 | 0 | 791 | 10668 | 971 | 17396 | 49076 | 508028 | 926326 | 201717 | 1729750 | 6608632 | 2264569 | 140572 | 31846879 | 3905064 | 16931737 | 3776287 | 25722526 | 34007 | 0 | 0 | 423399 | 0 | 430 | 14653 | 0 | 1880229 | 0 | 2422 | 41689 | 0 | 158777 | 45387 | 7226 | 45122 | 16065 | 4579 | 0 | 5469 | 66451 | 0 | 74762 |  | Unknown
WO-23003-63-A-1 | 0 | 36004 | 67696 | 8968 | 45451 | 0 | 3829429 | 0 | 56945 | 1382 | 681 | 9425 | 1753 | 19475 | 60563 | 589679 | 1024970 | 245238 | 1889404 | 7221355 | 2596358 | 176897 | 33426157 | 4088866 | 19240240 | 4026837 | 27207511 | 47697 | 0 | 0 | 501018 | 0 | 392 | 15904 | 0 | 1522078 | 0 | 3484 | 51325 | 0 | 163453 | 42634 | 8271 | 52239 | 17863 | 4691 | 0 | 4038 | 76247 | 0 | 86294 |  | Unknown
WO-23003-62-A-1 | 0 | 37306 | 74484 | 8885 | 70423 | 0 | 3752867 | 0 | 69091 | 0 | 1788 | 9919 | 3103 | 20923 | 70215 | 625665 | 1079636 | 244470 | 2105867 | 8268549 | 2937539 | 211321 | 33327527 | 4027168 | 20501940 | 4020798 | 28012206 | 58574 | 0 | 0 | 533107 | 0 | 716 | 14347 | 0 | 980633 | 0 | 3715 | 30162 | 0 | 144824 | 46212 | 3461 | 66276 | 19679 | 5206 | 0 | 7169 | 86053 | 0 | 64381 |  | Unknown
WO-23003-61-A-1 | 0 | 38310 | 83062 | 16143 | 169893 | 0 | 3716879 | 0 | 49089 | 0 | 626 | 16747 | 0 | 24713 | 61495 | 614972 | 1086930 | 184116 | 1857673 | 7382427 | 2553636 | 173143 | 31165442 | 3882644 | 18165095 | 3818796 | 25807274 | 39650 | 1514 | 0 | 473310 | 0 | 445 | 24553 | 0 | 659319 | 0 | 1839 | 68951 | 0 | 130772 | 69055 | 14259 | 66110 | 15462 | 11978 | 0 | 6549 | 143976 | 0 | 108778 |  | Unknown
AB-230117-02 LFSM-H 3 | 0 | 56850 | 456647 | 643062 | 73707881 | 0 | 2726498 | 0 | 41829 | 578319 | 192546 | 3387854 | 1799709 | 3073156 | 55229 | 406774 | 600001 | 210395 | 1830236 | 7067342 | 2497078 | 206173 | 25558148 | 3377565 | 17167760 | 3212403 | 22029527 | 34135 | 0 | 0 | 468608 | 0 | 730 | 2287862 | 63012 | 25536709 | 0 | 3124 | 6921542 | 2362256 | 212301 | 20199923 | 3363898 | 16377660 | 60341 | 3573590 | 0 | 1098719 | 31569909 | 0 | 21796468 |  | Unknown
AB-230117-02 CCC-H 4 | 0 | 325555 | 670045 | 740961 | 120230961 | 0 | 2248001 | 0 | 258513 | 2951692 | 896575 | 5733532 | 4509367 | 4832420 | 135482 | 362711 | 364696 | 463618 | 2194755 | 11188051 | 5962507 | 545200 | 22953023 | 3114180 | 20043337 | 2555118 | 21827783 | 245197 | 0 | 0 | 1173296 | 0 | 353 | 3584230 | 451380 | 42633436 | 0 | 279 | 21725133 | 11117525 | 223974 | 38195355 | 4999985 | 37579983 | 72389 | 5846053 | 0 | 4734852 | 53125089 | 0 | 35077246 |  | Unknown
AB-230117-02 LFB-H 4 | 0 | 144101 | 454007 | 551792 | 68280380 | 0 | 2773144 | 0 | 252306 | 1504306 | 456291 | 3376106 | 2663538 | 2752725 | 142217 | 440656 | 502842 | 443262 | 2745742 | 13343664 | 6098605 | 582983 | 28066312 | 3332796 | 23171554 | 2951750 | 25512501 | 238833 | 0 | 0 | 1181283 | 904 | 232 | 2231474 | 208346 | 26451828 | 0 | 305 | 13253636 | 5693054 | 163076 | 23028556 | 2992807 | 21527843 | 38387 | 3138970 | 0 | 2572635 | 29132558 | 0 | 22970324 |  | Unknown
AB-230117-02 LD 3 | 0 | 40607 | 64993 | 4225 | 6501 | 0 | 3379100 | 0 | 37987 | 0 | 532 | 10460 | 1302 | 15109 | 58006 | 484238 | 841083 | 195206 | 1797806 | 6986338 | 2303037 | 178149 | 28505998 | 3264048 | 17701996 | 3312750 | 24386723 | 33217 | 0 | 0 | 398970 | 0 | 846 | 8487 | 25399 | 244830 | 0 | 1872 | 45214 | 0 | 57604 | 35649 | 1033 | 47772 | 9545 | 1098 | 0 | 8481 | 42795 | 0 | 30066 |  | Unknown
AB-230117-02 LRB 7 | 1263 | 15209 | 35579 | 2036 | 10032 | 0 | 192510 | 0 | 5418 | 0 | 530 | 4203 | 0 | 14687 | 4709 | 19612 | 44665 | 6978 | 78334 | 250544 | 78411 | 7836 | 1533140 | 146825 | 663256 | 154930 | 1096848 | 5123 | 0 | 0 | 12758 | 0 | 2152 | 2483 | 0 | 580373 | 0 | 26 | 0 | 0 | 18119 | 23992 | 1936 | 47828 | 6610 | 1567 | 0 | 13677 | 20833 | 1062 | 27134 |  | Unknown
LRB7-inj1 | 2038 | 53970 | 31236 | 5536 | 11304 | 0 | 3537374 | 0 | 253227 | 0 | 386 | 5817 | 1297 | 18694 | 145586 | 611887 | 735305 | 501100 | 1875354 | 13421045 | 6075374 | 445401 | 34977386 | 3432771 | 26555056 | 3619684 | 30394983 | 240978 | 0 | 0 | 1135717 | 0 | 2126 | 11251 | 0 | 4665299 | 0 | 174 | 36282 | 0 | 62873 | 30176 | 4439 | 63583 | 5504 | 1968 | 0 | 12392 | 53584 | 0 | 82383 |  | Unknown
WO-23003-60-A-1 | 0 | 40697 | 45819 | 7707 | 17520 | 0 | 2298348 | 0 | 5085 | 0 | 904 | 7987 | 1318 | 11980 | 26657 | 372629 | 275521 | 99288 | 1576405 | 5395806 | 1307468 | 55113 | 9685815 | 2699768 | 15734824 | 2805663 | 21786661 | 3676 | 0 | 0 | 229596 | 0 | 818 | 10683 | 0 | 202156810 | 0 | 2020 | 108396 | 0 | 54824 | 288133 | 4453 | 152927 | 8818 | 1370 | 0 | 4011 | 55533 | 950 | 94826 |  | Unknown
WO-23003-59-A-1 | 0 | 43976 | 68020 | 3092 | 35296 | 0 | 2815865 | 46 | 45359 | 0 | 0 | 4105 | 2083 | 15373 | 52962 | 564596 | 532379 | 267109 | 2051000 | 7946611 | 2062473 | 143849 | 20147347 | 2954149 | 18992116 | 3414709 | 25958296 | 35406 | 0 | 0 | 404907 | 0 | 492 | 11901 | 0 | 65216052 | 0 | 2958 | 28261 | 0 | 69472 | 90322 | 4792 | 76702 | 11863 | 4632 | 0 | 9008 | 71072 | 0 | 92753 |  | Unknown
WO-23003-58-A-1 | 0 | 53117 | 66607 | 12998 | 60583 | 0 | 2967392 | 39 | 50012 | 0 | 0 | 12661 | 2596 | 21062 | 58400 | 505157 | 733475 | 191825 | 1777737 | 6905818 | 2336952 | 230470 | 27314869 | 2849235 | 16830488 | 3252455 | 24137122 | 43031 | 0 | 0 | 419100 | 0 | 543 | 12625 | 0 | 8133209 | 0 | 3131 | 54950 | 0 | 63211 | 137152 | 9962 | 58849 | 11751 | 4166 | 0 | 6244 | 86593 | 0 | 100105 |  | Unknown
WO-23003-138-A-1 | 0 | 22640 | 71618 | 4963 | 0 | 0 | 5714048 | 0 | 113300 | 0 | 90863 | 5267 | 4079 | 10255 | 81370 | 959218 | 1543897 | 379201 | 2401875 | 9433895 | 3349023 | 365920 | 35416341 | 4321380 | 22679273 | 4161902 | 28595325 | 96973 | 0 | 0 | 583191 | 0 | 602 | 14789 | 0 | 66502 | 0 | 3461 | 39268 | 0 | 159969 | 25178 | 1400 | 41085 | 15305 | 1720 | 0 | 6585 | 48708 | 0 | 28011 |  | Unknown
WO-23003-137-A-1 | 0 | 26465 | 77300 | 8334 | 0 | 0 | 5683525 | 0 | 151711 | 0 | 79951 | 4894 | 0 | 8338 | 89840 | 955129 | 1479764 | 456434 | 2499664 | 9881285 | 3548669 | 431510 | 35718603 | 4543633 | 22937395 | 4280646 | 29179604 | 132259 | 1113 | 0 | 629946 | 0 | 419 | 14611 | 0 | 70238 | 0 | 2918 | 0 | 0 | 165375 | 22172 | 921 | 54106 | 17720 | 1285 | 0 | 6417 | 44177 | 0 | 32367 |  | Unknown
WO-23003-136-A-1 | 0 | 22127 | 70840 | 8207 | 0 | 0 | 5377131 | 0 | 14069 | 1707 | 32484 | 4723 | 85 | 7985 | 35638 | 557647 | 1336808 | 156718 | 1386247 | 4975276 | 1406570 | 77642 | 30705787 | 3865038 | 13518015 | 3699614 | 22565939 | 10602 | 0 | 0 | 255224 | 0 | 517 | 14024 | 0 | 68033 | 0 | 1991 | 0 | 0 | 167303 | 16462 | 1316 | 27299 | 17368 | 741 | 0 | 3780 | 51889 | 0 | 25694 |  | Unknown
WO-23003-135-A-1 | 0 | 24477 | 71095 | 5038 | 0 | 0 | 4832413 | 0 | 119113 | 0 | 68330 | 2864 | 704 | 11655 | 72970 | 663416 | 1431970 | 351604 | 1568439 | 6554848 | 2814769 | 355541 | 30085287 | 3942866 | 14745212 | 3590842 | 22759139 | 103340 | 896 | 0 | 508344 | 0 | 165 | 10799 | 0 | 67432 | 0 | 2608 | 48168 | 0 | 165981 | 25240 | 1028 | 41601 | 14571 | 972 | 0 | 5583 | 46396 | 1138 | 21265 |  | Unknown
WO-23003-134-A-1 | 0 | 24527 | 88605 | 4900 | 0 | 0 | 5533954 | 0 | 116705 | 624 | 81449 | 5063 | 1024 | 8854 | 73249 | 829767 | 1748752 | 365967 | 2043914 | 7994462 | 3030173 | 367619 | 34656130 | 4726989 | 19418592 | 4282981 | 26244060 | 98231 | 0 | 0 | 543832 | 0 | 281 | 15000 | 0 | 72886 | 352 | 3441 | 19754 | 0 | 170367 | 20337 | 3026 | 50856 | 14721 | 1043 | 0 | 8274 | 47341 | 0 | 34518 |  | Unknown
WO-23003-133-A-1 | 0 | 3457 | 78269 | 9854 | 0 | 2045 | 4892994 | 0 | 73186 | 3078 | 52812 | 4138 | 0 | 9430 | 60825 | 582022 | 1242523 | 274414 | 1214269 | 5539638 | 2242924 | 245807 | 27027843 | 3543979 | 12101965 | 3200615 | 19143175 | 61954 | 0 | 0 | 406838 | 0 | 559 | 11276 | 0 | 71217 | 726 | 2926 | 0 | 0 | 161791 | 17824 | 741 | 31440 | 16882 | 1176 | 0 | 1346 | 44741 | 1353 | 31651 |  | Unknown
WO-23003-132-A-1 | 0 | 24412 | 72305 | 5768 | 0 | 0 | 5513665 | 0 | 38092 | 0 | 21350 | 3411 | 0 | 8120 | 48919 | 590867 | 1312655 | 237812 | 1478371 | 6017671 | 2055522 | 158852 | 33892710 | 4002897 | 15305524 | 3707534 | 24372583 | 31432 | 0 | 0 | 362798 | 0 | 327 | 13049 | 0 | 130371 | 0 | 2874 | 0 | 0 | 182212 | 155153 | 2668 | 96747 | 18465 | 1034 | 0 | 5428 | 48194 | 0 | 42759 |  | Unknown
WO-23003-131-A-1 | 0 | 0 | 58576 | 6062 | 0 | 0 | 6015082 | 0 | 111380 | 1054 | 27311 | 6142 | 0 | 7067 | 68496 | 655395 | 1271623 | 347480 | 1809424 | 7525095 | 2844506 | 341053 | 34747393 | 4159293 | 18207144 | 3956493 | 26549325 | 94356 | 0 | 0 | 530790 | 0 | 810 | 11666 | 0 | 129982 | 0 | 3308 | 22775 | 0 | 157890 | 26148 | 1692 | 47243 | 15283 | 786 | 0 | 3864 | 45513 | 0 | 27020 |  | Unknown
AB-230117-02 LRB 14 | 1050 | 26295 | 33956 | 1614 | 0 | 0 | 4987787 | 0 | 534411 | 0 | 0 | 1923 | 1453 | 8398 | 185531 | 654226 | 738541 | 984231 | 3160520 | 16995537 | 7990089 | 1061146 | 40930026 | 4078276 | 28640931 | 3665930 | 33465281 | 515255 | 0 | 0 | 1427915 | 0 | 415 | 13041 | 0 | 136180 | 0 | 0 | 27874 | 32432 | 123799 | 27972 | 1310 | 104436 | 14037 | 1130 | 0 | 7346 | 56038 | 0 | 32790 |  | Unknown
LRB-14-inj1 | 0 | 25334 | 33745 | 3472 | 0 | 0 | 5546533 | 0 | 552859 | 0 | 1337 | 3994 | 1962 | 8087 | 189896 | 728109 | 934100 | 970248 | 3645393 | 18338564 | 7939528 | 1092118 | 44684916 | 4437394 | 31459571 | 4127495 | 36416657 | 523742 | 1572 | 0 | 1423067 | 0 | 966 | 17661 | 0 | 152220 | 0 | 62 | 57116 | 30452 | 171266 | 25896 | 2350 | 100379 | 17426 | 1708 | 0 | 8293 | 53273 | 897 | 34496 |  | Unknown
WO-23003-130-A-1 | 0 | 24966 | 72673 | 6479 | 0 | 0 | 5013821 | 0 | 48031 | 1162 | 19435 | 2603 | 0 | 8552 | 56984 | 759322 | 1376809 | 263797 | 1830424 | 6867673 | 2212515 | 195110 | 32804320 | 4018962 | 17821794 | 3863138 | 25163894 | 37183 | 0 | 0 | 392233 | 0 | 313 | 15547 | 0 | 127522 | 0 | 1383 | 44304 | 0 | 166665 | 23579 | 2074 | 56465 | 15340 | 750 | 0 | 6646 | 45548 | 0 | 24257 |  | Unknown
WO-23003-129-A-1 | 0 | 23507 | 66874 | 7996 | 0 | 0 | 5134153 | 0 | 75829 | 0 | 18985 | 3631 | 637 | 11667 | 58825 | 467697 | 1141750 | 277829 | 1148923 | 5448219 | 2274633 | 257650 | 28034179 | 3583773 | 11565885 | 3174631 | 18874520 | 60467 | 0 | 0 | 408787 | 0 | 826 | 10124 | 0 | 161030 | 1329 | 2917 | 21013 | 0 | 164716 | 21990 | 1899 | 34916 | 16158 | 1122 | 0 | 2010 | 50498 | 1068 | 31022 |  | Unknown
WO-23003-128-A-1 | 0 | 18812 | 45827 | 5026 | 0 | 0 | 4521838 | 0 | 55745 | 1461 | 11438 | 4285 | 531 | 6493 | 50172 | 269409 | 704893 | 237811 | 934843 | 4814475 | 1982972 | 202431 | 23495238 | 2804990 | 9105334 | 2260065 | 14623735 | 45411 | 0 | 0 | 350118 | 0 | 556 | 12780 | 0 | 182487 | 1140 | 1431 | 19144 | 0 | 168441 | 17179 | 1111 | 33477 | 18839 | 1100 | 0 | 1166 | 54024 | 721 | 36177 |  | Unknown
WO-23003-127-A-1 | 0 | 21473 | 57751 | 7152 | 0 | 0 | 5124173 | 0 | 96709 | 0 | 14519 | 2172 | 1078 | 7406 | 60889 | 372386 | 977979 | 289169 | 1117778 | 5581262 | 2451430 | 307105 | 28731300 | 3410322 | 11281829 | 2945807 | 19103573 | 79591 | 0 | 0 | 446521 | 0 | 232 | 10830 | 0 | 267880 | 0 | 1747 | 0 | 0 | 163671 | 24903 | 1790 | 34051 | 17623 | 690 | 0 | 5109 | 51149 | 0 | 25144 |  | Unknown
WO-23003-126-A-1 | 0 | 18768 | 55411 | 5999 | 0 | 0 | 4323259 | 0 | 36613 | 0 | 9455 | 4677 | 661 | 7609 | 38435 | 272297 | 734369 | 184292 | 846387 | 4204367 | 1591971 | 142551 | 19273236 | 2465688 | 8130357 | 1976140 | 12752173 | 30509 | 0 | 0 | 292620 | 0 | 362 | 11958 | 0 | 447069 | 0 | 1539 | 0 | 0 | 156849 | 18785 | 1683 | 0 | 19381 | 1552 | 0 | 1334 | 57323 | 1523 | 24456 |  | Unknown
WO-23003-125-A-1 | 0 | 23650 | 78437 | 10344 | 3747 | 0 | 4651750 | 0 | 40013 | 0 | 16372 | 5374 | 0 | 10272 | 56069 | 799273 | 1335919 | 256722 | 1690241 | 6368431 | 2125778 | 159849 | 28262592 | 3845229 | 15625123 | 3598788 | 22372578 | 32552 | 0 | 0 | 367291 | 0 | 489 | 10809 | 0 | 6511456 | 0 | 2964 | 0 | 0 | 149126 | 33089 | 1446 | 51902 | 16838 | 1273 | 0 | 6853 | 50670 | 0 | 23675 |  | Unknown
WO-23003-121-A-1B | 0 | 23009 | 50176 | 6308 | 6891 | 0 | 3437692 | 0 | 207188 | 1268 | 0 | 3327 | 0 | 8109 | 92009 | 697381 | 390908 | 507402 | 2728083 | 11420592 | 4054371 | 548877 | 13605851 | 3981506 | 24419468 | 3747443 | 30006898 | 191988 | 0 | 0 | 745087 | 0 | 266 | 12075 | 0 | 169342682 | 1139 | 1767 | 178871 | 29376 | 104584 | 1086943 | 3317 | 1059285 | 10819 | 1245 | 0 | 9294 | 48739 | 784 | 42559 |  | Unknown
WO-23003-111-A-1B | 0 | 22987 | 215961 | 5057 | 8746 | 0 | 4416627 | 0 | 56235 | 0 | 5967 | 6483 | 1554 | 18040 | 58277 | 573944 | 598968 | 285769 | 2043807 | 7810096 | 2481005 | 228257 | 24339869 | 3970597 | 19681184 | 3769791 | 27636877 | 52232 | 0 | 0 | 445432 | 0 | 747 | 12013 | 0 | 88156947 | 0 | 1702 | 106086 | 21224 | 77852 | 291450 | 1371 | 460268 | 8147 | 2138 | 0 | 7769 | 55371 | 1239 | 35912 |  | Unknown
AB-230117-02 LRB 13 20230123105625 | 0 | 48 | 31003 | 2695 | 12976 | 0 | 4969519 | 0 | 500247 | 1324 | 0 | 4883 | 2644 | 8344 | 169339 | 618027 | 739289 | 899201 | 3126958 | 16592397 | 7254703 | 977390 | 37983588 | 3999587 | 27657980 | 3550476 | 32002182 | 477597 | 0 | 0 | 1377450 | 0 | 613 | 13969 | 0 | 622893 | 0 | 394 | 38693 | 28051 | 103917 | 27822 | 3275 | 105403 | 9754 | 1106 | 0 | 9059 | 60145 | 547 | 41851 |  | Unknown
AB-230117-02 LRB 12 20230123102918 | 0 | 51924 | 39787 | 1712 | 11402 | 0 | 38195 | 0 | 9141 | 0 | 0 | 1854 | 1260 | 8123 | 2499 | 5368 | 6186 | 7032 | 22189 | 107309 | 50717 | 10281 | 273976 | 25458 | 176137 | 24467 | 226469 | 7694 | 0 | 0 | 9823 | 0 | 661 | 700 | 0 | 1267000 | 1504 | 72 | 21364 | 0 | 11025 | 27721 | 3603 | 43623 | 6560 | 892 | 0 | 9575 | 29345 | 614 | 28837 | C | Unknown
AB-230117-02 QCS 2 | 0 | 204253 | 441359 | 603646 | 81574396 | 0 | 4420463 | 368 | 422987 | 1618969 | 947855 | 3470683 | 2709675 | 3151001 | 150030 | 363559 | 496802 | 802983 | 2333913 | 11682604 | 6179804 | 951336 | 29133809 | 3171892 | 20001675 | 2572196 | 22790999 | 401060 | 0 | 2142 | 1130478 | 1272 | 754 | 457901 | 425937 | 31453513 | 0 | 0 | 13269483 | 6529999 | 217306 | 24509136 | 429956 | 21781698 | 63781 | 378738 | 1042298 | 2843790 | 7168617 | 0 | 30110560 |  | Unknown
AB-230117-02 CCC-H 12 | 482 | 289264 | 554433 | 603771 | 109505664 | 1561 | 3107982 | 0 | 417924 | 2739489 | 1648536 | 4821065 | 3840892 | 8282863 | 135023 | 282032 | 290018 | 796744 | 1940837 | 9825469 | 5657886 | 970177 | 22029282 | 2691385 | 17254104 | 1870799 | 17173406 | 397349 | 0 | 0 | 1095778 | 0 | 314 | 4711087 | 752635 | 40232440 | 0 | 296 | 19466371 | 10779089 | 223959 | 29681762 | 4255560 | 32324222 | 80667 | 5001092 | 1884785 | 4503255 | 47603263 | 1165 | 33908553 |  | Unknown
WO-23003-212-A-1 | 0 | 19419 | 0 | 3222 | 5466 | 1194 | 5070144 | 0 | 477435 | 1700 | 8131 | 3080891 | 94447 | 31515941 | 164522 | 411196 | 580215 | 937794 | 3311931 | 16646158 | 7037727 | 1087278 | 37972771 | 1891630 | 26368237 | 818293 | 7639842 | 454164 | 0 | 0 | 1295538 | 0 | 300 | 13453 | 0 | 55514 | 1046 | 319 | 50215 | 27137 | 145798 | 0 | 194759 | 171241 | 14140 | 413 | 0 | 8545 | 52427 | 1310 | 29615 |  | Unknown
AB-230117-02 LRB 22 | 0 | 20776 | 21580 | 3009 | 5542 | 0 | 5414775 | 0 | 486154 | 1201 | 0 | 196503 | 7930 | 6682665 | 171590 | 576160 | 666014 | 956316 | 3315056 | 16895262 | 7312818 | 1124938 | 41046510 | 3648200 | 28520880 | 2760987 | 27771783 | 462212 | 0 | 0 | 1317309 | 0 | 519 | 13221 | 0 | 56623 | 0 | 0 | 30966 | 26495 | 145642 | 11911 | 17080 | 103592 | 6599 | 1402 | 0 | 8452 | 60896 | 728 | 34669 |  | Unknown
WO-23003-211-A-1 | 0 | 18241 | 0 | 3545 | 6295 | 0 | 5074136 | 0 | 476536 | 0 | 11928 | 3254453 | 95711 | 32536957 | 166175 | 415258 | 567747 | 925945 | 3336733 | 16561750 | 7069050 | 1128506 | 38808812 | 1924035 | 26980041 | 851737 | 7497363 | 453256 | 0 | 0 | 1336035 | 0 | 83 | 14156 | 0 | 53221 | 0 | 491 | 50261 | 31206 | 160001 | 0 | 203630 | 165269 | 14465 | 1071 | 0 | 5640 | 54935 | 1134 | 32574 |  | Unknown
WO-23003-210-A-1 | 1268 | 20087 | 0 | 2656 | 6746 | 0 | 4789669 | 0 | 438982 | 1066 | 12992 | 4235526 | 156049 | 32563042 | 159055 | 375087 | 561719 | 883799 | 3225027 | 16267255 | 6958152 | 1056920 | 37852841 | 1807117 | 25194154 | 765246 | 6993795 | 417165 | 0 | 0 | 1243565 | 0 | 496 | 12246 | 0 | 56277 | 0 | 0 | 55223 | 25410 | 156773 | 0 | 195119 | 207478 | 12164 | 542 | 0 | 6156 | 55722 | 405 | 39044 |  | Unknown
WO-23003-209-A-1 | 0 | 16521 | 55666 | 4033 | 6840 | 0 | 4838962 | 0 | 206665 | 1093 | 282685 | 17813585 | 2359020 | 44410429 | 91140 | 123396 | 606109 | 511993 | 2358730 | 10502171 | 3913333 | 622713 | 34154426 | 1368453 | 11487087 | 490738 | 3580538 | 189791 | 0 | 0 | 695284 | 0 | 287 | 13231 | 0 | 54296 | 0 | 241 | 504604 | 86410 | 145035 | 0 | 63903 | 1463923 | 10388 | 738 | 10985 | 20916 | 50729 | 1283 | 39549 |  | Unknown
WO-23003-208-A-1 | 0 | 15930 | 69997 | 3365 | 8842 | 0 | 4611754 | 0 | 150711 | 1534 | 237609 | 17234114 | 2163204 | 44379838 | 82548 | 114085 | 595119 | 463464 | 2194316 | 9255556 | 3484935 | 516429 | 34058358 | 1430752 | 11019572 | 534977 | 3671026 | 137549 | 0 | 0 | 641950 | 0 | 324 | 12584 | 0 | 45890 | 0 | 501 | 455636 | 84758 | 166427 | 0 | 61337 | 1375569 | 13978 | 491 | 0 | 19779 | 52614 | 1616 | 30435 |  | Unknown
WO-23003-207-A-1 | 0 | 17046 | 102659 | 6844 | 7940 | 0 | 4351378 | 0 | 149658 | 0 | 253584 | 17174020 | 2223808 | 45579388 | 82367 | 113003 | 600897 | 452265 | 2228845 | 9434078 | 3493807 | 522053 | 33979020 | 1387819 | 10652995 | 526947 | 3472443 | 133800 | 0 | 0 | 613027 | 0 | 344 | 13622 | 0 | 45362 | 0 | 1301 | 459228 | 78642 | 169542 | 0 | 61520 | 1423802 | 15579 | 977 | 0 | 15266 | 55974 | 722 | 40990 |  | Unknown
WO-23003-206-A-1 | 0 | 15853 | 94194 | 2851 | 11424 | 1393 | 4380456 | 0 | 171368 | 1987 | 279986 | 20102645 | 2573993 | 52157495 | 88220 | 112528 | 579466 | 503077 | 2217176 | 9988571 | 3682234 | 597765 | 31615078 | 1158911 | 9117678 | 396341 | 2571564 | 157626 | 1160 | 0 | 674686 | 0 | 457 | 13305 | 0 | 50005 | 0 | 486 | 591642 | 95677 | 142309 | 0 | 84171 | 1609566 | 12972 | 841 | 0 | 19419 | 54321 | 1124 | 41815 |  | Unknown
WO-23003-204-A-1 | 0 | 19091 | 232622 | 8059 | 14706 | 0 | 4457283 | 0 | 54720 | 1285 | 352169 | 21427404 | 3471206 | 49172875 | 55157 | 86133 | 896529 | 289048 | 1553806 | 6530989 | 2327833 | 244962 | 24559170 | 1094035 | 6360974 | 409960 | 2250114 | 44987 | 0 | 0 | 400465 | 0 | 289 | 9865 | 0 | 39040 | 0 | 944 | 373127 | 60363 | 147629 | 0 | 58207 | 941185 | 10345 | 628 | 0 | 29248 | 52330 | 1052 | 38913 |  | Unknown
WO-23003-203-A-1 | 0 | 15339 | 83037 | 4319 | 21224 | 0 | 4408325 | 0 | 41609 | 2572 | 297646 | 21399204 | 3196260 | 49723305 | 45012 | 81385 | 914523 | 247272 | 1473104 | 6136837 | 2082824 | 190796 | 24116360 | 1042863 | 5989206 | 363290 | 2065247 | 36289 | 0 | 0 | 361261 | 0 | 565 | 12044 | 0 | 44092 | 1092 | 437 | 448061 | 68990 | 151079 | 0 | 61176 | 1188738 | 9120 | 3290 | 0 | 23298 | 57431 | 0 | 36132 |  | Unknown
WO-23003-202-A-1 | 0 | 3310 | 40998 | 5869 | 52014 | 0 | 5499305 | 26 | 93189 | 0 | 304201 | 11494993 | 1379644 | 32216989 | 68950 | 190759 | 1033057 | 348026 | 1808625 | 7524275 | 2742275 | 354772 | 31378239 | 1780291 | 13140881 | 854523 | 6111114 | 80081 | 0 | 0 | 481883 | 0 | 387 | 15766 | 0 | 63395 | 0 | 1015 | 256110 | 55309 | 153111 | 0 | 31367 | 738795 | 13702 | 1891 | 0 | 19085 | 73874 | 0 | 56083 |  | Unknown
WO-23003-201-A-1 | 0 | 19192 | 34513 | 4056 | 123587 | 0 | 5353869 | 0 | 79493 | 4089 | 306194 | 11732335 | 1307838 | 32425999 | 62493 | 187388 | 989081 | 344747 | 1756825 | 7393348 | 2646422 | 334233 | 30882941 | 1754895 | 12833795 | 824947 | 6282357 | 65972 | 1654 | 0 | 454091 | 0 | 509 | 19440 | 0 | 93716 | 1277 | 1143 | 260033 | 48550 | 132485 | 20799 | 42662 | 728548 | 12476 | 5597 | 0 | 17364 | 124888 | 1086 | 112740 |  | Unknown
AB-230117-02 LFSM-H 10 | 0 | 70953 | 204025 | 630315 | 57277270 | 0 | 4038662 | 0 | 103082 | 615482 | 688288 | 12265304 | 2920422 | 31792958 | 69780 | 178413 | 615366 | 383532 | 1727810 | 6980925 | 2898810 | 392898 | 22408400 | 1452498 | 12232708 | 831195 | 5712771 | 88543 | 0 | 0 | 525551 | 0 | 349 | 3130516 | 94035 | 21394431 | 0 | 1773 | 6873541 | 2715295 | 205138 | 5076872 | 2779753 | 11647483 | 60066 | 1700508 | 360721 | 1145766 | 28175259 | 2351 | 19999415 |  | Unknown
AB-230117-02 CCC-H 11 | 0 | 268222 | 513983 | 529486 | 112284316 | 0 | 3566690 | 0 | 483672 | 2857702 | 1774053 | 4539835 | 3798387 | 4060758 | 140379 | 263608 | 279671 | 906117 | 2011273 | 10596408 | 6258063 | 1091221 | 22917705 | 3086199 | 17894682 | 2175740 | 19829498 | 460334 | 0 | 0 | 1221504 | 0 | 491 | 4986840 | 784632 | 40559193 | 907 | 50 | 19343685 | 11049893 | 231831 | 32857948 | 3977206 | 31793768 | 74378 | 5518390 | 2004953 | 4345397 | 49147562 | 0 | 37518346 |  | Unknown
AB-230117-02 LFB-H 11 | 0 | 170024 | 371829 | 420457 | 64756087 | 0 | 4883870 | 0 | 442155 | 1391869 | 863970 | 3366270 | 2279066 | 5924030 | 149140 | 357736 | 428283 | 839595 | 2303348 | 12023295 | 6137861 | 982929 | 26941243 | 3014461 | 20602763 | 2379547 | 21417114 | 409354 | 0 | 0 | 1116067 | 0 | 1025 | 3695826 | 364703 | 24627744 | 0 | 143 | 11601376 | 5689178 | 168253 | 18369737 | 2356765 | 18917556 | 36302 | 2918014 | 926579 | 2537970 | 26231482 | 783 | 22540694 |  | Unknown
AB-230117-02 LD 10 | 0 | 18272 | 30213 | 12258 | 0 | 0 | 4836993 | 0 | 13709 | 1457 | 119322 | 9963744 | 832213 | 32349720 | 32101 | 136780 | 936647 | 153826 | 1171678 | 4465287 | 1269400 | 92056 | 26862747 | 1353161 | 9817354 | 825097 | 5515946 | 11231 | 0 | 0 | 219738 | 0 | 560 | 11785 | 0 | 37934 | 2735 | 1936 | 159365 | 21506 | 146587 | 0 | 27535 | 536224 | 11910 | 1299 | 0 | 11112 | 45455 | 823 | 34180 |  | Unknown
AB-230117-02 LRB 21 | 448 | 19044 | 29776 | 4928 | 0 | 0 | 4661309 | 0 | 368862 | 348 | 920 | 41098 | 2931 | 153907 | 147842 | 474214 | 601928 | 869335 | 2577534 | 14108098 | 6373522 | 929339 | 32716061 | 3282778 | 23046320 | 2926896 | 26813713 | 354881 | 0 | 0 | 1185440 | 862 | 203 | 11024 | 0 | 54481 | 1646 | 0 | 33752 | 22000 | 101044 | 17796 | 1172 | 87678 | 4398 | 733 | 0 | 9934 | 41623 | 0 | 33136 |  | Unknown
LRB-21-inj1 | 0 | 22881 | 25148 | 8435 | 0 | 0 | 5480285 | 0 | 459951 | 0 | 0 | 598753 | 33889 | 4144237 | 167395 | 582188 | 734722 | 904832 | 3249198 | 17016118 | 7134991 | 1115417 | 42486830 | 3895662 | 28451576 | 3086215 | 29645595 | 437337 | 0 | 0 | 1252977 | 0 | 172 | 11208 | 0 | 63670 | 967 | 55 | 30678 | 31970 | 152410 | 19148 | 2979 | 114533 | 6702 | 707 | 0 | 7995 | 53048 | 559 | 33275 |  | Unknown
WO-23003-200-A-1 | 0 | 15533 | 20406 | 2924 | 2840 | 0 | 5430894 | 0 | 91589 | 0 | 305820 | 11327929 | 1305896 | 32182845 | 66944 | 171366 | 853529 | 366120 | 1842376 | 7731473 | 2804146 | 357210 | 34422132 | 1947721 | 14015330 | 896167 | 7062188 | 75617 | 0 | 0 | 489517 | 0 | 215 | 12377 | 0 | 37273 | 948 | 403 | 321363 | 63442 | 158357 | 0 | 27306 | 956748 | 15421 | 192 | 0 | 19489 | 46140 | 0 | 28762 |  | Unknown
WO-23003-199-A-1 | 604 | 19323 | 26449 | 5155 | 0 | 1006 | 5919865 | 0 | 117446 | 2140 | 380486 | 8530076 | 1156529 | 24121155 | 76961 | 324737 | 1088913 | 401948 | 2049307 | 8530274 | 3228719 | 462322 | 33610321 | 2375162 | 16534708 | 1399210 | 10973119 | 103257 | 0 | 0 | 569404 | 0 | 339 | 10843 | 0 | 43221 | 0 | 1241 | 232757 | 49058 | 165054 | 7386 | 14390 | 564085 | 13400 | 670 | 0 | 19712 | 45117 | 1115 | 31587 |  | Unknown
WO-23003-198-A-1 | 1236 | 16277 | 31790 | 5987 | 0 | 962 | 5025130 | 0 | 29450 | 1209 | 208593 | 12431155 | 1529477 | 33122706 | 39771 | 161211 | 1077263 | 200546 | 1312846 | 5297880 | 1671541 | 130518 | 27215197 | 1603065 | 9768776 | 821874 | 5348640 | 22061 | 0 | 3670 | 280313 | 0 | 632 | 9453 | 0 | 33696 | 1215 | 1133 | 224934 | 36742 | 168458 | 8460 | 24677 | 724754 | 15985 | 274 | 0 | 16796 | 42085 | 1744 | 22934 |  | Unknown
WO-23003-197-A-1 | 1526 | 15549 | 34712 | 6383 | 0 | 821 | 5331614 | 0 | 80004 | 3691 | 366496 | 14037953 | 2120324 | 33654683 | 65794 | 133657 | 806799 | 356339 | 1753471 | 7263200 | 2828714 | 328986 | 33394941 | 1835139 | 11764041 | 818531 | 6191372 | 68455 | 0 | 0 | 480511 | 0 | 274 | 12520 | 0 | 41671 | 0 | 1091 | 288228 | 53340 | 169628 | 0 | 26434 | 817122 | 15311 | 206 | 0 | 20853 | 47597 | 0 | 27613 |  | Unknown
WO-23003-196-A-1 | 0 | 16346 | 46070 | 6034 | 3682 | 0 | 5840777 | 0 | 39933 | 2787 | 62666 | 992694 | 86302 | 7464271 | 48782 | 380038 | 879971 | 250865 | 1581333 | 6057717 | 2018624 | 201306 | 34051812 | 3422189 | 15600340 | 2732450 | 20086012 | 35397 | 0 | 0 | 346963 | 0 | 366 | 11800 | 0 | 33805 | 0 | 2977 | 96976 | 10320 | 163803 | 15138 | 4401 | 74703 | 14399 | 800 | 0 | 5753 | 50055 | 1180 | 41679 |  | Unknown
WO-23003-195-A-1 | 0 | 17344 | 43781 | 5562 | 0 | 1386 | 5883393 | 0 | 77675 | 875 | 80332 | 1123101 | 100803 | 7419197 | 60371 | 467396 | 951624 | 345615 | 1913407 | 7573851 | 2680389 | 325336 | 35999627 | 3446401 | 18682395 | 2779978 | 21631474 | 64274 | 0 | 0 | 483473 | 0 | 184 | 12991 | 0 | 38050 | 0 | 2028 | 44059 | 14517 | 170354 | 0 | 3804 | 94285 | 14539 | 1121 | 0 | 4900 | 44098 | 768 | 34210 |  | Unknown
WO-23003-194-A-1 | 0 | 19255 | 45450 | 7060 | 3255 | 0 | 6166172 | 0 | 97597 | 3990 | 105819 | 1192217 | 107192 | 7836621 | 69391 | 538330 | 1092333 | 383637 | 2008907 | 8124924 | 2968645 | 362662 | 35050320 | 3469696 | 19273079 | 2755911 | 21045555 | 81652 | 0 | 0 | 509360 | 0 | 395 | 11051 | 0 | 38492 | 1613 | 1421 | 63898 | 11957 | 160526 | 17871 | 4215 | 101817 | 13435 | 665 | 0 | 4061 | 45384 | 740 | 29874 |  | Unknown
WO-23003-193-A-1 | 0 | 18269 | 43431 | 4115 | 0 | 0 | 5875707 | 0 | 77905 | 578 | 72626 | 842913 | 77144 | 6135963 | 62469 | 509278 | 967140 | 353054 | 1916812 | 7536871 | 2720983 | 328437 | 36116897 | 3554282 | 18691242 | 2918698 | 22666720 | 64886 | 0 | 0 | 481457 | 0 | 530 | 13755 | 0 | 31500 | 0 | 1700 | 48981 | 10963 | 168995 | 17257 | 4506 | 73375 | 16234 | 897 | 0 | 4577 | 48748 | 0 | 23420 |  | Unknown
WO-23003-192-A-1 | 0 | 17008 | 40866 | 7763 | 0 | 0 | 5956394 | 0 | 81339 | 0 | 169397 | 1071079 | 129371 | 5888696 | 61431 | 464571 | 1038094 | 358394 | 1938047 | 7685883 | 2784636 | 330002 | 35628495 | 3587671 | 18266146 | 2910447 | 22132711 | 67594 | 0 | 0 | 491857 | 0 | 380 | 10311 | 0 | 33166 | 0 | 1869 | 61063 | 15170 | 181493 | 23894 | 3653 | 95026 | 14357 | 1484 | 0 | 6308 | 44811 | 0 | 40344 |  | Unknown
WO-23003-191-A-1 | 0 | 15805 | 45192 | 5134 | 0 | 1019 | 5305642 | 0 | 20738 | 2244 | 81713 | 854251 | 91272 | 5402853 | 34592 | 375770 | 979005 | 182248 | 1279172 | 4983417 | 1507161 | 116649 | 31907415 | 3289420 | 13583557 | 2775798 | 19573895 | 15693 | 0 | 0 | 258299 | 0 | 786 | 10671 | 0 | 28528 | 2216 | 1477 | 91502 | 11261 | 168909 | 18004 | 2251 | 67846 | 16098 | 790 | 0 | 3713 | 51335 | 952 | 35571 |  | Unknown
AB-230117-02 LRB 20 | 0 | 21860 | 24698 | 1993 | 7853 | 0 | 5313269 | 0 | 452410 | 1315 | 0 | 4327 | 1558 | 14661 | 167592 | 588226 | 680975 | 968936 | 3149172 | 16405925 | 7324266 | 1109714 | 39648533 | 4020201 | 28240850 | 3481649 | 31973298 | 437246 | 0 | 0 | 1374949 | 0 | 218 | 15117 | 0 | 57752 | 0 | 243 | 45513 | 29633 | 152191 | 21965 | 1083 | 91655 | 12865 | 1027 | 0 | 8236 | 54712 | 0 | 30553 |  | Unknown
LRB-20-inj1 | 883 | 18912 | 30241 | 2231 | 7216 | 0 | 5681547 | 0 | 461591 | 0 | 0 | 37585 | 5775 | 251705 | 169333 | 586673 | 727335 | 957130 | 3272710 | 17037185 | 7502454 | 1143863 | 41119870 | 4158859 | 28880873 | 3544377 | 32546918 | 439978 | 0 | 0 | 1357430 | 0 | 432 | 14621 | 0 | 64856 | 807 | 237 | 36366 | 27426 | 169109 | 22722 | 2129 | 97222 | 13852 | 1899 | 0 | 8333 | 58129 | 746 | 40901 |  | Unknown
WO-23003-190-A-1 | 0 | 17753 | 43830 | 5521 | 9986 | 0 | 5312904 | 0 | 48964 | 0 | 147146 | 1107758 | 128322 | 6160498 | 51364 | 423956 | 940632 | 293861 | 1684141 | 6778091 | 2251107 | 229402 | 33587846 | 3588782 | 17055709 | 2896586 | 21148930 | 39881 | 0 | 0 | 394071 | 0 | 349 | 12122 | 0 | 65791 | 0 | 2595 | 46970 | 17026 | 171524 | 23060 | 4369 | 106894 | 14640 | 1391 | 0 | 5463 | 52839 | 568 | 35349 |  | Unknown
WO-23003-189-A-1 | 0 | 18329 | 51966 | 2645 | 9870 | 0 | 5798846 | 0 | 100311 | 2484 | 202440 | 1502818 | 190058 | 7520299 | 66707 | 573473 | 1151973 | 382234 | 2040648 | 8009665 | 2905790 | 379687 | 34075321 | 3460972 | 18794074 | 2918972 | 21142994 | 85580 | 0 | 0 | 512194 | 0 | 226 | 11627 | 0 | 65867 | 0 | 2347 | 112033 | 23597 | 167181 | 37239 | 5788 | 107755 | 14500 | 1383 | 0 | 13440 | 45746 | 2041 | 27851 |  | Unknown
WO-23003-188-A-1 | 0 | 17461 | 53949 | 5467 | 9911 | 0 | 5461830 | 0 | 82363 | 0 | 11997 | 107694 | 8947 | 1246448 | 62805 | 537514 | 956384 | 351928 | 1994692 | 7913775 | 2742557 | 339425 | 36804811 | 4020680 | 19244736 | 3607841 | 26529526 | 69449 | 0 | 0 | 484668 | 0 | 364 | 12425 | 0 | 39256 | 1138 | 2753 | 70396 | 0 | 170338 | 24493 | 1327 | 47615 | 16149 | 1639 | 0 | 2516 | 57931 | 0 | 46490 |  | Unknown
WO-23003-187-A-1 | 0 | 16312 | 49524 | 7741 | 10080 | 1442 | 5811317 | 0 | 82884 | 0 | 7269 | 91078 | 8413 | 1089742 | 68118 | 524589 | 999776 | 355746 | 1924367 | 7665300 | 2762067 | 344564 | 36518213 | 3997412 | 18966931 | 3637597 | 26285689 | 67936 | 0 | 0 | 474748 | 0 | 539 | 11240 | 0 | 43855 | 0 | 2566 | 18145 | 0 | 162875 | 17240 | 6008 | 40138 | 14766 | 1954 | 0 | 5789 | 54127 | 0 | 39627 |  | Unknown
WO-23003-186-A-1 | 0 | 20291 | 56233 | 2637 | 22171 | 0 | 5875287 | 0 | 89518 | 746 | 8764 | 94955 | 5730 | 1107808 | 72680 | 612699 | 1063101 | 361916 | 2032877 | 8345047 | 2949424 | 367737 | 35994281 | 3993570 | 19807118 | 3634040 | 25912912 | 74934 | 0 | 0 | 529083 | 0 | 395 | 12949 | 0 | 42108 | 0 | 1796 | 16895 | 0 | 171658 | 19701 | 3031 | 51814 | 15437 | 2015 | 0 | 7334 | 54059 | 1871 | 34912 |  | Unknown
WO-23003-185-A-1 | 0 | 19636 | 53015 | 5749 | 20589 | 0 | 5742435 | 0 | 90247 | 1589 | 26003 | 78216 | 8677 | 758587 | 71130 | 575329 | 1005008 | 380600 | 1990488 | 8085328 | 2904735 | 378090 | 35863933 | 4113732 | 19844099 | 3640898 | 26407794 | 77847 | 0 | 0 | 506374 | 0 | 341 | 15714 | 0 | 45017 | 0 | 1768 | 0 | 5834 | 161258 | 33578 | 3663 | 52530 | 15543 | 2521 | 0 | 8199 | 65089 | 871 | 51651 |  | Unknown
WO-23003-184-A-1 | 0 | 21220 | 54208 | 5408 | 26340 | 0 | 5555022 | 0 | 89785 | 2103 | 29062 | 98810 | 12248 | 861360 | 67579 | 614094 | 983345 | 403204 | 2344458 | 9300913 | 3094685 | 379749 | 35514802 | 4197163 | 21936335 | 3691711 | 27768915 | 76146 | 0 | 0 | 537967 | 0 | 165 | 14145 | 0 | 45911 | 1748 | 2047 | 0 | 0 | 162701 | 26832 | 2284 | 62005 | 14646 | 4062 | 0 | 8487 | 61122 | 0 | 42122 |  | Unknown
WO-23003-183-A-1 | 0 | 18983 | 53834 | 6023 | 40390 | 1044 | 5873467 | 0 | 96783 | 1902 | 31949 | 95078 | 11393 | 905281 | 70533 | 532112 | 1022261 | 379271 | 1920242 | 7749766 | 2915834 | 379708 | 36206525 | 4082820 | 19124889 | 3641152 | 26049925 | 82736 | 741 | 0 | 520642 | 0 | 73 | 11767 | 0 | 60280 | 0 | 2378 | 0 | 0 | 171342 | 24340 | 7859 | 51851 | 13878 | 4523 | 0 | 5706 | 72525 | 0 | 47103 |  | Unknown
WO-23003-182-A-1 | 0 | 19964 | 54736 | 4336 | 70544 | 0 | 5591705 | 0 | 87411 | 1022 | 0 | 29536 | 3616 | 299928 | 70023 | 607631 | 1037320 | 376204 | 2053163 | 8094669 | 2889862 | 345390 | 35678436 | 4108076 | 19655372 | 3699468 | 25915073 | 73792 | 0 | 0 | 498288 | 0 | 956 | 16708 | 0 | 57922 | 0 | 1758 | 21207 | 0 | 160905 | 35686 | 3837 | 60630 | 14609 | 5292 | 0 | 3763 | 84746 | 0 | 57621 |  | Unknown
WO-23003-181-A-1 | 0 | 19135 | 48323 | 2369 | 166338 | 0 | 5675313 | 38 | 92596 | 2822 | 0 | 30185 | 4652 | 321537 | 66458 | 546008 | 946983 | 387068 | 2055317 | 8126718 | 2959952 | 373765 | 35900244 | 3985933 | 19786204 | 3623051 | 26881852 | 79367 | 0 | 0 | 520554 | 0 | 568 | 25276 | 0 | 95090 | 1456 | 1605 | 0 | 0 | 155041 | 53566 | 7224 | 75533 | 16701 | 10481 | 0 | 5263 | 158450 | 0 | 99874 |  | Unknown
AB-230117-02 LFSM-H 9 | 0 | 32384 | 267992 | 448033 | 66906671 | 0 | 3856330 | 0 | 17742 | 284813 | 166334 | 2410899 | 1153267 | 2745748 | 30225 | 243236 | 438694 | 165991 | 1255344 | 4647554 | 1389125 | 100715 | 25494037 | 3190817 | 12329974 | 2649459 | 18646924 | 14244 | 0 | 0 | 242322 | 0 | 183 | 3088604 | 13110 | 23688789 | 701 | 1972 | 4399134 | 1239693 | 240014 | 16456350 | 2940918 | 11241430 | 62525 | 3180136 | 88281 | 659251 | 32990263 | 0 | 23187317 |  | Unknown
AB-230117-02 CCC-H 10 | 0 | 290172 | 523343 | 526720 | 113452906 | 0 | 4211550 | 0 | 519751 | 2910868 | 1806060 | 4429739 | 3951852 | 4203307 | 153719 | 289518 | 301436 | 952125 | 2100146 | 11053440 | 6632334 | 1163934 | 23743526 | 3304165 | 18727431 | 2348444 | 20658508 | 493383 | 0 | 0 | 1279916 | 0 | 207 | 5748985 | 819684 | 39821091 | 871 | 0 | 19088423 | 11549089 | 256409 | 33231224 | 4163821 | 31886679 | 89069 | 5653111 | 2099382 | 4562258 | 53680169 | 0 | 37459508 |  | Unknown
AB-230117-02 LFB-H 10 | 0 | 164708 | 338197 | 371151 | 65671209 | 0 | 4243422 | 0 | 497557 | 1516092 | 909153 | 2857946 | 2190806 | 2212065 | 158320 | 363146 | 417744 | 985252 | 2433931 | 13098976 | 7015382 | 1154004 | 29407136 | 3513343 | 22259117 | 2675897 | 24459732 | 474655 | 0 | 0 | 1319412 | 0 | 302 | 2927973 | 390545 | 24982872 | 696 | 39 | 11813877 | 5923381 | 204849 | 19498312 | 2520733 | 19062464 | 45627 | 2994708 | 1035459 | 2469719 | 28294699 | 1207 | 23566734 |  | Unknown
AB-230117-02 LD 9 | 0 | 17897 | 53592 | 7994 | 0 | 0 | 5428884 | 0 | 57753 | 1662 | 0 | 23938 | 0 | 306434 | 56838 | 568906 | 957577 | 317353 | 1834832 | 6972579 | 2457924 | 264953 | 33893145 | 4223106 | 18049984 | 3726813 | 25240310 | 49183 | 0 | 0 | 423819 | 1153 | 310 | 14358 | 0 | 41624 | 601 | 1514 | 0 | 0 | 152562 | 20532 | 3695 | 42711 | 7181 | 1031 | 0 | 4257 | 46272 | 645 | 29160 |  | Unknown
AB-230117-02 LRB 19 | 0 | 20459 | 28787 | 4756 | 0 | 0 | 5586287 | 0 | 494139 | 0 | 0 | 2365 | 1417 | 4320 | 172302 | 598899 | 703044 | 979270 | 3210571 | 16560517 | 7657476 | 1200068 | 39049072 | 4146167 | 27607743 | 3510414 | 31221498 | 475072 | 0 | 0 | 1364602 | 0 | 361 | 12275 | 0 | 57681 | 0 | 87 | 28202 | 30422 | 142486 | 26466 | 2099 | 91748 | 6970 | 850 | 0 | 8529 | 48293 | 0 | 35267 |  | Unknown
LRB-19-inj1 | 0 | 19613 | 29780 | 4606 | 0 | 668 | 5404909 | 0 | 484961 | 1366 | 0 | 2748 | 897 | 20673 | 159754 | 613166 | 881002 | 975650 | 3402141 | 16798319 | 7363776 | 1179640 | 40849895 | 4040698 | 28767903 | 3416155 | 32008275 | 462613 | 0 | 0 | 1333367 | 0 | 48 | 12497 | 0 | 59145 | 0 | 461 | 37173 | 32143 | 177734 | 23030 | 1270 | 96571 | 13930 | 1280 | 0 | 9718 | 46409 | 0 | 28132 |  | Unknown
WO-23003-180-A-1 | 0 | 20313 | 70873 | 5818 | 0 | 0 | 4942557 | 0 | 51502 | 0 | 0 | 20820 | 0 | 276511 | 53407 | 468558 | 1209231 | 266808 | 1270011 | 5428612 | 2117294 | 227330 | 26608435 | 3586635 | 12516657 | 3073816 | 18800561 | 41414 | 0 | 0 | 351343 | 0 | 338 | 9519 | 0 | 31590 | 0 | 1490 | 0 | 0 | 168473 | 16426 | 1484 | 30908 | 14922 | 1588 | 0 | 3428 | 38074 | 857 | 28712 |  | Unknown
WO-23003-179-A-1 | 0 | 13037 | 44232 | 6409 | 0 | 0 | 3958668 | 0 | 50030 | 1029 | 0 | 6214 | 0 | 81028 | 41144 | 176462 | 497361 | 235710 | 797813 | 4088557 | 1727550 | 209111 | 19752521 | 2471074 | 7635678 | 1761842 | 11761625 | 38951 | 0 | 0 | 306279 | 0 | 360 | 11686 | 0 | 25500 | 1291 | 1754 | 58461 | 0 | 164177 | 0 | 1599 | 20186 | 18132 | 589 | 0 | 1162 | 49869 | 966 | 24049 |  | Unknown
WO-23003-178-A-1 | 0 | 14936 | 57427 | 10542 | 0 | 0 | 5124041 | 0 | 61954 | 435 | 0 | 11370 | 1528 | 145760 | 50033 | 369986 | 961501 | 287900 | 1080273 | 5047723 | 2095557 | 244182 | 26034693 | 3434464 | 10682658 | 2808552 | 17221198 | 49961 | 1353 | 1450 | 378797 | 0 | 694 | 10818 | 0 | 28199 | 544 | 1793 | 65457 | 0 | 162124 | 13111 | 942 | 17024 | 18038 | 823 | 0 | 1599 | 39643 | 808 | 29092 |  | Unknown
WO-23003-177-A-1 | 0 | 17204 | 57413 | 8704 | 0 | 0 | 5755279 | 0 | 58342 | 0 | 0 | 19165 | 2039 | 233231 | 54269 | 467704 | 946314 | 304495 | 1708034 | 6642716 | 2401343 | 249877 | 35158130 | 4133434 | 16951043 | 3653857 | 24435144 | 47934 | 0 | 0 | 412953 | 0 | 228 | 14686 | 0 | 44867 | 0 | 2174 | 34101 | 0 | 171980 | 21597 | 2029 | 44959 | 14592 | 718 | 0 | 3868 | 53026 | 485 | 33476 |  | Unknown
WO-23003-176-A-1 | 0 | 15472 | 49559 | 2410 | 0 | 0 | 5743114 | 0 | 77044 | 3672 | 0 | 9862 | 0 | 109010 | 56337 | 453338 | 912424 | 334777 | 1633243 | 6725419 | 2623095 | 329025 | 34980436 | 4075895 | 16601625 | 3523101 | 24917770 | 62970 | 0 | 0 | 475018 | 0 | 180 | 11734 | 0 | 33908 | 0 | 1737 | 0 | 0 | 169438 | 17931 | 2979 | 43738 | 15268 | 611 | 0 | 4074 | 50657 | 987 | 42125 |  | Unknown
WO-23003-175-A-1 | 0 | 18904 | 94111 | 5666 | 0 | 1315 | 5819420 | 43 | 90208 | 1055 | 0 | 11159 | 1581 | 111070 | 66209 | 517187 | 987646 | 369722 | 1999689 | 7824120 | 2959853 | 363998 | 35315191 | 4183782 | 19082738 | 3626223 | 25711792 | 74350 | 0 | 0 | 511880 | 0 | 373 | 14006 | 0 | 35492 | 0 | 2312 | 11451 | 0 | 154626 | 22955 | 2801 | 46917 | 12396 | 1045 | 0 | 5371 | 49377 | 0 | 33902 |  | Unknown
WO-23003-174-A-1 | 0 | 20565 | 181962 | 3794 | 0 | 0 | 5982507 | 0 | 72251 | 2472 | 0 | 10316 | 2055 | 103769 | 61088 | 493560 | 987552 | 351613 | 1943904 | 7512317 | 2663445 | 312279 | 35919221 | 4338605 | 18862711 | 3765508 | 25850470 | 59001 | 0 | 0 | 468510 | 0 | 462 | 10536 | 0 | 42895 | 0 | 2285 | 16682 | 0 | 163282 | 21217 | 2378 | 46328 | 13249 | 1406 | 0 | 5654 | 54091 | 0 | 33864 |  | Unknown
WO-23003-173-A-1 | 0 | 21769 | 483769 | 6252 | 0 | 0 | 5668810 | 0 | 19145 | 508 | 0 | 8296 | 1649 | 104701 | 33451 | 317791 | 883639 | 167549 | 1318094 | 4821696 | 1406071 | 99326 | 33014680 | 3907669 | 13489980 | 3306529 | 21850161 | 14126 | 0 | 0 | 258711 | 0 | 298 | 12253 | 0 | 45647 | 902 | 1670 | 63934 | 0 | 158053 | 21392 | 2801 | 34207 | 14080 | 660 | 0 | 3935 | 51009 | 833 | 37383 |  | Unknown
WO-23003-172-A-1 | 0 | 20481 | 132575 | 9242 | 0 | 0 | 6051473 | 0 | 74304 | 2307 | 0 | 15109 | 1152 | 439001 | 70487 | 615949 | 1187780 | 372477 | 2023323 | 7784524 | 2828001 | 324177 | 33821093 | 4304216 | 19258794 | 3705559 | 25363940 | 61777 | 0 | 0 | 493198 | 0 | 644 | 13679 | 0 | 49182 | 1415 | 2554 | 34009 | 0 | 165134 | 23035 | 2734 | 31121 | 13503 | 687 | 0 | 6400 | 44404 | 1471 | 31103 |  | Unknown
WO-23003-171-A-1 | 0 | 20613 | 82471 | 7442 | 0 | 0 | 5805586 | 0 | 111450 | 0 | 0 | 15558 | 0 | 436563 | 80798 | 704555 | 1313062 | 417810 | 2132079 | 8466266 | 3174860 | 440752 | 32821029 | 4227071 | 20015689 | 3679650 | 24823141 | 97618 | 0 | 0 | 558713 | 0 | 983 | 9877 | 0 | 46321 | 0 | 3687 | 0 | 0 | 145423 | 20636 | 1852 | 53941 | 13538 | 802 | 0 | 8338 | 44892 | 1256 | 29169 |  | Unknown
AB-230117-02 LRB 18 | 0 | 19947 | 32473 | 3654 | 6477 | 0 | 5589208 | 0 | 502462 | 0 | 0 | 1231 | 1412 | 2376 | 170856 | 620570 | 720985 | 1016943 | 3250951 | 16438593 | 7600592 | 1207701 | 38512113 | 4161213 | 27888958 | 3514841 | 30466100 | 480642 | 1323 | 0 | 1380572 | 0 | 33 | 10016 | 0 | 61101 | 0 | 137 | 31803 | 25851 | 128273 | 20447 | 1103 | 91758 | 10133 | 1213 | 0 | 9227 | 52721 | 0 | 37121 |  | Unknown
LRB-18-inj1 | 0 | 22623 | 29539 | 4078 | 5787 | 0 | 5854183 | 0 | 489437 | 0 | 0 | 1988 | 1200 | 8296 | 172865 | 602302 | 753951 | 999773 | 3231707 | 16830928 | 7666312 | 1214027 | 40339726 | 4371654 | 28983106 | 3593951 | 31732010 | 465299 | 0 | 0 | 1407079 | 0 | 1099 | 12160 | 0 | 59669 | 881 | 198 | 0 | 26257 | 156584 | 20160 | 1767 | 96280 | 12830 | 1129 | 0 | 9391 | 58781 | 888 | 32785 |  | Unknown
WO-23003-170-A-1 | 0 | 16960 | 105726 | 8968 | 8496 | 0 | 5917529 | 0 | 58044 | 2020 | 0 | 7900 | 117 | 79463 | 57206 | 512977 | 978720 | 321024 | 1918476 | 7331982 | 2560727 | 269131 | 35655967 | 4409533 | 18979309 | 3701497 | 25869636 | 48476 | 3108 | 0 | 441216 | 0 | 201 | 14523 | 0 | 41185 | 0 | 2113 | 30190 | 0 | 157365 | 18150 | 3464 | 44012 | 12588 | 1473 | 0 | 5799 | 48169 | 862 | 33062 |  | Unknown
WO-23003-169-A-1 | 0 | 19089 | 117183 | 3395 | 8208 | 0 | 5645860 | 0 | 16061 | 705 | 0 | 9547 | 2200 | 102070 | 29215 | 457559 | 1085844 | 151875 | 1289295 | 4800114 | 1260188 | 87573 | 31308682 | 4052739 | 13380648 | 3482955 | 22045140 | 12005 | 0 | 0 | 230313 | 0 | 415 | 10630 | 0 | 41679 | 1004 | 1413 | 12200 | 0 | 147720 | 18126 | 4124 | 34607 | 13523 | 1592 | 0 | 1864 | 50188 | 717 | 25370 |  | Unknown
WO-23003-168-A-1 | 0 | 52389 | 282364 | 8779 | 8653 | 0 | 3206745 | 0 | 124553 | 2839 | 17220004 | 4564 | 0 | 19896 | 107894 | 1204166 | 1483243 | 338226 | 2389505 | 8665500 | 3560271 | 477327 | 30011092 | 4283816 | 19229820 | 3746978 | 22941001 | 107661 | 0 | 0 | 612076 | 0 | 433 | 9604 | 0 | 96993 | 0 | 3078 | 402077 | 15243 | 145602 | 46946 | 1276 | 62983 | 8701 | 1348 | 0 | 6217 | 42678 | 1122 | 16648 |  | Unknown
WO-23003-167-A-1 | 0 | 56219 | 969237 | 8778 | 12161 | 0 | 3076346 | 0 | 80943 | 0 | 15089772 | 6156 | 2911 | 20072 | 93333 | 916137 | 1450492 | 283386 | 2079503 | 7744759 | 2920180 | 336671 | 29415023 | 4147141 | 17747007 | 3630087 | 22371425 | 66326 | 0 | 0 | 515902 | 0 | 302 | 7574 | 0 | 104484 | 0 | 3714 | 371567 | 23801 | 146247 | 48884 | 2077 | 56837 | 14462 | 2577 | 0 | 7677 | 40186 | 936 | 15278 |  | Unknown
WO-23003-166-A-1 | 0 | 41102 | 559975 | 7125 | 14664 | 935 | 4871720 | 0 | 152127 | 0 | 18877176 | 8835 | 4833 | 20268 | 106965 | 996926 | 1396697 | 362297 | 2539693 | 9476008 | 3810986 | 563524 | 31878795 | 4264002 | 21890359 | 3901813 | 25407205 | 136720 | 0 | 0 | 681771 | 0 | 189 | 10612 | 0 | 100409 | 0 | 2168 | 453809 | 22010 | 158691 | 48274 | 3629 | 59592 | 14832 | 2415 | 0 | 6983 | 49384 | 0 | 24545 |  | Unknown
WO-23003-165-A-1 | 0 | 33724 | 82375 | 8942 | 17715 | 1363 | 5041041 | 0 | 85476 | 3530 | 14718012 | 9303 | 2541 | 20635 | 74201 | 890067 | 1408951 | 284546 | 2077889 | 7691472 | 2977046 | 357214 | 31033433 | 4129061 | 18802474 | 3702482 | 23474981 | 69809 | 0 | 0 | 520027 | 0 | 233 | 12497 | 0 | 77862 | 0 | 1829 | 371353 | 11265 | 165660 | 42184 | 4983 | 82315 | 15342 | 2908 | 0 | 6463 | 50118 | 0 | 32573 |  | Unknown
WO-23003-164-A-1 | 0 | 24282 | 77223 | 6590 | 24362 | 0 | 4415184 | 0 | 38712 | 2239 | 3876149 | 4691 | 1819 | 7928 | 52071 | 835361 | 1384845 | 222523 | 1648049 | 6145104 | 2117720 | 188997 | 28647728 | 3911273 | 15546125 | 3467978 | 20805999 | 33077 | 0 | 0 | 361499 | 0 | 651 | 11223 | 0 | 44639 | 843 | 2610 | 57158 | 0 | 161888 | 23630 | 2578 | 49552 | 10304 | 2873 | 0 | 5238 | 53165 | 0 | 42032 |  | Unknown
WO-23003-163-A-1 | 0 | 17339 | 65770 | 5194 | 39764 | 0 | 5348519 | 298 | 18531 | 2813 | 2732124 | 7087 | 1326 | 7003 | 32663 | 436918 | 1218578 | 168104 | 1319432 | 4843303 | 1435424 | 105682 | 30799128 | 3829879 | 13691119 | 3292263 | 21426804 | 16334 | 0 | 0 | 245372 | 0 | 355 | 15498 | 0 | 47181 | 874 | 1218 | 0 | 0 | 168267 | 24993 | 6727 | 41189 | 16016 | 2765 | 0 | 2306 | 70557 | 0 | 43443 |  | Unknown
WO-23003-162-A-1 | 0 | 21825 | 66990 | 7666 | 62632 | 0 | 5232667 | 0 | 61833 | 2326 | 5285361 | 9612 | 1860 | 10652 | 62408 | 753630 | 1325620 | 295461 | 1993288 | 7401610 | 2617183 | 290613 | 32557192 | 4343109 | 18864513 | 3892407 | 24824767 | 53537 | 0 | 0 | 465424 | 0 | 418 | 12234 | 0 | 103152 | 1237 | 3603 | 100563 | 0 | 167683 | 42326 | 4663 | 57217 | 15538 | 6911 | 0 | 5193 | 76178 | 0 | 55215 |  | Unknown
WO-23003-161-A-1 | 0 | 22787 | 76440 | 3586 | 152694 | 0 | 5088481 | 81 | 85480 | 2682 | 9288137 | 13156 | 3464 | 12946 | 80928 | 965444 | 1378803 | 320634 | 2091007 | 7897875 | 3033931 | 353953 | 30116754 | 4135400 | 18538399 | 3656138 | 23473716 | 69693 | 0 | 0 | 541326 | 0 | 118 | 22271 | 0 | 101707 | 0 | 2191 | 110312 | 0 | 149565 | 50707 | 9207 | 65287 | 18017 | 10847 | 0 | 4824 | 137458 | 0 | 83798 |  | Unknown
AB-230117-02 LFSM-H 8 | 0 | 74011 | 639663 | 840003 | 63919630 | 0 | 3612855 | 0 | 80115 | 582113 | 9507546 | 2988031 | 1650643 | 2593001 | 70695 | 614889 | 842819 | 314744 | 1896632 | 6956418 | 2890177 | 338215 | 21881749 | 3362595 | 15740897 | 2939093 | 18646299 | 67281 | 0 | 0 | 497745 | 0 | 630 | 2793355 | 65112 | 20508600 | 0 | 2162 | 6422955 | 2449174 | 221852 | 16017906 | 2846015 | 14060570 | 60033 | 3259519 | 296262 | 1153197 | 30150560 | 0 | 20674873 |  | Unknown
AB-230117-02 CCC-H 9 | 0 | 306484 | 565672 | 574419 | 114967475 | 0 | 3303441 | 0 | 530533 | 3108414 | 1879967 | 4782937 | 3958558 | 4171529 | 160801 | 299670 | 311580 | 997179 | 2169590 | 11147682 | 7007399 | 1220309 | 24904792 | 3456287 | 19631655 | 2417403 | 21703755 | 502223 | 2366 | 0 | 1384656 | 0 | 499 | 4606806 | 837726 | 42220888 | 1012 | 26 | 19688885 | 12187522 | 253031 | 34274579 | 4348208 | 33047919 | 79735 | 5797848 | 2165743 | 4761844 | 53697048 | 0 | 39196538 |  | Unknown
AB-230117-02 LFB-H 9 | 0 | 225283 | 415022 | 539995 | 71579149 | 0 | 4175842 | 0 | 575944 | 1944739 | 1170389 | 3159892 | 2554066 | 2666954 | 193231 | 422993 | 531671 | 1168510 | 2731349 | 14133441 | 8495625 | 1408230 | 32099934 | 3846696 | 24747417 | 3098356 | 26967994 | 546872 | 0 | 0 | 1664664 | 0 | 712 | 2957691 | 479426 | 27930035 | 0 | 177 | 12982365 | 7655089 | 206678 | 22393229 | 2947254 | 22038961 | 46676 | 3429393 | 1295998 | 2963244 | 31311722 | 886 | 26967713 |  | Unknown
AB-230117-02 LD 8 | 1611 | 25954 | 70832 | 8331 | 10559 | 912 | 4748298 | 0 | 46632 | 0 | 4681698 | 3760 | 0 | 5032 | 64320 | 701205 | 1287837 | 244540 | 1713812 | 6467909 | 2240979 | 220661 | 31064596 | 4233622 | 16793031 | 3731604 | 23159551 | 41236 | 0 | 0 | 387823 | 0 | 364 | 11429 | 0 | 76846 | 740 | 1760 | 55817 | 0 | 144366 | 30724 | 1657 | 47573 | 13061 | 1612 | 0 | 3577 | 45043 | 0 | 26800 |  | Unknown
AB-230117-02 LRB 17 | 0 | 24826 | 41324 | 3212 | 0 | 0 | 5462917 | 0 | 490612 | 0 | 0 | 2966 | 2972 | 5781 | 206021 | 788146 | 894505 | 1036892 | 3197912 | 16600735 | 8001582 | 1224859 | 38781650 | 4352404 | 28252225 | 3737957 | 32070431 | 472207 | 0 | 0 | 1408963 | 0 | 1094 | 15323 | 0 | 65644 | 0 | 135 | 60173 | 34258 | 132417 | 20214 | 1007 | 109993 | 12076 | 781 | 0 | 10364 | 49486 | 879 | 26536 |  | Unknown
LRB-17-inj1 | 0 | 27496 | 48338 | 5217 | 0 | 0 | 5912452 | 0 | 509488 | 0 | 0 | 4445 | 1791 | 6345 | 207061 | 872961 | 1025191 | 1050190 | 3379366 | 16961166 | 7902759 | 1252637 | 40057113 | 4798619 | 29850546 | 4215275 | 32933569 | 484708 | 0 | 0 | 1439985 | 0 | 547 | 11256 | 0 | 66179 | 0 | 318 | 36308 | 29797 | 177830 | 24722 | 2891 | 110538 | 13454 | 1097 | 0 | 7832 | 52510 | 880 | 32420 |  | Unknown
WO-23003-160-A-1 | 0 | 24192 | 97428 | 6833 | 0 | 0 | 4869433 | 0 | 69393 | 0 | 8007913 | 6435 | 1656 | 6165 | 75084 | 1114633 | 1656464 | 282179 | 1772186 | 6610913 | 2716039 | 292414 | 30127225 | 4210310 | 15260217 | 3469802 | 20522506 | 54321 | 0 | 0 | 451646 | 0 | 353 | 8607 | 0 | 74133 | 1499 | 2498 | 101063 | 23874 | 179169 | 20829 | 1123 | 39208 | 13684 | 1718 | 0 | 4194 | 45985 | 510 | 24830 |  | Unknown
WO-23003-159-A-1 | 0 | 18199 | 54811 | 8685 | 0 | 0 | 5813320 | 0 | 47925 | 0 | 7473240 | 5898 | 0 | 6418 | 58561 | 510795 | 1121185 | 260875 | 1670059 | 6774462 | 2368282 | 236671 | 35695339 | 4425844 | 17583008 | 3888346 | 25763344 | 38069 | 0 | 0 | 410528 | 0 | 661 | 14204 | 0 | 62563 | 891 | 1700 | 118306 | 0 | 178283 | 27650 | 3048 | 39182 | 13153 | 1343 | 0 | 4567 | 50806 | 0 | 29580 |  | Unknown
WO-23003-158-A-1 | 0 | 18669 | 54091 | 4822 | 0 | 0 | 5839522 | 0 | 108189 | 2999 | 3044505 | 6001 | 821 | 4161 | 75476 | 624740 | 1187744 | 386892 | 1987225 | 8038532 | 3121432 | 425090 | 36461695 | 4599970 | 19224864 | 3959551 | 26801827 | 92849 | 0 | 0 | 556363 | 1826 | 467 | 13345 | 0 | 318373 | 0 | 2067 | 85261 | 0 | 189222 | 21466 | 1523 | 44288 | 14856 | 946 | 0 | 5690 | 49863 | 1703 | 29593 |  | Unknown
WO-23003-157-A-1 | 0 | 16778 | 70621 | 3441 | 0 | 0 | 4730005 | 0 | 23847 | 638 | 1244427 | 4039 | 0 | 8379 | 36092 | 598311 | 1308848 | 165777 | 1257699 | 4863416 | 1513836 | 113209 | 27568727 | 3801571 | 12335038 | 3244749 | 19264744 | 15191 | 0 | 0 | 255885 | 0 | 339 | 9467 | 0 | 35938 | 1593 | 1734 | 10591 | 0 | 172878 | 17168 | 1255 | 32881 | 16553 | 1275 | 0 | 2634 | 46457 | 924 | 23984 |  | Unknown
WO-23003-156-A-1 | 0 | 8465 | 73909 | 5430 | 0 | 0 | 4607077 | 0 | 96035 | 0 | 2858671 | 8114 | 0 | 7770 | 84509 | 933303 | 1433286 | 377701 | 2226147 | 8307874 | 3178765 | 404392 | 31947500 | 4308448 | 19713169 | 3762476 | 24511254 | 82072 | 0 | 0 | 554647 | 0 | 487 | 10384 | 0 | 41384 | 0 | 2522 | 74313 | 0 | 162504 | 18913 | 1191 | 50223 | 12842 | 927 | 0 | 6944 | 42391 | 1730 | 25579 |  | Unknown
WO-23003-155-A-1 | 0 | 18984 | 56806 | 5907 | 0 | 1434 | 5745681 | 0 | 80135 | 3658 | 4609996 | 5391 | 1190 | 0 | 65083 | 540690 | 1078858 | 311653 | 1783363 | 7339596 | 2785854 | 334619 | 36089654 | 4463528 | 17934940 | 3803417 | 26608371 | 64013 | 0 | 0 | 488040 | 1636 | 456 | 12988 | 0 | 49169 | 0 | 2308 | 34952 | 0 | 177305 | 34513 | 1786 | 42917 | 17140 | 1505 | 0 | 6405 | 48169 | 1180 | 35323 |  | Unknown
WO-23003-154-A-1 | 0 | 19292 | 69099 | 5652 | 0 | 0 | 5591321 | 0 | 77779 | 3250 | 3987003 | 4813 | 1992 | 5138 | 66763 | 587517 | 1297606 | 310399 | 1614797 | 6881998 | 2742671 | 324559 | 33684806 | 4330562 | 16827970 | 3772794 | 24393253 | 63488 | 0 | 0 | 471730 | 0 | 219 | 11812 | 0 | 50777 | 0 | 2314 | 0 | 0 | 174011 | 173823 | 1430 | 46819 | 16439 | 1069 | 0 | 2600 | 47230 | 1098 | 26725 |  | Unknown
WO-23003-153-A-1 | 0 | 15489 | 48111 | 3590 | 4898 | 0 | 4749892 | 0 | 55849 | 0 | 2042070 | 5116 | 0 | 4158 | 50120 | 274975 | 837140 | 256724 | 969843 | 4818959 | 2031301 | 230281 | 25804341 | 3521984 | 9743480 | 2604457 | 15843986 | 45525 | 0 | 0 | 358758 | 0 | 452 | 11485 | 0 | 52909 | 628 | 1828 | 0 | 0 | 183103 | 18557 | 1791 | 24967 | 18237 | 1113 | 0 | 1894 | 55324 | 1332 | 29217 |  | Unknown
WO-23003-152-A-1 | 0 | 15366 | 46886 | 5006 | 0 | 1069 | 5371058 | 0 | 41256 | 2505 | 720672 | 4646 | 0 | 3685 | 46744 | 354280 | 912526 | 235400 | 1233668 | 5391073 | 1952820 | 189287 | 33483385 | 3767259 | 13012060 | 3345514 | 22355921 | 35316 | 0 | 0 | 337422 | 0 | 15 | 11140 | 0 | 46667 | 0 | 2135 | 0 | 0 | 180310 | 17818 | 4413 | 35985 | 16314 | 1304 | 0 | 3974 | 55874 | 0 | 28553 |  | Unknown
WO-23003-151-A-1 | 0 | 19257 | 47544 | 6173 | 24246 | 0 | 5270755 | 0 | 79245 | 0 | 921625 | 2622 | 0 | 9689 | 61429 | 391388 | 871021 | 328737 | 1307811 | 6057087 | 2494156 | 331070 | 34276127 | 3530487 | 13904521 | 3304266 | 23442267 | 64985 | 0 | 0 | 439616 | 0 | 247 | 14239 | 0 | 176764 | 0 | 1533 | 0 | 0 | 181496 | 24927 | 1650 | 50373 | 17180 | 1808 | 0 | 2479 | 66138 | 0 | 38360 |  | Unknown
AB-230117-02 LRB 16 | 0 | 25617 | 25258 | 3181 | 6149 | 0 | 4772346 | 0 | 329369 | 0 | 0 | 1611 | 1748 | 7824 | 152551 | 505547 | 581921 | 823464 | 2838192 | 15611997 | 6549995 | 848175 | 38439401 | 3360589 | 26693963 | 3350159 | 31526459 | 319065 | 0 | 0 | 1158028 | 0 | 787 | 12880 | 0 | 84008 | 0 | 784 | 25142 | 23480 | 132719 | 25155 | 3709 | 102717 | 9155 | 1412 | 0 | 7397 | 57002 | 739 | 38475 |  | Unknown
LRB-16-inj1 | 0 | 34633 | 26579 | 3279 | 5744 | 0 | 5131704 | 0 | 371747 | 5594 | 0 | 2333 | 589 | 7384 | 154118 | 579847 | 729806 | 785930 | 3153054 | 16541082 | 6366458 | 807509 | 43038656 | 3538721 | 28695723 | 3752352 | 34424377 | 356839 | 0 | 0 | 1105469 | 0 | 370 | 13662 | 0 | 82227 | 0 | 175 | 33639 | 22574 | 169831 | 28595 | 1780 | 95879 | 16268 | 1527 | 0 | 10960 | 61780 | 1178 | 47942 |  | Unknown
WO-23003-150-A-1 | 0 | 25999 | 48550 | 4610 | 6739 | 0 | 5402224 | 0 | 41987 | 0 | 772102 | 3593 | 2229 | 12904 | 54961 | 476900 | 1045410 | 226839 | 1624270 | 6436548 | 2039599 | 175331 | 37504622 | 3583959 | 17346607 | 3851752 | 27896754 | 31830 | 0 | 0 | 344916 | 0 | 205 | 17239 | 0 | 67162 | 0 | 2774 | 0 | 0 | 168347 | 29764 | 2597 | 50465 | 18144 | 961 | 0 | 3584 | 55651 | 0 | 31969 |  | Unknown
WO-23003-149-A-1 | 0 | 31756 | 74306 | 7934 | 0 | 0 | 4995140 | 0 | 42190 | 2371 | 714003 | 7170 | 455 | 0 | 53871 | 729774 | 1462005 | 233992 | 1703351 | 6632741 | 2007245 | 173026 | 33375145 | 3332020 | 17551789 | 3793251 | 26469640 | 33126 | 0 | 0 | 340903 | 0 | 451 | 12412 | 0 | 62182 | 0 | 2708 | 22728 | 0 | 165988 | 30888 | 4574 | 34272 | 15205 | 1579 | 0 | 5554 | 44561 | 0 | 33342 |  | Unknown
WO-23003-148-A-1 | 0 | 33242 | 71676 | 4193 | 7807 | 0 | 4922542 | 0 | 52706 | 3259 | 822590 | 6784 | 0 | 7467 | 60120 | 741259 | 1422064 | 242650 | 1794748 | 6992145 | 2224399 | 215368 | 33859088 | 3427961 | 18475522 | 3756816 | 26373254 | 42946 | 0 | 0 | 376745 | 1045 | 176 | 11789 | 0 | 62022 | 1517 | 549 | 26655 | 0 | 165293 | 32288 | 1776 | 49955 | 15657 | 1175 | 0 | 5834 | 50330 | 988 | 26221 |  | Unknown
WO-23003-147-A-1 | 0 | 30924 | 71512 | 4472 | 10219 | 0 | 5142605 | 0 | 64852 | 0 | 921139 | 8907 | 0 | 8118 | 65484 | 824335 | 1416512 | 273863 | 1906985 | 7518003 | 2453173 | 242382 | 35745303 | 3448670 | 19525177 | 3832871 | 27348352 | 49981 | 0 | 0 | 429523 | 0 | 174 | 13148 | 0 | 66734 | 0 | 3297 | 0 | 0 | 170796 | 29791 | 2720 | 60509 | 16224 | 1115 | 0 | 4259 | 55962 | 2079 | 26723 |  | Unknown
WO-23003-146-A-1 | 0 | 41385 | 82883 | 9759 | 13771 | 0 | 4132950 | 0 | 73875 | 0 | 1275663 | 7409 | 1396 | 9969 | 68778 | 1053999 | 1545661 | 295815 | 2056685 | 7941116 | 2652987 | 273558 | 32156990 | 3404050 | 20046846 | 3906282 | 25700237 | 61502 | 0 | 0 | 443958 | 0 | 778 | 10733 | 0 | 66596 | 0 | 2056 | 86259 | 0 | 160014 | 27436 | 3431 | 68942 | 14760 | 1095 | 0 | 4417 | 46385 | 0 | 36219 |  | Unknown
WO-23003-145-A-1 | 0 | 33084 | 75418 | 18477 | 18471 | 977 | 4784068 | 0 | 85705 | 0 | 1189624 | 8559 | 2345 | 9466 | 72646 | 815428 | 1493470 | 317013 | 2031906 | 8174231 | 2684127 | 285286 | 34635365 | 3507343 | 20234548 | 3890902 | 27672780 | 72268 | 0 | 0 | 458870 | 0 | 39 | 13539 | 0 | 73664 | 0 | 1682 | 0 | 0 | 168834 | 30509 | 2001 | 56555 | 15471 | 2486 | 0 | 6089 | 63483 | 0 | 33928 |  | Unknown
WO-23003-144-A-1 | 0 | 32643 | 50119 | 9115 | 26776 | 0 | 5449959 | 14 | 71134 | 0 | 220632 | 6891 | 0 | 10689 | 64226 | 512297 | 997226 | 297135 | 1800955 | 7520719 | 2550021 | 259757 | 40148666 | 3600614 | 19453429 | 4008874 | 30179345 | 56921 | 0 | 1296 | 430289 | 0 | 989 | 16926 | 0 | 94491 | 0 | 1947 | 33772 | 0 | 177620 | 34045 | 4223 | 57860 | 18222 | 3894 | 0 | 7134 | 70804 | 0 | 48991 |  | Unknown
WO-23003-143-A-1 | 0 | 30072 | 49826 | 3394 | 40609 | 0 | 5507041 | 0 | 67125 | 3961 | 175736 | 9048 | 1431 | 11692 | 63220 | 535788 | 1024663 | 295991 | 1912046 | 7728237 | 2412260 | 258218 | 40160262 | 3553445 | 19532034 | 3855879 | 29438075 | 54078 | 0 | 0 | 411477 | 0 | 835 | 13077 | 0 | 135017 | 0 | 2482 | 9602 | 0 | 184242 | 36691 | 4147 | 56924 | 17961 | 2552 | 0 | 5635 | 78639 | 0 | 51888 |  | Unknown
WO-23003-142-A-1 | 0 | 40119 | 105636 | 5778 | 69571 | 0 | 4303242 | 0 | 53273 | 0 | 238490 | 9350 | 0 | 9341 | 68546 | 1213708 | 1857096 | 269680 | 1970255 | 7378311 | 2269744 | 226409 | 33259350 | 3742151 | 18317393 | 3757708 | 23460827 | 43086 | 1484 | 0 | 393507 | 0 | 724 | 15336 | 0 | 152427 | 0 | 2769 | 41147 | 16232 | 173756 | 42580 | 9146 | 55478 | 16163 | 4753 | 0 | 4451 | 87557 | 0 | 51764 |  | Unknown
WO-23003-141-A-1 | 0 | 33862 | 73455 | 7331 | 202024 | 0 | 4012351 | 0 | 15985 | 0 | 98896 | 13572 | 0 | 13210 | 36234 | 609994 | 1365222 | 149370 | 1267875 | 4997696 | 1333669 | 75888 | 29934648 | 2967222 | 13563890 | 3333604 | 22187134 | 12317 | 0 | 0 | 225755 | 0 | 376 | 15725 | 0 | 137878 | 0 | 3093 | 60250 | 0 | 159210 | 64726 | 9620 | 66169 | 17581 | 8538 | 0 | 2978 | 121292 | 0 | 68336 |  | Unknown
AB-230117-02 LFSM-H 7 | 0 | 77033 | 574253 | 845422 | 64946442 | 0 | 3055770 | 0 | 35356 | 364126 | 356101 | 2931477 | 1413249 | 2739312 | 46555 | 564917 | 817266 | 219237 | 1618655 | 6120918 | 1886454 | 151982 | 24540149 | 2659010 | 15199777 | 3058586 | 20777930 | 29375 | 0 | 0 | 326061 | 1576 | 600 | 2370424 | 30329 | 22725019 | 0 | 1579 | 5600255 | 1626107 | 207792 | 18101405 | 3019053 | 13862555 | 63331 | 2684124 | 0 | 828444 | 25706787 | 0 | 18496056 |  | Unknown
AB-230117-02 CCC-H 8 | 0 | 283634 | 540834 | 571652 | 112230191 | 0 | 3094138 | 0 | 397317 | 2619326 | 1402414 | 4929560 | 4048473 | 4314371 | 142821 | 294559 | 299509 | 793697 | 2157964 | 12328264 | 6025403 | 809826 | 28531073 | 2863416 | 21121392 | 2383178 | 24578455 | 380122 | 0 | 841 | 1095340 | 1343 | 531 | 4369011 | 641387 | 48461870 | 0 | 158 | 22191139 | 10407726 | 256513 | 40138544 | 4324048 | 36438473 | 90069 | 5011331 | 0 | 4411997 | 55940087 | 0 | 44911171 |  | Unknown
AB-230117-02 LFB-H 8 | 0 | 188692 | 370761 | 477303 | 65860546 | 0 | 4175234 | 0 | 393779 | 1257647 | 702162 | 2974567 | 2374829 | 2524445 | 160540 | 401129 | 522896 | 766022 | 2647643 | 14439857 | 6265642 | 844149 | 35223208 | 3148685 | 24434056 | 3035404 | 29190429 | 373163 | 0 | 0 | 1103741 | 0 | 352 | 2898271 | 304173 | 29820113 | 1037 | 508 | 12869591 | 5444796 | 220698 | 23541569 | 2822259 | 21261067 | 48319 | 2692329 | 0 | 2391425 | 30845357 | 0 | 26687083 |  | Unknown
AB-230117-02 LD 7 | 0 | 37694 | 99384 | 4779 | 7757 | 0 | 4265080 | 0 | 25222 | 0 | 151098 | 6204 | 0 | 0 | 44592 | 919188 | 1734024 | 171470 | 1522140 | 5446739 | 1583020 | 103016 | 30179486 | 3487801 | 14327453 | 3377280 | 20794686 | 17781 | 0 | 0 | 258314 | 0 | 434 | 10583 | 0 | 179349 | 1198 | 629 | 20106 | 13626 | 150196 | 23302 | 2912 | 40808 | 13323 | 1198 | 0 | 2389 | 47071 | 0 | 34956 |  | Unknown
AB-230117-02 LRB 15 | 0 | 36561 | 29706 | 3350 | 0 | 900 | 4501277 | 0 | 311627 | 4308 | 0 | 2372 | 0 | 7481 | 156236 | 493514 | 607056 | 725349 | 2627401 | 14095219 | 5947417 | 712231 | 35095576 | 2922875 | 23444026 | 3163740 | 29548719 | 302709 | 0 | 0 | 1054039 | 0 | 270 | 13232 | 0 | 223796 | 0 | 319 | 0 | 25794 | 119012 | 28448 | 2146 | 79462 | 13375 | 1174 | 0 | 6263 | 48352 | 496 | 37811 |  | Unknown
LRB-15-inj1 | 0 | 35677 | 24942 | 2977 | 8069 | 0 | 5368014 | 0 | 369245 | 2553 | 0 | 677 | 0 | 7071 | 169071 | 603724 | 806236 | 776399 | 3163272 | 16402500 | 6363611 | 833703 | 42072343 | 3527219 | 28246415 | 3750497 | 33904567 | 353297 | 0 | 0 | 1122844 | 968 | 664 | 14291 | 0 | 295261 | 0 | 580 | 28963 | 27899 | 165345 | 33895 | 2643 | 106634 | 17126 | 1247 | 0 | 8153 | 59646 | 0 | 41605 |  | Unknown
WO-23003-140-A-1 | 0 | 26785 | 58847 | 3020 | 0 | 919 | 5629488 | 0 | 60590 | 2213 | 239678 | 7245 | 0 | 7782 | 58816 | 612534 | 1176296 | 269934 | 1776710 | 7167007 | 2241844 | 228036 | 36342071 | 3473055 | 18299776 | 3732851 | 27809792 | 50264 | 0 | 2062 | 386869 | 0 | 111 | 12385 | 0 | 360866 | 0 | 2617 | 0 | 0 | 164935 | 31433 | 2528 | 46615 | 13623 | 1394 | 0 | 4089 | 55989 | 955 | 33577 |  | Unknown
WO-23003-139-A-1B | 0 | 29012 | 51109 | 3933 | 13611 | 0 | 5394826 | 0 | 56577 | 6464 | 211700 | 7344 | 1874 | 25846 | 55362 | 526887 | 918806 | 247705 | 1842386 | 7243909 | 2194892 | 203238 | 35046837 | 3279201 | 18277425 | 3553271 | 27124738 | 47777 | 0 | 0 | 393619 | 0 | 654 | 12674 | 0 | 1076143 | 0 | 2592 | 0 | 0 | 126606 | 31596 | 6511 | 48481 | 14010 | 1439 | 0 | 5891 | 55714 | 0 | 51714 |  | Unknown
WO-23003-139-A-1 | 0 | 31677 | 78099 | 2887 | 56842 | 0 | 5594901 | 0 | 88805 | 9909 | 165071 | 4129 | 0 | 23235 | 55825 | 434235 | 818506 | 200900 | 1893044 | 7669561 | 2565629 | 223068 | 31886772 | 3793322 | 18606048 | 3544326 | 26342001 | 75063 | 0 | 0 | 501825 | 0 | 296 | 17674 | 0 | 20941273 | 1125 | 665 | 33655 | 0 | 133668 | 58013 | 7003 | 78779 | 7423 | 5750 | 0 | 11430 | 73011 | 745 | 64482 |  | Unknown
Rinse12 | 0 | 69108 | 38289 | 1611 | 44362 | 526 | 403553 | 0 | 14699 | 4833 | 0 | 3768 | 3206 | 8803 | 3058 | 24669 | 54053 | 6697 | 124329 | 448804 | 100284 | 8211 | 2431795 | 293562 | 1287286 | 265578 | 2143667 | 12916 | 1743 | 0 | 18818 | 0 | 112 | 4649 | 0 | 46587122 | 0 | 0 | 38972 | 0 | 29102 | 59619 | 4323 | 72330 | 8497 | 6410 | 0 | 12964 | 37454 | 1177 | 48108 | C | Unknown
WO-23003-203-A-1B | 1217 | 47274 | 26042 | 10615 | 21734 | 0 | 5035644 | 0 | 520821 | 1827 | 54088 | 4246293 | 312484 | 17162755 | 189706 | 375335 | 719115 | 965643 | 3062842 | 16186626 | 7261125 | 1091433 | 38456781 | 1705287 | 25456947 | 1662963 | 17509722 | 500203 | 0 | 0 | 1342376 | 0 | 46 | 11755 | 0 | 73952 | 0 | 0 | 141859 | 30330 | 152455 | 18826 | 9494 | 589947 | 15896 | 20671 | 0 | 12169 | 62247 | 796 | 51875 |  | Unknown
WO-23003-208-A-1B | 0 | 50776 | 30813 | 5820 | 6824 | 1010 | 5449984 | 0 | 468018 | 5653 | 53138 | 4509247 | 323889 | 17762006 | 173990 | 381752 | 744333 | 936372 | 3246724 | 16632516 | 7138306 | 1048049 | 39806308 | 1792928 | 26469318 | 1709606 | 18251157 | 449514 | 0 | 0 | 1264936 | 0 | 78 | 12573 | 0 | 64591 | 0 | 512 | 137623 | 25757 | 145826 | 16946 | 9986 | 532428 | 15794 | 18041 | 0 | 8482 | 53437 | 1376 | 41182 |  | Unknown
WO-23003-207-A-1B | 1199 | 63726 | 43149 | 11019 | 12344 | 0 | 5229263 | 0 | 488640 | 0 | 57334 | 4736971 | 339119 | 18623889 | 169531 | 369012 | 714456 | 955947 | 3170272 | 16481551 | 7203777 | 1064852 | 39965297 | 1736488 | 26394783 | 1638880 | 17579970 | 465880 | 0 | 1658 | 1261198 | 0 | 125 | 12554 | 0 | 69410 | 0 | 8 | 136627 | 28185 | 170601 | 22137 | 10276 | 544293 | 15528 | 20264 | 0 | 7930 | 60092 | 1346 | 35236 |  | Unknown
WO-23003-206-A-1B | 0 | 46003 | 28949 | 4921 | 9961 | 0 | 5277615 | 0 | 517899 | 2866 | 61629 | 6158966 | 391191 | 22867936 | 180495 | 328000 | 692831 | 982724 | 3234755 | 17195748 | 7396823 | 1065695 | 38969513 | 1474619 | 25605917 | 1336272 | 14106716 | 495795 | 1481 | 0 | 1330125 | 0 | 352 | 13542 | 0 | 69182 | 1332 | 210 | 146637 | 28185 | 170843 | 17544 | 14533 | 644270 | 17091 | 16978 | 0 | 5924 | 59071 | 917 | 50468 |  | Unknown
WO-23003-205-A-1B | 0 | 50932 | 131176 | 14087 | 13320 | 0 | 5034118 | 0 | 315303 | 1835 | 163867 | 11344895 | 1223008 | 29055816 | 124021 | 231073 | 844695 | 646357 | 2759228 | 13341459 | 4933200 | 775866 | 35260483 | 1129694 | 18925067 | 1005913 | 9179316 | 295207 | 0 | 0 | 865937 | 0 | 173 | 11906 | 0 | 61895 | 999 | 882 | 183530 | 33918 | 149655 | 0 | 22286 | 741288 | 14662 | 17489 | 0 | 12026 | 48120 | 481 | 35573 |  | Unknown
WO-23003-204-A-1B | 0 | 37706 | 43674 | 10397 | 16622 | 1286 | 5304768 | 0 | 330524 | 1732 | 171614 | 11563289 | 1224814 | 30353372 | 133075 | 223653 | 887806 | 668503 | 2797550 | 13671055 | 5203349 | 807996 | 36292475 | 1261782 | 19003081 | 973200 | 9018587 | 317187 | 0 | 0 | 918572 | 0 | 127 | 14053 | 0 | 58739 | 833 | 544 | 251246 | 44746 | 168461 | 0 | 17724 | 899930 | 14707 | 15121 | 0 | 14517 | 46377 | 1058 | 36795 |  | Unknown
LRB-22-reprep-inj2 | 0 | 38167 | 47778 | 8019 | 28191 | 534 | 4117270 | 0 | 586108 | 1678 | 0 | 6515 | 1812 | 54856 | 185799 | 487393 | 547399 | 1013485 | 2515105 | 13811734 | 7210561 | 1106935 | 30900382 | 1775117 | 23303576 | 2794056 | 25322946 | 572685 | 0 | 0 | 1356743 | 0 | 249 | 12724 | 0 | 63395 | 0 | 126 | 47965 | 16745 | 108165 | 40019 | 1732 | 278282 | 13739 | 22361 | 0 | 11880 | 49133 | 807 | 59138 |  | Unknown
LRB-22-reprep | 0 | 46681 | 48010 | 10821 | 46088 | 0 | 4861937 | 0 | 602141 | 0 | 0 | 7036 | 2998 | 65300 | 193159 | 517580 | 582015 | 1038861 | 2908422 | 15865842 | 7867451 | 1143671 | 35010831 | 2026559 | 26885937 | 3077562 | 28422575 | 587046 | 0 | 0 | 1501253 | 0 | 191 | 14361 | 0 | 74148 | 0 | 231 | 55189 | 15968 | 126285 | 45443 | 2004 | 333638 | 16239 | 23096 | 0 | 8553 | 72411 | 611 | 64068 |  | Unknown
217288-d4-reprep | 0 | 25116 | 90144 | 8774 | 5972 | 924 | 5841501 | 0 | 173166 | 2389 | 10483422 | 14464 | 2776 | 106021 | 84346 | 776789 | 1276571 | 393326 | 2579520 | 10548765 | 3761331 | 546421 | 33642603 | 2810267 | 24106931 | 3658609 | 27874542 | 157464 | 0 | 0 | 650768 | 0 | 270 | 8955 | 0 | 97569 | 0 | 1022 | 225873 | 13843 | 149657 | 40414 | 1117 | 107384 | 16273 | 11139 | 0 | 8231 | 46063 | 1332 | 27299 |  | Unknown
217289-d4-reprep | 984 | 25491 | 102093 | 6685 | 6408 | 0 | 6090198 | 0 | 183091 | 0 | 11259640 | 12745 | 2776 | 63669 | 93373 | 829328 | 1237734 | 421121 | 2727214 | 11279977 | 3904277 | 606199 | 37427258 | 3215678 | 25204572 | 3967468 | 28154739 | 168253 | 0 | 0 | 695262 | 0 | 95 | 13646 | 0 | 70233 | 0 | 1819 | 225909 | 11936 | 165897 | 32560 | 3128 | 101759 | 18200 | 10049 | 0 | 7395 | 51852 | 927 | 39689 |  | Unknown
217295-d1000-LD-reprep | 696 | 20883 | 37335 | 4391 | 0 | 0 | 6272970 | 0 | 665964 | 1827 | 0 | 593721 | 19486 | 13785228 | 219005 | 605430 | 812553 | 1318561 | 3518786 | 18814017 | 9513170 | 1431177 | 42625232 | 3040714 | 30393908 | 1986992 | 21408323 | 641965 | 0 | 0 | 1798305 | 0 | 685 | 12102 | 0 | 230246 | 0 | 371 | 41469 | 13264 | 171099 | 18479 | 41942 | 68628 | 16383 | 1046 | 0 | 9763 | 52855 | 0 | 44600 |  | Unknown
217296-d1000-LFSM-200-reprep | 1991 | 199052 | 330968 | 416262 | 66925428 | 0 | 4567584 | 0 | 576390 | 1843876 | 1114612 | 3209818 | 2358648 | 14344332 | 180419 | 341907 | 428412 | 1113075 | 2384164 | 13926006 | 7752940 | 1181855 | 28143191 | 2543040 | 21920330 | 1563970 | 16594272 | 556150 | 0 | 0 | 1449487 | 0 | 261 | 3316007 | 503731 | 26331193 | 0 | 348 | 13244952 | 7308955 | 199059 | 14122256 | 2465503 | 19783496 | 48641 | 2442402 | 1153771 | 2706163 | 27604180 | 684 | 24509778 |  | Unknown
217315-d2-reprep | 0 | 20725 | 71577 | 6322 | 6503 | 0 | 5917056 | 0 | 63208 | 0 | 0 | 7349 | 1369 | 40863 | 56351 | 508385 | 1029380 | 315346 | 1848991 | 7099238 | 2492778 | 267490 | 35013426 | 2894322 | 19174601 | 3730780 | 25701650 | 52733 | 0 | 0 | 398883 | 0 | 547 | 10980 | 39223 | 41662 | 934 | 2365 | 79003 | 0 | 156992 | 27076 | 3305 | 66067 | 14281 | 8693 | 0 | 5990 | 50405 | 1053 | 52289 |  | Unknown
217335-d2-LD-reprep | 1252 | 21156 | 71679 | 10688 | 0 | 0 | 5853497 | 0 | 49101 | 0 | 0 | 17196 | 523 | 1083104 | 62092 | 534687 | 960509 | 305429 | 1888781 | 7420324 | 2367644 | 225857 | 33019176 | 2852319 | 19509902 | 3588274 | 25665879 | 39079 | 0 | 0 | 398711 | 0 | 637 | 8444 | 93208 | 165617 | 1780 | 3306 | 19139 | 0 | 171100 | 27821 | 3533 | 70810 | 10500 | 8026 | 0 | 4155 | 46380 | 830 | 32456 |  | Unknown
217336-d2-LFSM-200-reprep | 0 | 45298 | 331633 | 530673 | 59872634 | 0 | 4068641 | 0 | 12256 | 238309 | 138888 | 2291264 | 1034837 | 3114858 | 26388 | 286281 | 528893 | 135649 | 1150272 | 4144848 | 1164473 | 79394 | 21729644 | 1999570 | 11821768 | 2616570 | 16744072 | 9796 | 0 | 0 | 201239 | 0 | 302 | 3077124 | 32977 | 20732272 | 803 | 1314 | 3911811 | 1069112 | 209474 | 14521447 | 2675781 | 10344228 | 55875 | 2598330 | 73737 | 583376 | 26119912 | 981 | 20291816 |  | Unknown
217356-d2-reprep | 0 | 20069 | 82053 | 15334 | 3814 | 0 | 5870232 | 8 | 76680 | 875 | 0 | 4803 | 1412 | 31725 | 68036 | 653299 | 1069240 | 362507 | 2077732 | 8096076 | 2734706 | 308952 | 35078474 | 2845001 | 21287780 | 3725669 | 26439273 | 64045 | 0 | 0 | 488862 | 0 | 559 | 12986 | 0 | 736551 | 946 | 4141 | 18245 | 6113 | 151352 | 28435 | 3570 | 74310 | 17170 | 11041 | 0 | 4985 | 45274 | 613 | 34940 |  | Unknown
217357-d2-reprep | 0 | 32052 | 75356 | 44566 | 5801 | 0 | 5244048 | 0 | 67015 | 2362 | 0 | 4165 | 1693 | 32992 | 65595 | 604243 | 1001574 | 337621 | 1990548 | 7773840 | 2671781 | 296061 | 33193268 | 2849331 | 19919468 | 3765206 | 26011235 | 54807 | 0 | 0 | 453180 | 0 | 230 | 11121 | 0 | 7868034 | 0 | 3284 | 18296 | 0 | 164534 | 100436 | 1901 | 75649 | 15712 | 8492 | 0 | 5553 | 42301 | 0 | 41112 |  | Unknown
217358-d2-reprep | 0 | 23677 | 72136 | 24997 | 0 | 0 | 4729105 | 0 | 79584 | 0 | 0 | 10085 | 2848 | 24194 | 65171 | 566785 | 342994 | 358746 | 1989970 | 7845774 | 2800127 | 333002 | 11191331 | 2870189 | 20363549 | 3693069 | 26309173 | 66583 | 0 | 0 | 489764 | 0 | 798 | 10327 | 0 | 141319505 | 621 | 4149 | 0 | 0 | 163871 | 134025 | 3024 | 90072 | 15353 | 7193 | 0 | 4477 | 41984 | 1018 | 69786 |  | Unknown
217359-d2-LD | 971 | 17870 | 66847 | 8947 | 0 | 0 | 3414232 | 0 | 65162 | 0 | 0 | 3112 | 837 | 12195 | 61072 | 440804 | 214548 | 310530 | 1734908 | 7024500 | 2495623 | 280112 | 5267128 | 3279024 | 17329609 | 3248796 | 23961101 | 52260 | 0 | 0 | 423000 | 0 | 461 | 11501 | 0 | 225954508 | 0 | 2149 | 20654 | 0 | 141407 | 296928 | 595 | 96978 | 17439 | 744 | 0 | 4925 | 47573 | 1027 | 134731 |  | Unknown
217359-d2-reprep | 0 | 18373 | 65981 | 17265 | 0 | 0 | 3297853 | 0 | 80302 | 0 | 0 | 5443 | 1556 | 22622 | 63996 | 496533 | 254931 | 371808 | 1898308 | 7682663 | 2777479 | 336158 | 5216648 | 2893285 | 19783128 | 3700322 | 27355959 | 65076 | 0 | 0 | 490242 | 0 | 443 | 13916 | 0 | 253317030 | 0 | 2051 | 40984 | 0 | 162467 | 369067 | 2577 | 135285 | 18232 | 8939 | 0 | 5175 | 47985 | 1078 | 153944 |  | Unknown
217360-d2-LFSM-200 | 0 | 51737 | 386323 | 499128 | 65617849 | 0 | 4506121 | 0 | 50386 | 464645 | 293350 | 2807931 | 1420577 | 2570668 | 49901 | 339709 | 493857 | 299353 | 1637116 | 6241613 | 2189551 | 228744 | 21867262 | 2969746 | 15306684 | 2838369 | 20095160 | 41412 | 0 | 0 | 387900 | 0 | 774 | 3289355 | 41315 | 41189890 | 0 | 2566 | 5937270 | 2024299 | 209091 | 17229779 | 2461164 | 14150233 | 61962 | 2941876 | 218097 | 917692 | 28022948 | 937 | 22066640 |  | Unknown
217360-d2-reprep | 860 | 20584 | 86535 | 9732 | 0 | 0 | 5448112 | 70 | 47413 | 0 | 0 | 4708 | 1281 | 21560 | 54102 | 585753 | 788245 | 289713 | 1758120 | 6538550 | 2278169 | 216977 | 24984103 | 2875732 | 17853311 | 3663221 | 23347612 | 36267 | 967 | 0 | 381007 | 0 | 661 | 12630 | 0 | 46315496 | 791 | 2628 | 33153 | 0 | 155210 | 27133 | 1220 | 70856 | 13699 | 7236 | 0 | 5296 | 45437 | 730 | 42640 |  | Unknown
217375-d2-LD-reprep | 0 | 25384 | 390526 | 26969 | 3790 | 0 | 5411978 | 0 | 39026 | 0 | 0 | 3022 | 777 | 33629 | 50156 | 401069 | 942737 | 252033 | 1396354 | 5781199 | 2050127 | 199392 | 30356496 | 2682278 | 15040827 | 3283182 | 21663026 | 32562 | 0 | 0 | 353343 | 0 | 436 | 12187 | 0 | 4245281 | 560 | 3347 | 18647 | 9205 | 157693 | 27408 | 1103 | 75167 | 13059 | 7257 | 0 | 4041 | 47983 | 0 | 31763 |  | Unknown
217376-d2-LFSM-200-reprep | 0 | 65564 | 700419 | 558954 | 71558837 | 0 | 4773089 | 0 | 56985 | 567072 | 355833 | 3050272 | 1975287 | 2746035 | 63670 | 437084 | 567229 | 384978 | 2245190 | 9294029 | 2947540 | 272169 | 25146296 | 2431843 | 20267034 | 3180995 | 22022785 | 47681 | 0 | 0 | 513604 | 0 | 461 | 3400935 | 51364 | 32891947 | 0 | 4066 | 8276010 | 2544314 | 198470 | 18045981 | 2684779 | 16832591 | 52201 | 3037869 | 245830 | 1343814 | 26256051 | 0 | 21917322 |  | Unknown
217393-d4-reprep | 0 | 24587 | 100701 | 24049 | 0 | 0 | 5640495 | 0 | 183676 | 0 | 0 | 2823 | 1524 | 19025 | 99243 | 812677 | 590817 | 543343 | 2761566 | 11319101 | 4115075 | 582291 | 18289200 | 2968573 | 25726568 | 3815020 | 29222054 | 167116 | 0 | 1465 | 714264 | 0 | 1021 | 13864 | 0 | 95234437 | 0 | 1593 | 103059 | 15287 | 165174 | 385799 | 1237 | 500710 | 14230 | 9251 | 0 | 7239 | 46953 | 0 | 33570 |  | Unknown
217394-d4-reprep | 0 | 23761 | 78999 | 15420 | 0 | 0 | 5491444 | 0 | 179021 | 2159 | 0 | 3568 | 1171 | 15388 | 91400 | 664419 | 474956 | 526280 | 2660302 | 11534145 | 3908243 | 571202 | 16523189 | 3013250 | 26064711 | 3844497 | 30325157 | 164144 | 0 | 0 | 686028 | 0 | 386 | 15099 | 0 | 114609128 | 0 | 810 | 136162 | 23132 | 167818 | 552495 | 1636 | 615378 | 15394 | 9407 | 0 | 7571 | 44258 | 0 | 37063 |  | Unknown
217395-d4-LD-reprep | 838 | 22985 | 75246 | 8174 | 0 | 0 | 5095457 | 0 | 174845 | 0 | 0 | 7080 | 616 | 15865 | 93252 | 704026 | 377820 | 546714 | 2715983 | 11584406 | 3962948 | 566562 | 11471986 | 2855096 | 26488218 | 3864173 | 30315119 | 162917 | 0 | 0 | 686795 | 1317 | 681 | 12491 | 0 | 146680128 | 442 | 1469 | 180267 | 29448 | 165754 | 1165027 | 1345 | 1180013 | 14110 | 11149 | 0 | 8521 | 50782 | 0 | 38085 |  | Unknown
217395-d4-reprep | 0 | 20327 | 60736 | 10125 | 0 | 0 | 4851454 | 0 | 183884 | 515 | 0 | 2363 | 1137 | 14219 | 97274 | 671870 | 354150 | 528791 | 2784250 | 11700873 | 4108574 | 581720 | 11213632 | 3077095 | 26148752 | 3838100 | 30556450 | 167429 | 0 | 0 | 707509 | 0 | 235 | 12535 | 0 | 149269121 | 945 | 1328 | 211675 | 30759 | 163359 | 1167012 | 1060 | 1198362 | 15868 | 8321 | 0 | 7380 | 44033 | 745 | 40267 |  | Unknown
217396-d200-LFSM-200 | 768 | 206007 | 356184 | 237947 | 65496904 | 0 | 3821865 | 0 | 597045 | 1739541 | 1101227 | 2817359 | 2304057 | 2357196 | 198069 | 363751 | 199307 | 1163052 | 2649285 | 14478970 | 8146487 | 1222092 | 8403033 | 2527858 | 23886278 | 2716368 | 25767847 | 574760 | 0 | 0 | 1554299 | 0 | 195 | 2508156 | 479586 | 166477566 | 0 | 322 | 12954911 | 7106511 | 224430 | 21266355 | 970107 | 20081097 | 56253 | 3028653 | 1070801 | 2659873 | 30255459 | 651 | 24175109 |  | Unknown
217399-d2-reprep | 436 | 24780 | 79046 | 16361 | 2733 | 0 | 4784404 | 0 | 73560 | 0 | 23475 | 3918 | 0 | 14221 | 75456 | 867722 | 1302752 | 368700 | 1942365 | 7573260 | 2773002 | 304273 | 27407650 | 2628133 | 18007646 | 3282239 | 21440348 | 60835 | 1499 | 0 | 445606 | 0 | 560 | 8737 | 0 | 581372 | 0 | 3394 | 84409 | 0 | 133264 | 21631 | 900 | 74854 | 15246 | 10291 | 0 | 6211 | 34634 | 0 | 20985 |  | Unknown
217400-d2-reprep | 632 | 22957 | 81470 | 13331 | 0 | 0 | 5367379 | 0 | 68656 | 0 | 24767 | 9145 | 0 | 12985 | 85595 | 1056779 | 1451839 | 387634 | 2705573 | 10805928 | 3286410 | 308810 | 31039101 | 2832429 | 23929613 | 3662110 | 24811018 | 57369 | 0 | 0 | 543005 | 0 | 501 | 10580 | 0 | 298312 | 0 | 1818 | 74290 | 0 | 143622 | 26057 | 1539 | 102749 | 13292 | 9435 | 0 | 9733 | 35047 | 0 | 20493 |  | Unknown
217415-d2-LD-reprep | 808 | 25656 | 78343 | 4050 | 0 | 0 | 4155058 | 0 | 46103 | 1320 | 206464 | 7465 | 0 | 11063 | 60766 | 805962 | 1242937 | 291605 | 1720166 | 6871050 | 2263857 | 205106 | 26651054 | 3222867 | 16258879 | 3296585 | 20903479 | 37833 | 0 | 0 | 374812 | 0 | 443 | 6241 | 0 | 1381459 | 0 | 3889 | 94703 | 0 | 133017 | 23718 | 1170 | 51414 | 13381 | 628 | 0 | 5660 | 31225 | 0 | 16381 |  | Unknown
217416-d2-LFSM-200-reprep | 0 | 83241 | 788735 | 1003138 | 62131940 | 0 | 3782519 | 0 | 47699 | 456242 | 550463 | 2852845 | 1598869 | 2454321 | 53730 | 741540 | 986325 | 301510 | 1743168 | 6376900 | 2212125 | 223481 | 20335293 | 2868447 | 13558401 | 2693800 | 15469898 | 38456 | 0 | 0 | 389577 | 0 | 752 | 2787936 | 42059 | 19451561 | 0 | 2641 | 6051152 | 2028848 | 200643 | 13403418 | 2584366 | 12631617 | 60323 | 2844100 | 210802 | 1013276 | 25704393 | 0 | 17533860 |  | Unknown
CCC-H-12-reinject | 0 | 264141 | 458136 | 485437 | 113179950 | 0 | 3427960 | 0 | 398600 | 2607896 | 1687348 | 4313477 | 3777218 | 3813315 | 131248 | 244035 | 255173 | 865786 | 1929331 | 10935415 | 5985086 | 907564 | 23410432 | 2706727 | 17713375 | 2071302 | 20159351 | 382165 | 0 | 0 | 1120831 | 0 | 240 | 4717789 | 644587 | 40550901 | 0 | 277 | 20000465 | 10397662 | 125859 | 33049742 | 3520165 | 31103722 | 19847 | 4778743 | 1641406 | 4193316 | 42349328 | 0 | 38429312 |  | Unknown
LFB-H-11-reinject | 0 | 212962 | 385844 | 480726 | 72657344 | 0 | 4968740 | 973 | 616911 | 2088516 | 1278132 | 3202712 | 2659566 | 2642671 | 183445 | 331273 | 520881 | 1096484 | 2263951 | 13426920 | 7447833 | 1184394 | 28409833 | 1931638 | 22092585 | 2412550 | 23966959 | 590402 | 1534 | 0 | 1449455 | 0 | 485 | 363552 | 628744 | 29701807 | 0 | 35 | 14702218 | 8102199 | 226466 | 24864372 | 276486 | 22656358 | 66662 | 276542 | 1304200 | 3049501 | 5261060 | 0 | 33033199 |  | Unknown
QCS-2-Reinject | 0 | 215077 | 377764 | 370620 | 66660321 | 0 | 3917737 | 0 | 598270 | 1895423 | 1188016 | 2895607 | 2411660 | 2438116 | 182022 | 328270 | 349367 | 1055790 | 2267959 | 13391104 | 7455815 | 1167081 | 26824835 | 1872018 | 21990461 | 2463527 | 23190728 | 579051 | 0 | 0 | 1405041 | 0 | 611 | 3250291 | 552812 | 26975487 | 1918 | 129 | 13787266 | 7434792 | 200321 | 21660094 | 2580785 | 20857326 | 51320 | 2953398 | 1194124 | 2787749 | 30139689 | 417 | 25655151 |  | Unknown
LRB-22-reprep-inj2 | 0 | 24277 | 47797 | 6246 | 138063 | 1479 | 5467133 | 0 | 653772 | 2285 | 0 | 10235 | 4227 | 14820 | 211547 | 539032 | 635159 | 1182442 | 3194747 | 17482559 | 8603870 | 1266632 | 39102975 | 2783438 | 28518683 | 3287098 | 31593070 | 629245 | 0 | 0 | 1611581 | 0 | 616 | 22288 | 0 | 316513 | 0 | 150 | 75425 | 12330 | 145128 | 64827 | 12603 | 130659 | 15483 | 14691 | 0 | 7181 | 135995 | 845 | 90870 |  | Unknown
LRB-22-reprep | 0 | 23031 | 63576 | 31461 | 4232912 | 0 | 5989035 | 47 | 602175 | 0 | 0 | 114178 | 49182 | 137205 | 217265 | 611764 | 779377 | 1194069 | 3444078 | 18194704 | 8922549 | 1268080 | 40081974 | 2997945 | 30780928 | 3558349 | 32350728 | 580530 | 0 | 0 | 1595323 | 0 | 252 | 208061 | 0 | 1903816 | 0 | 719 | 174189 | 15791 | 186971 | 856870 | 207949 | 576952 | 25336 | 150619 | 0 | 12235 | 3503220 | 1184 | 2025680 |  | Unknown
WO-23003-57-A-1 | 0 | 40256 | 73763 | 4948 | 0 | 0 | 4214765 | 0 | 52953 | 0 | 296 | 7653 | 0 | 19400 | 63711 | 517283 | 1003054 | 169211 | 1718804 | 6773288 | 2450418 | 194259 | 34045107 | 3234054 | 16898084 | 3724831 | 25858895 | 44556 | 0 | 0 | 455739 | 0 | 303 | 13174 | 0 | 819985 | 0 | 1374 | 42898 | 0 | 164590 | 32501 | 1900 | 41419 | 18099 | 629 | 0 | 5073 | 61928 | 1558 | 37403 |  | Unknown

## Chain of Custody
# | Time Stamp: | Sample ID: | Pruned Sample IDs: | Sample IDs | Submission IDs | Collection Date: | Holding Time: | Match?:
20 | 2023-01-18 20:55 | 217275-d2 | 217275 | 121422PFAS_IS152_04A | 217275 | 2023-01-12 | 6.8715 | Match
21 | 2023-01-18 21:22 | 217276-d2 | 217276 | 121422PFAS_IS152_04B | 217276 | 2023-01-12 | 6.8903 | Match
22 | 2023-01-18 21:49 | 217277-d2 | 217277 | 121422PFAS_IS152_04C | 217277 | 2023-01-12 | 6.9091 | Match
23 | 2023-01-18 22:16 | 217278-d2 | 217278 | 121422PFAS_IS152_04AF | 217278 | 2023-01-12 | 6.9279 | Match
24 | 2023-01-18 22:43 | 217279-d2 | 217279 | 121422PFAS_IS152_04BF | 217279 | 2023-01-12 | 6.9467 | Match
25 | 2023-01-18 23:10 | 217280-d2 | 217280 | 121422PFAS_IS152_04CF | 217280 | 2023-01-12 | 6.9656 | Match
26 | 2023-01-18 23:37 | 217281-d200 | 217281 | 121422PFAS_IS152_06A | 217281 | 2023-01-12 | 6.9844 | Match
27 | 2023-01-19 00:05 | 217282-d200 | 217282 | 121422PFAS_IS152_06B | 217282 | 2023-01-12 | 7.0032 | Match
28 | 2023-01-19 00:32 | 217283-d200 | 217283 | 121422PFAS_IS152_06C | 217283 | 2023-01-12 | 7.022 | Match
29 | 2023-01-19 00:59 | 217284-d200 | 217284 | 121422PFAS_IS152_06C DUP | 217284 | 2023-01-12 | 7.0408 | Match
32 | 2023-01-19 02:20 | 217285-d200 | 217285 | 121422PFAS_IS152_06AF | 217285 | 2023-01-12 | 7.0972 | Match
33 | 2023-01-19 02:47 | 217286-d200 | 217286 | 121422PFAS_IS152_06BF | 217286 | 2023-01-12 | 7.116 | Match
34 | 2023-01-19 03:14 | 217287-d200 | 217287 | 121422PFAS_IS152_06CF | 217287 | 2023-01-12 | 7.1348 | Match
336 | 2023-01-30 15:56 | 217288-d4-reprep | 217288 | 121422PFAS_IS152_08A | 217288 | 2023-01-12 | 18.664 | Match
337 | 2023-01-30 16:23 | 217289-d4-reprep | 217289 | 121422PFAS_IS152_08B | 217289 | 2023-01-12 | 18.6828 | Match
37 | 2023-01-19 04:35 | 217290-d4 | 217290 | 121422PFAS_IS152_08C | 217290 | 2023-01-12 | 7.1912 | Match
38 | 2023-01-19 05:02 | 217291-d4 | 217291 | 121422PFAS_IS152_08AF | 217291 | 2023-01-12 | 7.21 | Match
39 | 2023-01-19 05:29 | 217292-d4 | 217292 | 121422PFAS_IS152_08BF | 217292 | 2023-01-12 | 7.2288 | Match
40 | 2023-01-19 05:57 | 217293-d4 | 217293 | 121422PFAS_IS152_08CF | 217293 | 2023-01-12 | 7.2476 | Match
41 | 2023-01-19 06:24 | 217294-d1000 | 217294 | 121422PFAS_IS152_10A | 217294 | 2023-01-12 | 7.2664 | Match
47 | 2023-01-19 09:33 | 217295-d1000 | 217295 | 121422PFAS_IS152_10B | 217295 | 2023-01-12 | 7.398 | Match
48 | 2023-01-19 10:00 | 217296-d1000 | 217296 | 121422PFAS_IS152_10C | 217296 | 2023-01-12 | 7.4168 | Match
49 | 2023-01-19 10:27 | 217297-d1000 | 217297 | 121422PFAS_IS152_10AF | 217297 | 2023-01-12 | 7.4356 | Match
50 | 2023-01-19 10:54 | 217298-d1000 | 217298 | 121422PFAS_IS152_10BF | 217298 | 2023-01-12 | 7.4544 | Match
51 | 2023-01-19 11:21 | 217299-d1000 | 217299 | 121422PFAS_IS152_10CF | 217299 | 2023-01-12 | 7.4733 | Match
95 | 2023-01-20 10:01 | 217300-d2-inj2 | 217300 | 011223PFAS_IS152_02A | 217300 | 2023-01-12 | 8.4172 | Match
53 | 2023-01-19 12:16 | 217301-d2 | 217301 | 011223PFAS_IS152_02B | 217301 | 2023-01-12 | 7.5109 | Match
54 | 2023-01-19 12:43 | 217302-d2 | 217302 | 011223PFAS_IS152_02C | 217302 | 2023-01-12 | 7.5297 | Match
55 | 2023-01-19 13:10 | 217303-d2 | 217303 | 011223PFAS_IS152_03A | 217303 | 2023-01-12 | 7.5485 | Match
56 | 2023-01-19 13:37 | 217304-d2 | 217304 | 011223PFAS_IS152_03B | 217304 | 2023-01-12 | 7.5673 | Match
59 | 2023-01-19 14:58 | 217305-d2 | 217305 | 011223PFAS_IS152_03C | 217305 | 2023-01-12 | 7.6237 | Match
60 | 2023-01-19 15:25 | 217306-d2 | 217306 | 011223PFAS_IS152_03C DUP | 217306 | 2023-01-12 | 7.6425 | Match
61 | 2023-01-19 15:52 | 217307-d2 | 217307 | 011223PFAS_IS152_01A | 217307 | 2023-01-12 | 7.6613 | Match
62 | 2023-01-19 16:19 | 217308-d2 | 217308 | 011223PFAS_IS152_01B | 217308 | 2023-01-12 | 7.6801 | Match
63 | 2023-01-19 16:46 | 217309-d2 | 217309 | 011223PFAS_IS152_01C | 217309 | 2023-01-12 | 7.6989 | Match
64 | 2023-01-19 17:14 | 217310-d2 | 217310 | 011223PFAS_IS152_01AF | 217310 | 2023-01-12 | 7.7177 | Match
65 | 2023-01-19 17:41 | 217311-d2 | 217311 | 011223PFAS_IS152_01BF | 217311 | 2023-01-12 | 7.7365 | Match
66 | 2023-01-19 18:08 | 217312-d2 | 217312 | 011223PFAS_IS152_01CF | 217312 | 2023-01-12 | 7.7553 | Match
67 | 2023-01-19 18:35 | 217313-d2 | 217313 | 010923PFAS_IS152_12A | 217313 | 2023-01-12 | 7.7741 | Match
68 | 2023-01-19 19:02 | 217314-d2 | 217314 | 010923PFAS_IS152_12B | 217314 | 2023-01-12 | 7.7929 | Match
338 | 2023-01-30 16:50 | 217315-d2-reprep | 217315 | 010923PFAS_IS152_12C | 217315 | 2023-01-12 | 18.7016 | Match
77 | 2023-01-19 23:06 | 217316-d2 | 217316 | 010923PFAS_IS152_12C DUP | 217316 | 2023-01-12 | 7.9622 | Match
78 | 2023-01-19 23:33 | 217317-d2 | 217317 | 010923PFAS_IS152_18A | 217317 | 2023-01-12 | 7.981 | Match
79 | 2023-01-20 | 217318-d2 | 217318 | 010923PFAS_IS152_18B | 217318 | 2023-01-12 | 7.9998 | Match
80 | 2023-01-20 00:27 | 217319-d2 | 217319 | 010923PFAS_IS152_18C | 217319 | 2023-01-12 | 8.0186 | Match
81 | 2023-01-20 00:54 | 217320-d2 | 217320 | 010923PFAS_IS152_18C DUP | 217320 | 2023-01-12 | 8.0374 | Match
82 | 2023-01-20 01:21 | 217321-d2 | 217321 | 010923PFAS_IS152_13A | 217321 | 2023-01-12 | 8.0563 | Match
83 | 2023-01-20 01:48 | 217322-d2 | 217322 | 010923PFAS_IS152_13B | 217322 | 2023-01-12 | 8.0751 | Match
84 | 2023-01-20 02:15 | 217323-d2 | 217323 | 010923PFAS_IS152_13C | 217323 | 2023-01-12 | 8.0939 | Match
85 | 2023-01-20 02:42 | 217324-d2 | 217324 | 010923PFAS_IS152_19A | 217324 | 2023-01-12 | 8.1127 | Match
129 | 2023-01-21 02:16 | 217325-d2 | 217325 | 010923PFAS_IS152_19B | 217325 | 2023-01-12 | 9.0947 | Match
130 | 2023-01-21 02:43 | 217326-d2 | 217326 | 010923PFAS_IS152_19C | 217326 | 2023-01-12 | 9.1135 | Match
131 | 2023-01-21 03:11 | 217327-d2 | 217327 | 010923PFAS_IS152_14A | 217327 | 2023-01-12 | 9.1323 | Match
132 | 2023-01-21 03:38 | 217328-d2 | 217328 | 010923PFAS_IS152_14B | 217328 | 2023-01-12 | 9.1511 | Match
133 | 2023-01-21 04:05 | 217329-d2 | 217329 | 010923PFAS_IS152_14C | 217329 | 2023-01-12 | 9.1699 | Match
140 | 2023-01-21 07:41 | 217330-d2 | 217330 | 010923PFAS_IS152_20A | 217330 | 2023-01-12 | 9.3204 | Match
141 | 2023-01-21 08:08 | 217331-d2 | 217331 | 010923PFAS_IS152_20B | 217331 | 2023-01-12 | 9.3392 | Match
142 | 2023-01-21 08:36 | 217332-d2 | 217332 | 010923PFAS_IS152_20C | 217332 | 2023-01-12 | 9.358 | Match
143 | 2023-01-21 09:03 | 217333-d2 | 217333 | 010923PFAS_IS152_15A | 217333 | 2023-01-12 | 9.3768 | Match
144 | 2023-01-21 09:30 | 217334-d2 | 217334 | 010923PFAS_IS152_15B | 217334 | 2023-01-12 | 9.3956 | Match
145 | 2023-01-21 09:57 | 217335-d2 | 217335 | 010923PFAS_IS152_15C | 217335 | 2023-01-12 | 9.4144 | Match
146 | 2023-01-21 10:24 | 217336-d2 | 217336 | 010923PFAS_IS152_15C DUP | 217336 | 2023-01-12 | 9.4332 | Match
147 | 2023-01-21 10:51 | 217337-d2 | 217337 | 010923PFAS_IS152_21A | 217337 | 2023-01-12 | 9.452 | Match
148 | 2023-01-21 11:18 | 217338-d2 | 217338 | 010923PFAS_IS152_21B | 217338 | 2023-01-12 | 9.4708 | Match
149 | 2023-01-21 11:45 | 217339-d2 | 217339 | 010923PFAS_IS152_21C | 217339 | 2023-01-12 | 9.4896 | Match
152 | 2023-01-21 13:06 | 217340-d2 | 217340 | 010923PFAS_IS152_21C DUP | 217340 | 2023-01-12 | 9.546 | Match
153 | 2023-01-21 13:33 | 217341-d2 | 217341 | 010923PFAS_IS152_16A | 217341 | 2023-01-12 | 9.5648 | Match
154 | 2023-01-21 14:00 | 217342-d2 | 217342 | 010923PFAS_IS152_16B | 217342 | 2023-01-12 | 9.5836 | Match
155 | 2023-01-21 14:27 | 217343-d2 | 217343 | 010923PFAS_IS152_16C | 217343 | 2023-01-12 | 9.6024 | Match
156 | 2023-01-21 14:54 | 217344-d2 | 217344 | 010923PFAS_IS152_22A | 217344 | 2023-01-12 | 9.6212 | Match
157 | 2023-01-21 15:22 | 217345-d2 | 217345 | 010923PFAS_IS152_22B | 217345 | 2023-01-12 | 9.64 | Match
158 | 2023-01-21 15:49 | 217346-d2 | 217346 | 010923PFAS_IS152_22C | 217346 | 2023-01-12 | 9.6588 | Match
159 | 2023-01-21 16:16 | 217347-d2 | 217347 | 010923PFAS_IS152_17A | 217347 | 2023-01-12 | 9.6776 | Match
160 | 2023-01-21 16:43 | 217348-d2 | 217348 | 010923PFAS_IS152_17B | 217348 | 2023-01-12 | 9.6964 | Match
161 | 2023-01-21 17:10 | 217349-d2 | 217349 | 010923PFAS_IS152_17C | 217349 | 2023-01-12 | 9.7152 | Match
88 | 2023-01-20 04:04 | 217350-d2 | 217350 | 010923PFAS_IS152_23A | 217350 | 2023-01-12 | 8.1692 | Match
89 | 2023-01-20 04:31 | 217351-d2 | 217351 | 010923PFAS_IS152_23B | 217351 | 2023-01-12 | 8.188 | Match
90 | 2023-01-20 04:58 | 217352-d2 | 217352 | 010923PFAS_IS152_23C | 217352 | 2023-01-12 | 8.2068 | Match
91 | 2023-01-20 05:25 | 217353-d2 | 217353 | 010923PFAS_IS152_05A | 217353 | 2023-01-12 | 8.2256 | Match
92 | 2023-01-20 05:52 | 217354-d2 | 217354 | 010923PFAS_IS152_05B | 217354 | 2023-01-12 | 8.2444 | Match
93 | 2023-01-20 06:19 | 217355-d2 | 217355 | 010923PFAS_IS152_05C | 217355 | 2023-01-12 | 8.2632 | Match
94 | 2023-01-20 06:46 | 217356-d2 | 217356 | 011023PFAS_IS152_24A | 217356 | 2023-01-12 | 8.282 | Match
96 | 2023-01-20 10:28 | 217357-d2-inj2 | 217357 | 011023PFAS_IS152_24B | 217357 | 2023-01-12 | 8.4363 | Match
97 | 2023-01-20 10:55 | 217358-d2-inj2 | 217358 | 011023PFAS_IS152_24C | 217358 | 2023-01-12 | 8.4552 | Match
98 | 2023-01-20 11:23 | 217359-d2-inj2 | 217359 | 011023PFAS_IS152_24C DUP | 217359 | 2023-01-12 | 8.474 | Match
112 | 2023-01-20 18:36 | 217360-d2 | 217360 | 011023PFAS_IS152_30A | 217360 | 2023-01-12 | 8.775 | Match
113 | 2023-01-20 19:03 | 217361-d2 | 217361 | 011023PFAS_IS152_30B | 217361 | 2023-01-12 | 8.7939 | Match
114 | 2023-01-20 19:30 | 217362-d2 | 217362 | 011023PFAS_IS152_30C | 217362 | 2023-01-12 | 8.8127 | Match
115 | 2023-01-20 19:57 | 217363-d2 | 217363 | 011023PFAS_IS152_30C DUP | 217363 | 2023-01-12 | 8.8315 | Match
116 | 2023-01-20 20:24 | 217364-d2 | 217364 | 011023PFAS_IS152_25A | 217364 | 2023-01-12 | 8.8503 | Match
117 | 2023-01-20 20:51 | 217365-d2 | 217365 | 011023PFAS_IS152_25B | 217365 | 2023-01-12 | 8.8691 | Match
118 | 2023-01-20 21:19 | 217366-d2 | 217366 | 011023PFAS_IS152_25C | 217366 | 2023-01-12 | 8.8879 | Match
119 | 2023-01-20 21:46 | 217367-d2 | 217367 | 011023PFAS_IS152_31A | 217367 | 2023-01-12 | 8.9067 | Match
120 | 2023-01-20 22:13 | 217368-d2 | 217368 | 011023PFAS_IS152_31B | 217368 | 2023-01-12 | 8.9255 | Match
121 | 2023-01-20 22:40 | 217369-d2 | 217369 | 011023PFAS_IS152_31C | 217369 | 2023-01-12 | 8.9443 | Match
124 | 2023-01-21 00:01 | 217370-d2 | 217370 | 011023PFAS_IS152_26A | 217370 | 2023-01-12 | 9.0007 | Match
125 | 2023-01-21 00:28 | 217371-d2 | 217371 | 011023PFAS_IS152_26B | 217371 | 2023-01-12 | 9.0195 | Match
126 | 2023-01-21 00:55 | 217372-d2 | 217372 | 011023PFAS_IS152_26C | 217372 | 2023-01-12 | 9.0383 | Match
127 | 2023-01-21 01:22 | 217373-d2 | 217373 | 011023PFAS_IS152_32A | 217373 | 2023-01-12 | 9.0571 | Match
128 | 2023-01-21 01:49 | 217374-d2 | 217374 | 011023PFAS_IS152_32B | 217374 | 2023-01-12 | 9.0759 | Match
168 | 2023-01-21 20:47 | 217375-d2 | 217375 | 011023PFAS_IS152_32C | 217375 | 2023-01-12 | 9.8657 | Match
169 | 2023-01-21 21:14 | 217376-d2 | 217376 | 011023PFAS_IS152_27A | 217376 | 2023-01-12 | 9.8845 | Match
170 | 2023-01-21 21:41 | 217377-d2 | 217377 | 011023PFAS_IS152_27B | 217377 | 2023-01-12 | 9.9033 | Match
171 | 2023-01-21 22:08 | 217378-d2 | 217378 | 011023PFAS_IS152_27C | 217378 | 2023-01-12 | 9.9221 | Match
172 | 2023-01-21 22:35 | 217379-d2 | 217379 | 011023PFAS_IS152_27C DUP | 217379 | 2023-01-12 | 9.9409 | Match
173 | 2023-01-21 23:02 | 217380-d2 | 217380 | 011023PFAS_IS152_33A | 217380 | 2023-01-12 | 9.9597 | Match
174 | 2023-01-21 23:29 | 217381-d2 | 217381 | 011023PFAS_IS152_33B | 217381 | 2023-01-12 | 9.9785 | Match
175 | 2023-01-21 23:56 | 217382-d2 | 217382 | 011023PFAS_IS152_33C | 217382 | 2023-01-12 | 9.9973 | Match
176 | 2023-01-22 00:23 | 217383-d2 | 217383 | 011023PFAS_IS152_33C DUP | 217383 | 2023-01-12 | 10.0161 | Match
177 | 2023-01-22 00:50 | 217384-d2 | 217384 | 011023PFAS_IS152_28A | 217384 | 2023-01-12 | 10.0349 | Match
203 | 2023-01-23 11:29 | 217385-d2-inj2 | 217385 | 011023PFAS_IS152_28B | 217385 | 2023-01-12 | 11.4786 | Match
181 | 2023-01-22 02:39 | 217386-d2 | 217386 | 011023PFAS_IS152_28C | 217386 | 2023-01-12 | 10.1102 | Match
182 | 2023-01-22 03:06 | 217387-d2 | 217387 | 011023PFAS_IS152_34A | 217387 | 2023-01-12 | 10.129 | Match
183 | 2023-01-22 03:33 | 217388-d2 | 217388 | 011023PFAS_IS152_34B | 217388 | 2023-01-12 | 10.1478 | Match
184 | 2023-01-22 04:00 | 217389-d2 | 217389 | 011023PFAS_IS152_34C | 217389 | 2023-01-12 | 10.1666 | Match
185 | 2023-01-22 04:27 | 217390-d4 | 217390 | 011023PFAS_IS152_29A | 217390 | 2023-01-12 | 10.1854 | Match
186 | 2023-01-22 04:54 | 217391-d4 | 217391 | 011023PFAS_IS152_29B | 217391 | 2023-01-12 | 10.2042 | Match
187 | 2023-01-22 05:21 | 217392-d4 | 217392 | 011023PFAS_IS152_29C | 217392 | 2023-01-12 | 10.223 | Match
344 | 2023-01-30 19:33 | 217393-d4-reprep | 217393 | 011023PFAS_IS152_35A | 217393 | 2023-01-12 | 18.8144 | Match
345 | 2023-01-30 20:00 | 217394-d4-reprep | 217394 | 011023PFAS_IS152_35B | 217394 | 2023-01-12 | 18.8332 | Match
346 | 2023-01-30 21:21 | 217395-d4-reprep | 217395 | 011023PFAS_IS152_35C | 217395 | 2023-01-12 | 18.8897 | Match
197 | 2023-01-22 10:19 | 217396-d200 | 217396 | 011023PFAS_IS152_07A | 217396 | 2023-01-12 | 10.43 | Match
198 | 2023-01-22 10:46 | 217397-d200 | 217397 | 011023PFAS_IS152_07B | 217397 | 2023-01-12 | 10.4488 | Match
199 | 2023-01-22 11:13 | 217398-d200 | 217398 | 011023PFAS_IS152_07C | 217398 | 2023-01-12 | 10.4676 | Match
347 | 2023-01-30 22:15 | 217399-d2-reprep | 217399 | 011123PFAS_IS152_36A | 217399 | 2023-01-12 | 18.9273 | Match
348 | 2023-01-30 22:42 | 217400-d2-reprep | 217400 | 011123PFAS_IS152_36B | 217400 | 2023-01-12 | 18.9461 | Match
207 | 2023-01-23 13:18 | 217401-d2 | 217401 | 011123PFAS_IS152_36C | 217401 | 2023-01-12 | 11.554 | Match
208 | 2023-01-23 13:45 | 217402-d2 | 217402 | 011123PFAS_IS152_36C DUP | 217402 | 2023-01-12 | 11.5728 | Match
209 | 2023-01-23 14:12 | 217403-d2 | 217403 | 011123PFAS_IS152_42A | 217403 | 2023-01-12 | 11.5916 | Match
210 | 2023-01-23 14:39 | 217404-d2 | 217404 | 011123PFAS_IS152_42B | 217404 | 2023-01-12 | 11.6104 | Match
213 | 2023-01-23 16:00 | 217405-d2 | 217405 | 011123PFAS_IS152_42C | 217405 | 2023-01-12 | 11.6668 | Match
214 | 2023-01-23 16:27 | 217406-d2 | 217406 | 011123PFAS_IS152_42C DUP | 217406 | 2023-01-12 | 11.6856 | Match
215 | 2023-01-23 16:54 | 217407-d2 | 217407 | 011123PFAS_IS152_37A | 217407 | 2023-01-12 | 11.7044 | Match
216 | 2023-01-23 17:21 | 217408-d2 | 217408 | 011123PFAS_IS152_37B | 217408 | 2023-01-12 | 11.7232 | Match
217 | 2023-01-23 17:49 | 217409-d2 | 217409 | 011123PFAS_IS152_37C | 217409 | 2023-01-12 | 11.742 | Match
218 | 2023-01-23 18:16 | 217410-d2 | 217410 | 011123PFAS_IS152_43A | 217410 | 2023-01-12 | 11.7608 | Match
219 | 2023-01-23 18:43 | 217411-d2 | 217411 | 011123PFAS_IS152_43B | 217411 | 2023-01-12 | 11.7796 | Match
220 | 2023-01-23 19:10 | 217412-d2 | 217412 | 011123PFAS_IS152_43C | 217412 | 2023-01-12 | 11.7984 | Match
222 | 2023-01-27 08:54 | 217413-d2-inj2 | 217413 | 011123PFAS_IS152_38A | 217413 | 2023-01-12 | 15.3709 | Match
223 | 2023-01-27 09:21 | 217414-d2 | 217414 | 011123PFAS_IS152_38B | 217414 | 2023-01-12 | 15.3896 | Match
230 | 2023-01-27 12:58 | 217415-d2 | 217415 | 011123PFAS_IS152_38C | 217415 | 2023-01-12 | 15.5401 | Match
231 | 2023-01-27 13:25 | 217416-d2 | 217416 | 011123PFAS_IS152_44A | 217416 | 2023-01-12 | 15.5589 | Match
232 | 2023-01-27 13:52 | 217417-d2 | 217417 | 011123PFAS_IS152_44B | 217417 | 2023-01-12 | 15.5777 | Match
233 | 2023-01-27 14:19 | 217418-d2 | 217418 | 011123PFAS_IS152_44C | 217418 | 2023-01-12 | 15.5965 | Match
234 | 2023-01-27 14:46 | 217419-d2 | 217419 | 011123PFAS_IS152_39A | 217419 | 2023-01-12 | 15.6153 | Match
235 | 2023-01-27 15:13 | 217420-d2 | 217420 | 011123PFAS_IS152_39B | 217420 | 2023-01-12 | 15.6341 | Match
236 | 2023-01-27 15:40 | 217421-d2 | 217421 | 011123PFAS_IS152_39C | 217421 | 2023-01-12 | 15.6529 | Match
237 | 2023-01-27 16:07 | 217422-d2 | 217422 | 011123PFAS_IS152_39C DUP | 217422 | 2023-01-12 | 15.6717 | Match
238 | 2023-01-27 16:34 | 217423-d2 | 217423 | 011123PFAS_IS152_45A | 217423 | 2023-01-12 | 15.6905 | Match
239 | 2023-01-27 17:01 | 217424-d2 | 217424 | 011123PFAS_IS152_45B | 217424 | 2023-01-12 | 15.7093 | Match
242 | 2023-01-27 18:23 | 217425-d2 | 217425 | 011123PFAS_IS152_45C | 217425 | 2023-01-12 | 15.7657 | Match
243 | 2023-01-27 18:50 | 217426-d2 | 217426 | 011123PFAS_IS152_45C DUP | 217426 | 2023-01-12 | 15.7845 | Match
244 | 2023-01-27 19:17 | 217427-d2 | 217427 | 011123PFAS_IS152_40A | 217427 | 2023-01-12 | 15.8033 | Match
245 | 2023-01-27 19:44 | 217428-d2 | 217428 | 011123PFAS_IS152_40B | 217428 | 2023-01-12 | 15.8221 | Match
246 | 2023-01-27 20:11 | 217429-d2 | 217429 | 011123PFAS_IS152_40C | 217429 | 2023-01-12 | 15.8409 | Match
247 | 2023-01-27 20:38 | 217430-d2 | 217430 | 011123PFAS_IS152_46A | 217430 | 2023-01-12 | 15.8597 | Match
248 | 2023-01-27 21:05 | 217431-d2 | 217431 | 011123PFAS_IS152_46B | 217431 | 2023-01-12 | 15.8785 | Match
249 | 2023-01-27 21:32 | 217432-d2 | 217432 | 011123PFAS_IS152_46C | 217432 | 2023-01-12 | 15.8973 | Match
250 | 2023-01-27 21:59 | 217433-d2 | 217433 | 011123PFAS_IS152_41A | 217433 | 2023-01-12 | 15.9161 | Match
251 | 2023-01-27 22:26 | 217434-d2 | 217434 | 011123PFAS_IS152_41B | 217434 | 2023-01-12 | 15.9349 | Match
258 | 2023-01-28 02:03 | 217435-d2 | 217435 | 011123PFAS_IS152_41C | 217435 | 2023-01-12 | 16.0854 | Match
259 | 2023-01-28 02:30 | 217436-d2 | 217436 | 011123PFAS_IS152_47A | 217436 | 2023-01-12 | 16.1042 | Match
260 | 2023-01-28 02:57 | 217437-d2 | 217437 | 011123PFAS_IS152_47B | 217437 | 2023-01-12 | 16.123 | Match
261 | 2023-01-28 03:24 | 217438-d2 | 217438 | 011123PFAS_IS152_47C | 217438 | 2023-01-12 | 16.1418 | Match
262 | 2023-01-28 03:51 | 217439-d2 | 217439 | 011123PFAS_IS152_09A | 217439 | 2023-01-12 | 16.1606 | Match
263 | 2023-01-28 04:18 | 217440-d2 | 217440 | 011123PFAS_IS152_09B | 217440 | 2023-01-12 | 16.1794 | Match
264 | 2023-01-28 04:45 | 217441-d2 | 217441 | 011123PFAS_IS152_09C | 217441 | 2023-01-12 | 16.1982 | Match
265 | 2023-01-28 05:12 | 217442-d2 | 217442 | 011123PFAS_IS152_09C DUP | 217442 | 2023-01-12 | 16.217 | Match
266 | 2023-01-28 05:40 | 217443-d2 | 217443 | 011223PFAS_IS152_48A | 217443 | 2023-01-12 | 16.2358 | Match
267 | 2023-01-28 06:07 | 217444-d2 | 217444 | 011223PFAS_IS152_48B | 217444 | 2023-01-12 | 16.2546 | Match
270 | 2023-01-28 07:28 | 217445-d2 | 217445 | 011223PFAS_IS152_48C | 217445 | 2023-01-12 | 16.311 | Match
271 | 2023-01-28 07:55 | 217446-d2 | 217446 | 011223PFAS_IS152_48C DUP | 217446 | 2023-01-12 | 16.3298 | Match
272 | 2023-01-28 08:22 | 217447-d2 | 217447 | 011223PFAS_IS152_54A | 217447 | 2023-01-12 | 16.3487 | Match
273 | 2023-01-28 08:49 | 217448-d2 | 217448 | 011223PFAS_IS152_54B | 217448 | 2023-01-12 | 16.3675 | Match
274 | 2023-01-28 09:16 | 217449-d2 | 217449 | 011223PFAS_IS152_54C | 217449 | 2023-01-12 | 16.3863 | Match
275 | 2023-01-28 09:43 | 217450-d2 | 217450 | 011223PFAS_IS152_54C DUP | 217450 | 2023-01-12 | 16.4051 | Match
276 | 2023-01-28 10:10 | 217451-d2 | 217451 | 011223PFAS_IS152_49A | 217451 | 2023-01-12 | 16.4239 | Match
277 | 2023-01-28 10:38 | 217452-d2 | 217452 | 011223PFAS_IS152_49B | 217452 | 2023-01-12 | 16.4427 | Match
278 | 2023-01-28 11:05 | 217453-d2 | 217453 | 011223PFAS_IS152_49C | 217453 | 2023-01-12 | 16.4615 | Match
279 | 2023-01-28 11:32 | 217454-d2 | 217454 | 011223PFAS_IS152_55A | 217454 | 2023-01-12 | 16.4803 | Match
286 | 2023-01-28 15:08 | 217455-d2 | 217455 | 011223PFAS_IS152_55B | 217455 | 2023-01-12 | 16.6308 | Match
287 | 2023-01-28 15:35 | 217456-d2 | 217456 | 011223PFAS_IS152_55C | 217456 | 2023-01-12 | 16.6496 | Match
288 | 2023-01-28 16:03 | 217457-d2 | 217457 | 011223PFAS_IS152_50A | 217457 | 2023-01-12 | 16.6684 | Match
289 | 2023-01-28 16:30 | 217458-d2 | 217458 | 011223PFAS_IS152_50B | 217458 | 2023-01-12 | 16.6873 | Match
290 | 2023-01-28 16:57 | 217459-d2 | 217459 | 011223PFAS_IS152_50C | 217459 | 2023-01-12 | 16.7061 | Match
291 | 2023-01-28 17:24 | 217460-d2 | 217460 | 011223PFAS_IS152_56A | 217460 | 2023-01-12 | 16.7249 | Match
292 | 2023-01-28 17:51 | 217461-d2 | 217461 | 011223PFAS_IS152_56B | 217461 | 2023-01-12 | 16.7437 | Match
293 | 2023-01-28 18:18 | 217462-d2 | 217462 | 011223PFAS_IS152_56C | 217462 | 2023-01-12 | 16.7625 | Match
294 | 2023-01-28 18:45 | 217463-d2 | 217463 | 011223PFAS_IS152_51A | 217463 | 2023-01-12 | 16.7813 | Match
295 | 2023-01-28 19:12 | 217464-d2 | 217464 | 011223PFAS_IS152_51B | 217464 | 2023-01-12 | 16.8002 | Match
298 | 2023-01-28 20:33 | 217465-d2 | 217465 | 011223PFAS_IS152_51C | 217465 | 2023-01-12 | 16.8566 | Match
299 | 2023-01-28 21:01 | 217466-d2 | 217466 | 011223PFAS_IS152_51C DUP | 217466 | 2023-01-12 | 16.8754 | Match
300 | 2023-01-28 21:28 | 217467-d2 | 217467 | 011223PFAS_IS152_57A | 217467 | 2023-01-12 | 16.8942 | Match
301 | 2023-01-28 21:55 | 217468-d2 | 217468 | 011223PFAS_IS152_57B | 217468 | 2023-01-12 | 16.913 | Match
302 | 2023-01-28 22:22 | 217469-d2 | 217469 | 011223PFAS_IS152_57C | 217469 | 2023-01-12 | 16.9318 | Match
303 | 2023-01-28 22:49 | 217470-d2 | 217470 | 011223PFAS_IS152_57C DUP | 217470 | 2023-01-12 | 16.9506 | Match
304 | 2023-01-28 23:16 | 217471-d2 | 217471 | 011223PFAS_IS152_52A | 217471 | 2023-01-12 | 16.9694 | Match
305 | 2023-01-28 23:43 | 217472-d2 | 217472 | 011223PFAS_IS152_52B | 217472 | 2023-01-12 | 16.9883 | Match
306 | 2023-01-29 00:10 | 217473-d2 | 217473 | 011223PFAS_IS152_52C | 217473 | 2023-01-12 | 17.0071 | Match
307 | 2023-01-29 00:37 | 217474-d2 | 217474 | 011223PFAS_IS152_58A | 217474 | 2023-01-12 | 17.0259 | Match
314 | 2023-01-29 04:14 | 217475-d2 | 217475 | 011223PFAS_IS152_58B | 217475 | 2023-01-12 | 17.1763 | Match
315 | 2023-01-29 04:41 | 217476-d2 | 217476 | 011223PFAS_IS152_58C | 217476 | 2023-01-12 | 17.1951 | Match
332 | 2023-01-30 13:03 | 217477-d10 | 217477 | 011223PFAS_IS152_53A | 217477 | 2023-01-12 | 18.544 | Match
331 | 2023-01-30 12:36 | 217478-d10 | 217478 | 011223PFAS_IS152_53B | 217478 | 2023-01-12 | 18.5252 | Match
330 | 2023-01-30 12:09 | 217479-d20 | 217479 | 011223PFAS_IS152_59A | 217479 | 2023-01-12 | 18.5064 | Match
333 | 2023-01-30 13:30 | 217480-d20 | 217480 | 011223PFAS_IS152_59B | 217480 | 2023-01-12 | 18.5628 | Match
334 | 2023-01-30 13:58 | 217481-d20 | 217481 | 011223PFAS_IS152_59C | 217481 | 2023-01-12 | 18.5816 | Match
335 | 2023-01-30 14:25 | 217482-d20 | 217482 | 011223PFAS_IS152_59C DUP | 217482 | 2023-01-12 | 18.6004 | Match
322 | 2023-01-29 07:50 | 217483-d200 | 217483 | 011223PFAS_IS152_11A | 217483 | 2023-01-12 | 17.3267 | Match
323 | 2023-01-29 08:18 | 217484-d200 | 217484 | 011223PFAS_IS152_11B | 217484 | 2023-01-12 | 17.3455 | Match
325 | 2023-01-29 09:12 | 217485-d200 | 217485 | 011223PFAS_IS152_11C | 217485 | 2023-01-12 | 17.3831 | Match

## Internal Standard Recovery
| F-PFBA | F-PFBS | F-HFPO-DA | F-PFHxA | F-PFHxS | F-PFOA | F-PFNA | F-PFOS | F-PFDA | F-PFBA | F-PFBS | F-HFPO-DA | F-PFHxA | F-PFHxS | F-PFOA | F-PFNA | F-PFOS | F-PFDA | F-PFBA | F-PFBS | F-HFPO-DA | F-PFHxA | F-PFHxS | F-PFOA | F-PFNA | F-PFOS | F-PFDA
AB-230117-02 Cal Blank 1 | 1.7338 | 1.0688 | 0.81626 | 0.90701 | 1.0251 | 1.0063 | 0.95362 | 1.079 | 0.92396 | IS1 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS1 |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 Cal Std 1 1 | 1.0258 | 0.64147 | 1.154 | 1.0667 | 1.1021 | 1.1286 | 1.1243 | 1.1475 | 1.2331 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 Cal Std 2 1 | 0.91686 | 1.235 | 1.0622 | 1.021 | 1.129 | 1.1513 | 1.1018 | 1.1223 | 1.1893 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 Cal Std 3 1 | 1.1869 | 1.3032 | 1.0391 | 1.0767 | 1.1097 | 1.0027 | 1.0821 | 1.054 | 1.2746 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 Cal Std 4 1 | 1.0664 | 1.0478 | 1.0507 | 1.0091 | 1.111 | 1.1544 | 1.1669 | 1.1542 | 1.3105 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 Cal Std 5 1 | 1.0526 | 0.7245 | 1.0487 | 1.0256 | 1.1004 | 0.99731 | 1.066 | 1.0549 | 1.2736 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 Cal Std 6 1 | 1.1283 | 0.96658 | 0.96762 | 0.99698 | 1.0649 | 1.1102 | 1.0318 | 0.98506 | 1.1892 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 Cal Std 7 1 | 0.73111 | 0.99636 | 0.82234 | 0.83743 | 0.97495 | 0.98004 | 1.021 | 1.0149 | 1.0195 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 Cal Std 8 1 | 0.97416 | 1.2736 | 1.061 | 1.0497 | 0.96455 | 1.0072 | 0.97895 | 0.92118 | 0.98053 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 Cal Std 9 1 | 0.84125 | 1.1126 | 1.0324 | 1.003 | 0.79635 | 0.85721 | 0.8062 | 0.73092 | 0.79332 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 Cal Std 10 1 | 0.97325 | 0.99288 | 0.95319 | 0.86523 | 0.54738 | 0.58667 | 0.5212 | 0.47743 | 0.49089 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | IS3 | #N/A |  |  | #N/A |  |  | #N/A |  | IS3 | IS3
AB-230117-02 Cal Std 11 1 | 1.0949 | 1.0036 | 0.76046 | 0.68571 | 0.35002 | 0.38649 | 0.32002 | 0.28216 | 0.3047 | FALSE | FALSE | #N/A | FALSE | IS3 | #N/A | IS3 | IS3 | #N/A |  |  | #N/A |  | IS3 | IS3 | IS3 | IS3 | IS3
AB-230117-02 Cal Std 12 1 | 0.94279 | 0.844 | 0.53112 | 0.47122 | 0.23727 | 0.22057 | 0.1709 | 0.18127 | 0.16189 | FALSE | FALSE | #N/A | IS3 | IS3 | #N/A | IS3 | IS3 | #N/A |  |  | #N/A | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
AB-230117-02 Cal Std 13 1 | 0.84208 | 0.53842 | 0.27537 | 0.23672 | 0.1369 | 0.096773 | 0.082578 | 0.10344 | 0.077361 | FALSE | FALSE | #N/A | IS3 | IS3 | #N/A | IS3 | IS3 | #N/A |  |  | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 | IS3
AB-230117-02 LRB 1 | 1.0335 | 1.1903 | 1.0588 | 1.0487 | 1.085 | 1.1158 | 1.1228 | 1.062 | 1.135 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  | b | b | b | b | b
AB-230117-02 LFB-H 1 | 0.9436 | 0.84419 | 0.99213 | 0.9719 | 0.94869 | 0.98889 | 0.93137 | 0.85355 | 0.95419 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 QCS 1 | 1.0096 | 1.0281 | 1.0456 | 1.0165 | 0.90507 | 0.9542 | 0.86521 | 0.75474 | 1.0558 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 IBC 1 | 0.93972 | 0.93766 | 1.0343 | 1.0044 | 0.93149 | 0.96612 | 0.94633 | 0.89023 | 1.0366 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 CCC-H 1 | 0.9883 | 1.0565 | 1.0195 | 0.96058 | 0.77031 | 0.83187 | 0.72342 | 0.62481 | 0.74877 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-01-A-1 @10 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | IS3 | IS3 | IS3 | IS4 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
WO-23003-02-A-1 | 0.10794 | 0.35019 | 0.3657 | 0.36132 | 0.64846 | 0.70169 | 0.81632 | 1.0264 | 0.95805 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-03-A-1 | 0.26389 | 0.58852 | 0.54143 | 0.50633 | 0.86765 | 0.88001 | 0.87997 | 1.0206 | 0.86525 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-04-A-1 | 0.023178 | 0.23011 | 0.18923 | 0.20178 | 0.4844 | 0.59164 | 0.81568 | 1.001 | 1.0525 | IS3 | IS3 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-05-A-1 | 0.044176 | 0.27983 | 0.22873 | 0.23117 | 0.45259 | 0.54964 | 0.7712 | 0.9745 | 1.0329 | IS3 | IS3 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-06-A-1 | 0.02798 | 0.1628 | 0.21188 | 0.2022 | 0.38049 | 0.47413 | 0.7047 | 0.88927 | 0.97679 | IS3 | IS3 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-07-A-1 | 0.90605 | 1.1032 | 0.9329 | 0.82914 | 1.0019 | 1.0671 | 1.0636 | 0.97073 | 0.34276 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | IS4
WO-23003-08-A-1 | 0.93898 | 1.0305 | 0.96609 | 0.89284 | 1.0189 | 1.0613 | 1.0619 | 0.9563 | 0.33254 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | IS4
WO-23003-09-A-1 | 0.7807 | 0.81764 | 0.9646 | 0.61708 | 1.0063 | 1.0794 | 1.0721 | 0.98311 | 0.35382 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | IS4
WO-23003-10-A-1 | 0.84329 | 0.9153 | 0.93716 | 0.96162 | 0.99254 | 1.0753 | 1.0709 | 0.97895 | 0.32758 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | IS4
LRB2-Inj1 | 1.107 | 0.9587 | 1.0125 | 0.94495 | 1.0515 | 1.0887 | 1.083 | 1.0309 | 1.1624 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | b
AB-230117-02 LRB 2 | 0.78776 | 0.78227 | 1.022 | 1.0268 | 0.94346 | 0.96205 | 1.0113 | 0.93234 | 1.1693 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | b
WO-23003-11-A-1 | 0.80866 | 0.95429 | 0.94052 | 0.94709 | 0.98213 | 1.0159 | 1.0017 | 0.92656 | 0.33484 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | IS4
WO-23003-12-A-1 | 0.78423 | 0.84624 | 0.85658 | 0.82265 | 0.97977 | 1.0326 | 1.0274 | 0.96402 | 0.33735 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | IS4
WO-23003-13-A-1 | 0.94114 | 0.90697 | 0.92845 | 0.93372 | 0.97955 | 1.0342 | 1.0402 | 0.94229 | 0.33152 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | IS4
WO-23003-14-A-1 | 0.36658 | 0.44305 | 0.52524 | 0.50144 | 0.75921 | 0.82678 | 0.91816 | 1.0136 | 0.8816 | IS3 | IS4 | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | #N/A |  |  | #N/A |  |  | C
WO-23003-15-A-1 | 0.39846 | 0.35141 | 0.34674 | 0.28463 | 0.79792 | 0.84334 | 0.90457 | 1.0187 | 1.0795 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | C
WO-23003-16-A-1 | 0.37639 | 0.47994 | 0.54566 | 0.5269 | 0.77378 | 0.83539 | 0.96451 | 1.0071 | 1.121 | IS3 | IS4 | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-17-A-1 | 0.34843 | 0.34488 | 0.49577 | 0.48838 | 0.74851 | 0.81446 | 0.93245 | 0.99354 | 1.1147 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-18-A-1 | 0.33157 | 0.42408 | 0.48528 | 0.4634 | 0.69158 | 0.78444 | 0.91327 | 0.98184 | 1.009 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-19-A-1 | 0.39297 | 0.46318 | 0.32537 | 0.3714 | 0.66104 | 0.74863 | 0.74434 | 0.80994 | 1.0108 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-20-A-1 | 0.89513 | 0.91274 | 0.58625 | 0.49129 | 0.86292 | 0.92691 | 0.64654 | 0.5246 | 1.1572 | FALSE | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A | IS4 |  | #N/A |  |  | #N/A
AB-230117-02 LRB 3 | 0.90346 | 0.82201 | 0.96449 | 0.95866 | 0.98197 | 1.0163 | 1.0206 | 0.96731 | 1.1911 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  | b | #N/A
AB-230117-02 LD 1 | 0.25337 | 0.71887 | 0.79214 | 0.83467 | 1.0533 | 0.98799 | 0.87607 | 1.0304 | 1.0051 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LFB-H 2 | 0.91143 | 1.0617 | 0.93291 | 0.89663 | 0.79902 | 0.85444 | 0.8508 | 0.77511 | 0.95126 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 CCC-H 2 | 0.90277 | 0.92794 | 0.89034 | 0.84951 | 0.67875 | 0.73317 | 0.71245 | 0.61578 | 0.76578 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LFSM-H 1 @5 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE! | IS3 | IS3 | IS3 | IS3 | #VALUE! | #VALUE! | #VALUE! | #VALUE! | #VALUE!
WO-23003-21-A-1 | 0.76764 | 0.94203 | 0.91018 | 0.92284 | 0.94741 | 0.97334 | 0.66846 | 0.52112 | 1.1122 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-22-A-1 | 0.90655 | 0.73232 | 0.93309 | 0.93339 | 0.98047 | 0.99144 | 0.69396 | 0.55083 | 1.2004 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-23-A-1 | 0.77651 | 1.0419 | 0.88149 | 0.86229 | 0.98482 | 0.98241 | 0.66573 | 0.5262 | 1.1709 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-24-A-1 | 0.88864 | 0.8292 | 0.86318 | 0.79567 | 0.94635 | 0.98597 | 0.68963 | 0.53868 | 1.1983 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-25-A-1 | 0.90222 | 0.90075 | 0.91134 | 0.87179 | 0.92069 | 0.98125 | 0.66478 | 0.51261 | 1.1796 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-26-A-1 | 0.18101 | 0.38383 | 0.3501 | 0.36814 | 0.58 | 0.71025 | 0.84767 | 0.91299 | 1.0678 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  | C | #N/A
WO-23003-27-A-1 | 0.23486 | 0.52283 | 0.46328 | 0.45992 | 0.66658 | 0.7583 | 0.8386 | 0.89849 | 0.93621 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-28-A-1 | 0.26298 | 0.34176 | 0.43923 | 0.4381 | 0.6214 | 0.74755 | 0.89386 | 0.97404 | 1.0791 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-29-A-1 | 0.233 | 0.49899 | 0.46392 | 0.4515 | 0.63041 | 0.76283 | 0.92869 | 0.98649 | 1.1358 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-30-A-1 | 0.19328 | 0.46141 | 0.40958 | 0.41689 | 0.58934 | 0.69861 | 0.82402 | 0.94464 | 0.96918 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
LRB4-inj1 | 0.90679 | 1.0584 | 0.9542 | 0.91282 | 0.93586 | 0.98592 | 1.0562 | 0.98747 | 1.2331 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LRB 4 | 0.69979 | 1.0983 | 0.59404 | 0.63465 | 0.92346 | 0.98867 | 1.0144 | 0.93617 | 1.1928 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-31-A-1 | 0.15099 | 0.40957 | 0.26239 | 0.21739 | 0.47659 | 0.61876 | 0.8159 | 0.87022 | 1.0376 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-32-A-1 | 0.11229 | 0.38674 | 0.25173 | 0.21324 | 0.49313 | 0.63515 | 0.83488 | 0.91307 | 1.0453 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-33-A-1 | 0.25112 | 0.34951 | 0.24367 | 0.42437 | 0.50282 | 0.64738 | 0.84687 | 0.94375 | 1.0689 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-34-A-1 | 0.21971 | 0.49242 | 0.4189 | 0.41921 | 0.63026 | 0.71355 | 0.89874 | 0.98139 | 1.0965 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-35-A-1 | 0.22672 | 0.44227 | 0.4266 | 0.42824 | 0.61797 | 0.72039 | 0.87532 | 0.98614 | 1.0604 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-36-A-1 | 0.24525 | 0.42762 | 0.43394 | 0.4337 | 0.63077 | 0.7346 | 0.88519 | 0.98257 | 1.0702 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-37-A-1 | 0.20439 | 0.41959 | 0.44105 | 0.42098 | 0.60216 | 0.74133 | 0.90223 | 0.99181 | 1.0801 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-38-A-1 | 0.1908 | 0.52859 | 0.45839 | 0.45156 | 0.6546 | 0.75211 | 0.89285 | 0.98497 | 1.1044 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-39-A-1 | 0.093571 | 0.32518 | 0.30249 | 0.29706 | 0.47314 | 0.58983 | 0.81743 | 0.91491 | 1.0504 | IS3 | IS3 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-40-A-1 | 0.20351 | 0.46595 | 0.40965 | 0.40862 | 0.58096 | 0.70571 | 0.86496 | 0.94137 | 1.0757 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
LRB5-inj1 | 0.91357 | 1.0676 | 0.94652 | 0.96133 | 0.9388 | 0.98624 | 1.0366 | 0.96637 | 1.2095 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LRB 5 | 0.77971 | 0.83441 | 0.92262 | 0.86698 | 0.88337 | 0.94559 | 0.96666 | 0.89469 | 1.1282 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LD 2 | 0.86036 | 1.0225 | 0.9312 | 0.94399 | 0.95973 | 0.92412 | 0.29189 | 0.22668 | 1.12 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS4 | IS4 | #N/A |  |  | #N/A |  |  | #N/A | IS4 | IS4 | #N/A
AB-230117-02 LFB-H 3 | 0.92298 | 0.98421 | 0.89772 | 0.90201 | 0.78507 | 0.85782 | 0.78432 | 0.64046 | 0.97084 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  | C | #N/A
AB-230117-02 CCC-H 3 | 0.97279 | 0.92063 | 0.61065 | 0.50907 | 0.62971 | 0.72537 | 0.71377 | 0.61325 | 0.76271 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LFSM-H 2 | 1.0025 | 0.95041 | 0.9082 | 0.91434 | 0.78244 | 0.79296 | 0.27413 | 0.22003 | 0.82936 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | #N/A |  |  | #N/A |  |  | #N/A | IS3 | IS3 | #N/A
Rinse5 | 0.034118 | 0.022571 | 0.019129 | 0.017238 | 0.043142 | 0.047402 | 0.056009 | 0.055422 | 0.085556 | IS3 | IS4 | #N/A | IS4 | IS3 | #N/A | IS3 | IS3 | #N/A | IS3 | IS4 | IS4 | IS4 | IS3 | IS3 | IS3 | IS3 | IS3
WO-23003-41-A-1 | 0.15553 | 0.30669 | 0.3108 | 0.3852 | 0.5691 | 0.69745 | 0.85336 | 0.92074 | 1.0656 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  | C | #N/A
WO-23003-42-A-1 | 0.12104 | 0.31127 | 0.37445 | 0.32059 | 0.56539 | 0.68377 | 0.84789 | 0.94975 | 1.088 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-43-A-1 | 0.16488 | 0.33908 | 0.31101 | 0.368 | 0.52644 | 0.64384 | 0.74991 | 0.91464 | 0.94711 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-44-A-1 | 0.20396 | 0.29823 | 0.30162 | 0.24523 | 0.58132 | 0.69795 | 0.88679 | 1.0046 | 1.0618 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-45-A-1 | 0.22915 | 0.33706 | 0.29845 | 0.26248 | 0.57404 | 0.70859 | 0.86042 | 0.94068 | 1.0628 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-46-A-1 | 0.12049 | 0.38581 | 0.28588 | 0.35719 | 0.4077 | 0.56576 | 0.81065 | 0.95152 | 1.0268 | IS3 | IS3 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-47-A-1 | 0.025227 | 0.19582 | 0.22257 | 0.23142 | 0.4129 | 0.49935 | 0.67835 | 0.8743 | 0.84082 | IS3 | IS3 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-48-A-1 | 0.18127 | 0.28519 | 0.41561 | 0.40025 | 0.54708 | 0.68665 | 0.86109 | 0.91507 | 1.0787 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-49-A-1 | 0.19158 | 0.35483 | 0.41522 | 0.36648 | 0.59453 | 0.69837 | 0.8751 | 0.98336 | 1.0696 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-50-A-1 | 0.028547 | 0.13434 | 0.10604 | 0.16457 | 0.28928 | 0.4117 | 0.66611 | 0.79132 | 0.95875 | IS3 | IS3 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 | IS4 | IS4 |  |  | #N/A
LRB6-inj1 | 0.70934 | 0.66323 | 0.89154 | 0.95386 | 0.98872 | 1.0173 | 1.0382 | 0.96745 | 1.0412 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LRB 6 | 0.78974 | 0.71042 | 0.62449 | 0.84881 | 0.68472 | 0.84348 | 0.91341 | 0.84872 | 1.1215 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-51-A-1 | 0.11008 | 0.28979 | 0.19731 | 0.17122 | 0.486 | 0.64782 | 0.84995 | 0.93071 | 1.0683 | IS3 | IS3 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-52-A-1 | 0.23688 | 0.35006 | 0.23143 | 0.29722 | 0.52446 | 0.67703 | 0.79544 | 0.95666 | 0.97635 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-53-A-1 | 0.24954 | 0.35916 | 0.49063 | 0.4927 | 0.76419 | 0.83895 | 0.94125 | 1.0497 | 1.1139 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-54-A-1 | 0.13228 | 0.27514 | 0.38622 | 0.37035 | 0.51759 | 0.61995 | 0.83273 | 0.94078 | 1.0467 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-55-A-1 | 0.14635 | 0.32293 | 0.46266 | 0.47794 | 0.84181 | 0.88645 | 0.91061 | 0.99993 | 1.0239 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-56-A-1 | 0.14624 | 0.27744 | 0.34619 | 0.33634 | 0.47991 | 0.58698 | 0.78107 | 0.93051 | 0.9626 | IS3 | IS3 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS3 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-57-A-1 | 0.12204 | 0.22852 | 0.37233 | 0.3603 | 0.50898 | 0.61214 | 0.81088 | 0.94057 | 1.0381 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-26-A-1B | 0.11529 | 0.29962 | 0.3159 | 0.3031 | 0.4994 | 0.57347 | 0.71266 | 0.79001 | 0.7051 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-58-A-1B | 0.24476 | 0.43077 | 0.41048 | 0.34452 | 0.54005 | 0.63891 | 0.773 | 0.889 | 0.95771 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-59-A-1B | 0.19 | 0.32875 | 0.43084 | 0.41115 | 0.65486 | 0.75558 | 0.85859 | 0.94888 | 0.71685 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-60-A-1B | 0.02669 | 0.15105 | 0.2313 | 0.22385 | 0.46363 | 0.567 | 0.7732 | 0.84364 | 0.27092 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | #N/A |  |  | IS4
LRB-7-inj2 | 0.77239 | 0.59912 | 0.96056 | 0.87749 | 0.96185 | 1.0018 | 0.904 | 0.85305 | 1.0556 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | b
LRB-7-inj3 | 0.13895 | 0.30258 | 0.39967 | 0.39282 | 0.58771 | 0.68323 | 0.84364 | 0.94808 | 1.0705 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-14-A-1B | 0.26859 | 0.25835 | 0.47948 | 0.47465 | 0.70094 | 0.75532 | 0.83091 | 0.88953 | 0.96581 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
LFB-H-3-Inj2 | 0.72282 | 0.56219 | 0.9096 | 0.84365 | 0.73069 | 0.65524 | 0.69989 | 0.62395 | 0.8163 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-58-A-1 | 0.11787 | 0.25906 | 0.3424 | 0.34362 | 0.52643 | 0.60969 | 0.75689 | 0.82129 | 0.8329 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | b
WO-23003-59-A-1 | 0.096981 | 0.36072 | 0.33081 | 0.30326 | 0.60735 | 0.688 | 0.814 | 0.86226 | 0.61435 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-60-A-1 | 0.010069 | 0.13409 | 0.18758 | 0.19225 | 0.46681 | 0.57 | 0.68318 | 0.70847 | 0.29535 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | #N/A |  |  | IS4
LRB7-inj1 | 0.66006 | 0.67672 | 0.92787 | 0.89331 | 0.55533 | 0.96197 | 0.95312 | 0.91402 | 1.0666 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | b
AB-230117-02 LRB 7 | 0.014032 | 0.0094236 | 0.010423 | 0.011529 | 0.023196 | 0.024027 | 0.034395 | 0.039122 | 0.04675 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | IS3 | IS4 | #N/A | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | b
AB-230117-02 LD 3 | 0.090985 | 0.26362 | 0.32595 | 0.33863 | 0.53237 | 0.64126 | 0.76472 | 0.83652 | 0.86922 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
AB-230117-02 LFB-H 4 | 0.65419 | 0.59861 | 0.9651 | 0.89672 | 0.81307 | 0.8394 | 0.80002 | 0.74536 | 0.85582 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 CCC-H 4 | 0.67162 | 0.62611 | 0.95857 | 0.87671 | 0.64991 | 0.72608 | 0.68447 | 0.6452 | 0.6999 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LFSM-H 3 | 0.093499 | 0.28413 | 0.38285 | 0.36716 | 0.54197 | 0.62191 | 0.6908 | 0.81118 | 0.77934 | IS3 | IS3 | #N/A | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS3 | IS3 |  | #N/A |  |  | #N/A
WO-23003-61-A-1 | 0.10861 | 0.24864 | 0.38669 | 0.37548 | 0.5501 | 0.65804 | 0.80926 | 0.9643 | 0.95032 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-62-A-1 | 0.16044 | 0.33015 | 0.43554 | 0.43193 | 0.62359 | 0.74269 | 0.8784 | 1.0153 | 1.0162 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | b
WO-23003-63-A-1 | 0.13065 | 0.33119 | 0.40933 | 0.38176 | 0.55949 | 0.69699 | 0.85317 | 1.0168 | 1.0193 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | b
WO-23003-64-A-1 | 0.093149 | 0.27241 | 0.34591 | 0.33298 | 0.51222 | 0.61336 | 0.8066 | 0.95357 | 0.9711 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | b
WO-23003-65-A-1 | 0.10928 | 0.30531 | 0.36685 | 0.35251 | 0.56956 | 0.65723 | 0.79148 | 0.98554 | 0.93464 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | b
WO-23003-66-A-1 | 0.22887 | 0.40659 | 0.4701 | 0.44894 | 0.6503 | 0.76403 | 0.83654 | 1.0465 | 0.90982 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | b
WO-23003-67-A-1 | 0.19819 | 0.42661 | 0.46438 | 0.43849 | 0.65109 | 0.74132 | 0.87757 | 1.0156 | 1.0314 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | b
WO-23003-68-A-1 | 0.19798 | 0.42312 | 0.44182 | 0.42549 | 0.62702 | 0.70552 | 0.76289 | 0.9642 | 0.88727 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | b
WO-23003-69-A-1 | 0.18831 | 0.43701 | 0.41271 | 0.4113 | 0.59616 | 0.71307 | 0.84807 | 1.0088 | 1.0165 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | b
WO-23003-70-A-1 | 0.1779 | 0.43942 | 0.42758 | 0.42103 | 0.61742 | 0.71285 | 0.85048 | 0.96027 | 1.0158 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | b
LRB-8-inj1 | 0.94336 | 0.99135 | 0.94367 | 0.93922 | 0.98037 | 1.031 | 1.0336 | 1.0424 | 1.1633 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LRB 8 | 0.9502 | 0.99486 | 0.99638 | 0.94369 | 0.92097 | 0.98301 | 0.98263 | 0.95729 | 1.0743 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-71-A-1 | 0.23674 | 0.46199 | 0.42616 | 0.42705 | 0.6248 | 0.7375 | 0.86029 | 0.9911 | 0.97791 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-72-A-1 | 0.18526 | 0.42712 | 0.40373 | 0.39443 | 0.61115 | 0.67544 | 0.73853 | 0.92232 | 0.86289 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-73-A-1 | 0.18863 | 0.43646 | 0.39616 | 0.41174 | 0.58628 | 0.70952 | 0.8503 | 0.99393 | 0.99028 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-74-A-1 | 0.25784 | 0.49034 | 0.45381 | 0.4435 | 0.6658 | 0.75774 | 0.88233 | 1.0064 | 1.0112 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-75-A-1 | 0.19233 | 0.4674 | 0.41945 | 0.40576 | 0.60339 | 0.69081 | 0.78564 | 0.98638 | 0.87596 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-76-A-1 | 0.22648 | 0.47847 | 0.452 | 0.44783 | 0.65122 | 0.76637 | 0.8961 | 1.0263 | 1.0854 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-77-A-1 | 0.27105 | 0.49677 | 0.48472 | 0.46548 | 0.71213 | 0.82189 | 0.94235 | 1.0514 | 1.1037 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-78-A-1 | 0.23923 | 0.47872 | 0.4632 | 0.44803 | 0.66903 | 0.7598 | 0.89378 | 1.0278 | 1.0358 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-79-A-1 | 0.20813 | 0.47653 | 0.43129 | 0.41717 | 0.64786 | 0.74864 | 0.8576 | 1.0186 | 1.0047 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-80-A-1 | 0.11886 | 0.36957 | 0.34632 | 0.34217 | 0.55445 | 0.65362 | 0.83492 | 1.0048 | 1.0158 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
AB-230117-02 LRB 9 | 1.0599 | 1.1304 | 1.0611 | 1.04 | 1.0059 | 1.0954 | 1.086 | 1.0176 | 1.2648 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
LD4-inj1 | 0.10396 | 0.33251 | 0.31109 | 0.31358 | 0.52664 | 0.61551 | 0.77318 | 0.91474 | 0.94795 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
AB-230117-02 LD 4 | 0.057294 | 0.26504 | 0.25569 | 0.25892 | 0.48871 | 0.57588 | 0.73187 | 0.86257 | 0.8885 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | #N/A |  |  | #N/A
AB-230117-02 LFB-H 5 | 1.0936 | 1.1213 | 1.0856 | 1.0047 | 0.84261 | 0.9254 | 0.89924 | 0.7869 | 0.99251 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 CCC-H 5 | 1.2107 | 1.1581 | 1.0704 | 1.0004 | 0.69954 | 0.79274 | 0.75212 | 0.66533 | 0.80091 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LFSM-H 4 | 0.24001 | 0.48026 | 0.44907 | 0.44585 | 0.61376 | 0.68363 | 0.72919 | 0.7979 | 0.79485 | IS3 | IS3 | #N/A | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS3 | IS3 |  | #N/A |  |  | #N/A
WO-23003-81-A-1 | 0.25696 | 0.48703 | 0.45709 | 0.4654 | 0.66544 | 0.78853 | 0.93705 | 1.0525 | 1.1405 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-82-A-1 | 0.14994 | 0.27103 | 0.25724 | 0.24961 | 0.38166 | 0.45141 | 0.55696 | 0.62248 | 0.64973 | IS3 | IS3 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5
WO-23003-83-A-1 | 0.24304 | 0.48053 | 0.45011 | 0.43497 | 0.65708 | 0.74708 | 0.91941 | 1.0382 | 1.097 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-84-A-1 | 0.26668 | 0.53316 | 0.4574 | 0.45236 | 0.69607 | 0.77583 | 0.82663 | 1.0145 | 0.94681 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-85-A-1 | 0.26243 | 0.50732 | 0.46384 | 0.47216 | 0.71459 | 0.80426 | 0.87902 | 1.0322 | 1.0028 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-86-A-1 | 0.1245 | 0.38601 | 0.35254 | 0.3746 | 0.59548 | 0.71942 | 0.8805 | 1.0194 | 1.061 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-87-A-1 | 0.082883 | 0.29738 | 0.29335 | 0.29461 | 0.50713 | 0.61517 | 0.80924 | 0.97989 | 1.0237 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-88-A-1 | 0.26865 | 0.5279 | 0.46859 | 0.46743 | 0.70107 | 0.78633 | 0.92774 | 1.0567 | 1.1117 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-89-A-1 | 0.22461 | 0.47307 | 0.45529 | 0.43844 | 0.67159 | 0.79121 | 0.9377 | 1.0708 | 1.1252 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-90-A-1 | 0.1892 | 0.41762 | 0.3997 | 0.40702 | 0.63907 | 0.74508 | 0.88426 | 1.0576 | 1.0271 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
LRB-10-inj1 | 1.0708 | 1.1535 | 1.0771 | 1.0684 | 1.044 | 1.1088 | 1.1256 | 1.0184 | 1.3283 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LRB 10 | 1.0787 | 1.1092 | 1.0752 | 1.0541 | 0.98743 | 1.0373 | 1.0563 | 0.95098 | 1.2271 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-91-A-1 | 0.13325 | 0.39037 | 0.37433 | 0.37513 | 0.614 | 0.74065 | 0.88914 | 1.0347 | 1.0651 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-92-A-1 | 0.30654 | 0.56074 | 0.51553 | 0.54735 | 0.78587 | 0.78787 | 0.8364 | 1.0748 | 1.0186 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-93-A-1 | 0.16996 | 0.4559 | 0.43393 | 0.4226 | 0.65628 | 0.78707 | 0.92348 | 1.0867 | 1.084 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-94-A-1 | 0.2479 | 0.54024 | 0.47993 | 0.47121 | 0.72026 | 0.84308 | 1.0328 | 1.1324 | 1.2398 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-95-A-1 | 0.23952 | 0.55931 | 0.48355 | 0.48376 | 0.72198 | 0.85678 | 1.035 | 1.1669 | 1.2682 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-96-A-1 | 0.23716 | 0.4993 | 0.46601 | 0.46135 | 0.69174 | 0.82018 | 0.95554 | 1.1129 | 1.1825 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-97-A-1 | 0.23383 | 0.52408 | 0.43197 | 0.45755 | 0.67644 | 0.7847 | 0.9024 | 1.1131 | 1.0875 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-98-A-1 | 0.24854 | 0.50389 | 0.47101 | 0.46009 | 0.66362 | 0.82472 | 1.0018 | 1.1163 | 1.2258 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-99-A-1 | 0.11351 | 0.39997 | 0.35558 | 0.35581 | 0.57339 | 0.69604 | 0.90444 | 1.0552 | 1.1695 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-100-A-1 | 0.20689 | 0.46555 | 0.44099 | 0.43576 | 0.62481 | 0.77041 | 0.97425 | 1.093 | 1.1965 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
AB-230117-02 LRB 11 | 1.0026 | 1.1517 | 1.0295 | 1.0265 | 1.0282 | 1.1109 | 1.1147 | 1.0308 | 1.3219 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
LD-5-inj1 | 0.22591 | 0.51425 | 0.47753 | 0.48396 | 0.78603 | 0.91028 | 1.0329 | 1.0836 | 1.1621 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | D
AB-230117-02 LD 5 | 0.19173 | 0.45067 | 0.44438 | 0.44157 | 0.7307 | 0.8439 | 0.95049 | 1.0096 | 1.0544 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | D
AB-230117-02 LFB-H 6 | 0.99583 | 1.1101 | 1.0207 | 0.98699 | 0.82201 | 0.9173 | 0.89842 | 0.77323 | 1.0236 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 CCC-H 6 | 1.0397 | 1.0945 | 0.99629 | 0.95244 | 0.7049 | 0.79886 | 0.75524 | 0.64155 | 0.82709 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LFSM-H 5 | 0.24552 | 0.61502 | 0.59417 | 0.5887 | 0.83758 | 0.87586 | 0.82823 | 0.86552 | 0.88374 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-101-A-1 | 0.19539 | 0.50849 | 0.45604 | 0.44772 | 0.68385 | 0.81108 | 0.99933 | 1.0648 | 1.1822 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-102-A-1 | 0.16456 | 0.37101 | 0.33914 | 0.33104 | 0.31273 | 0.38331 | 0.56493 | 0.68068 | 0.84696 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-103-A-1 | 0.17245 | 0.41654 | 0.36668 | 0.36262 | 0.45973 | 0.56166 | 0.82529 | 0.9413 | 1.0907 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-104-A-1 | 0.32626 | 0.60482 | 0.56988 | 0.54567 | 0.82223 | 0.96639 | 1.0686 | 1.1559 | 1.1482 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-105-A-1 | 0.12138 | 0.37207 | 0.35331 | 0.34921 | 0.50265 | 0.63551 | 0.85485 | 0.96776 | 1.0472 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-106-A-1 | 0.1688 | 0.37819 | 0.34449 | 0.34393 | 0.38682 | 0.48537 | 0.70744 | 0.84398 | 0.93676 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-107-A-1 | 0.21911 | 0.45481 | 0.38757 | 0.39483 | 0.54749 | 0.68265 | 0.89406 | 1.0009 | 1.0614 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-108-A-1 | 0.11133 | 0.32827 | 0.30325 | 0.29703 | 0.3619 | 0.45192 | 0.70429 | 0.85864 | 0.95729 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-109-A-1 | 0.22043 | 0.50014 | 0.45706 | 0.46745 | 0.74023 | 0.87683 | 1.0433 | 1.0961 | 1.1559 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-110-A-1 | 0.23447 | 0.4848 | 0.44383 | 0.45441 | 0.68069 | 0.83336 | 0.97751 | 1.075 | 0.78881 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
LRB-12-inj1 | 1.0082 | 1.1305 | 1.0324 | 1.0252 | 1.0356 | 1.1409 | 1.1782 | 1.0284 | 1.329 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | b
AB-230117-02 LRB 12 | 0.99239 | 1.1434 | 1.0296 | 1.0335 | 0.97684 | 1.0643 | 1.0939 | 0.94776 | 1.2797 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-111-A-1 | 0.1861 | 0.44387 | 0.41569 | 0.4229 | 0.64267 | 0.77831 | 0.97876 | 1.0541 | 0.78603 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-112-A-1 | 0.20356 | 0.49196 | 0.43143 | 0.42677 | 0.66268 | 0.80457 | 0.96421 | 1.0713 | 0.83891 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | b
WO-23003-113-A-1 | 0.19377 | 0.47696 | 0.42165 | 0.43629 | 0.66598 | 0.82826 | 0.99813 | 1.0928 | 0.82816 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | b
WO-23003-114-A-1 | 0.20394 | 0.46974 | 0.43287 | 0.44121 | 0.6722 | 0.8329 | 0.98357 | 1.0911 | 0.83724 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | b
WO-23003-115-A-1 | 0.24285 | 0.47836 | 0.44961 | 0.44166 | 0.68656 | 0.8402 | 0.96728 | 1.0911 | 0.76245 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-116-A-1 | 0.38402 | 0.56178 | 0.53455 | 0.53236 | 0.76383 | 0.88534 | 1.0393 | 0.9931 | 0.45825 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | IS4
WO-23003-117-A-1 | 0.37675 | 0.5866 | 0.53656 | 0.53663 | 0.78134 | 0.93377 | 1.069 | 1.043 | 0.46371 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | IS4
WO-23003-118-A-1 | 0.43256 | 0.61546 | 0.57391 | 0.58004 | 0.85134 | 0.97981 | 1.0841 | 1.0284 | 0.41873 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | IS4
WO-23003-119-A-1 | 0.41772 | 0.63582 | 0.57606 | 0.57154 | 0.8441 | 0.99999 | 1.0963 | 1.0993 | 0.69518 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-120-A-1 | 0.38128 | 0.58707 | 0.55626 | 0.56626 | 0.78306 | 0.95352 | 1.0832 | 1.0896 | 0.65899 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LRB 13 | 0.89631 | 1.0917 | 0.97255 | 0.95758 | 0.97965 | 1.0972 | 1.1489 | 0.97935 | 1.2976 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | b
LD6-inj1 | 0.94859 | 1.1027 | 0.96232 | 0.97896 | 0.9908 | 1.1329 | 1.1626 | 0.97682 | 0.36405 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | IS4
AB-230117-02 LD 6 | 0.93271 | 1.0759 | 0.97867 | 0.9776 | 0.96125 | 1.0758 | 1.0888 | 0.95214 | 0.34882 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | IS4
AB-230117-02 LFB-H 7 | 0.97648 | 1.0739 | 0.9758 | 0.94311 | 0.83002 | 0.96958 | 0.95758 | 0.79143 | 1.0061 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 CCC-H 7 | 0.96801 | 1.0119 | 0.94612 | 0.90538 | 0.70233 | 0.84009 | 0.83371 | 0.66376 | 0.89554 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LFSM-H 6 | 0.39098 | 0.55887 | 0.51799 | 0.51974 | 0.7028 | 0.82336 | 0.88762 | 0.82832 | 0.39936 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | IS3
WO-23003-121-A-1 | 0.4512 | 0.64471 | 0.573 | 0.57526 | 0.80666 | 0.95877 | 1.0597 | 1.0184 | 0.45426 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | IS4
WO-23003-122-A-1 | 0.8697 | 1.08 | 0.95769 | 0.95769 | 0.96562 | 1.1113 | 1.1233 | 0.96562 | 0.35966 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | IS4
WO-23003-123-A-1 | 0.93459 | 1.0641 | 0.97373 | 0.97824 | 1.021 | 1.1513 | 1.151 | 0.98852 | 0.35862 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | IS4
WO-23003-124-A-1 | 0.94436 | 1.0564 | 0.9372 | 0.96239 | 1.0107 | 1.1426 | 1.1698 | 1.0172 | 0.35518 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | IS4
Rinse12 | 0.035378 | 0.0090441 | 0.015374 | 0.014745 | 0.036817 | 0.046633 | 0.067221 | 0.067062 | 0.074152 | IS3 | IS4 | #N/A | IS3 | IS4 | #N/A | IS3 | IS3 | #N/A | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5
AB-230117-02 LRB 12 20230123102918 | 0.021075 | 0.0094965 | 0.0080253 | 0.0074573 | 0.0065706 | 0.0063807 | 0.0071016 | 0.0061783 | 0.0083543 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | IS3 | IS3 | #N/A | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5
AB-230117-02 LRB 13 20230123105625 | 1.3082 | 1.2143 | 1.1254 | 1.0667 | 0.92596 | 1.0019 | 1.0035 | 0.89655 | 1.1582 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | b
WO-23003-111-A-1B | 0.14307 | 0.38592 | 0.36391 | 0.3648 | 0.60522 | 0.71296 | 0.86664 | 0.95193 | 0.74219 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-121-A-1B | 0.52588 | 0.68523 | 0.60873 | 0.59614 | 0.80784 | 0.88461 | 0.94095 | 0.94628 | 0.41488 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | IS4
WO-23003-125-A-1 | 0.089163 | 0.3467 | 0.30007 | 0.31257 | 0.50052 | 0.56603 | 0.70156 | 0.90874 | 0.8618 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | C
WO-23003-126-A-1 | 0.083567 | 0.24888 | 0.23907 | 0.23408 | 0.25063 | 0.29453 | 0.39988 | 0.499 | 0.58769 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | IS4 | IS4 | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | #N/A
WO-23003-127-A-1 | 0.21801 | 0.39052 | 0.3648 | 0.36045 | 0.331 | 0.40869 | 0.59905 | 0.74386 | 0.87609 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-128-A-1 | 0.12439 | 0.32116 | 0.28604 | 0.29157 | 0.27683 | 0.32985 | 0.45857 | 0.5707 | 0.71643 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | IS4 | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 |  | #N/A
WO-23003-129-A-1 | 0.16563 | 0.3752 | 0.33398 | 0.33446 | 0.34022 | 0.41898 | 0.59187 | 0.80164 | 0.85484 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-130-A-1 | 0.10185 | 0.35625 | 0.32045 | 0.32532 | 0.54203 | 0.6456 | 0.78909 | 0.9755 | 1.0003 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
LRB-14-inj1 | 1.4346 | 1.3103 | 1.1626 | 1.1674 | 1.0795 | 1.1396 | 1.142 | 1.0423 | 1.3626 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LRB 14 | 1.4113 | 1.3292 | 1.1666 | 1.1748 | 0.9359 | 1.0375 | 1.0494 | 0.9257 | 1.2481 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-131-A-1 | 0.25845 | 0.46926 | 0.43365 | 0.41825 | 0.53581 | 0.65956 | 0.83253 | 0.99907 | 1.0595 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-132-A-1 | 0.086096 | 0.32116 | 0.2964 | 0.30224 | 0.43778 | 0.55445 | 0.76427 | 0.9362 | 1.0335 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-133-A-1 | 0.1697 | 0.37059 | 0.33238 | 0.32979 | 0.35957 | 0.4384 | 0.60029 | 0.8082 | 0.82415 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-134-A-1 | 0.26907 | 0.49423 | 0.44431 | 0.44555 | 0.60525 | 0.70345 | 0.82296 | 1.0815 | 1.0568 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-135-A-1 | 0.28306 | 0.47483 | 0.41531 | 0.41388 | 0.46445 | 0.53415 | 0.71368 | 0.90674 | 0.91738 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-136-A-1 | 0.02904 | 0.21164 | 0.20852 | 0.20682 | 0.4105 | 0.4897 | 0.70762 | 0.9342 | 0.9363 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-137-A-1 | 0.36227 | 0.6164 | 0.51466 | 0.52179 | 0.7402 | 0.83092 | 0.91501 | 1.0809 | 1.0892 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-138-A-1 | 0.26562 | 0.5121 | 0.47646 | 0.49243 | 0.71125 | 0.82157 | 0.89669 | 1.0509 | 1.0799 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-139-A-1 | 0.20561 | 0.27131 | 0.40999 | 0.37724 | 0.56057 | 0.67401 | 0.82603 | 0.89499 | 0.97231 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | C
WO-23003-139-A-1B | 0.13087 | 0.33452 | 0.32158 | 0.32273 | 0.54557 | 0.66211 | 0.85058 | 0.89725 | 1.0687 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | C
WO-23003-140-A-1 | 0.13768 | 0.36454 | 0.31607 | 0.32963 | 0.52612 | 0.66292 | 0.87206 | 0.9426 | 1.1082 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
LRB-15-inj1 | 0.96772 | 1.0485 | 0.91735 | 0.93569 | 0.93671 | 1.0232 | 1.0632 | 0.94705 | 1.2829 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LRB 15 | 0.82915 | 0.97957 | 0.86114 | 0.87449 | 0.77803 | 0.84927 | 0.92659 | 0.79889 | 1.0702 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LD 7 | 0.048704 | 0.23157 | 0.21104 | 0.23276 | 0.45074 | 0.51902 | 0.65208 | 0.85281 | 0.92025 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | #N/A |  |  | #N/A
AB-230117-02 LFB-H 8 | 1.0221 | 1.0345 | 0.90175 | 0.92128 | 0.78402 | 0.88514 | 0.91535 | 0.76648 | 1.074 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 CCC-H 8 | 1.0412 | 1.0719 | 0.89488 | 0.88596 | 0.63902 | 0.76513 | 0.77073 | 0.60179 | 0.86999 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LFSM-H 7 | 0.080461 | 0.29607 | 0.26639 | 0.27738 | 0.47932 | 0.55062 | 0.65155 | 0.77234 | 0.74829 | IS3 | IS3 | #N/A | IS3 | IS3 | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS3 | IS3 | IS3 | #N/A |  |  | #N/A
WO-23003-141-A-1 | 0.033738 | 0.20172 | 0.18444 | 0.1961 | 0.37545 | 0.49136 | 0.69574 | 0.84178 | 0.91279 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-142-A-1 | 0.11802 | 0.3642 | 0.32149 | 0.33374 | 0.58344 | 0.66356 | 0.73568 | 0.94887 | 1.0142 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-143-A-1 | 0.14813 | 0.39973 | 0.33617 | 0.35469 | 0.5662 | 0.70756 | 0.92312 | 0.97366 | 1.2246 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-144-A-1 | 0.15591 | 0.40127 | 0.35154 | 0.37495 | 0.5333 | 0.70471 | 0.94636 | 1.0123 | 1.2242 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-145-A-1 | 0.19795 | 0.42812 | 0.37489 | 0.39467 | 0.60169 | 0.73301 | 0.86776 | 0.98251 | 1.0561 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-146-A-1 | 0.16846 | 0.39949 | 0.36271 | 0.39009 | 0.60903 | 0.72621 | 0.80591 | 0.98639 | 0.98055 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-147-A-1 | 0.1369 | 0.36985 | 0.35092 | 0.36071 | 0.5647 | 0.70731 | 0.85759 | 0.96785 | 1.09 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-148-A-1 | 0.11763 | 0.32769 | 0.3078 | 0.32707 | 0.53146 | 0.66929 | 0.82701 | 0.94865 | 1.0325 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-149-A-1 | 0.090736 | 0.316 | 0.27851 | 0.29514 | 0.5044 | 0.63582 | 0.83003 | 0.95785 | 1.0177 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-150-A-1 | 0.087186 | 0.30634 | 0.28179 | 0.2999 | 0.48098 | 0.62839 | 0.87478 | 0.97262 | 1.1436 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | #N/A |  |  | #N/A
LRB-16-inj1 | 0.97742 | 1.0614 | 0.90316 | 0.93611 | 0.93369 | 1.0395 | 1.0795 | 0.94752 | 1.3124 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LRB 16 | 0.87395 | 1.1121 | 0.9461 | 0.96309 | 0.84045 | 0.967 | 0.9886 | 0.84596 | 1.1721 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-151-A-1 | 0.178 | 0.44395 | 0.35916 | 0.36673 | 0.38727 | 0.5037 | 0.7351 | 0.83437 | 1.0452 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-152-A-1 | 0.096734 | 0.3179 | 0.27567 | 0.28714 | 0.36532 | 0.47137 | 0.70104 | 0.84479 | 1.021 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-153-A-1 | 0.1247 | 0.3467 | 0.2931 | 0.29868 | 0.28719 | 0.35296 | 0.49683 | 0.65766 | 0.78684 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | IS4 | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 | IS4 |  | #N/A
WO-23003-154-A-1 | 0.1739 | 0.41919 | 0.3854 | 0.40327 | 0.47818 | 0.6096 | 0.76492 | 0.95268 | 1.0271 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-155-A-1 | 0.17534 | 0.42088 | 0.39872 | 0.40962 | 0.52809 | 0.6497 | 0.83438 | 0.96042 | 1.1005 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-156-A-1 | 0.2248 | 0.51008 | 0.45314 | 0.4674 | 0.65921 | 0.71412 | 0.76862 | 0.95008 | 0.97416 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-157-A-1 | 0.04161 | 0.22388 | 0.20906 | 0.22259 | 0.37243 | 0.44684 | 0.6041 | 0.81934 | 0.84064 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-158-A-1 | 0.25432 | 0.52249 | 0.45454 | 0.45897 | 0.58846 | 0.69643 | 0.84045 | 0.99984 | 1.1118 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-159-A-1 | 0.10428 | 0.35231 | 0.3354 | 0.34823 | 0.49454 | 0.63695 | 0.80788 | 0.98186 | 1.0884 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-160-A-1 | 0.14879 | 0.38108 | 0.36899 | 0.39936 | 0.52478 | 0.55281 | 0.64354 | 0.87617 | 0.91866 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
LRB-17-inj1 | 1.3277 | 1.4183 | 1.1765 | 1.162 | 1.0007 | 1.0814 | 1.0327 | 1.0644 | 1.2214 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LRB 17 | 1.2934 | 1.4003 | 1.1511 | 1.1765 | 0.94697 | 1.0235 | 1.0057 | 0.94389 | 1.1826 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LD 8 | 0.11295 | 0.33025 | 0.31685 | 0.32951 | 0.5075 | 0.60834 | 0.72624 | 0.94228 | 0.94724 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
AB-230117-02 LFB-H 9 | 1.4979 | 1.578 | 1.36 | 1.2492 | 0.80881 | 0.89649 | 0.84566 | 0.78238 | 0.97881 | FALSE | IS1 | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  | IS1 | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 CCC-H 9 | 1.3756 | 1.3467 | 1.1313 | 1.0303 | 0.64246 | 0.71117 | 0.68058 | 0.61043 | 0.75941 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LFSM-H 8 | 0.18429 | 0.42505 | 0.40665 | 0.42496 | 0.56163 | 0.57022 | 0.58471 | 0.74216 | 0.66723 | IS3 | IS3 | #N/A | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS3 | IS3 |  | #N/A |  |  | #N/A
WO-23003-161-A-1 | 0.1909 | 0.43301 | 0.44226 | 0.4461 | 0.61919 | 0.67156 | 0.73609 | 0.92323 | 0.91834 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-162-A-1 | 0.14664 | 0.39901 | 0.38025 | 0.38482 | 0.59026 | 0.68338 | 0.77845 | 0.98289 | 0.99276 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-163-A-1 | 0.044741 | 0.22702 | 0.20047 | 0.21106 | 0.39071 | 0.49597 | 0.6719 | 0.83134 | 0.93915 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-164-A-1 | 0.090601 | 0.30051 | 0.29534 | 0.31138 | 0.48802 | 0.56317 | 0.65243 | 0.87571 | 0.87355 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-165-A-1 | 0.19121 | 0.38427 | 0.42486 | 0.43774 | 0.61531 | 0.68113 | 0.73613 | 0.93493 | 0.94629 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-166-A-1 | 0.37449 | 0.48927 | 0.557 | 0.56036 | 0.75206 | 0.79299 | 0.79672 | 0.98526 | 0.97207 | IS3 | IS4 | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-167-A-1 | 0.18167 | 0.38271 | 0.42149 | 0.42938 | 0.61579 | 0.64289 | 0.70152 | 0.91665 | 0.89694 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-168-A-1 | 0.29489 | 0.45677 | 0.50006 | 0.52349 | 0.70758 | 0.69661 | 0.71938 | 0.94616 | 0.91512 | IS3 | IS4 | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-169-A-1 | 0.032883 | 0.2051 | 0.18816 | 0.18529 | 0.38179 | 0.48472 | 0.69129 | 0.8795 | 0.95468 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-170-A-1 | 0.13278 | 0.43354 | 0.36047 | 0.37652 | 0.5681 | 0.68754 | 0.81122 | 0.93468 | 1.0872 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
LRB-18-inj1 | 1.2745 | 1.3502 | 1.1496 | 1.1272 | 0.95698 | 1.0499 | 0.99505 | 0.90752 | 1.2301 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LRB 18 | 1.3165 | 1.3734 | 1.1279 | 1.1176 | 0.96268 | 1.0103 | 0.95535 | 0.88755 | 1.1743 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-171-A-1 | 0.26739 | 0.56424 | 0.45646 | 0.46682 | 0.63135 | 0.72508 | 0.7784 | 0.92916 | 1.0008 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-172-A-1 | 0.16921 | 0.50302 | 0.40294 | 0.41582 | 0.59915 | 0.69766 | 0.79536 | 0.93571 | 1.0313 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-173-A-1 | 0.038693 | 0.22627 | 0.21136 | 0.20674 | 0.39032 | 0.48868 | 0.68518 | 0.83495 | 1.0067 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-174-A-1 | 0.16161 | 0.47485 | 0.38277 | 0.39163 | 0.57563 | 0.68331 | 0.81062 | 0.95084 | 1.0953 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-175-A-1 | 0.20365 | 0.4993 | 0.4182 | 0.43521 | 0.59215 | 0.69128 | 0.80627 | 0.91567 | 1.0769 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-176-A-1 | 0.17248 | 0.45211 | 0.38809 | 0.38569 | 0.48364 | 0.6014 | 0.78137 | 0.88963 | 1.0666 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-177-A-1 | 0.1313 | 0.41121 | 0.33738 | 0.35309 | 0.50579 | 0.61406 | 0.76624 | 0.92265 | 1.0721 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-178-A-1 | 0.13685 | 0.3888 | 0.30947 | 0.30813 | 0.31989 | 0.38698 | 0.54002 | 0.7092 | 0.79387 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  | #N/A
WO-23003-179-A-1 | 0.10669 | 0.31832 | 0.25023 | 0.25401 | 0.23625 | 0.27661 | 0.36882 | 0.44489 | 0.60231 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | IS4 | IS4 | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | IS4 | #N/A
WO-23003-180-A-1 | 0.11344 | 0.36032 | 0.28704 | 0.31132 | 0.37608 | 0.45342 | 0.58955 | 0.77618 | 0.81136 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | IS4 |  |  | #N/A
LRB-19-inj1 | 1.2671 | 1.3176 | 1.0893 | 1.0827 | 1.0074 | 1.0421 | 1.0037 | 0.86263 | 1.2456 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LRB 19 | 1.3013 | 1.3225 | 1.1149 | 1.1259 | 0.95072 | 1.0001 | 0.97904 | 0.88643 | 1.1907 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LD 9 | 0.13472 | 0.42858 | 0.34626 | 0.36141 | 0.54333 | 0.65387 | 0.79148 | 0.94107 | 1.0335 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
AB-230117-02 LFB-H 10 | 1.3001 | 1.3306 | 1.0779 | 1.0315 | 0.72074 | 0.80635 | 0.76701 | 0.6757 | 0.8967 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 CCC-H 10 | 1.3514 | 1.2858 | 1.0457 | 0.9752 | 0.6219 | 0.67841 | 0.64781 | 0.59302 | 0.724 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LFSM-H 9 | 0.039016 | 0.22417 | 0.19797 | 0.20425 | 0.37173 | 0.44666 | 0.58473 | 0.66903 | 0.77738 | IS3 | IS3 | #N/A | IS3 | IS3 | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  |  | #N/A
WO-23003-181-A-1 | 0.21739 | 0.52273 | 0.42529 | 0.43522 | 0.60862 | 0.71677 | 0.84296 | 0.91487 | 1.0947 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-182-A-1 | 0.20212 | 0.50805 | 0.4071 | 0.42492 | 0.60799 | 0.71203 | 0.81264 | 0.93417 | 1.0879 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-183-A-1 | 0.22662 | 0.5122 | 0.42536 | 0.42874 | 0.56863 | 0.69281 | 0.81687 | 0.91944 | 1.104 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-184-A-1 | 0.20857 | 0.54452 | 0.43951 | 0.45503 | 0.69424 | 0.79466 | 0.87078 | 0.93221 | 1.0829 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-185-A-1 | 0.21323 | 0.51399 | 0.4137 | 0.4271 | 0.58943 | 0.71886 | 0.82809 | 0.91938 | 1.0936 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-186-A-1 | 0.20525 | 0.48876 | 0.43226 | 0.43368 | 0.60198 | 0.71752 | 0.81258 | 0.91765 | 1.0976 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-187-A-1 | 0.18608 | 0.48043 | 0.38786 | 0.40613 | 0.56985 | 0.68709 | 0.82426 | 0.91854 | 1.1135 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
WO-23003-188-A-1 | 0.19023 | 0.47527 | 0.39597 | 0.40326 | 0.59067 | 0.69715 | 0.83191 | 0.91103 | 1.1223 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-189-A-1 | 0.23441 | 0.5162 | 0.41846 | 0.42726 | 0.60428 | 0.68083 | 0.663 | 0.73708 | 1.039 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-190-A-1 | 0.10924 | 0.39685 | 0.32195 | 0.331 | 0.49871 | 0.61785 | 0.66319 | 0.73143 | 1.0242 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | #N/A |  |  | #N/A
LRB-20-inj1 | 1.2051 | 1.2926 | 1.109 | 1.1031 | 0.96912 | 1.0462 | 1.0206 | 0.89501 | 1.2539 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  | b | #N/A
AB-230117-02 LRB 20 | 1.1977 | 1.3085 | 1.1233 | 1.0769 | 0.93254 | 1.023 | 1.0026 | 0.87917 | 1.209 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
WO-23003-191-A-1 | 0.042985 | 0.24612 | 0.21103 | 0.22161 | 0.37879 | 0.49207 | 0.6138 | 0.70093 | 0.97294 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | IS3 |  |  | #N/A
WO-23003-192-A-1 | 0.18515 | 0.484 | 0.40184 | 0.40945 | 0.5739 | 0.6617 | 0.69404 | 0.73493 | 1.0864 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-193-A-1 | 0.17773 | 0.47679 | 0.39335 | 0.40009 | 0.56761 | 0.6771 | 0.71078 | 0.73701 | 1.1013 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-194-A-1 | 0.22365 | 0.51809 | 0.41614 | 0.4365 | 0.59488 | 0.69818 | 0.65994 | 0.69591 | 1.0688 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-195-A-1 | 0.17605 | 0.46674 | 0.39499 | 0.39412 | 0.5666 | 0.67678 | 0.67832 | 0.70198 | 1.0977 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
WO-23003-196-A-1 | 0.096956 | 0.33879 | 0.28347 | 0.29681 | 0.46827 | 0.56513 | 0.62986 | 0.68998 | 1.0383 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | #N/A |  |  | #N/A
WO-23003-197-A-1 | 0.18751 | 0.48123 | 0.39257 | 0.41593 | 0.51924 | 0.42616 | 0.19415 | 0.20669 | 1.0183 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | IS3 | IS4 | #N/A | IS3 | IS4 | IS3 | IS4 |  | IS3 | IS3 | IS4 | #N/A
WO-23003-198-A-1 | 0.060427 | 0.27083 | 0.22901 | 0.24578 | 0.38876 | 0.35388 | 0.16772 | 0.20753 | 0.82986 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | IS3 | IS4 | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | #N/A
WO-23003-199-A-1 | 0.28283 | 0.54282 | 0.4652 | 0.47474 | 0.60684 | 0.59898 | 0.34409 | 0.35332 | 1.0249 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | IS3 | IS4 | #N/A | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5 | IS5
WO-23003-200-A-1 | 0.20712 | 0.49444 | 0.39993 | 0.41231 | 0.54557 | 0.50771 | 0.22146 | 0.22629 | 1.0496 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | IS3 | IS4 | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A | IS3 | IS4 | #N/A
LRB-21-inj1 | 1.1979 | 1.222 | 1.0237 | 1.0491 | 0.96216 | 1.0307 | 0.92962 | 0.77931 | 1.2955 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  | b | #N/A
AB-230117-02 LRB 21 | 0.97206 | 1.174 | 0.96849 | 0.93715 | 0.76326 | 0.83486 | 0.84082 | 0.73908 | 0.9976 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  | b | #N/A
AB-230117-02 LD 10 | 0.030763 | 0.20774 | 0.17952 | 0.18665 | 0.34696 | 0.35564 | 0.17297 | 0.20835 | 0.81912 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | IS3 | IS4 | #N/A | IS3 | IS4 | IS3 | IS4 | IS4 | IS3 | IS3 | IS4 | #N/A
AB-230117-02 LFB-H 11 | 1.1213 | 1.1339 | 0.91182 | 0.90249 | 0.68207 | 0.74635 | 0.6716 | 0.60087 | 0.82151 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 CCC-H 11 | 1.2609 | 1.2237 | 0.99796 | 0.92017 | 0.59558 | 0.64824 | 0.62181 | 0.5494 | 0.69882 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
AB-230117-02 LFSM-H 10 | 0.24253 | 0.51795 | 0.42937 | 0.42623 | 0.51164 | 0.44314 | 0.17914 | 0.20989 | 0.68329 | IS3 | FALSE | #N/A | IS3 | FALSE | #N/A | IS3 | IS3 | #N/A | IS3 |  | IS3 | IS3 |  | IS3 | IS3 | IS3 | #N/A
WO-23003-201-A-1 | 0.1807 | 0.46557 | 0.37099 | 0.38912 | 0.52023 | 0.46491 | 0.197 | 0.20831 | 0.9417 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | IS3 | IS4 | #N/A | IS3 | IS4 | IS3 | IS4 |  | IS3 | IS3 | IS4 | #N/A
WO-23003-202-A-1 | 0.21935 | 0.47 | 0.39369 | 0.40322 | 0.53557 | 0.47604 | 0.19163 | 0.21578 | 0.95681 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | IS3 | IS4 | #N/A | IS3 | IS4 | IS3 | IS4 |  | IS3 | IS3 | IS4 | #N/A
WO-23003-203-A-1 | 0.099399 | 0.33393 | 0.29515 | 0.30625 | 0.43622 | 0.21696 | 0.064762 | 0.091736 | 0.73537 | IS3 | IS4 | #N/A | IS4 | IS3 | #N/A | IS3 | IS4 | #N/A | IS3 | IS4 | IS3 | IS4 | IS3 | IS3 | IS3 | UL | #N/A
WO-23003-204-A-1 | 0.12322 | 0.39035 | 0.32718 | 0.34228 | 0.46012 | 0.23043 | 0.070559 | 0.10352 | 0.74887 | IS3 | IS4 | #N/A | IS4 | IS3 | #N/A | IS3 | IS4 | #N/A | IS3 | IS4 | IS3 | IS4 | IS3 | IS3 | IS3 | UL | #N/A
WO-23003-206-A-1 | 0.43175 | 0.67939 | 0.55121 | 0.54143 | 0.65655 | 0.33029 | 0.080639 | 0.10008 | 0.96403 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | #N/A | IS3 |  | #N/A |  |  | IS3 | IS3 | UL | #N/A
WO-23003-207-A-1 | 0.36649 | 0.61077 | 0.50084 | 0.51372 | 0.66001 | 0.38591 | 0.10889 | 0.13306 | 1.0361 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | #N/A | IS3 |  | #N/A |  |  | IS3 | IS3 | UL | #N/A
WO-23003-208-A-1 | 0.37676 | 0.6259 | 0.52447 | 0.51242 | 0.64978 | 0.39919 | 0.11512 | 0.13509 | 1.0385 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | #N/A | IS3 |  | #N/A |  |  | IS3 | IS3 | UL | #N/A
WO-23003-209-A-1 | 0.51986 | 0.69143 | 0.56804 | 0.57541 | 0.69847 | 0.41613 | 0.11228 | 0.12392 | 1.0415 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | #N/A |  |  | #N/A |  |  | IS3 | IS3 | UL | #N/A
WO-23003-210-A-1 | 1.1427 | 1.1935 | 1.016 | 1.0231 | 0.955 | 0.91267 | 0.21931 | 0.19324 | 1.1542 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS4 | IS4 | #N/A |  |  | #N/A |  |  | #N/A | IS4 | IS4 | #N/A
WO-23003-211-A-1 | 1.2415 | 1.2505 | 1.0915 | 1.0394 | 0.98808 | 0.97737 | 0.2351 | 0.21508 | 1.1834 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS4 | IS4 | #N/A |  |  | #N/A |  |  | #N/A | IS4 | IS4 | #N/A
AB-230117-02 LRB 22 | 1.266 | 1.2915 | 1.0762 | 1.0753 | 0.98166 | 1.0332 | 0.87087 | 0.69719 | 1.2516 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  | b | #N/A
WO-23003-212-A-1 | 1.244 | 1.2665 | 1.0584 | 1.0348 | 0.98073 | 0.9552 | 0.23957 | 0.20663 | 1.1579 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS4 | IS4 | #N/A |  |  | #N/A |  |  | #N/A | IS4 | IS4 | #N/A
AB-230117-02 CCC-H 12 | 1.0884 | 1.076 | 0.89524 | 0.83192 | 0.57472 | 0.62504 | 0.53852 | 0.4724 | 0.67173 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | IS3 | #N/A |  |  | #N/A |  |  | #N/A |  | C | #N/A
AB-230117-02 QCS 2 | 1.0985 | 1.0844 | 0.92359 | 0.90866 | 0.69112 | 0.72457 | 0.71468 | 0.64952 | 0.88837 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A
LRB-22-reprep | 1.608 | 1.403 | 1.2265 | 1.1568 | 0.86125 | 0.97396 | 0.89127 | 0.77713 | 1.0676 | IS1 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS1 |  | #N/A |  |  | #N/A |  | b | #N/A
LRB-22-reprep-inj2 | 1.5686 | 1.3687 | 1.1084 | 1.0602 | 0.74478 | 0.84418 | 0.79407 | 0.70554 | 0.94223 | IS1 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS1 |  | #N/A |  |  | #N/A |  | b | #N/A
WO-23003-203-A-1B | 1.3701 | 1.3041 | 1.0967 | 1.0677 | 0.90697 | 0.92219 | 0.54907 | 0.41992 | 1.1726 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | IS4 | #N/A |  |  | #N/A |  |  | #N/A |  | IS4 | #N/A
WO-23003-204-A-1B | 0.86881 | 0.9028 | 0.75046 | 0.76509 | 0.82842 | 0.6884 | 0.2828 | 0.24575 | 1.1067 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | #N/A |  |  | #N/A |  |  | #N/A | IS3 | IS4 | #N/A
WO-23003-205-A-1B | 0.8086 | 0.87289 | 0.70746 | 0.72536 | 0.81707 | 0.68557 | 0.28784 | 0.25401 | 1.0752 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | #N/A |  |  | #N/A |  |  | #N/A | IS3 | IS4 | #N/A
WO-23003-206-A-1B | 1.358 | 1.3271 | 1.0867 | 1.0876 | 0.95788 | 0.92759 | 0.44236 | 0.33743 | 1.1883 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS4 | IS4 | #N/A |  |  | #N/A |  |  | #N/A | IS4 | IS4 | #N/A
WO-23003-207-A-1B | 1.2761 | 1.291 | 1.0304 | 1.0592 | 0.93879 | 0.95616 | 0.55127 | 0.41384 | 1.2186 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | IS4 | #N/A |  |  | #N/A |  |  | #N/A |  | IS4 | #N/A
WO-23003-208-A-1B | 1.2313 | 1.2645 | 1.0334 | 1.0496 | 0.96143 | 0.95886 | 0.57232 | 0.4317 | 1.2138 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | IS4 | #N/A |  |  | #N/A |  |  | #N/A |  | IS4 | #N/A
217288-d4-reprep | 0.43131 | 0.53118 | 0.53167 | 0.55306 | 0.76385 | 0.87329 | 0.87409 | 0.92385 | 1.0259 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | #N/A
217289-d4-reprep | 0.46086 | 0.56871 | 0.56802 | 0.57407 | 0.80759 | 0.91305 | 0.88287 | 1.0018 | 1.1413 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | #N/A
217315-d2-reprep | 0.14444 | 0.42587 | 0.32588 | 0.36653 | 0.54753 | 0.69461 | 0.80595 | 0.94207 | 1.0677 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
217356-d2-reprep | 0.17543 | 0.48956 | 0.3994 | 0.4021 | 0.61526 | 0.77116 | 0.82908 | 0.94078 | 1.0696 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
217357-d2-reprep | 0.15012 | 0.45595 | 0.37024 | 0.39285 | 0.58944 | 0.72159 | 0.81566 | 0.95077 | 1.0122 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | #N/A
217358-d2-reprep | 0.18238 | 0.48448 | 0.40013 | 0.41172 | 0.58927 | 0.73768 | 0.825 | 0.93255 | 0.34125 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | IS4
217359-d2-reprep | 0.17825 | 0.50212 | 0.40052 | 0.40839 | 0.56213 | 0.71665 | 0.85783 | 0.93438 | 0.15907 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS4 | IS4 |  | #N/A |  |  | IS4
217360-d2-reprep | 0.099339 | 0.39125 | 0.31128 | 0.33498 | 0.52062 | 0.64675 | 0.73213 | 0.92501 | 0.76183 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | C
217393-d4-reprep | 0.45775 | 0.73377 | 0.58355 | 0.60507 | 0.81776 | 0.93196 | 0.91634 | 0.96335 | 0.55769 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | #N/A
217394-d4-reprep | 0.44961 | 0.71073 | 0.56048 | 0.57466 | 0.78777 | 0.94421 | 0.95093 | 0.97079 | 0.50384 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | #N/A
217395-d4-reprep | 0.45861 | 0.71412 | 0.57803 | 0.60411 | 0.82448 | 0.94725 | 0.95819 | 0.96917 | 0.34193 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | IS4
217399-d2-reprep | 0.16663 | 0.49792 | 0.36406 | 0.40773 | 0.57518 | 0.65234 | 0.67232 | 0.82881 | 0.83573 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
217400-d2-reprep | 0.15714 | 0.52349 | 0.44363 | 0.48322 | 0.80118 | 0.86686 | 0.77802 | 0.92473 | 0.94646 | IS3 | FALSE | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS3 | IS4 |  | #N/A |  |  | #N/A
217295-d1000-LD-reprep | 1.7584 | 1.7807 | 1.4692 | 1.3988 | 1.042 | 1.101 | 0.67132 | 0.50174 | 1.2998 | IS1 | IS2 | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS1 | IS2 | #N/A |  |  | #N/A |  |  | #N/A
217335-d2-LD-reprep | 0.10704 | 0.41247 | 0.32574 | 0.34813 | 0.55931 | 0.70676 | 0.80483 | 0.90609 | 1.0068 | IS3 | IS3 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS4 | IS4 |  | #N/A |  |  | #N/A
217375-d2-LD-reprep | 0.089191 | 0.34036 | 0.28868 | 0.30145 | 0.41349 | 0.54486 | 0.67931 | 0.82905 | 0.92565 | IS3 | IS4 | #N/A | IS4 | IS4 | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 | IS4 | #N/A |  |  | #N/A
217395-d4-LD-reprep | 0.44625 | 0.73832 | 0.5611 | 0.5827 | 0.80426 | 0.95955 | 0.95062 | 0.97576 | 0.34981 | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | #N/A |  |  | #N/A |  |  | IS4
217359-d2-LD | 0.14315 | 0.41936 | 0.34559 | 0.36695 | 0.51374 | 0.62777 | 0.75137 | 0.82037 | 0.16061 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS4 | IS4 |  | #N/A |  |  | IS4
217415-d2-LD-reprep | 0.10363 | 0.39381 | 0.30622 | 0.33287 | 0.50938 | 0.58899 | 0.65549 | 0.83243 | 0.81266 | IS3 | IS4 | #N/A | IS4 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 | IS4 | IS3 | IS4 |  | #N/A |  |  | #N/A
217296-d1000-LFSM-200-reprep | 1.5234 | 1.5032 | 1.1842 | 1.14 | 0.706 | 0.79408 | 0.52036 | 0.39492 | 0.85816 | IS1 | IS1 | #N/A | FALSE | FALSE | #N/A | FALSE | IS3 | #N/A | IS1 | IS1 | #N/A |  |  | #N/A |  | IS3 | #N/A
217336-d2-LFSM-200-reprep | 0.026832 | 0.18319 | 0.16441 | 0.17122 | 0.34062 | 0.42825 | 0.52506 | 0.66072 | 0.66259 | IS3 | IS3 | #N/A | IS3 | IS3 | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS3 | IS3 | IS3 | IS3 |  |  | #N/A
217376-d2-LFSM-200-reprep | 0.1306 | 0.5199 | 0.41961 | 0.4334 | 0.66485 | 0.73418 | 0.69059 | 0.80325 | 0.76678 | IS3 | FALSE | #N/A | IS3 | FALSE | #N/A | FALSE | FALSE | #N/A | IS3 |  | IS3 | IS3 |  | #N/A |  |  | #N/A
217396-d200-LFSM-200 | 1.5743 | 1.5707 | 1.2698 | 1.1978 | 0.78451 | 0.86529 | 0.80803 | 0.68592 | 0.25623 | IS1 | IS1 | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS1 | IS1 | #N/A |  |  | #N/A |  |  | IS3
217360-d2-LFSM-200 | 0.11343 | 0.40427 | 0.31691 | 0.32195 | 0.48479 | 0.55449 | 0.63014 | 0.71673 | 0.66679 | IS3 | IS3 | #N/A | IS3 | IS3 | #N/A | FALSE | FALSE | #N/A | IS3 | IS3 | IS3 | IS3 | IS3 | #N/A |  |  | C
217416-d2-LFSM-200-reprep | 0.10534 | 0.40718 | 0.31828 | 0.32526 | 0.51619 | 0.49116 | 0.4851 | 0.68022 | 0.62008 | IS3 | IS3 | #N/A | IS3 | FALSE | #N/A | IS3 | FALSE | #N/A | IS3 | IS3 | IS3 | IS3 |  | IS3 | IS3 |  | #N/A
LRB-22-reprep | 1.608 | 1.403 | 1.2265 | 1.1568 | 0.86125 | 0.97396 | 0.89127 | 0.77713 | 1.0676 | IS1 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS1 | IS2 | #N/A |  | b | b | b | b | b
LRB-22-reprep-inj2 | 1.5686 | 1.3687 | 1.1084 | 1.0602 | 0.74478 | 0.84418 | 0.79407 | 0.70554 | 0.94223 | IS1 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS1 | IS2 | #N/A |  |  | #N/A |  |  | #N/A
QCS-2-Reinject | 1.5861 | 1.4258 | 1.1479 | 1.0963 | 0.67159 | 0.79662 | 0.72721 | 0.62208 | 0.81796 | IS1 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS1 |  | #N/A |  |  | #N/A |  |  | #N/A
LFB-H-11-reinject | 1.6172 | 1.4808 | 1.1842 | 1.0951 | 0.6704 | 0.80032 | 0.75155 | 0.6092 | 0.86629 | IS1 | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | IS1 |  | #N/A |  |  | #N/A |  |  | #N/A
CCC-H-12-reinject | 1.0468 | 1.1692 | 0.91571 | 0.88003 | 0.57132 | 0.64168 | 0.63216 | 0.52303 | 0.71385 | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A | FALSE | FALSE | #N/A |  |  | #N/A |  |  | #N/A |  |  | #N/A

## Processing Narrative
1. Lots of manual integration for PFBA needed through the whole batch.
2. Two samples had seemingly bad internal standard spikes, either 2-fold or it looks like half, and will be reanalyzed.
3. Lots of carryover again. Some samples were rerun as a result.
4. Due to the carryover issues and general concerns about it I reprepared numerous extra LDs/LFSMS. These are on the QAQC overflow along with the extra LRBs and LFBs